## PowerPivot
Channel | Percent Diff | Avg Landscape Centred | Avg Landscape Repeated Centred
Blue | -0.03709% | 25026.2829 | 25016.9998
Green | -0.1459% | 40410.2298 | 40351.2767
Red | -0.1904% | 33158.2855 | 33095.1494

## Original
Description | Average | StdDev
Landscape repeated; centred
Red | 33095.1494 | 31.784
Green | 40351.2767 | 21.1264
Blue | 25016.9998 | 43.0854
Landscape; centred
Red | 33158.2855 | 50.1356
Green | 40410.2298 | 57.1823
Blue | 25026.2829 | 56.7549
Landscape; centred; flipped
Red | 32348.7649 | 45.0438
Green | 39609.7264 | 31.053
Blue | 24410.6999 | 27.1845
Landscape; to left
Red | 32472.833 | 96.7878
Green | 40196.8478 | 75.3202 | `
Blue | 24645.7675 | 80.0833
Landscape; to right
Red | 33665.4801 | 7.7633
Green | 41321.8139 | 10.0794
Blue | 25786.1772 | 18.3795
Portrait; centred
Red | 30375.9717 | 75.7898
Green | 38285.6412 | 85.7415
Blue | 24236.5178 | 65.5581
Portrait; centred; flipped
Red | 29688.5443 | 49.4165
Green | 37474.9884 | 39.1699
Blue | 23446.7256 | 18.0469

## orientation and uniformity
Orientation | Channel | pixel value | Repeat number
Landscape; centred | Red | 33125 | 1 |  | 50.1356 | 33158.2855
Landscape; centred | Red | 33133.9082 | 2
Landscape; centred | Red | 33215.9482 | 3
Landscape; centred | Green | 40365 | 1
Landscape; centred | Green | 40391.1852 | 2
Landscape; centred | Green | 40474.5043 | 3
Landscape; centred | Blue | 24973 | 1
Landscape; centred | Blue | 25019.8815 | 2
Landscape; centred | Blue | 25085.9671 | 3
Landscape; to left | Red | 32361.3591 | 1
Landscape; to left | Red | 32521.6381 | 2
Landscape; to left | Red | 32535.5018 | 3
Landscape; to left | Green | 40110.1726 | 1
Landscape; to left | Green | 40233.9661 | 2
Landscape; to left | Green | 40246.4049 | 3
Landscape; to left | Blue | 24555.689 | 1
Landscape; to left | Blue | 24672.703 | 2
Landscape; to left | Blue | 24708.9104 | 3
Landscape; to right | Red | 33674.4007 | 1
Landscape; to right | Red | 33661.7854 | 2
Landscape; to right | Red | 33660.2544 | 3
Landscape; to right | Green | 41330.1835 | 1
Landscape; to right | Green | 41310.625 | 2
Landscape; to right | Green | 41324.6331 | 3
Landscape; to right | Blue | 25774.2282 | 1
Landscape; to right | Blue | 25776.9622 | 2
Landscape; to right | Blue | 25807.3413 | 3
Landscape repeated; centred | Red | 33116.4371 | 1 | 31.784
Landscape repeated; centred | Red | 33110.3966 | 2
Landscape repeated; centred | Red | 33058.6144 | 3
Landscape repeated; centred | Green | 40357.0248 | 1
Landscape repeated; centred | Green | 40368.9341 | 2
Landscape repeated; centred | Green | 40327.8711 | 3
Landscape repeated; centred | Blue | 24967.2986 | 1
Landscape repeated; centred | Blue | 25039.9278 | 2
Landscape repeated; centred | Blue | 25043.7729 | 3
Portrait; centred | Red | 30288.4591 | 1
Portrait; centred | Red | 30419.2159 | 2
Portrait; centred | Red | 30420.24 | 3
Portrait; centred | Green | 38186.6372 | 1
Portrait; centred | Green | 38335.6396 | 2
Portrait; centred | Green | 38334.6469 | 3
Portrait; centred | Blue | 24161.5055 | 1
Portrait; centred | Blue | 24265.2073 | 2
Portrait; centred | Blue | 24282.8406 | 3
Landscape; centred; flipped | Red | 32374.9474 | 1
Landscape; centred; flipped | Red | 32296.7532 | 2
Landscape; centred; flipped | Red | 32374.5941 | 3
Landscape; centred; flipped | Green | 39628.4528 | 1
Landscape; centred; flipped | Green | 39573.8815 | 2
Landscape; centred; flipped | Green | 39626.8449 | 3
Landscape; centred; flipped | Blue | 24396.0178 | 1
Landscape; centred; flipped | Blue | 24394.0134 | 2
Landscape; centred; flipped | Blue | 24442.0685 | 3
Portrait; centred; flipped | Red | 29717.5421 | 1
Portrait; centred; flipped | Red | 29631.4856 | 2
Portrait; centred; flipped | Red | 29716.6052 | 3
Portrait; centred; flipped | Green | 37498.3605 | 1
Portrait; centred; flipped | Green | 37429.7674 | 2
Portrait; centred; flipped | Green | 37496.8372 | 3
Portrait; centred; flipped | Blue | 23435.1251 | 1
Portrait; centred; flipped | Blue | 23437.5338 | 2
Portrait; centred; flipped | Blue | 23467.5179 | 3

## linearity Vidar
OD_FILENAME=N:\rtp\RIT\Vidar\RIT Calib. Files\ODcalibration_jan2014_stepwedge4118.cal
OD from stepwedge | SV | OD linearity of vidar | straight line
0.04 | 65535 | 0 | 0 | 0
0.16 | 51680 | 0.10315 | 3 | 3
0.27 | 40061 | 0.21375
0.38 | 30778 | 0.32823
0.49 | 24018 | 0.43594
0.61 | 18031 | 0.56045
0.72 | 13960 | 0.67159
0.87 | 9934 | 0.81935
0.97 | 7870 | 0.9205
1.08 | 5998 | 1.0385
1.21 | 4472 | 1.166
1.31 | 3597 | 1.2605
1.46 | 2515 | 1.4159
1.57 | 1936 | 1.5296
1.69 | 1449 | 1.6554
1.8 | 1114 | 1.7696
1.92 | 853 | 1.8855
2.03 | 650 | 2.0036
2.14 | 503 | 2.1149
2.27 | 382 | 2.2344
2.39 | 286 | 2.3601
2.5 | 222 | 2.4701
2.59 | 180 | 2.5612
2.71 | 137 | 2.6798
2.83 | 108 | 2.783
2.94 | 88 | 2.872

## linearity Scanner A
OD_FILENAME=N:\rtp\RIT\11000XL\calibration files\optical density\scanner A OD calibration\scanner A OD calibration 18 Feb 2014.cal
| OD from stepwedge | SV | OD linearity of Scanner A | straight line
| 0.04 | 61702 | 0.026174 | 0 | 0
| 0.16 | 50910 | 0.10967 | 3 | 3
| 0.27 | 42702 | 0.18603
| 0.38 | 35897 | 0.26142
| 0.49 | 30610 | 0.33061
| 0.61 | 25393 | 0.41176 |  |  | 0.81291
| 0.72 | 21514 | 0.48375
| 0.87 | 17344 | 0.57732
| 0.97 | 15011 | 0.64006
| 1.08 | 12696 | 0.71281
| 1.21 | 10610 | 0.79076
| 1.31 | 9257 | 0.85
| 1.46 | 7451 | 0.94426
| 1.57 | 6350 | 1.0137
| 1.69 | 5284 | 1.0935
| 1.8 | 4516 | 1.1617
| 1.92 | 3864 | 1.2294
| 2.03 | 3529 | 1.2688
| 2.14 | 3089 | 1.3267
| 2.27 | 2664 | 1.3909
| 2.39 | 2306 | 1.4536
| 2.5 | 2030 | 1.509
| 2.59 | 1826 | 1.555
| 2.71 | 1610 | 1.6096
| 2.83 | 1454 | 1.6539
| 2.94 | 1344 | 1.6881
| 3.05 | 1245 | 1.7213
| 3.18 | 1162 | 1.7513
| 3.28 | 1107 | 1.7723
| 3.36 | 1064 | 1.7895
| 3.48 | 1026 | 1.8053
| 3.7 | 1006 | 1.8139

## linearity Scanner B
OD_FILENAME=N:\rtp\RIT\11000XL\calibration files\optical density\scanner B OD calibration\scanner B OD calibration 18 Feb 2014.cal
| OD from stepwedge | SV | OD linearity of Scanner A | straight line
| 0.04 | 61910 | 0.024712 | 0 | 0
| 0.16 | 50961 | 0.10924 | 3 | 3
| 0.27 | 42723 | 0.18581
| 0.38 | 35866 | 0.26179
| 0.49 | 30542 | 0.33158
| 0.61 | 25345 | 0.41258
| 0.72 | 21489 | 0.48426
| 0.87 | 17283 | 0.57885
| 0.97 | 14929 | 0.64244
| 1.08 | 12621 | 0.71538
| 1.21 | 10532 | 0.79396
| 1.31 | 9170 | 0.8541
| 1.46 | 7383 | 0.94824
| 1.57 | 6306 | 1.0167
| 1.69 | 5244 | 1.0968
| 1.8 | 4485 | 1.1647
| 1.92 | 3843 | 1.2318
| 2.03 | 3528 | 1.2689
| 2.14 | 3078 | 1.3282
| 2.27 | 2655 | 1.3924
| 2.39 | 2313 | 1.4523
| 2.5 | 2024 | 1.5103
| 2.59 | 1840 | 1.5517
| 2.71 | 1637 | 1.6024
| 2.83 | 1488 | 1.6439
| 2.94 | 1383 | 1.6757
| 3.05 | 1287 | 1.7069
| 3.18 | 1219 | 1.7305
| 3.28 | 1164 | 1.7505
| 3.36 | 1133 | 1.7622
| 3.48 | 1074 | 1.7855

## Vidar dose curves
FILE_NAME=N:\rtp\RIT\Vidar\RIT Calib. Files\March2012NoSleeve178Res.CAL
dose cGy | scanner value | optical density | net optical density | log Dose
647.796 | 17162.46 | 0.58189 | 0.44997 | 2.8114
587.916 | 18233.6 | 0.5556 | 0.42368 | 2.7693
528.036 | 19314.94 | 0.53058 | 0.39865 | 2.7227
478.265 | 20262.58 | 0.50978 | 0.37785 | 2.6797
418.385 | 21848.48 | 0.47705 | 0.34513 | 2.6216
378.724 | 22983.02 | 0.45507 | 0.32314 | 2.5783
339.063 | 24215.45 | 0.43238 | 0.30046 | 2.5303 | 0.39243 | 26549 | 0.2605 | 0.0067592
308.734 | 25325.4 | 0.41292 | 0.28099 | 2.4896 | 0.39918 | 26139 | 0.26726
258.963 | 27140.38 | 0.38286 | 0.25093 | 2.4132
219.302 | 29208.83 | 0.35096 | 0.21903 | 2.341
179.641 | 31240.51 | 0.32176 | 0.18983 | 2.2544
149.312 | 33167.53 | 0.29576 | 0.16383 | 2.1741
119.761 | 35364.92 | 0.2679 | 0.13598 | 2.0783
69.99 | 39754.3 | 0.21709 | 0.085164 | 1.845
29.551 | 44288.96 | 0.17018 | 0.038252 | 1.4706
0 | 48366.86 | 0.13193 | 0 | #NUM!

## autosequence
Doses and Mus to be used in the gafchromic dose calibration: |  |  |  | Autosequence
dose Gy | Mus | rounded Mus | actual dose Gy | field | X1 (cm) | X2 | Y1 | Y2 | MU
0 | 0 | 0 | 0 | 1 | -8 | 0 | 5 | 10 | 23
0.2 | 22.7273 | 23 | 0.2024 | 2 | 0 | 8 | 5 | 10 | 136
0.7 | 79.5455 | 80 | 0.704 | 3 | -8 | 0 | 0 | 5 | 250
1.2 | 136.3636 | 136 | 1.1968 | 4 | 0 | 8 | 0 | 5 | 364
1.7 | 193.1818 | 193 | 1.6984 | 5 | -8 | 0 | -5 | 0 | 477
2.2 | 250 | 250 | 2.2 | 6 | 0 | 8 | -5 | 0 | 591
2.7 | 306.8182 | 307 | 2.7016 | 7 | -8 | 0 | -10 | -5 | 705
3.2 | 363.6364 | 364 | 3.2032 | 8 | 0 | 8 | -10 | -5 | 818
3.7 | 420.4545 | 420 | 3.696
4.2 | 477.2727 | 477 | 4.1976
4.7 | 534.0909 | 534 | 4.6992
5.2 | 590.9091 | 591 | 5.2008
5.7 | 647.7273 | 648 | 5.7024
6.2 | 704.5455 | 705 | 6.204
6.7 | 761.3636 | 761 | 6.6968
7.2 | 818.1818 | 818 | 7.1984
7.7 | 875 | 875 | 7.7

## scanner warm up
x3 | -2e-11
x2 | 3e-06
x | -0.113
c | 1823.5
positioned the film in the portrait orientation centrally to the scan array
repeat measurements taken
scanner turned on at
epson software used 72dpi; 48bit tiff format
preview scan taken; however, when in clinical use there will be no need for a preview scan if the template is in place; however, the preview scans could be used as a warm up perhaps; this could be before the film is placed on the plate.
or perhaps if there is a warm up effect then we should be doing preview scans before placeing it on the scanner. Is the preview scans the same as the real scan? Can 3 previews stop any warm up effect. Or should the real scan be an average of 3 scans? To remove the repeatability uncertainty?
I shall test this :
1 preview scan is required to get the positioning tof the scan area.
then reapeat scans can be carried out ( the previews can't be used for that because there is no image formed form the scan)
then if a warm up effect is observed then the scanner can be switched off and then 1 preview scan given - then a real scan - see if the real scan is the same as a fully settled scanner
then if 1 preview is not enough - the nallow to cool then repeat by doing 2 previews, and so on until the required amounts of warm up previews are done.
if repeatability is unsatisfactory then repeat scans can be done and the image averaged.
location of scan already determined before switched on (10 minutes ago)
time turned on 12.40
time first scan
time second scan
25th March scanner turned on at
11:28
set 1
scan name | time of scan | time since scanner was turned on | seconds since scanner turned on | minutes since turned on | imagej read out
[5 empty] | region 1 | region 2
scan1009 | 11:30 | 0:02 | 97 | 1.6167 | 26549.7076 | 44016.9394
scan1010 | 11:31 | 0:03 | 151 | 2.5167 | 26542.8156 | 44034.315
scan1011 | 11:32 | 0:04 | 212 | 3.5333 | 26403.9188 | 43853.7719
scan1012 | 11:33 | 0:05 | 274 | 4.5667 | 26442.9156 | 43955.6231
scan1013 | 11:34 | 0:06 | 332 | 5.5333 | 26430.1628 | 43923.1956
scan1014 | 11:35 | 0:07 | 392 | 6.5333 | 26430.7984 | 43920.1489
scan1015 | 11:36 | 0:08 | 452 | 7.5333 | 26437.6652 | 43922.2875
scan1016 | 11:37 | 0:09 | 512 | 8.5333 | 26431.9068 | 43917.6383
scan1017 | 11:38 | 0:10 | 574 | 9.5667 | 26430.9268 | 43927.7806
scan1018 | 11:39 | 0:11 | 631 | 10.5167 | 26429.944 | 43923.8364
set 2
One scan taken directly after scanner switched on
scan1008 | 26557.4372
set 3
scanner was left for an hour then scanning resumed. This is after set 1
however, I turned the scanner off then immediately on before this set - perhaps this makes a difference
scan name | time of scan | time since scanner was turned on | seconds since scanner turned on | minutes since turned on | imagej read out
[5 empty] | region 1 | region 2
scan1019 | 12:50 | 1:22 | 4937 | 82.2833 | 26318.3308 | 43726.2978
scan1020 | 12:51 | 1:23 | 4959 | 82.65 | 26352.192 | 43792.2456
scan1021 | 12:51 | 1:23 | 4980 | 83 | 26355.2984 | 43813.2322
scan1022 | 12:51 | 1:23 | 5007 | 83.45 | 26324.4524 | 43795.5989
scan1023 | 12:52 | 1:24 | 5029 | 83.8167 | 26361.5988 | 43823.0917
scan1024 | 12:52 | 1:24 | 5052 | 84.2 | 26327.7432 | 43792.9833
scan1025 | 12:53 | 1:25 | 5076 | 84.6 | 26324.9236 | 43787.8083
set 4
then went for lunch and then did a set of scanning without turning scanner off
[5 empty] | region 1 | region 2
scan1026 | 13:47 | 2:19 | 8369 | 139.4833 | 26343.7548 | 43783.7139
scan1027 | 13:48 | 2:20 | 8390 | 139.8333 | 26330.0492 | 43768.9119
scan1028 | 13:48 | 2:20 | 8412 | 140.2 | 26325.0412 | 43769.5253
scan1029 | 13:49 | 2:21 | 8435 | 140.5833 | 26300.588 | 43721.4097
scan1030 | 13:49 | 2:21 | 8457 | 140.95 | 26313.3156 | 43749.5819
scan1031 | 13:49 | 2:21 | 8475 | 141.25 | 26318.6504 | 43748.5961
scan1032 | 13:50 | 2:22 | 8502 | 141.7 | 26250.9672 | 43653.1731
scan1033 | 13:50 | 2:22 | 8523 | 142.05 | 26306.1532 | 43714.94
set 5
then went to the group meeting without turning the scanner off:
[5 empty] | region 1 | region 2
scan1034 | 15:01 | 3:33 | 12754 | 212.5667 | 26254.2072 | 43662.4114
scan1035 | 15:01 | 3:33 | 12775 | 212.9167 | 26185.036 | 43581.7264
scan1036 | 15:01 | 3:33 | 12796 | 213.2667 | 26184.8224 | 43580.9506
scan1037 | 15:02 | 3:34 | 12817 | 213.6167 | 26245.2572 | 43670.0381
scan1038 | 15:02 | 3:34 | 12838 | 213.9667 | 26251.8804 | 43681.5411
scan1039 | 15:02 | 3:34 | 12861 | 214.35 | 26294.1868 | 43771.3803
then the scanner was turned off then back on again immediately and scanning resumed
scan1040 | 15:04 | 3:36 | 12938 | 215.6333 | 26165.1944 | 43532.8014
scan1041 | 15:04 | 3:36 | 12959 | 215.9833 | 26240.0012 | 43656.1833
scan1042 | 15:04 | 3:36 | 12980 | 216.3333 | 26195.4492 | 43609.7667
scan1043 | 15:05 | 3:37 | 13001 | 216.6833 | 26216.3248 | 43637.8997
scan1044 | 15:05 | 3:37 | 13022 | 217.0333 | 26210.8464 | 43628.2767
scan1045 | 15:05 | 3:37 | 13043 | 217.3833 | 26204.0624 | 43623.8753

## scanner warm up day 1 0Gy
a | -2e-11
b | 3e-06
c | -0.113
d | 1823.5
positioned the film in the portrait orientation centrally to the scan array
repeat measurements taken
scanner turned on at
epson software used 72dpi; 48bit tiff format
preview scan taken; however, when in clinical use there will be no need for a preview scan if the template is in place; however, the preview scans could be used as a warm up perhaps; this could be before the film is placed on the plate.
or perhaps if there is a warm up effect then we should be doing preview scans before placeing it on the scanner. Is the preview scans the same as the real scan? Can 3 previews stop any warm up effect. Or should the real scan be an average of 3 scans? To remove the repeatability uncertainty?
I shall test this :
1 preview scan is required to get the positioning tof the scan area.
then reapeat scans can be carried out ( the previews can't be used for that because there is no image formed form the scan)
then if a warm up effect is observed then the scanner can be switched off and then 1 preview scan given - then a real scan - see if the real scan is the same as a fully settled scanner
then if 1 preview is not enough - the nallow to cool then repeat by doing 2 previews, and so on until the required amounts of warm up previews are done.
if repeatability is unsatisfactory then repeat scans can be done and the image averaged.
location of scan already determined before switched on (10 minutes ago)
time turned on 12.40
time first scan
time second scan
25th March scanner turned on at
11:28
set 1
scan name | time of scan | time since scanner was turned on | seconds since scanner turned on | minutes since turned on | imagej read out
[5 empty] | region 1 | region 2 | cGy
scan1009 | 11:30 | 0:02 | 97 | 1.6167 | 26549.7076 | 44016.9394 | 956.4103
scan1010 | 11:31 | 0:03 | 151 | 2.5167 | 26542.8156 | 44034.315 | 957.016
scan1011 | 11:32 | 0:04 | 212 | 3.5333 | 26403.9188 | 43853.7719 | 950.7332
scan1012 | 11:33 | 0:05 | 274 | 4.5667 | 26442.9156 | 43955.6231 | 954.2746
scan1013 | 11:34 | 0:06 | 332 | 5.5333 | 26430.1628 | 43923.1956 | 953.1463
scan1014 | 11:35 | 0:07 | 392 | 6.5333 | 26430.7984 | 43920.1489 | 953.0403
scan1015 | 11:36 | 0:08 | 452 | 7.5333 | 26437.6652 | 43922.2875 | 953.1147
scan1016 | 11:37 | 0:09 | 512 | 8.5333 | 26431.9068 | 43917.6383 | 952.953
scan1017 | 11:38 | 0:10 | 574 | 9.5667 | 26430.9268 | 43927.7806 | 953.3058
scan1018 | 11:39 | 0:11 | 631 | 10.5167 | 26429.944 | 43923.8364 | 953.1686
set 2
One scan taken directly after scanner switched on
scan1008 | 26557.4372
set 3
scanner was left for an hour then scanning resumed. This is after set 1
however, I turned the scanner off then immediately on before this set - perhaps this makes a difference
scan name | time of scan | time since scanner was turned on | seconds since scanner turned on | minutes since turned on | imagej read out
[5 empty] | region 1 | region 2
scan1019 | 12:50 | 1:22 | 4937 | 82.2833 | 26318.3308 | 43726.2978
scan1020 | 12:51 | 1:23 | 4959 | 82.65 | 26352.192 | 43792.2456
scan1021 | 12:51 | 1:23 | 4980 | 83 | 26355.2984 | 43813.2322
scan1022 | 12:51 | 1:23 | 5007 | 83.45 | 26324.4524 | 43795.5989
scan1023 | 12:52 | 1:24 | 5029 | 83.8167 | 26361.5988 | 43823.0917
scan1024 | 12:52 | 1:24 | 5052 | 84.2 | 26327.7432 | 43792.9833
scan1025 | 12:53 | 1:25 | 5076 | 84.6 | 26324.9236 | 43787.8083
set 4
then went for lunch and then did a set of scanning without turning scanner off
[5 empty] | region 1 | region 2
scan1026 | 13:47 | 2:19 | 8369 | 139.4833 | 26343.7548 | 43783.7139
scan1027 | 13:48 | 2:20 | 8390 | 139.8333 | 26330.0492 | 43768.9119
scan1028 | 13:48 | 2:20 | 8412 | 140.2 | 26325.0412 | 43769.5253
scan1029 | 13:49 | 2:21 | 8435 | 140.5833 | 26300.588 | 43721.4097
scan1030 | 13:49 | 2:21 | 8457 | 140.95 | 26313.3156 | 43749.5819
scan1031 | 13:49 | 2:21 | 8475 | 141.25 | 26318.6504 | 43748.5961
scan1032 | 13:50 | 2:22 | 8502 | 141.7 | 26250.9672 | 43653.1731
scan1033 | 13:50 | 2:22 | 8523 | 142.05 | 26306.1532 | 43714.94
set 5
then went to the group meeting without turning the scanner off:
[5 empty] | region 1 | region 2
scan1034 | 15:01 | 3:33 | 12754 | 212.5667 | 26254.2072 | 43662.4114
scan1035 | 15:01 | 3:33 | 12775 | 212.9167 | 26185.036 | 43581.7264
scan1036 | 15:01 | 3:33 | 12796 | 213.2667 | 26184.8224 | 43580.9506
scan1037 | 15:02 | 3:34 | 12817 | 213.6167 | 26245.2572 | 43670.0381
scan1038 | 15:02 | 3:34 | 12838 | 213.9667 | 26251.8804 | 43681.5411
scan1039 | 15:02 | 3:34 | 12861 | 214.35 | 26294.1868 | 43771.3803
then the scanner was turned off then back on again immediately and scanning resumed
scan1040 | 15:04 | 3:36 | 12938 | 215.6333 | 26165.1944 | 43532.8014
scan1041 | 15:04 | 3:36 | 12959 | 215.9833 | 26240.0012 | 43656.1833
scan1042 | 15:04 | 3:36 | 12980 | 216.3333 | 26195.4492 | 43609.7667
scan1043 | 15:05 | 3:37 | 13001 | 216.6833 | 26216.3248 | 43637.8997
scan1044 | 15:05 | 3:37 | 13022 | 217.0333 | 26210.8464 | 43628.2767
scan1045 | 15:05 | 3:37 | 13043 | 217.3833 | 26204.0624 | 43623.8753

## scanner warm up day 1 0Gy zoom
positioned the film in the portrait orientation centrally to the scan array
repeat measurements taken
scanner turned on at
epson software used 72dpi; 48bit tiff format
preview scan taken; however, when in clinical use there will be no need for a preview scan if the template is in place; however, the preview scans could be used as a warm up perhaps; this could be before the film is placed on the plate.
or perhaps if there is a warm up effect then we should be doing preview scans before placeing it on the scanner. Is the preview scans the same as the real scan? Can 3 previews stop any warm up effect. Or should the real scan be an average of 3 scans? To remove the repeatability uncertainty?
I shall test this :
1 preview scan is required to get the positioning tof the scan area.
then reapeat scans can be carried out ( the previews can't be used for that because there is no image formed form the scan)
then if a warm up effect is observed then the scanner can be switched off and then 1 preview scan given - then a real scan - see if the real scan is the same as a fully settled scanner
then if 1 preview is not enough - the nallow to cool then repeat by doing 2 previews, and so on until the required amounts of warm up previews are done.
if repeatability is unsatisfactory then repeat scans can be done and the image averaged.
location of scan already determined before switched on (10 minutes ago)
time turned on 12.40
time first scan
time second scan
set 1
scan name | time of scan | time since scanner was turned on | seconds since scanner turned on | minutes since turned on | imagej read out
[5 empty] | region 1 | region 2
scan1009 | 11:30 | 0:02 | 97 | 1.6167 | 26549.7076 | 44016.9394 | 25th March scanner turned on at
scan1010 | 11:31 | 0:03 | 151 | 2.5167 | 26542.8156 | 44034.315
scan1011 | 11:32 | 0:04 | 212 | 3.5333 | 26403.9188 | 43853.7719
scan1012 | 11:33 | 0:05 | 274 | 4.5667 | 26442.9156 | 43955.6231
scan1013 | 11:34 | 0:06 | 332 | 5.5333 | 26430.1628 | 43923.1956
scan1014 | 11:35 | 0:07 | 392 | 6.5333 | 26430.7984 | 43920.1489
scan1015 | 11:36 | 0:08 | 452 | 7.5333 | 26437.6652 | 43922.2875
scan1016 | 11:37 | 0:09 | 512 | 8.5333 | 26431.9068 | 43917.6383
scan1017 | 11:38 | 0:10 | 574 | 9.5667 | 26430.9268 | 43927.7806
scan1018 | 11:39 | 0:11 | 631 | 10.5167 | 26429.944 | 43923.8364
scan1019 | 12:50 | 1:22 | 4937 | 82.2833 | 26318.3308 | 43726.2978 | scanner was left for an hour then scanning resumed. This is after set 1
scan1020 | 12:51 | 1:23 | 4959 | 82.65 | 26352.192 | 43792.2456 | however, I turned the scanner off then immediately on before this set - perhaps this makes a difference |  |  | 11:28
scan1021 | 12:51 | 1:23 | 4980 | 83 | 26355.2984 | 43813.2322 | scan name
scan1022 | 12:51 | 1:23 | 5007 | 83.45 | 26324.4524 | 43795.5989
scan1023 | 12:52 | 1:24 | 5029 | 83.8167 | 26361.5988 | 43823.0917
scan1024 | 12:52 | 1:24 | 5052 | 84.2 | 26327.7432 | 43792.9833
scan1025 | 12:53 | 1:25 | 5076 | 84.6 | 26324.9236 | 43787.8083
scan1026 | 13:47 | 2:19 | 8369 | 139.4833 | 26343.7548 | 43783.7139 | then went for lunch and then did a set of scanning without turning scanner off
scan1027 | 13:48 | 2:20 | 8390 | 139.8333 | 26330.0492 | 43768.9119
scan1028 | 13:48 | 2:20 | 8412 | 140.2 | 26325.0412 | 43769.5253
scan1029 | 13:49 | 2:21 | 8435 | 140.5833 | 26300.588 | 43721.4097
scan1030 | 13:49 | 2:21 | 8457 | 140.95 | 26313.3156 | 43749.5819
scan1031 | 13:49 | 2:21 | 8475 | 141.25 | 26318.6504 | 43748.5961
scan1032 | 13:50 | 2:22 | 8502 | 141.7 | 26250.9672 | 43653.1731 |  | time of scan | time since scanner was turned on | seconds since scanner turned on | minutes since turned on | imagej read out
scan1033 | 13:50 | 2:22 | 8523 | 142.05 | 26306.1532 | 43714.94
scan1034 | 15:01 | 3:33 | 12754 | 212.5667 | 26254.2072 | 43662.4114 | then went to the group meeting without turning the scanner off:
scan1035 | 15:01 | 3:33 | 12775 | 212.9167 | 26185.036 | 43581.7264
scan1036 | 15:01 | 3:33 | 12796 | 213.2667 | 26184.8224 | 43580.9506
scan1037 | 15:02 | 3:34 | 12817 | 213.6167 | 26245.2572 | 43670.0381
scan1038 | 15:02 | 3:34 | 12838 | 213.9667 | 26251.8804 | 43681.5411
scan1039 | 15:02 | 3:34 | 12861 | 214.35 | 26294.1868 | 43771.3803
scan1040 | 15:04 | 3:36 | 12938 | 215.6333 | 26165.1944 | 43532.8014 | then the scanner was turned off then back on again immediately and scanning resumed
scan1041 | 15:04 | 3:36 | 12959 | 215.9833 | 26240.0012 | 43656.1833
scan1042 | 15:04 | 3:36 | 12980 | 216.3333 | 26195.4492 | 43609.7667
scan1043 | 15:05 | 3:37 | 13001 | 216.6833 | 26216.3248 | 43637.8997
scan1044 | 15:05 | 3:37 | 13022 | 217.0333 | 26210.8464 | 43628.2767
scan1045 | 15:05 | 3:37 | 13043 | 217.3833 | 26204.0624 | 43623.8753

## scanner warmup day1 ~3Gy
positioned the film in the portrait orientation centrally to the scan array
repeat measurements taken
scanner turned on at
epson software used 72dpi; 48bit tiff format
preview scan taken; however, when in clinical use there will be no need for a preview scan if the template is in place; however, the preview scans could be used as a warm up perhaps; this could be before the film is placed on the plate.
or perhaps if there is a warm up effect then we should be doing preview scans before placeing it on the scanner. Is the preview scans the same as the real scan? Can 3 previews stop any warm up effect. Or should the real scan be an average of 3 scans? To remove the repeatability uncertainty?
I shall test this :
1 preview scan is required to get the positioning tof the scan area.
then reapeat scans can be carried out ( the previews can't be used for that because there is no image formed form the scan)
then if a warm up effect is observed then the scanner can be switched off and then 1 preview scan given - then a real scan - see if the real scan is the same as a fully settled scanner
then if 1 preview is not enough - the nallow to cool then repeat by doing 2 previews, and so on until the required amounts of warm up previews are done.
if repeatability is unsatisfactory then repeat scans can be done and the image averaged.
location of scan already determined before switched on (10 minutes ago)
time turned on 12.40
time first scan
time second scan
set 1
scan name | time of scan | time since scanner was turned on | seconds since scanner turned on | minutes since turned on | imagej read out
[5 empty] | region 1 | region 2
scan1009 | 11:30 | 0:02 | 97 | 1.6167 | 26549.7076 | 44016.9394 | 25th March scanner turned on at
scan1010 | 11:31 | 0:03 | 151 | 2.5167 | 26542.8156 | 44034.315
scan1011 | 11:32 | 0:04 | 212 | 3.5333 | 26403.9188 | 43853.7719
scan1012 | 11:33 | 0:05 | 274 | 4.5667 | 26442.9156 | 43955.6231
scan1013 | 11:34 | 0:06 | 332 | 5.5333 | 26430.1628 | 43923.1956
scan1014 | 11:35 | 0:07 | 392 | 6.5333 | 26430.7984 | 43920.1489
scan1015 | 11:36 | 0:08 | 452 | 7.5333 | 26437.6652 | 43922.2875
scan1016 | 11:37 | 0:09 | 512 | 8.5333 | 26431.9068 | 43917.6383
scan1017 | 11:38 | 0:10 | 574 | 9.5667 | 26430.9268 | 43927.7806
scan1018 | 11:39 | 0:11 | 631 | 10.5167 | 26429.944 | 43923.8364
scan1019 | 12:50 | 1:22 | 4937 | 82.2833 | 26318.3308 | 43726.2978 | scanner was left for an hour then scanning resumed. This is after set 1
scan1020 | 12:51 | 1:23 | 4959 | 82.65 | 26352.192 | 43792.2456 | however, I turned the scanner off then immediately on before this set - perhaps this makes a difference |  |  | 11:28
scan1021 | 12:51 | 1:23 | 4980 | 83 | 26355.2984 | 43813.2322 | scan name
scan1022 | 12:51 | 1:23 | 5007 | 83.45 | 26324.4524 | 43795.5989
scan1023 | 12:52 | 1:24 | 5029 | 83.8167 | 26361.5988 | 43823.0917
scan1024 | 12:52 | 1:24 | 5052 | 84.2 | 26327.7432 | 43792.9833
scan1025 | 12:53 | 1:25 | 5076 | 84.6 | 26324.9236 | 43787.8083
scan1026 | 13:47 | 2:19 | 8369 | 139.4833 | 26343.7548 | 43783.7139 | then went for lunch and then did a set of scanning without turning scanner off
scan1027 | 13:48 | 2:20 | 8390 | 139.8333 | 26330.0492 | 43768.9119
scan1028 | 13:48 | 2:20 | 8412 | 140.2 | 26325.0412 | 43769.5253
scan1029 | 13:49 | 2:21 | 8435 | 140.5833 | 26300.588 | 43721.4097
scan1030 | 13:49 | 2:21 | 8457 | 140.95 | 26313.3156 | 43749.5819
scan1031 | 13:49 | 2:21 | 8475 | 141.25 | 26318.6504 | 43748.5961
scan1032 | 13:50 | 2:22 | 8502 | 141.7 | 26250.9672 | 43653.1731 |  | time of scan | time since scanner was turned on | seconds since scanner turned on | minutes since turned on | imagej read out
scan1033 | 13:50 | 2:22 | 8523 | 142.05 | 26306.1532 | 43714.94
scan1034 | 15:01 | 3:33 | 12754 | 212.5667 | 26254.2072 | 43662.4114 | then went to the group meeting without turning the scanner off:
scan1035 | 15:01 | 3:33 | 12775 | 212.9167 | 26185.036 | 43581.7264
scan1036 | 15:01 | 3:33 | 12796 | 213.2667 | 26184.8224 | 43580.9506
scan1037 | 15:02 | 3:34 | 12817 | 213.6167 | 26245.2572 | 43670.0381
scan1038 | 15:02 | 3:34 | 12838 | 213.9667 | 26251.8804 | 43681.5411
scan1039 | 15:02 | 3:34 | 12861 | 214.35 | 26294.1868 | 43771.3803
scan1040 | 15:04 | 3:36 | 12938 | 215.6333 | 26165.1944 | 43532.8014 | then the scanner was turned off then back on again immediately and scanning resumed
scan1041 | 15:04 | 3:36 | 12959 | 215.9833 | 26240.0012 | 43656.1833
scan1042 | 15:04 | 3:36 | 12980 | 216.3333 | 26195.4492 | 43609.7667
scan1043 | 15:05 | 3:37 | 13001 | 216.6833 | 26216.3248 | 43637.8997
scan1044 | 15:05 | 3:37 | 13022 | 217.0333 | 26210.8464 | 43628.2767
scan1045 | 15:05 | 3:37 | 13043 | 217.3833 | 26204.0624 | 43623.8753
explanations for this reduction in pixel value with time during the day
temperature
the scanner light is causing some film darkening,

## scanner warmup d1 both film
positioned the film in the portrait orientation centrally to the scan array
repeat measurements taken
scanner turned on at
epson software used 72dpi; 48bit tiff format
preview scan taken; however, when in clinical use there will be no need for a preview scan if the template is in place; however, the preview scans could be used as a warm up perhaps; this could be before the film is placed on the plate.
or perhaps if there is a warm up effect then we should be doing preview scans before placeing it on the scanner. Is the preview scans the same as the real scan? Can 3 previews stop any warm up effect. Or should the real scan be an average of 3 scans? To remove the repeatability uncertainty?
I shall test this :
1 preview scan is required to get the positioning tof the scan area.
then reapeat scans can be carried out ( the previews can't be used for that because there is no image formed form the scan)
then if a warm up effect is observed then the scanner can be switched off and then 1 preview scan given - then a real scan - see if the real scan is the same as a fully settled scanner
then if 1 preview is not enough - the nallow to cool then repeat by doing 2 previews, and so on until the required amounts of warm up previews are done.
if repeatability is unsatisfactory then repeat scans can be done and the image averaged.
location of scan already determined before switched on (10 minutes ago)
time turned on 12.40
time first scan
time second scan
set 1
scan name | time of scan | time since scanner was turned on | seconds since scanner turned on | minutes since turned on | imagej read out
[5 empty] | region 1 | region 2
scan1009 | 11:30 | 0:02 | 97 | 1.6167 | 26549.7076 | 44016.9394 | 25th March scanner turned on at
scan1010 | 11:31 | 0:03 | 151 | 2.5167 | 26542.8156 | 44034.315
scan1011 | 11:32 | 0:04 | 212 | 3.5333 | 26403.9188 | 43853.7719
scan1012 | 11:33 | 0:05 | 274 | 4.5667 | 26442.9156 | 43955.6231
scan1013 | 11:34 | 0:06 | 332 | 5.5333 | 26430.1628 | 43923.1956
scan1014 | 11:35 | 0:07 | 392 | 6.5333 | 26430.7984 | 43920.1489
scan1015 | 11:36 | 0:08 | 452 | 7.5333 | 26437.6652 | 43922.2875
scan1016 | 11:37 | 0:09 | 512 | 8.5333 | 26431.9068 | 43917.6383
scan1017 | 11:38 | 0:10 | 574 | 9.5667 | 26430.9268 | 43927.7806
scan1018 | 11:39 | 0:11 | 631 | 10.5167 | 26429.944 | 43923.8364
scan1019 | 12:50 | 1:22 | 4937 | 82.2833 | 26318.3308 | 43726.2978 | scanner was left for an hour then scanning resumed. This is after set 1
scan1020 | 12:51 | 1:23 | 4959 | 82.65 | 26352.192 | 43792.2456 | however, I turned the scanner off then immediately on before this set - perhaps this makes a difference |  |  | 11:28
scan1021 | 12:51 | 1:23 | 4980 | 83 | 26355.2984 | 43813.2322 | scan name
scan1022 | 12:51 | 1:23 | 5007 | 83.45 | 26324.4524 | 43795.5989
scan1023 | 12:52 | 1:24 | 5029 | 83.8167 | 26361.5988 | 43823.0917
scan1024 | 12:52 | 1:24 | 5052 | 84.2 | 26327.7432 | 43792.9833
scan1025 | 12:53 | 1:25 | 5076 | 84.6 | 26324.9236 | 43787.8083
scan1026 | 13:47 | 2:19 | 8369 | 139.4833 | 26343.7548 | 43783.7139 | then went for lunch and then did a set of scanning without turning scanner off
scan1027 | 13:48 | 2:20 | 8390 | 139.8333 | 26330.0492 | 43768.9119
scan1028 | 13:48 | 2:20 | 8412 | 140.2 | 26325.0412 | 43769.5253
scan1029 | 13:49 | 2:21 | 8435 | 140.5833 | 26300.588 | 43721.4097
scan1030 | 13:49 | 2:21 | 8457 | 140.95 | 26313.3156 | 43749.5819
scan1031 | 13:49 | 2:21 | 8475 | 141.25 | 26318.6504 | 43748.5961
scan1032 | 13:50 | 2:22 | 8502 | 141.7 | 26250.9672 | 43653.1731 |  | time of scan | time since scanner was turned on | seconds since scanner turned on | minutes since turned on | imagej read out
scan1033 | 13:50 | 2:22 | 8523 | 142.05 | 26306.1532 | 43714.94
scan1034 | 15:01 | 3:33 | 12754 | 212.5667 | 26254.2072 | 43662.4114 | then went to the group meeting without turning the scanner off:
scan1035 | 15:01 | 3:33 | 12775 | 212.9167 | 26185.036 | 43581.7264
scan1036 | 15:01 | 3:33 | 12796 | 213.2667 | 26184.8224 | 43580.9506
scan1037 | 15:02 | 3:34 | 12817 | 213.6167 | 26245.2572 | 43670.0381
scan1038 | 15:02 | 3:34 | 12838 | 213.9667 | 26251.8804 | 43681.5411
scan1039 | 15:02 | 3:34 | 12861 | 214.35 | 26294.1868 | 43771.3803
scan1040 | 15:04 | 3:36 | 12938 | 215.6333 | 26165.1944 | 43532.8014 | then the scanner was turned off then back on again immediately and scanning resumed
scan1041 | 15:04 | 3:36 | 12959 | 215.9833 | 26240.0012 | 43656.1833
scan1042 | 15:04 | 3:36 | 12980 | 216.3333 | 26195.4492 | 43609.7667
scan1043 | 15:05 | 3:37 | 13001 | 216.6833 | 26216.3248 | 43637.8997
scan1044 | 15:05 | 3:37 | 13022 | 217.0333 | 26210.8464 | 43628.2767
scan1045 | 15:05 | 3:37 | 13043 | 217.3833 | 26204.0624 | 43623.8753
scan1046 | 16:07 | 4:39 | 16729 | 278.8167 | 26161.2924 | 43537.7833 | waited an hour with scanner off, scanner turned on and scanning commenced immediately
scan1047 | 16:07 | 4:39 | 16759 | 279.3167 | 26174.988 | 43572.17
scan1048 | 16:08 | 4:40 | 16790 | 279.8333 | 26173.3016 | 43574.3936
scan1049 | 16:08 | 4:40 | 16828 | 280.4667 | 26178.166 | 43567.5056
scan1050 | 16:09 | 4:41 | 16853 | 280.8833 | 26214.2188 | 43674.03
scan1051 | 16:09 | 4:41 | 16878 | 281.3 | 26219.288 | 43676.8286
scan1052 | 16:10 | 4:42 | 16902 | 281.7 | 26220.5576 | 43689.4356
scan1053 | 16:11 | 4:43 | 16952 | 282.5333 | 26202.2632 | 43647.4594
scan1054 | 16:11 | 4:43 | 16989 | 283.15 | 26216.8996 | 43662.1475
scan1055 | 16:12 | 4:44 | 17011 | 283.5167 | 26224.4272 | 43686.6844
scan1056 | 16:12 | 4:44 | 17033 | 283.8833 | 26228.3396 | 43697.1089
scan1057 | 16:24 | 4:56 | 17767 | 296.1167 | 26132.2256 | 43497.6036 | did some analysis, leaving scanner on then did some more scans
scan1058 | 16:25 | 4:57 | 17790 | 296.5 | 26164.0728 | 43563.8133
scan1059 | 16:25 | 4:57 | 17813 | 296.8833 | 26155.8496 | 43561.0397
scan1060 | 16:25 | 4:57 | 17835 | 297.25 | 26180.858 | 43590.3328
scan1061 | 16:26 | 4:58 | 17856 | 297.6 | 26174.1224 | 43589.3236
scan1062 | 16:26 | 4:58 | 17883 | 298.05 | 26151.198 | 43553.6975
scan1063 | 16:26 | 4:58 | 17905 | 298.4167 | 26146.8524 | 43550.8644
scan1064 | 16:27 | 4:59 | 17930 | 298.8333 | 26163.464 | 43571.0217
scan1065 | 16:27 | 4:59 | 17954 | 299.2333 | 26131.972 | 43524.9175
scan1066 | 16:28 | 5:00 | 17976 | 299.6 | 26139.4448 | 43536.84
scan1067 | 16:36 | 5:08 | 18495 | 308.25 | 26107.9404 | 43474.8567
scan1068 | 16:37 | 5:09 | 18516 | 308.6 | 26115.1596 | 43500.0822
scan1069 | 16:37 | 5:09 | 18545 | 309.0833 | 26135.654 | 43548.35
scan1070 | 16:38 | 5:10 | 18571 | 309.5167 | 26151.9988 | 43572.9044
scan1071 | 16:38 | 5:10 | 18593 | 309.8833 | 26129.3544 | 43547.8369
scan1072 | 16:38 | 5:10 | 18614 | 310.2333 | 26123.5556 | 43532.7878
scan1073 | 16:45 | 5:17 | 18994 | 316.5667 | 26052.6328 | 43413.5486
scan1074 | 16:45 | 5:17 | 19018 | 316.9667 | 26130.718 | 43551.58
scan1075 | 16:45 | 5:17 | 19039 | 317.3167 | 26117.3864 | 43537.1581
scan1076 | 16:46 | 5:18 | 19061 | 317.6833 | 26119.8804 | 43549.4567
scan1077 | 16:46 | 5:18 | 19087 | 318.1167 | 26120.2404 | 43542.5081
scan1078 | 16:46 | 5:18 | 19109 | 318.4833 | 26122.8924 | 43560.2892
explanations for this reduction in pixel value with time during the day
temperature
the scanner light is causing some film darkening,
scans taken every 1 minute; subsequently taken every 30seconds
scanner left for an hour switched on. turned it off then immediately back on prior to scanning the 2nd set
went for lunch; scanner left on
then went to the group meeting without turning the scanner off:
half way through this set, turned scanner off then immediately on again (see zoomed in image
waited an hour with scanner off, scanner turned on and scanning commenced immediately
did some analysis, leaving scanner on then did some more scans

## scanner warmup day 2
scanner turned on on the 26th at 9.06 | 9:06
scan1079 | 9:07 | 0:01 | 80 | 1.3333 | 26160.4788 | 43455.92
scan1080 | 9:08 | 0:02 | 101 | 1.6833 | 26232.1152 | 43575.6625
scan1081 | 9:08 | 0:02 | 125 | 2.0833 | 26217.4332 | 43553.2153
scan1082 | 9:08 | 0:02 | 146 | 2.4333 | 26215.1268 | 43553.6881
scan1083 | 9:09 | 0:03 | 169 | 2.8167 | 26229.2128 | 43576.5092
scan1084 | 9:09 | 0:03 | 193 | 3.2167 | 26108.9 | 43394.0336
scan1085 | 9:10 | 0:04 | 220 | 3.6667 | 26101.1028 | 43380.0444
scan1086 | 9:10 | 0:04 | 241 | 4.0167 | 26099.9284 | 43388.2608
scan1087 | 9:10 | 0:04 | 265 | 4.4167 | 26114.54 | 43413.4761
scan1088 | 9:11 | 0:05 | 287 | 4.7833 | 26112.9456 | 43414.0839
scan1089 | 9:11 | 0:05 | 308 | 5.1333 | 26126.5196 | 43436.5533
scan1090 | 9:12 | 0:06 | 339 | 5.65 | 26101.99 | 43385.1011
scan1091 | 9:35 | 0:29 | 1761 | 29.35 | 26034.0016 | 43298.155
scan1092 | 9:36 | 0:30 | 1782 | 29.7 | 26161.544 | 43496.3675
scan1093 | 9:36 | 0:30 | 1806 | 30.1 | 26166.7132 | 43528.8864
scan1094 | 9:36 | 0:30 | 1828 | 30.4667 | 26058.2552 | 43370.015
scan1095 | 9:37 | 0:31 | 1849 | 30.8167 | 26063.4156 | 43381.5661
scan1096 | 9:37 | 0:31 | 1871 | 31.1833 | 26160.2688 | 43509.5278
scan1097 | 10:57 | 1:51 | 6681 | 111.35 | 26086.626 | 43348.1856 | turned off at 9.41; the idea is to test to see if the readings return
scan1098 | 10:58 | 1:52 | 6702 | 111.7 | 26123.9804 | 43424.8394 | to their original start of the day values or whether this
scan1099 | 10:58 | 1:52 | 6723 | 112.05 | 26131.5724 | 43434.0569 | strange decreasing in signal over the day shows itself again
scan1100 | 10:58 | 1:52 | 6747 | 112.45 | 26145.0888 | 43472.8247
scan1101 | 10:59 | 1:53 | 6773 | 112.8833 | 26097.0436 | 43399.8508
scan1102 | 10:59 | 1:53 | 6797 | 113.2833 | 26127.2756 | 43458.3069
scan1103 | 11:00 | 1:54 | 6818 | 113.6333 | 26159.0016 | 43496.1131
scan1104 | 11:00 | 1:54 | 6838 | 113.9667 | 26133.3328 | 43466.5125
scan1105 | 12:10 | 3:04 | 11026 | 183.7667 | 26018.6464 | 43255.4444
scan1106 | 12:10 | 3:04 | 11046 | 184.1 | 26069.3132 | 43353.1461
scan1107 | 12:10 | 3:05 | 11070 | 184.5 | 26072.2328 | 43374.1275
scan1108 | 12:18 | 3:12 | 11545 | 192.4167 | 26052.9124 | 43333.9589
scan1109 | 12:36 | 3:30 | 12598 | 209.9667 | 26027.1104 | 43290.0886
scan1110 | 12:36 | 3:30 | 12627 | 210.45 | 26044.668 | 43345.06
scan1111 | 12:37 | 3:31 | 12651 | 210.85 | 26041.386 | 43336.7442
scan1112 | 12:51 | 3:45 | 13516 | 225.2667 | 25983.054 | 43224.3819
scan1113 | 12:52 | 3:46 | 13537 | 225.6167 | 25984.516 | 43229.6722
scan1114 | 12:52 | 3:46 | 13558 | 225.9667 | 25974.49 | 43227.5944
scan1115 | 12:58 | 3:52 | 13933 | 232.2167 | 25975.5028 | 43207.5142 |  | 12:04 | switched on then waited for 5 minutes, to see if this has any stabilizing effect, then 3 scans to be taken every 10 minutes
scan1116 | 13:17 | 4:11 | 15035 | 250.5833 | 25959.3832 | 43183.9706
scan1117 | 13:17 | 4:11 | 15057 | 250.95 | 25967.966 | 43214.285
scan1118 | 13:17 | 4:11 | 15083 | 251.3833 | 25938.862 | 43165.1947
scan1119 | 13:29 | 4:23 | 15800 | 263.3333 | 25938.4096 | 43161.6961
scan1120 | 13:36 | 4:30 | 16186 | 269.7667 | 25899.7064 | 43133.6436
scan1121 | 13:39 | 4:33 | 16396 | 273.2667 | 25912.914 | 43149.3083
scan1122 | 13:45 | 4:39 | 16727 | 278.7833 | 25885.378 | 43115.1336
scan1123 | 13:49 | 4:43 | 16959 | 282.65 | 25887.3288 | 43117.1675
scan1124 | 14:13 | 5:07 | 18429 | 307.15 | 25879.5996 | 43107.7422
scan1125 | 14:13 | 5:07 | 18450 | 307.5 | 25934.7724 | 43201.4475
scan1126 | 14:14 | 5:08 | 18473 | 307.8833 | 25931.2532 | 43202.2756
scan1127 | 14:14 | 5:08 | 18499 | 308.3167 | 25974.3696 | 43286.6561
scan1128 | 14:38 | 5:32 | 19933 | 332.2167 | 25932.2612 | 43205.2231
scan1129 | 14:39 | 5:33 | 19958 | 332.6333 | 25934.1056 | 43229.2178
scan1130 | 14:40 | 5:34 | 20044 | 334.0667 | 25906.3084 | 43169.5236
scan1131 | 15:04 | 5:58 | 21490 | 358.1667 | 25885.578 | 43145.8506
scan1132 | 15:05 | 5:59 | 21511 | 358.5167 | 25923.6984 | 43207.0633
scan1133 | 15:05 | 5:59 | 21536 | 358.9333 | 25952.3628 | 43256.0519
scan1134 | 15:18 | 6:12 | 22336 | 372.2667 | 25930.9368 | 43210.7672
scan1135 | 15:19 | 6:13 | 22358 | 372.6333 | 25957.1136 | 43248.3611
scan1136 | 15:19 | 6:13 | 22379 | 372.9833 | 25946.9172 | 43249.8769

## scanner warm up day1 dose
|  |  |  | dose=c+b/(X-a) | [7 empty] | pixels to cm conversion | 1px = | 1/ | 1 inch = 2.54 centimetres
a |  |  |  | -0.040473 | [10 empty] | dpi | ppcm | cm/px
b |  |  |  | 2.7151 | [10 empty] | 72 | 28.3465 | 0.035278
c |  |  |  | -3.8345
positioned the film in the portrait orientation centrally to the scan array
repeat measurements taken
scanner turned on at
epson software used 72dpi; 48bit tiff format
preview scan taken; however, when in clinical use there will be no need for a preview scan if the template is in place; however, the preview scans could be used as a warm up perhaps; this could be before the film is placed on the plate.
or perhaps if there is a warm up effect then we should be doing preview scans before placeing it on the scanner. Is the preview scans the same as the real scan? Can 3 previews stop any warm up effect. Or should the real scan be an average of 3 scans? To remove the repeatability uncertainty?
I shall test this :
1 preview scan is required to get the positioning tof the scan area.
then reapeat scans can be carried out ( the previews can't be used for that because there is no image formed form the scan)
then if a warm up effect is observed then the scanner can be switched off and then 1 preview scan given - then a real scan - see if the real scan is the same as a fully settled scanner
then if 1 preview is not enough - the nallow to cool then repeat by doing 2 previews, and so on until the required amounts of warm up previews are done.
if repeatability is unsatisfactory then repeat scans can be done and the image averaged.
location of scan already determined before switched on (10 minutes ago)
time turned on 12.40
time first scan
time second scan
11:28
set 1
scan name | time of scan | time since scanner was turned on | seconds since scanner turned on | minutes since turned on | imagej read out
[5 empty] | region 1 | region 2 | region 1 | region 2 | 1 dose | 2 dose
scan1009 | 11:30 | 0:02 | 97 | 1.6167 | 26549.7076 | 44016.9394 | 0.40512 | 0.67166 | 2.2586 | -0.021942 | [7 empty] | 25th March scanner turned on at
scan1010 | 11:31 | 0:03 | 151 | 2.5167 | 26542.8156 | 44034.315 | 0.40502 | 0.67192 | 2.26 | -0.023361
scan1011 | 11:32 | 0:04 | 212 | 3.5333 | 26403.9188 | 43853.7719 | 0.4029 | 0.66917 | 2.2891 | -0.0085655
scan1012 | 11:33 | 0:05 | 274 | 4.5667 | 26442.9156 | 43955.6231 | 0.40349 | 0.67072 | 2.2809 | -0.016926
scan1013 | 11:34 | 0:06 | 332 | 5.5333 | 26430.1628 | 43923.1956 | 0.4033 | 0.67023 | 2.2836 | -0.014268
scan1014 | 11:35 | 0:07 | 392 | 6.5333 | 26430.7984 | 43920.1489 | 0.40331 | 0.67018 | 2.2835 | -0.014018
scan1015 | 11:36 | 0:08 | 452 | 7.5333 | 26437.6652 | 43922.2875 | 0.40341 | 0.67021 | 2.282 | -0.014194
scan1016 | 11:37 | 0:09 | 512 | 8.5333 | 26431.9068 | 43917.6383 | 0.40333 | 0.67014 | 2.2832 | -0.013813
scan1017 | 11:38 | 0:10 | 574 | 9.5667 | 26430.9268 | 43927.7806 | 0.40331 | 0.67029 | 2.2835 | -0.014644
scan1018 | 11:39 | 0:11 | 631 | 10.5167 | 26429.944 | 43923.8364 | 0.4033 | 0.67023 | 2.2837 | -0.014321
scan1019 | 12:50 | 1:22 | 4937 | 82.2833 | 26318.3308 | 43726.2978 | 0.40159 | 0.66722 | 2.3072 | 0.0019504 | [7 empty] | scanner was left for an hour then scanning resumed. This is after set 1
scan1020 | 12:51 | 1:23 | 4959 | 82.65 | 26352.192 | 43792.2456 | 0.40211 | 0.66823 | 2.3001 | -0.0034972 | [7 empty] | however, I turned the scanner off then immediately on before this set - perhaps this makes a difference
scan1021 | 12:51 | 1:23 | 4980 | 83 | 26355.2984 | 43813.2322 | 0.40216 | 0.66855 | 2.2994 | -0.0052275 | [7 empty] | scan name
scan1022 | 12:51 | 1:23 | 5007 | 83.45 | 26324.4524 | 43795.5989 | 0.40169 | 0.66828 | 2.3059 | -0.0037737
scan1023 | 12:52 | 1:24 | 5029 | 83.8167 | 26361.5988 | 43823.0917 | 0.40225 | 0.6687 | 2.2981 | -0.0060399
scan1024 | 12:52 | 1:24 | 5052 | 84.2 | 26327.7432 | 43792.9833 | 0.40174 | 0.66824 | 2.3052 | -0.003558
scan1025 | 12:53 | 1:25 | 5076 | 84.6 | 26324.9236 | 43787.8083 | 0.40169 | 0.66816 | 2.3058 | -0.0031311
scan1026 | 13:47 | 2:19 | 8369 | 139.4833 | 26343.7548 | 43783.7139 | 0.40198 | 0.6681 | 2.3018 | -0.0027933 | [7 empty] | then went for lunch and then did a set of scanning without turning scanner off
scan1027 | 13:48 | 2:20 | 8390 | 139.8333 | 26330.0492 | 43768.9119 | 0.40177 | 0.66787 | 2.3047 | -0.0015715
scan1028 | 13:48 | 2:20 | 8412 | 140.2 | 26325.0412 | 43769.5253 | 0.40169 | 0.66788 | 2.3058 | -0.0016221
scan1029 | 13:49 | 2:21 | 8435 | 140.5833 | 26300.588 | 43721.4097 | 0.40132 | 0.66715 | 2.311 | 0.0023548
scan1030 | 13:49 | 2:21 | 8457 | 140.95 | 26313.3156 | 43749.5819 | 0.40152 | 0.66758 | 2.3083 | 2.5247e-05
scan1031 | 13:49 | 2:21 | 8475 | 141.25 | 26318.6504 | 43748.5961 | 0.4016 | 0.66756 | 2.3072 | 0.00010672
scan1032 | 13:50 | 2:22 | 8502 | 141.7 | 26250.9672 | 43653.1731 | 0.40056 | 0.6661 | 2.3215 | 0.0080089 | imagej read out
scan1033 | 13:50 | 2:22 | 8523 | 142.05 | 26306.1532 | 43714.94 | 0.40141 | 0.66705 | 2.3098 | 0.0028901
scan1034 | 15:01 | 3:33 | 12754 | 212.5667 | 26254.2072 | 43662.4114 | 0.40061 | 0.66625 | 2.3209 | 0.0072424 | [7 empty] | then went to the group meeting without turning the scanner off:
scan1035 | 15:01 | 3:33 | 12775 | 212.9167 | 26185.036 | 43581.7264 | 0.39956 | 0.66501 | 2.3356 | 0.013947
scan1036 | 15:01 | 3:33 | 12796 | 213.2667 | 26184.8224 | 43580.9506 | 0.39955 | 0.665 | 2.3357 | 0.014011
scan1037 | 15:02 | 3:34 | 12817 | 213.6167 | 26245.2572 | 43670.0381 | 0.40048 | 0.66636 | 2.3228 | 0.0066099
scan1038 | 15:02 | 3:34 | 12838 | 213.9667 | 26251.8804 | 43681.5411 | 0.40058 | 0.66654 | 2.3214 | 0.0056563
scan1039 | 15:02 | 3:34 | 12861 | 214.35 | 26294.1868 | 43771.3803 | 0.40122 | 0.66791 | 2.3124 | -0.0017753
scan1040 | 15:04 | 3:36 | 12938 | 215.6333 | 26165.1944 | 43532.8014 | 0.39926 | 0.66427 | 2.3399 | 0.018024 | [7 empty] | then the scanner was turned off then back on again immediately and scanning resumed
scan1041 | 15:04 | 3:36 | 12959 | 215.9833 | 26240.0012 | 43656.1833 | 0.4004 | 0.66615 | 2.3239 | 0.0077591
scan1042 | 15:04 | 3:36 | 12980 | 216.3333 | 26195.4492 | 43609.7667 | 0.39972 | 0.66544 | 2.3334 | 0.011614
scan1043 | 15:05 | 3:37 | 13001 | 216.6833 | 26216.3248 | 43637.8997 | 0.40004 | 0.66587 | 2.3289 | 0.0092767
scan1044 | 15:05 | 3:37 | 13022 | 217.0333 | 26210.8464 | 43628.2767 | 0.39995 | 0.66572 | 2.3301 | 0.010076
scan1045 | 15:05 | 3:37 | 13043 | 217.3833 | 26204.0624 | 43623.8753 | 0.39985 | 0.66566 | 2.3316 | 0.010442

## scanner warm up day 2 dose
|  |  |  | dose=c+b/(X-a) |  |  |  | pixels to cm conversion | 1px = |  | 1/ | 1 inch = 2.54 centimetres
a |  |  |  | -0.040473 | [7 empty] | dpi | ppcm | cm/px
b |  |  |  | 2.7151 | [7 empty] | 72 | 28.3465 | 0.035278
c |  |  |  | -3.8345
scanner turned on on the 26th at 9.06 | 9:06 |  |  |  | r1 | r2 | r1 | r2 | r1 | r2
scan1081 | 9:08 | 0:02 | 125 | 2.0833 | 26217.4332 | 43553.2153 | 0.40005 | 0.66458 | 2.3287 | 0.016322
scan1082 | 9:08 | 0:02 | 146 | 2.4333 | 26215.1268 | 43553.6881 | 0.40002 | 0.66459 | 2.3292 | 0.016282
scan1083 | 9:09 | 0:03 | 169 | 2.8167 | 26229.2128 | 43576.5092 | 0.40023 | 0.66493 | 2.3262 | 0.014381
scan1084 | 9:09 | 0:03 | 193 | 3.2167 | 26108.9 | 43394.0336 | 0.3984 | 0.66215 | 2.352 | 0.029634
scan1085 | 9:10 | 0:04 | 220 | 3.6667 | 26101.1028 | 43380.0444 | 0.39828 | 0.66194 | 2.3536 | 0.030808
scan1086 | 9:10 | 0:04 | 241 | 4.0167 | 26099.9284 | 43388.2608 | 0.39826 | 0.66206 | 2.3539 | 0.030118
scan1087 | 9:10 | 0:04 | 265 | 4.4167 | 26114.54 | 43413.4761 | 0.39848 | 0.66245 | 2.3507 | 0.028003
scan1088 | 9:11 | 0:05 | 287 | 4.7833 | 26112.9456 | 43414.0839 | 0.39846 | 0.66246 | 2.3511 | 0.027952
scan1089 | 9:11 | 0:05 | 308 | 5.1333 | 26126.5196 | 43436.5533 | 0.39867 | 0.6628 | 2.3482 | 0.026069
scan1090 | 9:12 | 0:06 | 339 | 5.65 | 26101.99 | 43385.1011 | 0.39829 | 0.66201 | 2.3534 | 0.030384
scan1091 | 9:35 | 0:29 | 1761 | 29.35 | 26034.0016 | 43298.155 | 0.39725 | 0.66069 | 2.3681 | 0.037697
scan1092 | 9:36 | 0:30 | 1782 | 29.7 | 26161.544 | 43496.3675 | 0.3992 | 0.66371 | 2.3407 | 0.021065
scan1093 | 9:36 | 0:30 | 1806 | 30.1 | 26166.7132 | 43528.8864 | 0.39928 | 0.66421 | 2.3395 | 0.01835
scan1094 | 9:36 | 0:30 | 1828 | 30.4667 | 26058.2552 | 43370.015 | 0.39762 | 0.66178 | 2.3629 | 0.031651
scan1095 | 9:37 | 0:31 | 1849 | 30.8167 | 26063.4156 | 43381.5661 | 0.3977 | 0.66196 | 2.3618 | 0.030681
scan1096 | 9:37 | 0:31 | 1871 | 31.1833 | 26160.2688 | 43509.5278 | 0.39918 | 0.66391 | 2.3409 | 0.019966
scan1097 | 10:57 | 1:51 | 6681 | 111.35 | 26086.626 | 43348.1856 | 0.39806 | 0.66145 | 2.3567 | 0.033485 |  |  |  |  | turned off at 9.41; the idea is to test to see if the readings return
scan1098 | 10:58 | 1:52 | 6702 | 111.7 | 26123.9804 | 43424.8394 | 0.39863 | 0.66262 | 2.3487 | 0.02705 |  |  |  |  | to their original start of the day values or whether this
scan1099 | 10:58 | 1:52 | 6723 | 112.05 | 26131.5724 | 43434.0569 | 0.39874 | 0.66276 | 2.3471 | 0.026278 |  |  |  |  | strange decreasing in signal over the day shows itself again
scan1100 | 10:58 | 1:52 | 6747 | 112.45 | 26145.0888 | 43472.8247 | 0.39895 | 0.66335 | 2.3442 | 0.023033
scan1101 | 10:59 | 1:53 | 6773 | 112.8833 | 26097.0436 | 43399.8508 | 0.39822 | 0.66224 | 2.3545 | 0.029146
scan1102 | 10:59 | 1:53 | 6797 | 113.2833 | 26127.2756 | 43458.3069 | 0.39868 | 0.66313 | 2.348 | 0.024248
scan1103 | 11:00 | 1:54 | 6818 | 113.6333 | 26159.0016 | 43496.1131 | 0.39916 | 0.66371 | 2.3412 | 0.021086
scan1104 | 11:00 | 1:54 | 6838 | 113.9667 | 26133.3328 | 43466.5125 | 0.39877 | 0.66326 | 2.3467 | 0.023561
scan1105 | 12:10 | 3:04 | 11026 | 183.7667 | 26018.6464 | 43255.4444 | 0.39702 | 0.66004 | 2.3714 | 0.041299
scan1106 | 12:10 | 3:04 | 11046 | 184.1 | 26069.3132 | 43353.1461 | 0.39779 | 0.66153 | 2.3605 | 0.033068
scan1107 | 12:10 | 3:05 | 11070 | 184.5 | 26072.2328 | 43374.1275 | 0.39784 | 0.66185 | 2.3599 | 0.031305
scan1108 | 12:18 | 3:12 | 11545 | 192.4167 | 26052.9124 | 43333.9589 | 0.39754 | 0.66123 | 2.364 | 0.034682
scan1109 | 12:36 | 3:30 | 12598 | 209.9667 | 26027.1104 | 43290.0886 | 0.39715 | 0.66056 | 2.3696 | 0.038377
scan1110 | 12:36 | 3:30 | 12627 | 210.45 | 26044.668 | 43345.06 | 0.39742 | 0.6614 | 2.3658 | 0.033748
scan1111 | 12:37 | 3:31 | 12651 | 210.85 | 26041.386 | 43336.7442 | 0.39737 | 0.66128 | 2.3665 | 0.034448
scan1112 | 12:51 | 3:45 | 13516 | 225.2667 | 25983.054 | 43224.3819 | 0.39648 | 0.65956 | 2.3791 | 0.043924
scan1113 | 12:52 | 3:46 | 13537 | 225.6167 | 25984.516 | 43229.6722 | 0.3965 | 0.65964 | 2.3788 | 0.043476
scan1114 | 12:52 | 3:46 | 13558 | 225.9667 | 25974.49 | 43227.5944 | 0.39635 | 0.65961 | 2.381 | 0.043652
scan1115 | 12:58 | 3:52 | 13933 | 232.2167 | 25975.5028 | 43207.5142 | 0.39636 | 0.6593 | 2.3808 | 0.04535 |  |  |  |  | 12:04 | switched on then waited for 5 minutes, to see if this has any stabilizing effect, then 3 scans to be taken every 10 minutes
scan1116 | 13:17 | 4:11 | 15035 | 250.5833 | 25959.3832 | 43183.9706 | 0.39611 | 0.65895 | 2.3843 | 0.047343
scan1117 | 13:17 | 4:11 | 15057 | 250.95 | 25967.966 | 43214.285 | 0.39625 | 0.65941 | 2.3824 | 0.044777
scan1118 | 13:17 | 4:11 | 15083 | 251.3833 | 25938.862 | 43165.1947 | 0.3958 | 0.65866 | 2.3887 | 0.048934
scan1119 | 13:29 | 4:23 | 15800 | 263.3333 | 25938.4096 | 43161.6961 | 0.39579 | 0.65861 | 2.3888 | 0.04923
scan1120 | 13:36 | 4:30 | 16186 | 269.7667 | 25899.7064 | 43133.6436 | 0.3952 | 0.65818 | 2.3973 | 0.05161
scan1121 | 13:39 | 4:33 | 16396 | 273.2667 | 25912.914 | 43149.3083 | 0.39541 | 0.65842 | 2.3944 | 0.050281
scan1122 | 13:45 | 4:39 | 16727 | 278.7833 | 25885.378 | 43115.1336 | 0.39499 | 0.65789 | 2.4004 | 0.053182
scan1123 | 13:49 | 4:43 | 16959 | 282.65 | 25887.3288 | 43117.1675 | 0.39502 | 0.65793 | 2.4 | 0.053009
scan1124 | 14:13 | 5:07 | 18429 | 307.15 | 25879.5996 | 43107.7422 | 0.3949 | 0.65778 | 2.4017 | 0.05381
scan1125 | 14:13 | 5:07 | 18450 | 307.5 | 25934.7724 | 43201.4475 | 0.39574 | 0.65921 | 2.3896 | 0.045864
scan1126 | 14:14 | 5:08 | 18473 | 307.8833 | 25931.2532 | 43202.2756 | 0.39569 | 0.65922 | 2.3904 | 0.045793
scan1127 | 14:14 | 5:08 | 18499 | 308.3167 | 25974.3696 | 43286.6561 | 0.39634 | 0.66051 | 2.381 | 0.038666
scan1128 | 14:38 | 5:32 | 19933 | 332.2167 | 25932.2612 | 43205.2231 | 0.3957 | 0.65927 | 2.3902 | 0.045544
scan1129 | 14:39 | 5:33 | 19958 | 332.6333 | 25934.1056 | 43229.2178 | 0.39573 | 0.65964 | 2.3898 | 0.043515
scan1130 | 14:40 | 5:34 | 20044 | 334.0667 | 25906.3084 | 43169.5236 | 0.3953 | 0.65872 | 2.3958 | 0.048567
scan1131 | 15:04 | 5:58 | 21490 | 358.1667 | 25885.578 | 43145.8506 | 0.39499 | 0.65836 | 2.4004 | 0.050574
scan1132 | 15:05 | 5:59 | 21511 | 358.5167 | 25923.6984 | 43207.0633 | 0.39557 | 0.6593 | 2.392 | 0.045388
scan1133 | 15:05 | 5:59 | 21536 | 358.9333 | 25952.3628 | 43256.0519 | 0.39601 | 0.66005 | 2.3858 | 0.041248
scan1134 | 15:18 | 6:12 | 22336 | 372.2667 | 25930.9368 | 43210.7672 | 0.39568 | 0.65935 | 2.3905 | 0.045075
scan1135 | 15:19 | 6:13 | 22358 | 372.6333 | 25957.1136 | 43248.3611 | 0.39608 | 0.65993 | 2.3848 | 0.041897
scan1136 | 15:19 | 6:13 | 22379 | 372.9833 | 25946.9172 | 43249.8769 | 0.39592 | 0.65995 | 2.387 | 0.041769

## warmup d3 and uniformity
This sheet is for looking at and recording, processing the results from the scans that were taken to discount film darkening as a reason for the dip in response seen over time.
It is also used for looking at the scans taken to investigate non-uniformity of scanner corresponding to a dose greater than what was used previously; in fact previously a non-irradiated film was used. This was instructuive in that it showed that there was change in response over the area but that this corresponded ot a dose of less than 0.1% The area alluded to here is the area with the size of film sheet in portrait orientation, aligned with the template.
It is hypothesised that, although the non-uniformity drops off quite severly near to the edge of the scanner, that this can be mitigated by simply not scanning in this region. Provided that the film is placed on the template, the uncertainty in dose is reduced to below 0.1%. The non-uniformity in the scan drection is lower than the non-uniformity in the lateral direction. One paper fits a parabola to this lateral scan artefact and uses this fit to correct the readings from the scanner. This could be adopted in our work but given that the dose variation is low it can be deemed to not be necessary; this depends upon the required accuracy of dose of course.
So firstly I shall look at the change in response over time. In this test I cut two pieces of film from the same film sheet, scanned them both (using 3 consecuteive scan each time) initially at time zero; I then scanned one piece of film multiple amount of times over the proceeding 2 hours. Then the two pieces were scanned in the same fashion as at time zero. The results will hopefully show whether the change in response is due to to the film darkening over time - perhaps as a result of the scanner light - or whether it is a facet of the scanner itself - a warm-up effect. In practical terms what we are inspecting the films for is whether the responses are the same at the end or whether the film that scanned repeatedly has a lower reading on the scanner.
scanner switched on at | 9:55
film piece 1
name | pixel value | time taken | time since switch on | seconds | minutes
scan1174 | 43874.1832 | 9:57 | 0:02 | 110 | 1.8333
scan1175 | 43852.2614 | 9:57 | 0:02 | 144 | 2.4
scan1178 | 43744.4402 | 9:59 | 0:04 | 256 | 4.2667
scan1179 | 43746.5693 | 10:00 | 0:05 | 275 | 4.5833
scan1180 | 43571.3208 | 10:00 | 0:05 | 293 | 4.8833
scan1181 | 43568.6543 | 10:00 | 0:05 | 315 | 5.25
scan1182 | 43743.1632 | 10:01 | 0:06 | 331 | 5.5167
scan1183 | 43760.9685 | 10:01 | 0:06 | 345 | 5.75
scan1184 | 43764.0501 | 10:01 | 0:06 | 360 | 6
scan1185 | 43739.4139 | 10:15 | 0:20 | 1218 | 20.3
scan1186 | 43545.4705 | 10:22 | 0:27 | 1608 | 26.8
scan1187 | 43822.9776 | 10:22 | 0:27 | 1624 | 27.0667
scan1188 | 43841.3612 | 10:22 | 0:27 | 1649 | 27.4833
scan1189 | 43833.9317 | 10:23 | 0:28 | 1708 | 28.4667
scan1190 | 43812.2984 | 10:55 | 1:00 | 3572 | 59.5333
scan1191 | 43780.3707 | 10:55 | 1:00 | 3597 | 59.95
scan1192 | 43774.2974 | 10:56 | 1:01 | 3640 | 60.6667
scan1193 | 43749.7499 | 10:59 | 1:04 | 3811 | 63.5167
scan1194 | 43757.7067 | 10:59 | 1:04 | 3825 | 63.75
scan1195 | 43829.2028 | 10:59 | 1:04 | 3839 | 63.9833
scan1228 | 43859.9152 | 12:08 | 2:13 | 7966 | 132.7667
scan1229 | 43884.9945 | 12:08 | 2:13 | 7980 | 133
scan1230 | 43894.8636 | 12:08 | 2:13 | 7993 | 133.2167
film piece 2
name | pixel value | time taken | time since switch on | seconds | minutes
scan1231 | 43723.6139 | 9:58 | 0:03 | 186 | 3.1
scan1232 | 43751.8491 | 9:58 | 0:03 | 200 | 3.3333
scan1233 | 43757.7481 | 11:00 | 1:05 | 3882 | 64.7
scan1234 | 43744.777 | 11:00 | 1:05 | 3897 | 64.95
scan1235 | 43736.1939 | 11:00 | 1:05 | 3913 | 65.2167
scan1231 | 43742.984 | 12:09 | 2:14 | 8028 | 133.8
scan1232 | 43718.9051 | 12:09 | 2:14 | 8041 | 134.0167
scan1233 | 43708.324 | 12:09 | 2:14 | 8059 | 134.3167
scan1234 | 43701.3061 | 12:10 | 2:15 | 8071 | 134.5167
| position | distance from left edge
scan1199 | near side; left | 2 |  | 25180.1336
scan1200 | near side; left | 2 |  | 25177.7021
scan1201 | near side; left | 2 |  | 25178.9896
scan1202 | near side; centre | 10 |  | 24509.1314
scan1203 | near side; centre | 10 |  | 24540.2782
scan1204 | near side; centre | 10 |  | 24559.9417
scan1205 | near side; right | 18 |  | 23602.5179
scan1206 | near side; right | 18 |  | 23640.9946
scan1207 | near side; right | 18 |  | 23643.9554
scan1208 | far side; centre | 10 |  | 24299.0037
scan1209 | far side; centre | 10 |  | 24338.7142
scan1210 | far side; centre | 10 |  | 24347.5846
scan1211 | far side; left | 2 |  | 25111.9088
scan1212 | far side; left | 2 |  | 25068.5346
scan1213 | far side; left | 2 |  | 25097.4535
scan1214 | far side; right | 18 |  | 23985.3595
scan1215 | far side; right | 18 |  | 23981.1491
scan1216 | far side; right | 18 |  | 23986.1106
scan1217 | centre; centre | 10 |  | 24131.9721
scan1218 | centre; centre | 10 |  | 24172.2061
scan1219 | centre; centre | 10 |  | 24184.2255
scan1220 | centre; centre | 10 |  | 24181.511
scan1221 | centre; right | 18 |  | 23826.2875
scan1222 | centre; right | 18 |  | 23826.6385
scan1223 | centre; right | 18 |  | 23833.5583
scan1224 | centre; right | 18 |  | 23855.5115
scan1225 | centre; left | 2 |  | 25279.4969
scan1226 | centre; left | 2 |  | 25252.0385
scan1227 | centre; left | 2 |  | 25246.3792
43264.8407
43625.1003
43791.3336
43806.3743

## uniformity discrete
at scan array edge
centred
3cm left
6cm left
3cm right
6cm right
5cm from array edge
centred
3cm left
6cm left
3cm right
6cm right
10cm from array edge
centred
3cm left
6cm left
3cm right
6cm right
15cm from array edge
centred
3cm left
6cm left
3cm right
6cm right
150 is the first one where I decided to look at uniformity in the region that I will be using
154 starts the correct batch
as fraction of the max response of 65535
0.67285
0.66313
0.0097173
0.0097173
now, only looking at the uniformity of the region defined by the template
scan side
| centre | 44020.4672 | 44099.3126 | 10 | difference between centre and left side | 636.8226 | centre
| centre | 44107.8623 | 43554.0181 | 2 | difference between centre and right side | 208.766 | left
| centre | 44169.6083 | 43921.2454 | 18 | % drop in response from centre to left side | -0.014442 | difference
| left | 43434.3351
| left | 43673.7011
| rt | 43915.4637
| rt | 43927.027
centre
| centre | 44046.5276 | 44095.1147 | 10
| centre | 44143.7018 | 43458.2921 | 2
| left | 43465.0592 | 43886.3486 | 18
| left | 43451.5249
| rt | 43861.445
| rt | 43911.2523
far side
| centre | 44179.9943 | 44187.0499 | 10
| centre | 44194.1054 | 43459.2523 | 2
| left | 43429.6464 | 44005.1193 | 18
| left | 43488.8583
| rt | 43975.627
| rt | 44034.6116

## rational cal
| Calibration data, 16-bit # |  |  |  |  | scale by: | new film |  | original cal film
| Dose, cGy | Red | Green | Blue | k | l | 0 dose | other dose | 0 dose | other dose
| 300 | 22704.438 | 26601.2067 | 19960.2125 |  |  |  |  | 44441
| 200 | 26397.3885 | 30047.5799 | 21188.3132
| 100 | 32547.5085 | 35131.9428 | 22942.7761 | [8 empty] | 0.67285 | 0.019987 | these two are the difference in dose between the middle and left of the scan, for a unirradiated gafchromic film
| 0 | 44440.5987 | 41860.923 | 24778.5695 | [8 empty] | 0.66313 | 0.057634 |  | 0.037647
[11 empty] | 0 | 54073.3438 | 0.82511 | -0.43581
[11 empty] | 1 | 53979.8945 | 0.82368 | -0.43245
| Calibration data, 16-bit #/65535 | [9 empty] | 2 | 54002.2891 | 0.82402 | -0.43325
| Dose, Gy | Red | Green | Blue | [6 empty] | 3 | 54064.6328 | 0.82497 | -0.4355
| 3 | 0.34645 | 0.40591 | 0.30457 | [6 empty] | 4 | 53972.7109 | 0.82357 | -0.43219
| 2 | 0.4028 | 0.4585 | 0.32331 | [6 empty] | 5 | 54064.6328 | 0.82497 | -0.4355
| 1 | 0.49664 | 0.53608 | 0.35008 |  |  | k=1.22476 |  |  |  | 6 | 54001.1562 | 0.824 | -0.43321
| 0 | 0.67812 | 0.63876 | 0.3781 | [6 empty] | 7 | 54025.6562 | 0.82438 | -0.4341
[11 empty] | 8 | 54024.5508 | 0.82436 | -0.43406
| Calibration curve | [9 empty] | 9 | 54018.1836 | 0.82426 | -0.43383
|  | X(D) = A + B/(Dose -C) | [8 empty] | 10 | 54026.7891 | 0.8244 | -0.43414
where X(D) is response expressed as 16-bit #/65535, D is dose in Gy and A, B and C are constants | [10 empty] | 11 | 54049.7109 | 0.82475 | -0.43496
[11 empty] | 12 | 54091.9727 | 0.82539 | -0.43648
|  | Red | Green | Blue | [6 empty] | 13 | 53954.6836 | 0.8233 | -0.43154
| A | 0.11252 | 0.02256 | 0.063139 |  |  |  | -12.7365 |  |  | 14 | 54050.8672 | 0.82476 | -0.435
| B | 1.1953 | 3.0228 | 3.0654 | [6 empty] | 15 | 54027.6328 | 0.82441 | -0.43417
| C | -2.1132 | -4.903 | -9.7235 | [6 empty] | 16 | 54033.3945 | 0.8245 | -0.43437
[11 empty] | 17 | 54122.8438 | 0.82586 | -0.43759
[11 empty] | 18 | 54111.9219 | 0.8257 | -0.4372
[11 empty] | 19 | 54059.9727 | 0.8249 | -0.43533
[11 empty] | 20 | 54021.8438 | 0.82432 | -0.43396
[11 empty] | 21 | 53995.6055 | 0.82392 | -0.43301
| X(D) | Dose(Gy) | [8 empty] | 22 | 53971.5508 | 0.82355 | -0.43215
| 0.069654 | -30 | [8 empty] | 23 | 54014.1562 | 0.8242 | -0.43368
| 0.0695 | -29.9 | [8 empty] | 24 | 54008.6836 | 0.82412 | -0.43349
| 0.069344 | -29.8 | [8 empty] | 25 | 54071.7891 | 0.82508 | -0.43575
| 0.069188 | -29.7 | [8 empty] | 26 | 54019.9219 | 0.82429 | -0.43389
| 0.06903 | -29.6 | [8 empty] | 27 | 54004.4492 | 0.82406 | -0.43333
| 0.068871 | -29.5 | [8 empty] | 28 | 53978.7617 | 0.82366 | -0.43241
| 0.068711 | -29.4 | [8 empty] | 29 | 53951.3672 | 0.82325 | -0.43142
| 0.06855 | -29.3 | [8 empty] | 30 | 53953.1562 | 0.82327 | -0.43148
| 0.068388 | -29.2 | [8 empty] | 31 | 53965.0273 | 0.82345 | -0.43191
| 0.068224 | -29.1 | [8 empty] | 32 | 53953.5781 | 0.82328 | -0.4315
| 0.06806 | -29 | [8 empty] | 33 | 53967.6055 | 0.82349 | -0.432
| 0.067894 | -28.9 | [8 empty] | 34 | 53924.4492 | 0.82283 | -0.43045
| 0.067727 | -28.8 | [8 empty] | 35 | 54087.2383 | 0.82532 | -0.43631
| 0.067558 | -28.7 | [8 empty] | 36 | 53973.8672 | 0.82359 | -0.43223
| 0.067388 | -28.6 | [8 empty] | 37 | 53979.0273 | 0.82367 | -0.43242
| 0.067217 | -28.5 | [8 empty] | 38 | 53917.9492 | 0.82274 | -0.43021
| 0.067045 | -28.4 | [8 empty] | 39 | 54004.8672 | 0.82406 | -0.43335
| 0.066871 | -28.3 | 0.11 | [7 empty] | 40 | 53974.6836 | 0.8236 | -0.43226
| 0.066696 | -28.2 | [8 empty] | 41 | 53966.6055 | 0.82348 | -0.43197
| 0.06652 | -28.1 | [8 empty] | 42 | 53885.9219 | 0.82225 | -0.42905
| 0.066342 | -28 | [8 empty] | 43 | 53902.4219 | 0.8225 | -0.42965
| 0.066163 | -27.9 | [8 empty] | 44 | 53899.5 | 0.82245 | -0.42954
| 0.065983 | -27.8 | [8 empty] | 45 | 53879.4219 | 0.82215 | -0.42882
| 0.065801 | -27.7 | [8 empty] | 46 | 53806.7109 | 0.82104 | -0.42618
| 0.065618 | -27.6 | [8 empty] | 47 | 53850.4219 | 0.8217 | -0.42777
| 0.065433 | -27.5 | [8 empty] | 48 | 53807.7617 | 0.82105 | -0.42622
| 0.065247 | -27.4 | [8 empty] | 49 | 53850.3438 | 0.8217 | -0.42776
| 0.065059 | -27.3 | [8 empty] | 50 | 53822.5 | 0.82128 | -0.42675
| 0.06487 | -27.2 | [8 empty] | 51 | 53820.8672 | 0.82125 | -0.42669
| 0.064679 | -27.1 | [8 empty] | 52 | 53814.5781 | 0.82116 | -0.42646
| 0.064487 | -27 | [8 empty] | 53 | 53708.6562 | 0.81954 | -0.42261
| 0.064293 | -26.9 | [8 empty] | 54 | 53730.2617 | 0.81987 | -0.4234
| 0.064098 | -26.8 | [8 empty] | 55 | 53699.7891 | 0.81941 | -0.42228
| 0.063901 | -26.7 | [8 empty] | 56 | 53702.4492 | 0.81945 | -0.42238
| 0.063702 | -26.6 | [8 empty] | 57 | 53686.5 | 0.8192 | -0.4218
| 0.063502 | -26.5 | [8 empty] | 58 | 53669.6562 | 0.81895 | -0.42118
| 0.0633 | -26.4 | [8 empty] | 59 | 53693.4492 | 0.81931 | -0.42205
| 0.063097 | -26.3 | [8 empty] | 60 | 53655.2109 | 0.81873 | -0.42066
| 0.062892 | -26.2 | [8 empty] | 61 | 53602 | 0.81791 | -0.41871
| 0.062685 | -26.1 | [8 empty] | 62 | 53626.3672 | 0.81829 | -0.4196
| 0.062476 | -26 | [8 empty] | 63 | 53638.9492 | 0.81848 | -0.42006
| 0.062266 | -25.9 | [8 empty] | 64 | 53651.6328 | 0.81867 | -0.42053
| 0.062054 | -25.8 | [8 empty] | 65 | 53529.1836 | 0.8168 | -0.41603
| 0.06184 | -25.7 | [8 empty] | 66 | 53544.5781 | 0.81704 | -0.4166
| 0.061624 | -25.6 | [8 empty] | 67 | 53569 | 0.81741 | -0.4175
| 0.061406 | -25.5 | [8 empty] | 68 | 53402.9492 | 0.81488 | -0.41138
| 0.061187 | -25.4 | [8 empty] | 69 | 53414.5 | 0.81505 | -0.41181
| 0.060966 | -25.3 | [8 empty] | 70 | 53430.3945 | 0.8153 | -0.41239
| 0.060742 | -25.2 | [8 empty] | 71 | 53384.6328 | 0.8146 | -0.4107
| 0.060517 | -25.1 | [8 empty] | 72 | 53426.3438 | 0.81523 | -0.41224
| 0.06029 | -25 | [8 empty] | 73 | 53412.5508 | 0.81502 | -0.41174
| 0.060061 | -24.9 | [8 empty] | 74 | 53292.1328 | 0.81319 | -0.40727
| 0.059829 | -24.8 | [8 empty] | 75 | 53252.5 | 0.81258 | -0.4058
| 0.059596 | -24.7 | [8 empty] | 76 | 53218.1562 | 0.81206 | -0.40452
| 0.059361 | -24.6 | [8 empty] | 77 | 53146.3164 | 0.81096 | -0.40184
| 0.059123 | -24.5 | [8 empty] | 78 | 53114.6328 | 0.81048 | -0.40065
| 0.058884 | -24.4 | [8 empty] | 79 | 53031 | 0.8092 | -0.39752
| 0.058642 | -24.3 | [8 empty] | 80 | 52971.8672 | 0.8083 | -0.39529
| 0.058398 | -24.2 | [8 empty] | 81 | 52911.6562 | 0.80738 | -0.39302
| 0.058152 | -24.1 | [8 empty] | 82 | 52910.8945 | 0.80737 | -0.39299
| 0.057904 | -24 | [8 empty] | 83 | 52847.9219 | 0.80641 | -0.39061
| 0.057653 | -23.9 | [8 empty] | 84 | 52765.0273 | 0.80514 | -0.38746
| 0.0574 | -23.8 | [8 empty] | 85 | 52639.0781 | 0.80322 | -0.38266
| 0.057145 | -23.7 | [8 empty] | 86 | 52590.3945 | 0.80248 | -0.3808
| 0.056887 | -23.6 | [8 empty] | 87 | 52498.2891 | 0.80107 | -0.37726
| 0.056627 | -23.5 | [8 empty] | 88 | 52464.2109 | 0.80055 | -0.37595
| 0.056364 | -23.4 | [8 empty] | 89 | 52413.0273 | 0.79977 | -0.37398
| 0.056099 | -23.3 | [8 empty] | 90 | 52328.5 | 0.79848 | -0.37071
| 0.055832 | -23.2 | [8 empty] | 91 | 52300.6055 | 0.79806 | -0.36962
| 0.055562 | -23.1 | [8 empty] | 92 | 52175.8672 | 0.79615 | -0.36477
| 0.055289 | -23 | [8 empty] | 93 | 52020.7109 | 0.79379 | -0.35869
| 0.055014 | -22.9 | [8 empty] | 94 | 51950.3164 | 0.79271 | -0.35592
| 0.054736 | -22.8 | [8 empty] | 95 | 51779.5 | 0.7901 | -0.34916
| 0.054455 | -22.7 | [8 empty] | 96 | 51632.5 | 0.78786 | -0.3433
| 0.054172 | -22.6 | [8 empty] | 97 | 51481.9219 | 0.78556 | -0.33726
| 0.053885 | -22.5 | [8 empty] | 98 | 51302.4492 | 0.78283 | -0.33
| 0.053596 | -22.4 | [8 empty] | 99 | 51031.2891 | 0.77869 | -0.31893
| 0.053304 | -22.3 | [8 empty] | 100 | 50633.2383 | 0.77261 | -0.30242
| 0.05301 | -22.2 | [8 empty] | 101 | 50300.0781 | 0.76753 | -0.28837
| 0.052712 | -22.1 | [8 empty] | 102 | 49855.8672 | 0.76075 | -0.26928
| 0.052411 | -22 | [8 empty] | 103 | 49166.4727 | 0.75023 | -0.23887
| 0.052107 | -21.9 | [8 empty] | 104 | 48446.4492 | 0.73925 | -0.20601
| 0.051801 | -21.8 | [8 empty] | 105 | 47386.6836 | 0.72307 | -0.1555
| 0.051491 | -21.7 | [8 empty] | 106 | 46259.6055 | 0.70588 | -0.098752
| 0.051177 | -21.6 | [8 empty] | 107 | 44699.8945 | 0.68208 | -0.014576
| 0.050861 | -21.5 | [8 empty] | 108 | 43118.1055 | 0.65794 | 0.078295
| 0.050541 | -21.4 | [8 empty] | 109 | 41829.8945 | 0.63828 | 0.16023
| 0.050218 | -21.3 | [8 empty] | 110 | 40488.2109 | 0.61781 | 0.25234
| 0.049892 | -21.2 | [8 empty] | 111 | 39390.1328 | 0.60105 | 0.33347
| 0.049562 | -21.1 | [8 empty] | 112 | 38662.0508 | 0.58995 | 0.39041
| 0.049229 | -21 | [8 empty] | 113 | 38055.3164 | 0.58069 | 0.43992
| 0.048892 | -20.9 | [8 empty] | 114 | 37793.0273 | 0.57668 | 0.46193
| 0.048551 | -20.8 | [8 empty] | 115 | 37471.4727 | 0.57178 | 0.48944
| 0.048207 | -20.7 | [8 empty] | 116 | 37270.0508 | 0.5687 | 0.50698
| 0.047859 | -20.6 | [8 empty] | 117 | 37182.4492 | 0.56737 | 0.51468
| 0.047508 | -20.5 | [8 empty] | 118 | 37058.6328 | 0.56548 | 0.52564
| 0.047152 | -20.4 | [8 empty] | 119 | 37031.1055 | 0.56506 | 0.52809
| 0.046793 | -20.3 | [8 empty] | 120 | 36981.8164 | 0.56431 | 0.53249
| 0.04643 | -20.2 | [8 empty] | 121 | 36923.1328 | 0.56341 | 0.53774
| 0.046062 | -20.1 | [8 empty] | 122 | 36885.3438 | 0.56283 | 0.54114
| 0.045691 | -20 | [8 empty] | 123 | 36793.7891 | 0.56144 | 0.5494
| 0.045315 | -19.9 | [8 empty] | 124 | 36732.6328 | 0.5605 | 0.55494
| 0.044935 | -19.8 | [8 empty] | 125 | 36772.5 | 0.56111 | 0.55132
| 0.044551 | -19.7 | [8 empty] | 126 | 36809.3438 | 0.56167 | 0.54799
| 0.044162 | -19.6 | [8 empty] | 127 | 36825.8945 | 0.56193 | 0.54649
| 0.043769 | -19.5 | [8 empty] | 128 | 36720.2383 | 0.56031 | 0.55607
| 0.043371 | -19.4 | [8 empty] | 129 | 36721.7891 | 0.56034 | 0.55593
| 0.042969 | -19.3 | [8 empty] | 130 | 36732.7891 | 0.56051 | 0.55493
| 0.042562 | -19.2 | [8 empty] | 131 | 36692.1836 | 0.55989 | 0.55862
| 0.04215 | -19.1 | [8 empty] | 132 | 36700.7383 | 0.56002 | 0.55784
| 0.041733 | -19 | [8 empty] | 133 | 36722.8945 | 0.56036 | 0.55583
| 0.041312 | -18.9 | [8 empty] | 134 | 36639.9492 | 0.55909 | 0.56339
| 0.040885 | -18.8 | [8 empty] | 135 | 36582.0508 | 0.55821 | 0.5687
| 0.040453 | -18.7 | [8 empty] | 136 | 36633.5508 | 0.55899 | 0.56398
| 0.040016 | -18.6 | [8 empty] | 137 | 36702.8164 | 0.56005 | 0.55765
| 0.039574 | -18.5 | [8 empty] | 138 | 36661.4492 | 0.55942 | 0.56143
| 0.039126 | -18.4 | [8 empty] | 139 | 36655.6562 | 0.55933 | 0.56196
| 0.038672 | -18.3 | [8 empty] | 140 | 36686.9492 | 0.55981 | 0.5591
| 0.038213 | -18.2 | [8 empty] | 141 | 36671.8672 | 0.55958 | 0.56048
| 0.037748 | -18.1 | [8 empty] | 142 | 36657.7383 | 0.55936 | 0.56177
| 0.037278 | -18 | [8 empty] | 143 | 36646.6328 | 0.55919 | 0.56278
| 0.036801 | -17.9 | [8 empty] | 144 | 36646.7891 | 0.55919 | 0.56277
| 0.036319 | -17.8 | [8 empty] | 145 | 36633.4219 | 0.55899 | 0.56399
| 0.03583 | -17.7 | [8 empty] | 146 | 36602.5781 | 0.55852 | 0.56681
| 0.035335 | -17.6 | [8 empty] | 147 | 36652.4492 | 0.55928 | 0.56225
| 0.034833 | -17.5 | [8 empty] | 148 | 36607.2891 | 0.55859 | 0.56638
| 0.034325 | -17.4 | [8 empty] | 149 | 36595.4219 | 0.55841 | 0.56747
| 0.03381 | -17.3 | [8 empty] | 150 | 36621.1328 | 0.5588 | 0.56511
| 0.033288 | -17.2 | [8 empty] | 151 | 36637.7383 | 0.55906 | 0.56359
| 0.03276 | -17.1 | [8 empty] | 152 | 36608.3438 | 0.55861 | 0.56629
| 0.032224 | -17 | [8 empty] | 153 | 36626.2383 | 0.55888 | 0.56465
| 0.031681 | -16.9 | [8 empty] | 154 | 36658.1328 | 0.55937 | 0.56173
| 0.03113 | -16.8 | [8 empty] | 155 | 36628.8164 | 0.55892 | 0.56441
| 0.030572 | -16.7 | [8 empty] | 156 | 36616.5273 | 0.55873 | 0.56554
| 0.030007 | -16.6 | [8 empty] | 157 | 36602.2383 | 0.55851 | 0.56685
| 0.029433 | -16.5 | [8 empty] | 158 | 36597.7109 | 0.55845 | 0.56726
| 0.028852 | -16.4 | [8 empty] | 159 | 36635.2383 | 0.55902 | 0.56382
| 0.028262 | -16.3 | [8 empty] | 160 | 36659.6328 | 0.55939 | 0.56159
| 0.027664 | -16.2 | [8 empty] | 161 | 36671.9727 | 0.55958 | 0.56047
| 0.027057 | -16.1 | [8 empty] | 162 | 36670.5781 | 0.55956 | 0.56059
| 0.026442 | -16 | [8 empty] | 163 | 36699.3438 | 0.56 | 0.55797
| 0.025818 | -15.9 | [8 empty] | 164 | 36704.2891 | 0.56007 | 0.55752
| 0.025184 | -15.8 | [8 empty] | 165 | 36683.7383 | 0.55976 | 0.55939
| 0.024541 | -15.7 | [8 empty] | 166 | 36656.4219 | 0.55934 | 0.56189
| 0.023889 | -15.6 | [8 empty] | 167 | 36734.8945 | 0.56054 | 0.55474
| 0.023227 | -15.5 | [8 empty] | 168 | 36771.1328 | 0.56109 | 0.55145
| 0.022555 | -15.4 | [8 empty] | 169 | 36705.8672 | 0.5601 | 0.55738
| 0.021873 | -15.3 | [8 empty] | 170 | 36648.0781 | 0.55921 | 0.56265
| 0.02118 | -15.2 | [8 empty] | 171 | 36667.1836 | 0.55951 | 0.5609
| 0.020477 | -15.1 | [8 empty] | 172 | 36703.2383 | 0.56006 | 0.55762
| 0.019763 | -15 | [8 empty] | 173 | 36687.5273 | 0.55982 | 0.55905
| 0.019037 | -14.9 | [8 empty] | 174 | 36686.7383 | 0.5598 | 0.55912
| 0.0183 | -14.8 | [8 empty] | 175 | 36689.2617 | 0.55984 | 0.55889
| 0.017552 | -14.7 | [8 empty] | 176 | 36701.6055 | 0.56003 | 0.55776
| 0.016791 | -14.6 | [8 empty] | 177 | 36633.3945 | 0.55899 | 0.56399
| 0.016018 | -14.5 | [8 empty] | 178 | 36646.4219 | 0.55919 | 0.5628
| 0.015233 | -14.4 | [8 empty] | 179 | 36691.7617 | 0.55988 | 0.55866
| 0.014435 | -14.3 | [8 empty] | 180 | 36703.9727 | 0.56007 | 0.55755
| 0.013623 | -14.2 | [8 empty] | 181 | 36708.6836 | 0.56014 | 0.55712
| 0.012798 | -14.1 | [8 empty] | 182 | 36741.5 | 0.56064 | 0.55414
| 0.011959 | -14 | [8 empty] | 183 | 36694.9219 | 0.55993 | 0.55837
| 0.011106 | -13.9 | [8 empty] | 184 | 36660.1055 | 0.5594 | 0.56155
| 0.010239 | -13.8 | [8 empty] | 185 | 36628.4727 | 0.55891 | 0.56444
| 0.0093558 | -13.7 | [8 empty] | 186 | 36653.8438 | 0.5593 | 0.56212
| 0.0084578 | -13.6 | [8 empty] | 187 | 36659.5 | 0.55939 | 0.5616
| 0.0075439 | -13.5 | [8 empty] | 188 | 36663.5781 | 0.55945 | 0.56123
| 0.0066138 | -13.4 | [8 empty] | 189 | 36627.7109 | 0.5589 | 0.56451
| 0.0056672 | -13.3 | [8 empty] | 190 | 36667.4492 | 0.55951 | 0.56088
| 0.0047034 | -13.2 | [8 empty] | 191 | 36715.0273 | 0.56024 | 0.55654
| 0.0037221 | -13.1 | [8 empty] | 192 | 36809.0273 | 0.56167 | 0.54802
| 0.0027228 | -13 | [8 empty] | 193 | 36714.7109 | 0.56023 | 0.55657
| 0.0017049 | -12.9 | [8 empty] | 194 | 36690.1328 | 0.55986 | 0.55881
| 0.00066804 | -12.8 | [8 empty] | 195 | 36647.7891 | 0.55921 | 0.56267
| -0.00038845 | -12.7 | [8 empty] | 196 | 36689.1562 | 0.55984 | 0.5589
| -0.0014651 | -12.6 | [8 empty] | 197 | 36717.0273 | 0.56027 | 0.55636
| -0.0025625 | -12.5 | [8 empty] | 198 | 36627.4492 | 0.5589 | 0.56454
| -0.0036812 | -12.4 | [8 empty] | 199 | 36647.8164 | 0.55921 | 0.56267
| -0.0048218 | -12.3 | [8 empty] | 200 | 36658.1836 | 0.55937 | 0.56173
| -0.0059851 | -12.2 | [8 empty] | 201 | 36630.3672 | 0.55894 | 0.56427
| -0.0071717 | -12.1 | [8 empty] | 202 | 36636.6836 | 0.55904 | 0.56369
| -0.0083823 | -12 | [8 empty] | 203 | 36650.6055 | 0.55925 | 0.56242
| -0.0096176 | -11.9 | [8 empty] | 204 | 36638.4492 | 0.55907 | 0.56353
| -0.010878 | -11.8 | [8 empty] | 205 | 36636.8672 | 0.55904 | 0.56367
| -0.012166 | -11.7 | [8 empty] | 206 | 36706.2109 | 0.5601 | 0.55735
| -0.01348 | -11.6 | [8 empty] | 207 | 36657.5508 | 0.55936 | 0.56178
| -0.014822 | -11.5 | [8 empty] | 208 | 36666 | 0.55949 | 0.56101
| -0.016193 | -11.4 | [8 empty] | 209 | 36733.7617 | 0.56052 | 0.55484
| -0.017594 | -11.3 | [8 empty] | 210 | 36711.5273 | 0.56018 | 0.55686
| -0.019026 | -11.2 | [8 empty] | 211 | 36684.1055 | 0.55976 | 0.55936
| -0.02049 | -11.1 | [8 empty] | 212 | 36703.1055 | 0.56005 | 0.55763
| -0.021987 | -11 | [8 empty] | 213 | 36617.1562 | 0.55874 | 0.56548
| -0.023517 | -10.9 | [8 empty] | 214 | 36654.0781 | 0.55931 | 0.5621
| -0.025083 | -10.8 | [8 empty] | 215 | 36597.7109 | 0.55845 | 0.56726
| -0.026686 | -10.7 | [8 empty] | 216 | 36569.1328 | 0.55801 | 0.56988
| -0.028326 | -10.6 | [8 empty] | 217 | 36604.1055 | 0.55854 | 0.56667
| -0.030005 | -10.5 | [8 empty] | 218 | 36600.2383 | 0.55848 | 0.56703
| -0.031725 | -10.4 | [8 empty] | 219 | 36642.1836 | 0.55912 | 0.56319
| -0.033487 | -10.3 | [8 empty] | 220 | 36551.6836 | 0.55774 | 0.57149
| -0.035293 | -10.2 | [8 empty] | 221 | 36566.9492 | 0.55798 | 0.57008
| -0.037143 | -10.1 | [8 empty] | 222 | 36582.5 | 0.55821 | 0.56866
| -0.039041 | -10 | [8 empty] | 223 | 36557.8438 | 0.55784 | 0.57092
| -0.040987 | -9.9 | [8 empty] | 224 | 36627.2109 | 0.5589 | 0.56456
| -0.042984 | -9.8 | [8 empty] | 225 | 36590.8164 | 0.55834 | 0.56789
| -0.045034 | -9.7 | [8 empty] | 226 | 36648.2891 | 0.55922 | 0.56263
| -0.047138 | -9.6 | [8 empty] | 227 | 36632.4727 | 0.55898 | 0.56408
| -0.0493 | -9.5 | [8 empty] | 228 | 36555.0508 | 0.55779 | 0.57118
| -0.05152 | -9.4 | [8 empty] | 229 | 36630 | 0.55894 | 0.5643
| -0.053803 | -9.3 | [8 empty] | 230 | 36593.2891 | 0.55838 | 0.56767
| -0.05615 | -9.2 | [8 empty] | 231 | 36640.3164 | 0.5591 | 0.56336
| -0.058564 | -9.1 | [8 empty] | 232 | 36607.3438 | 0.55859 | 0.56638
| -0.061048 | -9 | [8 empty] | 233 | 36529.1836 | 0.5574 | 0.57356
| -0.063605 | -8.9 | [8 empty] | 234 | 36563.7109 | 0.55793 | 0.57038
| -0.066239 | -8.8 | [8 empty] | 235 | 36576.5 | 0.55812 | 0.56921
| -0.068953 | -8.7 | [8 empty] | 236 | 36557.2383 | 0.55783 | 0.57098
| -0.071751 | -8.6 | [8 empty] | 237 | 36540.9492 | 0.55758 | 0.57248
| -0.074636 | -8.5 | [8 empty] | 238 | 36554.8438 | 0.55779 | 0.5712
| -0.077613 | -8.4 | [8 empty] | 239 | 36590.6562 | 0.55834 | 0.56791
| -0.080686 | -8.3 | [8 empty] | 240 | 36581.2891 | 0.55819 | 0.56877
| -0.08386 | -8.2 | [8 empty] | 241 | 36585.1836 | 0.55825 | 0.56841
| -0.08714 | -8.1 | [8 empty] | 242 | 36622.6328 | 0.55883 | 0.56498
| -0.090532 | -8 | [8 empty] | 243 | 36641.7617 | 0.55912 | 0.56323
| -0.09404 | -7.9 | [8 empty] | 244 | 36531.6055 | 0.55744 | 0.57334
| -0.097673 | -7.8 | [8 empty] | 245 | 36557.3672 | 0.55783 | 0.57097
| -0.10143 | -7.7 | [8 empty] | 246 | 36523.9219 | 0.55732 | 0.57405
| -0.10533 | -7.6 | [8 empty] | 247 | 36502.5273 | 0.55699 | 0.57602
| -0.10938 | -7.5 | [8 empty] | 248 | 36561.2109 | 0.55789 | 0.57061
| -0.11358 | -7.4 | [8 empty] | 249 | 36570.6055 | 0.55803 | 0.56975
| -0.11793 | -7.3 | [8 empty] | 250 | 36549.8672 | 0.55772 | 0.57166
| -0.12247 | -7.2 | [8 empty] | 251 | 36495.9219 | 0.55689 | 0.57663
| -0.12718 | -7.1 | [8 empty] | 252 | 36493.6328 | 0.55686 | 0.57684
| -0.13208 | -7 | [8 empty] | 253 | 36553.2617 | 0.55777 | 0.57134
| -0.13719 | -6.9 | [8 empty] | 254 | 36578.8164 | 0.55816 | 0.56899
| -0.14252 | -6.8 | [8 empty] | 255 | 36605.1562 | 0.55856 | 0.56658
| -0.14808 | -6.7 | [8 empty] | 256 | 36647.3945 | 0.5592 | 0.56271
| -0.15389 | -6.6 | [8 empty] | 257 | 36687.6836 | 0.55982 | 0.55903
| -0.15996 | -6.5 | [8 empty] | 258 | 36715.7109 | 0.56025 | 0.55648
| -0.16632 | -6.4 | [8 empty] | 259 | 36748.2383 | 0.56074 | 0.55352
| -0.17298 | -6.3 | [8 empty] | 260 | 36824.4492 | 0.56191 | 0.54662
| -0.17996 | -6.2 | [8 empty] | 261 | 36925.0508 | 0.56344 | 0.53757
| -0.1873 | -6.1 | [8 empty] | 262 | 36982.7891 | 0.56432 | 0.5324
| -0.19501 | -6 | [8 empty] | 263 | 36996.0273 | 0.56452 | 0.53122
| -0.20313 | -5.9 | [8 empty] | 264 | 37065.0273 | 0.56558 | 0.52507
| -0.2117 | -5.8 | [8 empty] | 265 | 37149.1836 | 0.56686 | 0.51762
| -0.22073 | -5.7 | [8 empty] | 266 | 37270.4219 | 0.56871 | 0.50695
| -0.23029 | -5.6 | [8 empty] | 267 | 37391.3438 | 0.57056 | 0.49639
| -0.24041 | -5.5 | [8 empty] | 268 | 37622.7109 | 0.57409 | 0.47643
| -0.25115 | -5.4 | [8 empty] | 269 | 37820.3164 | 0.5771 | 0.45962
| -0.26256 | -5.3 | [8 empty] | 270 | 38043.5 | 0.58051 | 0.4409
| -0.27472 | -5.2 | [8 empty] | 271 | 38358.3438 | 0.58531 | 0.41495
| -0.28768 | -5.1 | [8 empty] | 272 | 38858.8945 | 0.59295 | 0.37476
| -0.30154 | -5 | [8 empty] | 273 | 39695.9219 | 0.60572 | 0.31033
| -0.3164 | -4.9 | [8 empty] | 274 | 40694.5781 | 0.62096 | 0.23769
| -0.33236 | -4.8 | [8 empty] | 275 | 42166.6562 | 0.64342 | 0.13822
| -0.34956 | -4.7 | [8 empty] | 276 | 43644.1055 | 0.66597 | 0.046513
| -0.36814 | -4.6 | [8 empty] | 277 | 45427.7891 | 0.69318 | -0.054719
| -0.38828 | -4.5 | [8 empty] | 278 | 47112.7891 | 0.7189 | -0.142
| -0.41018 | -4.4 | [8 empty] | 279 | 48211.3164 | 0.73566 | -0.19503
| -0.43409 | -4.3 | [8 empty] | 280 | 49127.2891 | 0.74963 | -0.23711
| -0.46028 | -4.2 | [8 empty] | 281 | 49706.2109 | 0.75847 | -0.26277
| -0.48911 | -4.1 | [8 empty] | 282 | 50057.8438 | 0.76383 | -0.27801
| -0.521 | -4 | [8 empty] | 283 | 50499.0273 | 0.77057 | -0.29678
| -0.55645 | -3.9 | [8 empty] | 284 | 50799.6055 | 0.77515 | -0.30936
| -0.59611 | -3.8 | [8 empty] | 285 | 51143.7891 | 0.7804 | -0.32354
| -0.64077 | -3.7 | [8 empty] | 286 | 51427.9492 | 0.78474 | -0.33508
| -0.69143 | -3.6 | [8 empty] | 287 | 51622.1836 | 0.7877 | -0.34289
| -0.74941 | -3.5 | [8 empty] | 288 | 51854.6836 | 0.79125 | -0.35214
| -0.81639 | -3.4 | [8 empty] | 289 | 51963.5273 | 0.79291 | -0.35644
| -0.89466 | -3.3 | [8 empty] | 290 | 52127.4492 | 0.79541 | -0.36288
| -0.98733 | -3.2 | [8 empty] | 291 | 52209.0508 | 0.79666 | -0.36606
| -1.0988 | -3.1 | [8 empty] | 292 | 52274.1328 | 0.79765 | -0.3686
| -1.2354 | -3 | [8 empty] | 293 | 52324.3164 | 0.79842 | -0.37054
| -1.4067 | -2.9 | [8 empty] | 294 | 52404.5508 | 0.79964 | -0.37365
| -1.6279 | -2.8 | [8 empty] | 295 | 52514.4727 | 0.80132 | -0.37788
| -1.9245 | -2.7 | [8 empty] | 296 | 52593.1562 | 0.80252 | -0.3809
| -2.343 | -2.6 | [8 empty] | 297 | 52610.5508 | 0.80279 | -0.38157
| -2.9778 | -2.5 | [8 empty] | 298 | 52676.4219 | 0.80379 | -0.38409
| -4.0554 | -2.4 | [8 empty] | 299 | 52761.1328 | 0.80508 | -0.38731
| -6.2868 | -2.3 | [8 empty] | 300 | 52820.8438 | 0.80599 | -0.38958
| -13.6605 | -2.2 | [8 empty] | 301 | 52914.5273 | 0.80742 | -0.39313
| 90.5663 | -2.1 | [8 empty] | 302 | 53014.9727 | 0.80896 | -0.39691
| 10.6704 | -2 | [8 empty] | 303 | 53023.8945 | 0.80909 | -0.39725
| 5.7186 | -1.9 | [8 empty] | 304 | 53069.8438 | 0.80979 | -0.39897
| 3.9288 | -1.8 | [8 empty] | 305 | 53126.0781 | 0.81065 | -0.40108
| 3.0052 | -1.7 | [8 empty] | 306 | 53248.6328 | 0.81252 | -0.40566
| 2.4416 | -1.6 | [8 empty] | 307 | 53291.5508 | 0.81318 | -0.40725
| 2.0618 | -1.5 | [8 empty] | 308 | 53369.3164 | 0.81436 | -0.41014
| 1.7885 | -1.4 | [8 empty] | 309 | 53309.0273 | 0.81344 | -0.4079
| 1.5824 | -1.3 | [8 empty] | 310 | 53322.7109 | 0.81365 | -0.40841
| 1.4214 | -1.2 | [8 empty] | 311 | 53490.5 | 0.81621 | -0.41461
| 1.2922 | -1.1 | [8 empty] | 312 | 53508.3945 | 0.81649 | -0.41527
| 1.1863 | -1 | [8 empty] | 313 | 53512.5 | 0.81655 | -0.41542
| 1.0978 | -0.9 | [8 empty] | 314 | 53507.2109 | 0.81647 | -0.41523
| 1.0227 | -0.8 | [8 empty] | 315 | 53619.0508 | 0.81817 | -0.41933
| 0.95832 | -0.7 | [8 empty] | 316 | 53637.1328 | 0.81845 | -0.41999
| 0.90242 | -0.6 | [8 empty] | 317 | 53532.2383 | 0.81685 | -0.41615
| 0.85346 | -0.5 | [8 empty] | 318 | 53670.8438 | 0.81896 | -0.42123
| 0.81021 | -0.4 | [8 empty] | 319 | 53642.2109 | 0.81853 | -0.42018
| 0.77173 | -0.3 | [8 empty] | 320 | 53598.9492 | 0.81787 | -0.4186
| 0.73728 | -0.2 | [8 empty] | 321 | 53682.3438 | 0.81914 | -0.42165
| 0.70624 | -0.1 | [8 empty] | 322 | 53741.5508 | 0.82004 | -0.42381
| 0.67815 | 4.0146e-13 | [8 empty] | 323 | 53726.1836 | 0.81981 | -0.42325
| 0.65259 | 0.1 | [8 empty] | 324 | 53760.2617 | 0.82033 | -0.42449
| 0.62924 | 0.2 | [8 empty] | 325 | 53780.1562 | 0.82063 | -0.42521
| 0.60783 | 0.3 | [8 empty] | 326 | 53720.2383 | 0.81972 | -0.42303
| 0.58812 | 0.4 | [8 empty] | 327 | 53827.1562 | 0.82135 | -0.42692
| 0.56992 | 0.5 | [8 empty] | 328 | 53813.7109 | 0.82114 | -0.42643
| 0.55306 | 0.6 | [8 empty] | 329 | 53745.1562 | 0.8201 | -0.42394
| 0.5374 | 0.7 | [8 empty] | 330 | 53771.2383 | 0.8205 | -0.42489
| 0.52282 | 0.8 | [8 empty] | 331 | 53808.8438 | 0.82107 | -0.42626
| 0.5092 | 0.9 | [8 empty] | 332 | 53844.3164 | 0.82161 | -0.42754
| 0.49646 | 1 | [8 empty] | 333 | 53839.5 | 0.82154 | -0.42737
| 0.48451 | 1.1 | [8 empty] | 334 | 53859.5273 | 0.82184 | -0.4281
| 0.47328 | 1.2 | [8 empty] | 335 | 53884.5508 | 0.82223 | -0.429
| 0.46271 | 1.3 | [8 empty] | 336 | 53935.2617 | 0.823 | -0.43084
| 0.45275 | 1.4 | [8 empty] | 337 | 53910.1562 | 0.82262 | -0.42993
| 0.44333 | 1.5 | [8 empty] | 338 | 53973.2617 | 0.82358 | -0.43221
| 0.43442 | 1.6 | [8 empty] | 339 | 53873.1055 | 0.82205 | -0.42859
| 0.42598 | 1.7 | [8 empty] | 340 | 53893.5508 | 0.82236 | -0.42933
| 0.41797 | 1.8 | [8 empty] | 341 | 53928.8164 | 0.8229 | -0.4306
| 0.41036 | 1.9 | [8 empty] | 342 | 53970.1836 | 0.82353 | -0.4321
| 0.40312 | 2 | [8 empty] | 343 | 53958.3672 | 0.82335 | -0.43167
| 0.39622 | 2.1 | [8 empty] | 344 | 53999.9727 | 0.82399 | -0.43317
| 0.38964 | 2.2 | [8 empty] | 345 | 53986.5273 | 0.82378 | -0.43269
| 0.38336 | 2.3 | [8 empty] | 346 | 53980.2383 | 0.82369 | -0.43246
| 0.37736 | 2.4 | [8 empty] | 347 | 53996.5273 | 0.82393 | -0.43305
| 0.37162 | 2.5 | [8 empty] | 348 | 54041.6328 | 0.82462 | -0.43467
| 0.36612 | 2.6 | [8 empty] | 349 | 54031.3945 | 0.82447 | -0.4343
| 0.36085 | 2.7 | [8 empty] | 350 | 53991.5781 | 0.82386 | -0.43287
| 0.3558 | 2.8 | [8 empty] | 351 | 54004.9492 | 0.82406 | -0.43335
| 0.35095 | 2.9 | [8 empty] | 352 | 54048.7617 | 0.82473 | -0.43493
| 0.34628 | 3 | [8 empty] | 353 | 54053.0508 | 0.8248 | -0.43508
| 0.3418 | 3.1 | [8 empty] | 354 | 54023.8164 | 0.82435 | -0.43403
| 0.33748 | 3.2 | [8 empty] | 355 | 54046.2383 | 0.82469 | -0.43484
| 0.33333 | 3.3 | [8 empty] | 356 | 54037.8672 | 0.82456 | -0.43454
| 0.32932 | 3.4 | [8 empty] | 357 | 54054.5 | 0.82482 | -0.43513
| 0.32546 | 3.5 | [8 empty] | 358 | 54111.2617 | 0.82568 | -0.43717
| 0.32173 | 3.6 | [8 empty] | 359 | 54061.5273 | 0.82493 | -0.43539
| 0.31813 | 3.7 | [8 empty] | 360 | 54115.5 | 0.82575 | -0.43732
| 0.31466 | 3.8 | [8 empty] | 361 | 54159.2383 | 0.82642 | -0.43889
| 0.3113 | 3.9 | [8 empty] | 362 | 54041.9727 | 0.82463 | -0.43468
| 0.30804 | 4 | [8 empty] | 363 | 54041.6836 | 0.82462 | -0.43467
| 0.3049 | 4.1 | [8 empty] | 364 | 54051.0781 | 0.82477 | -0.43501
| 0.30185 | 4.2 | [8 empty] | 365 | 54116.5 | 0.82576 | -0.43736
| 0.2989 | 4.3 | [8 empty] | 366 | 54069.3945 | 0.82505 | -0.43567
| 0.29604 | 4.4 | [8 empty] | 367 | 54071.6562 | 0.82508 | -0.43575
| 0.29326 | 4.5 | [8 empty] | 368 | 54055.2891 | 0.82483 | -0.43516
| 0.29057 | 4.6 | [8 empty] | 369 | 54016.1836 | 0.82423 | -0.43376
| 0.28796 | 4.7 | [8 empty] | 370 | 54085.3164 | 0.82529 | -0.43624
| 0.28542 | 4.8 | [8 empty] | 371 | 54081.7383 | 0.82523 | -0.43611
| 0.28295 | 4.9 | [8 empty] | 372 | 54079.3438 | 0.8252 | -0.43603
| 0.28056 | 5 | [8 empty] | 373 | 54120.6055 | 0.82583 | -0.43751
| 0.27823 | 5.1 | [8 empty] | 374 | 54054.5273 | 0.82482 | -0.43513
| 0.27596 | 5.2 | [8 empty] | 375 | 53992.6055 | 0.82387 | -0.43291
| 0.27376 | 5.3 | [8 empty] | 376 | 54037.6328 | 0.82456 | -0.43453
| 0.27161 | 5.4 | [8 empty] | 377 | 54032.1055 | 0.82448 | -0.43433
| 0.26952 | 5.5 | [8 empty] | 378 | 54046.2617 | 0.82469 | -0.43484
| 0.26748 | 5.6 | [8 empty] | 379 | 53990.5781 | 0.82384 | -0.43283
| 0.2655 | 5.7 | [8 empty] | 380 | 54003.1836 | 0.82404 | -0.43329
| 0.26357 | 5.8 | [8 empty] | 381 | 54021.1055 | 0.82431 | -0.43393
| 0.26168 | 5.9 | [8 empty] | 382 | 54010.0273 | 0.82414 | -0.43353
| 0.25984 | 6 | [8 empty] | 383 | 54047.4727 | 0.82471 | -0.43488
| 0.25805 | 6.1 | [8 empty] | 384 | 54056.5781 | 0.82485 | -0.43521
| 0.2563 | 6.2 | [8 empty] | 385 | 54087.3672 | 0.82532 | -0.43631
| 0.25459 | 6.3 | [8 empty] | 386 | 54060.6328 | 0.82491 | -0.43535
| 0.25292 | 6.4 | [8 empty] | 387 | 54073.2109 | 0.8251 | -0.43581
| 0.25129 | 6.5 | [8 empty] | 388 | 54083.6562 | 0.82526 | -0.43618
| 0.2497 | 6.6 | [8 empty] | 389 | 54126.6328 | 0.82592 | -0.43772
| 0.24814 | 6.7 | [8 empty] | 390 | 54091.2617 | 0.82538 | -0.43645
| 0.24662 | 6.8 | [8 empty] | 391 | 54070.3672 | 0.82506 | -0.4357
| 0.24513 | 6.9 | [8 empty] | 392 | 54144.3438 | 0.82619 | -0.43836
| 0.24368 | 7 | [8 empty] | 393 | 54130 | 0.82597 | -0.43784
| 0.24225 | 7.1 | [8 empty] | 394 | 54125.1055 | 0.8259 | -0.43767
| 0.24086 | 7.2 | [8 empty] | 395 | 54061.8164 | 0.82493 | -0.4354
| 0.2395 | 7.3 | [8 empty] | 396 | 54063.9219 | 0.82496 | -0.43547
| 0.23816 | 7.4 | [8 empty] | 397 | 54140.2383 | 0.82613 | -0.43821
| 0.23686 | 7.5 | [8 empty] | 398 | 54094.0273 | 0.82542 | -0.43655
| 0.23558 | 7.6 | [8 empty] | 399 | 54181.5781 | 0.82676 | -0.43969
| 0.23432 | 7.7 | [8 empty] | 400 | 54121.6328 | 0.82584 | -0.43754
| 0.23309 | 7.8 | [8 empty] | 401 | 54137.6328 | 0.82609 | -0.43812
| 0.23189 | 7.9 | [8 empty] | 402 | 54116.3945 | 0.82576 | -0.43736
| 0.23071 | 8 | [8 empty] | 403 | 54134.6055 | 0.82604 | -0.43801
| 0.22955 | 8.1 | [8 empty] | 404 | 54109.4492 | 0.82566 | -0.43711
| 0.22842 | 8.2 | [8 empty] | 405 | 54012.2617 | 0.82417 | -0.43361
| 0.2273 | 8.3 | [8 empty] | 406 | 54148.1328 | 0.82625 | -0.43849
| 0.22621 | 8.4 | [8 empty] | 407 | 54134 | 0.82603 | -0.43799
| 0.22514 | 8.5 | [8 empty] | 408 | 53913.9727 | 0.82267 | -0.43007
| 0.22409 | 8.6 | [8 empty] | 409 | 54083.6055 | 0.82526 | -0.43618
| 0.22306 | 8.7 | [8 empty] | 410 | 54125.6055 | 0.8259 | -0.43769
| 0.22204 | 8.8 | [8 empty] | 411 | 53990.1562 | 0.82384 | -0.43282
| 0.22105 | 8.9 | [8 empty] | 412 | 53989.7109 | 0.82383 | -0.4328
| 0.22007 | 9 | [8 empty] | 413 | 53986.6328 | 0.82378 | -0.43269
| 0.21911 | 9.1 | [8 empty] | 414 | 54103.9219 | 0.82557 | -0.43691
| 0.21817 | 9.2 | [8 empty] | 415 | 54184.1328 | 0.8268 | -0.43978
| 0.21725 | 9.3 | [8 empty] | 416 | 54061.2891 | 0.82492 | -0.43538
| 0.21634 | 9.4
| 0.21544 | 9.5
| 0.21456 | 9.6
| 0.2137 | 9.7
| 0.21285 | 9.8
| 0.21202 | 9.9
| 0.21119 | 10
| 0.21039 | 10.1
| 0.20959 | 10.2
| 0.20881 | 10.3
| 0.20804 | 10.4
| 0.20728 | 10.5
| 0.20654 | 10.6
| 0.2058 | 10.7
| 0.20508 | 10.8
| 0.20437 | 10.9
| 0.20367 | 11
| 0.20298 | 11.1
| 0.2023 | 11.2
| 0.20163 | 11.3
| 0.20097 | 11.4
| 0.20032 | 11.5
| 0.19968 | 11.6
| 0.19905 | 11.7
| 0.19843 | 11.8
| 0.19781 | 11.9
| 0.19721 | 12
| 0.19661 | 12.1
| 0.19603 | 12.2
| 0.19545 | 12.3
| 0.19488 | 12.4
| 0.19431 | 12.5
| 0.19376 | 12.6
| 0.19321 | 12.7
| 0.19267 | 12.8
| 0.19213 | 12.9
| 0.19161 | 13
| 0.19109 | 13.1
| 0.19057 | 13.2
| 0.19007 | 13.3
| 0.18957 | 13.4
| 0.18907 | 13.5
| 0.18859 | 13.6
| 0.18811 | 13.7
| 0.18763 | 13.8
| 0.18716 | 13.9
| 0.1867 | 14
| 0.18624 | 14.1
| 0.18579 | 14.2
| 0.18534 | 14.3
| 0.1849 | 14.4
| 0.18447 | 14.5
| 0.18403 | 14.6
| 0.18361 | 14.7
| 0.18319 | 14.8
| 0.18277 | 14.9
| 0.18236 | 15
| 0.18196 | 15.1
| 0.18156 | 15.2
| 0.18116 | 15.3
| 0.18077 | 15.4
| 0.18038 | 15.5
| 0.18 | 15.6
| 0.17962 | 15.7
| 0.17924 | 15.8
| 0.17887 | 15.9
| 0.17851 | 16
| 0.17814 | 16.1
| 0.17779 | 16.2
| 0.17743 | 16.3
| 0.17708 | 16.4
| 0.17673 | 16.5
| 0.17639 | 16.6
| 0.17605 | 16.7
| 0.17572 | 16.8
| 0.17538 | 16.9
| 0.17505 | 17
| 0.17473 | 17.1
| 0.17441 | 17.2
| 0.17409 | 17.3
| 0.17377 | 17.4
| 0.17346 | 17.5
| 0.17315 | 17.6
| 0.17284 | 17.7
| 0.17254 | 17.8
| 0.17224 | 17.9
| 0.17195 | 18
| 0.17165 | 18.1
| 0.17136 | 18.2
| 0.17107 | 18.3
| 0.17079 | 18.4
| 0.1705 | 18.5
| 0.17022 | 18.6
| 0.16995 | 18.7
| 0.16967 | 18.8
| 0.1694 | 18.9
| 0.16913 | 19
| 0.16886 | 19.1
| 0.1686 | 19.2
| 0.16834 | 19.3
| 0.16808 | 19.4
| 0.16782 | 19.5
| 0.16757 | 19.6
| 0.16731 | 19.7
| 0.16706 | 19.8
| 0.16682 | 19.9
| 0.16657 | 20
| 0.16633 | 20.1
| 0.16609 | 20.2
| 0.16585 | 20.3
| 0.16561 | 20.4
| 0.16538 | 20.5
| 0.16514 | 20.6
| 0.16491 | 20.7
| 0.16468 | 20.8
| 0.16446 | 20.9
| 0.16423 | 21
| 0.16401 | 21.1
| 0.16379 | 21.2
| 0.16357 | 21.3
| 0.16335 | 21.4
| 0.16314 | 21.5
| 0.16292 | 21.6
| 0.16271 | 21.7
| 0.1625 | 21.8
| 0.16229 | 21.9
| 0.16209 | 22
| 0.16188 | 22.1
| 0.16168 | 22.2
| 0.16148 | 22.3
| 0.16128 | 22.4
| 0.16108 | 22.5
| 0.16088 | 22.6
| 0.16069 | 22.7
| 0.1605 | 22.8
| 0.1603 | 22.9
| 0.16011 | 23
| 0.15992 | 23.1
| 0.15974 | 23.2
| 0.15955 | 23.3
| 0.15937 | 23.4
| 0.15918 | 23.5
| 0.159 | 23.6
| 0.15882 | 23.7
| 0.15864 | 23.8
| 0.15847 | 23.9
| 0.15829 | 24
| 0.15812 | 24.1
| 0.15794 | 24.2
| 0.15777 | 24.3
| 0.1576 | 24.4
| 0.15743 | 24.5
| 0.15726 | 24.6
| 0.1571 | 24.7
| 0.15693 | 24.8
| 0.15677 | 24.9
| 0.1566 | 25
| 0.15644 | 25.1
| 0.15628 | 25.2
| 0.15612 | 25.3
| 0.15596 | 25.4
| 0.1558 | 25.5
| 0.15565 | 25.6
| 0.15549 | 25.7
| 0.15534 | 25.8
| 0.15519 | 25.9
| 0.15503 | 26
| 0.15488 | 26.1
| 0.15473 | 26.2
| 0.15459 | 26.3
| 0.15444 | 26.4
| 0.15429 | 26.5
| 0.15415 | 26.6
| 0.154 | 26.7
| 0.15386 | 26.8
| 0.15372 | 26.9
| 0.15357 | 27
| 0.15343 | 27.1
| 0.15329 | 27.2
| 0.15315 | 27.3
| 0.15302 | 27.4
| 0.15288 | 27.5
| 0.15274 | 27.6
| 0.15261 | 27.7
| 0.15248 | 27.8
| 0.15234 | 27.9
| 0.15221 | 28
| 0.15208 | 28.1
| 0.15195 | 28.2
| 0.15182 | 28.3
| 0.15169 | 28.4
| 0.15156 | 28.5
| 0.15143 | 28.6
| 0.15131 | 28.7
| 0.15118 | 28.8
| 0.15106 | 28.9
| 0.15093 | 29
| 0.15081 | 29.1
| 0.15069 | 29.2
| 0.15057 | 29.3
| 0.15045 | 29.4
| 0.15033 | 29.5
| 0.15021 | 29.6
| 0.15009 | 29.7
| 0.14997 | 29.8
| 0.14985 | 29.9
| 0.14974 | 30
| 0.14962 | 30.1
| 0.14951 | 30.2
| 0.14939 | 30.3
| 0.14928 | 30.4
| 0.14917 | 30.5
| 0.14906 | 30.6
| 0.14894 | 30.7
| 0.14883 | 30.8
| 0.14872 | 30.9
| 0.14861 | 31
| 0.14851 | 31.1
| 0.1484 | 31.2
| 0.14829 | 31.3
| 0.14818 | 31.4
| 0.14808 | 31.5
| 0.14797 | 31.6
| 0.14787 | 31.7
| 0.14776 | 31.8
| 0.14766 | 31.9
| 0.14756 | 32
| 0.14745 | 32.1
| 0.14735 | 32.2
| 0.14725 | 32.3
| 0.14715 | 32.4
| 0.14705 | 32.5
| 0.14695 | 32.6
| 0.14685 | 32.7
| 0.14675 | 32.8
| 0.14666 | 32.9
| 0.14656 | 33
| 0.14646 | 33.1
| 0.14637 | 33.2
| 0.14627 | 33.3
| 0.14617 | 33.4
| 0.14608 | 33.5
| 0.14599 | 33.6
| 0.14589 | 33.7
| 0.1458 | 33.8
| 0.14571 | 33.9
| 0.14562 | 34
| 0.14552 | 34.1
| 0.14543 | 34.2
| 0.14534 | 34.3
| 0.14525 | 34.4
| 0.14516 | 34.5
| 0.14507 | 34.6
| 0.14499 | 34.7
| 0.1449 | 34.8
| 0.14481 | 34.9
| 0.14472 | 35
| 0.14464 | 35.1
| 0.14455 | 35.2
| 0.14447 | 35.3
| 0.14438 | 35.4
| 0.1443 | 35.5
| 0.14421 | 35.6
| 0.14413 | 35.7
| 0.14404 | 35.8
| 0.14396 | 35.9
| 0.14388 | 36
| 0.1438 | 36.1
| 0.14371 | 36.2
| 0.14363 | 36.3
| 0.14355 | 36.4
| 0.14347 | 36.5
| 0.14339 | 36.6
| 0.14331 | 36.7
| 0.14323 | 36.8
| 0.14315 | 36.9
| 0.14308 | 37
| 0.143 | 37.1
| 0.14292 | 37.2
| 0.14284 | 37.3
| 0.14277 | 37.4
| 0.14269 | 37.5
| 0.14261 | 37.6
| 0.14254 | 37.7
| 0.14246 | 37.8
| 0.14239 | 37.9
| 0.14231 | 38
| 0.14224 | 38.1
| 0.14217 | 38.2
| 0.14209 | 38.3
| 0.14202 | 38.4
| 0.14195 | 38.5
| 0.14188 | 38.6
| 0.1418 | 38.7
| 0.14173 | 38.8
| 0.14166 | 38.9
| 0.14159 | 39
| 0.14152 | 39.1
| 0.14145 | 39.2
| 0.14138 | 39.3
| 0.14131 | 39.4
| 0.14124 | 39.5
| 0.14117 | 39.6
| 0.1411 | 39.7
| 0.14103 | 39.8
| 0.14097 | 39.9
| 0.1409 | 40
| 0.14083 | 40.1
| 0.14077 | 40.2
| 0.1407 | 40.3
| 0.14063 | 40.4
| 0.14057 | 40.5
| 0.1405 | 40.6
| 0.14044 | 40.7
| 0.14037 | 40.8
| 0.14031 | 40.9
| 0.14024 | 41
| 0.14018 | 41.1
| 0.14011 | 41.2
| 0.14005 | 41.3
| 0.13999 | 41.4
| 0.13992 | 41.5
| 0.13986 | 41.6
| 0.1398 | 41.7
| 0.13974 | 41.8
| 0.13967 | 41.9
| 0.13961 | 42
| 0.13955 | 42.1
| 0.13949 | 42.2
| 0.13943 | 42.3
| 0.13937 | 42.4
| 0.13931 | 42.5
| 0.13925 | 42.6
| 0.13919 | 42.7
| 0.13913 | 42.8
| 0.13907 | 42.9
| 0.13901 | 43
| 0.13895 | 43.1
| 0.1389 | 43.2
| 0.13884 | 43.3
| 0.13878 | 43.4
| 0.13872 | 43.5
| 0.13866 | 43.6
| 0.13861 | 43.7
| 0.13855 | 43.8
| 0.13849 | 43.9
| 0.13844 | 44
| 0.13838 | 44.1
| 0.13833 | 44.2
| 0.13827 | 44.3
| 0.13821 | 44.4
| 0.13816 | 44.5
| 0.1381 | 44.6
| 0.13805 | 44.7
| 0.138 | 44.8
| 0.13794 | 44.9
| 0.13789 | 45
| 0.13783 | 45.1
| 0.13778 | 45.2
| 0.13773 | 45.3
| 0.13767 | 45.4
| 0.13762 | 45.5
| 0.13757 | 45.6
| 0.13752 | 45.7
| 0.13746 | 45.8
| 0.13741 | 45.9
| 0.13736 | 46
| 0.13731 | 46.1
| 0.13726 | 46.2
| 0.13721 | 46.3
| 0.13716 | 46.4
| 0.1371 | 46.5
| 0.13705 | 46.6
| 0.137 | 46.7
| 0.13695 | 46.8
| 0.1369 | 46.9
| 0.13685 | 47
| 0.1368 | 47.1
| 0.13676 | 47.2
| 0.13671 | 47.3
| 0.13666 | 47.4
| 0.13661 | 47.5
| 0.13656 | 47.6
| 0.13651 | 47.7
| 0.13646 | 47.8
| 0.13642 | 47.9
| 0.13637 | 48
| 0.13632 | 48.1
| 0.13627 | 48.2
| 0.13623 | 48.3
| 0.13618 | 48.4
| 0.13613 | 48.5
| 0.13609 | 48.6
| 0.13604 | 48.7
| 0.13599 | 48.8
| 0.13595 | 48.9
| 0.1359 | 49
| 0.13586 | 49.1
| 0.13581 | 49.2
| 0.13577 | 49.3
| 0.13572 | 49.4
| 0.13568 | 49.5
| 0.13563 | 49.6
| 0.13559 | 49.7
| 0.13554 | 49.8
| 0.1355 | 49.9
| 0.13545 | 50
| 0.13541 | 50.1
| 0.13537 | 50.2
| 0.13532 | 50.3
| 0.13528 | 50.4
| 0.13524 | 50.5
| 0.13519 | 50.6
| 0.13515 | 50.7
| 0.13511 | 50.8
| 0.13506 | 50.9
| 0.13502 | 51
| 0.13498 | 51.1
| 0.13494 | 51.2
| 0.13489 | 51.3
| 0.13485 | 51.4
| 0.13481 | 51.5
| 0.13477 | 51.6
| 0.13473 | 51.7
| 0.13469 | 51.8
| 0.13465 | 51.9
| 0.13461 | 52
| 0.13456 | 52.1
| 0.13452 | 52.2
| 0.13448 | 52.3
| 0.13444 | 52.4
| 0.1344 | 52.5
| 0.13436 | 52.6
| 0.13432 | 52.7
| 0.13428 | 52.8
| 0.13424 | 52.9
| 0.1342 | 53
| 0.13417 | 53.1
| 0.13413 | 53.2
| 0.13409 | 53.3
| 0.13405 | 53.4
| 0.13401 | 53.5
| 0.13397 | 53.6
| 0.13393 | 53.7
| 0.13389 | 53.8
| 0.13386 | 53.9
| 0.13382 | 54
| 0.13378 | 54.1
| 0.13374 | 54.2
| 0.1337 | 54.3
| 0.13367 | 54.4
| 0.13363 | 54.5
| 0.13359 | 54.6
| 0.13356 | 54.7
| 0.13352 | 54.8
| 0.13348 | 54.9
| 0.13345 | 55
| 0.13341 | 55.1
| 0.13337 | 55.2
| 0.13334 | 55.3
| 0.1333 | 55.4
| 0.13326 | 55.5
| 0.13323 | 55.6
| 0.13319 | 55.7
| 0.13316 | 55.8
| 0.13312 | 55.9
| 0.13309 | 56
| 0.13305 | 56.1
| 0.13301 | 56.2
| 0.13298 | 56.3
| 0.13294 | 56.4
| 0.13291 | 56.5
| 0.13287 | 56.6
| 0.13284 | 56.7
| 0.13281 | 56.8
| 0.13277 | 56.9
| 0.13274 | 57
| 0.1327 | 57.1
| 0.13267 | 57.2
| 0.13263 | 57.3
| 0.1326 | 57.4
| 0.13257 | 57.5
| 0.13253 | 57.6
| 0.1325 | 57.7
| 0.13247 | 57.8
| 0.13243 | 57.9
| 0.1324 | 58
| 0.13237 | 58.1
| 0.13233 | 58.2
| 0.1323 | 58.3
| 0.13227 | 58.4
| 0.13224 | 58.5
| 0.1322 | 58.6
| 0.13217 | 58.7
| 0.13214 | 58.8
| 0.13211 | 58.9
| 0.13208 | 59
| 0.13204 | 59.1
| 0.13201 | 59.2
| 0.13198 | 59.3
| 0.13195 | 59.4
| 0.13192 | 59.5
| 0.13189 | 59.6
| 0.13185 | 59.7
| 0.13182 | 59.8
| 0.13179 | 59.9
| 0.13176 | 60

## rational cal 2
Calibration data, 16-bit #
| Dose, cGy | Red | Green | Blue
| 300 | 22704.438 | 26601.2067 | 19960.2125
| 200 | 26397.3885 | 30047.5799 | 21188.3132
| 100 | 32547.5085 | 35131.9428 | 22942.7761
| 0 | 44440.5987 | 41860.923 | 24778.5695
Calibration data, 16-bit #/65535
| Dose, Gy | Red | Green | Blue
| 3 | 0.34645 | 0.40591 | 0.30457
| 2 | 0.4028 | 0.4585 | 0.32331
| 1 | 0.49664 | 0.53608 | 0.35008 |  |  |  |  | k=1.22476
| 0 | 0.67812 | 0.63876 | 0.3781
Calibration curve
X(D) = A + B/(Dose -C)
where X(D) is response expressed as 16-bit #/65535, D is dose in Gy and A, B and C are constants
|  | Red | Green | Blue
| A | 0.11252 | 0.02256 | 0.063139 | [5 empty] | -12.7365
| B | 1.1953 | 3.0228 | 3.0654
| C | -2.1132 | -4.903 | -9.7235
| X(D) | Dose(Gy)
| 0.069654 | -30
| 0.0695 | -29.9
| 0.069344 | -29.8
| 0.069188 | -29.7
| 0.06903 | -29.6
| 0.068871 | -29.5
| 0.068711 | -29.4
| 0.06855 | -29.3
| 0.068388 | -29.2
| 0.068224 | -29.1
| 0.06806 | -29
| 0.067894 | -28.9
| 0.067727 | -28.8
| 0.067558 | -28.7
| 0.067388 | -28.6
| 0.067217 | -28.5
| 0.067045 | -28.4
| 0.066871 | -28.3 | 0.11
| 0.066696 | -28.2
| 0.06652 | -28.1
| 0.066342 | -28
| 0.066163 | -27.9
| 0.065983 | -27.8
| 0.065801 | -27.7
| 0.065618 | -27.6
| 0.065433 | -27.5
| 0.065247 | -27.4
| 0.065059 | -27.3
| 0.06487 | -27.2
| 0.064679 | -27.1
| 0.064487 | -27 |  | X(D) | Dose(Gy)
| 0.064293 | -26.9 |  | 0.67815 | 4.0146e-13
| 0.064098 | -26.8 |  | 0.65259 | 0.1
| 0.063901 | -26.7 |  | 0.62924 | 0.2
| 0.063702 | -26.6 |  | 0.60783 | 0.3
| 0.063502 | -26.5 |  | 0.58812 | 0.4
| 0.0633 | -26.4 |  | 0.56992 | 0.5
| 0.063097 | -26.3 |  | 0.55306 | 0.6
| 0.062892 | -26.2 |  | 0.5374 | 0.7
| 0.062685 | -26.1 |  | 0.52282 | 0.8
| 0.062476 | -26 |  | 0.5092 | 0.9
| 0.062266 | -25.9 |  | 0.49646 | 1
| 0.062054 | -25.8 |  | 0.48451 | 1.1
| 0.06184 | -25.7 |  | 0.47328 | 1.2
| 0.061624 | -25.6 |  | 0.46271 | 1.3
| 0.061406 | -25.5 |  | 0.45275 | 1.4
| 0.061187 | -25.4 |  | 0.44333 | 1.5
| 0.060966 | -25.3 |  | 0.43442 | 1.6
| 0.060742 | -25.2 |  | 0.42598 | 1.7
| 0.060517 | -25.1 |  | 0.41797 | 1.8
| 0.06029 | -25 |  | 0.41036 | 1.9
| 0.060061 | -24.9 |  | 0.40312 | 2
| 0.059829 | -24.8 |  | 0.39622 | 2.1
| 0.059596 | -24.7 |  | 0.38964 | 2.2
| 0.059361 | -24.6 |  | 0.38336 | 2.3
| 0.059123 | -24.5 |  | 0.37736 | 2.4
| 0.058884 | -24.4 |  | 0.37162 | 2.5
| 0.058642 | -24.3 |  | 0.36612 | 2.6
| 0.058398 | -24.2 |  | 0.36085 | 2.7
| 0.058152 | -24.1 |  | 0.3558 | 2.8
| 0.057904 | -24 |  | 0.35095 | 2.9
| 0.057653 | -23.9 |  | 0.34628 | 3
| 0.0574 | -23.8 |  | 0.3418 | 3.1
| 0.057145 | -23.7 |  | 0.33748 | 3.2
| 0.056887 | -23.6 |  | 0.33333 | 3.3
| 0.056627 | -23.5 |  | 0.32932 | 3.4
| 0.056364 | -23.4 |  | 0.32546 | 3.5
| 0.056099 | -23.3 |  | 0.32173 | 3.6
| 0.055832 | -23.2 |  | 0.31813 | 3.7
| 0.055562 | -23.1 |  | 0.31466 | 3.8
| 0.055289 | -23 |  | 0.3113 | 3.9
| 0.055014 | -22.9 |  | 0.30804 | 4
| 0.054736 | -22.8 |  | 0.3049 | 4.1
| 0.054455 | -22.7 |  | 0.30185 | 4.2
| 0.054172 | -22.6 |  | 0.2989 | 4.3
| 0.053885 | -22.5 |  | 0.29604 | 4.4
| 0.053596 | -22.4 |  | 0.29326 | 4.5
| 0.053304 | -22.3 |  | 0.29057 | 4.6
| 0.05301 | -22.2 |  | 0.28796 | 4.7
| 0.052712 | -22.1 |  | 0.28542 | 4.8
| 0.052411 | -22 |  | 0.28295 | 4.9
| 0.052107 | -21.9 |  | 0.28056 | 5
| 0.051801 | -21.8 |  | 0.27823 | 5.1
| 0.051491 | -21.7 |  | 0.27596 | 5.2
| 0.051177 | -21.6 |  | 0.27376 | 5.3
| 0.050861 | -21.5 |  | 0.27161 | 5.4
| 0.050541 | -21.4 |  | 0.26952 | 5.5
| 0.050218 | -21.3 |  | 0.26748 | 5.6
| 0.049892 | -21.2 |  | 0.2655 | 5.7
| 0.049562 | -21.1 |  | 0.26357 | 5.8
| 0.049229 | -21 |  | 0.26168 | 5.9
| 0.048892 | -20.9 |  | 0.25984 | 6
| 0.048551 | -20.8 |  | 0.25805 | 6.1
| 0.048207 | -20.7 |  | 0.2563 | 6.2
| 0.047859 | -20.6 |  | 0.25459 | 6.3
| 0.047508 | -20.5 |  | 0.25292 | 6.4
| 0.047152 | -20.4 |  | 0.25129 | 6.5
| 0.046793 | -20.3 |  | 0.2497 | 6.6
| 0.04643 | -20.2 |  | 0.24814 | 6.7 |  | this equation, below, enables us to convert X(D) which is defined as 16 bit#/65535 , to dose; this can be used in the commissioning to ascertain what difference in response of the scanner across the scanner plate, or indeed what change in response over time, corresponds to which dose. The dose can then be assessed to see if the response of scanner is significant, if it needs to be corrected for during normal use or whether this difference can be subsumed into a considered allowed uncertainty budget.
| 0.046062 | -20.1 |  | 0.24662 | 6.8 |  |  |  |  | y = 547.95x4 - 1149.5x3 + 910.56x2 - 330.58x + 48.104
| 0.045691 | -20 |  | 0.24513 | 6.9 |  |  |  |  | y = 5688.1x6 - 17186x5 + 21549x4 - 14418x3 + 5477.3x2 - 1141.8x + 106.2
| 0.045315 | -19.9 |  | 0.24368 | 7
| 0.044935 | -19.8
| 0.044551 | -19.7
| 0.044162 | -19.6
| 0.043769 | -19.5
| 0.043371 | -19.4
| 0.042969 | -19.3
| 0.042562 | -19.2
| 0.04215 | -19.1
| 0.041733 | -19 | [5 empty] | change in X(D) |  |  | change in dose
| 0.041312 | -18.9 | [5 empty] | first X(D) | Second (XD) | first dose | second dose | change in dose
| 0.040885 | -18.8 | [5 empty] | 0.36011 | 0.38453 | 2.6739 | 2.2465 | 0.4274 | this one is the difference across the scan plate, using an already irradiated film to about 3Gy
| 0.040453 | -18.7 |  |  |  | 43900 | 43700 | 0.66987 | 0.66682 | 0.056687 | 0.055152 | 0.0015356 | this one is the difference measured over a day , it was the third day that I repeated these measurements, this time I used a non-irradiated film
| 0.040016 | -18.6 |  |  |  | 43850 | 43550 | 0.66911 | 0.66453 | 0.056173 | 0.054894 | 0.0012791 | this one is the variation in response or dose when the film repeatedly scanned, from the first batch of measurements on the third day scanning
| 0.039574 | -18.5 |  |  |  | 43460 | 44100 | 0.66316 | 0.67292 | 0.055097 | 0.059635 | -0.0045379 | this one is variation between middle and edge of scan plate when considering an unirradiated piece of film
| 0.039126 | -18.4 |  |  |  | 43900 | 43700 | 0.66987 | 0.66682 | 0.12973 | 0.13759 | -0.0078551 | this and the next two are the same as the previous 3 except they are using the order 6 polynomial ; the order of polynomial was increased because the 4th order one was appearing on the graph not to describe the low dose area very well
| 0.038672 | -18.3 |  |  |  | 43850 | 43550 | 0.66911 | 0.66453 | 0.13163 | 0.14394 | -0.012315
| 0.038213 | -18.2 |  |  |  | 43460 | 44100 | 0.66316 | 0.67292 | 0.14793 | 0.12264 | 0.025289
| 0.037748 | -18.1 |  |  |  | 43940 | 43520 | 0.67048 | 0.66407 | 0.12825 | 0.14526 | -0.017007 | first day looking at change over time
| 0.037278 | -18 |  |  |  | 26300 | 25900 | 0.40131 | 0.39521 | 2.0454 | 2.1343 | -0.088849 | second day over time
| 0.036801 | -17.9 |  |  |  | 46500 | 43250 | 0.70954 | 0.65995 | 0.12839 | 0.15773 | -0.029343
| 0.036319 | -17.8 |  |  |  | 45500 | 44000 | 0.69429 | 0.6714 | 0.10107 | 0.12609 | -0.025018
| 0.03583 | -17.7 |  |  |  | 46000 | 44000 | 0.70192 | 0.6714 | 0.10916 | 0.12609 | -0.01693
| 0.035335 | -17.6
| 0.034833 | -17.5
| 0.034325 | -17.4
| 0.03381 | -17.3
| 0.033288 | -17.2
| 0.03276 | -17.1
| 0.032224 | -17
| 0.031681 | -16.9
| 0.03113 | -16.8
| 0.030572 | -16.7
| 0.030007 | -16.6
| 0.029433 | -16.5
| 0.028852 | -16.4
| 0.028262 | -16.3
| 0.027664 | -16.2
| 0.027057 | -16.1
| 0.026442 | -16
| 0.025818 | -15.9
| 0.025184 | -15.8
| 0.024541 | -15.7
| 0.023889 | -15.6
| 0.023227 | -15.5
| 0.022555 | -15.4
| 0.021873 | -15.3
| 0.02118 | -15.2
| 0.020477 | -15.1
| 0.019763 | -15
| 0.019037 | -14.9
| 0.0183 | -14.8
| 0.017552 | -14.7
| 0.016791 | -14.6
| 0.016018 | -14.5
| 0.015233 | -14.4
| 0.014435 | -14.3
| 0.013623 | -14.2
| 0.012798 | -14.1
| 0.011959 | -14
| 0.011106 | -13.9
| 0.010239 | -13.8
| 0.0093558 | -13.7
| 0.0084578 | -13.6
| 0.0075439 | -13.5
| 0.0066138 | -13.4
| 0.0056672 | -13.3
| 0.0047034 | -13.2
| 0.0037221 | -13.1
| 0.0027228 | -13
| 0.0017049 | -12.9
| 0.00066804 | -12.8
| -0.00038845 | -12.7
| -0.0014651 | -12.6
| -0.0025625 | -12.5
| -0.0036812 | -12.4
| -0.0048218 | -12.3
| -0.0059851 | -12.2
| -0.0071717 | -12.1
| -0.0083823 | -12
| -0.0096176 | -11.9
| -0.010878 | -11.8
| -0.012166 | -11.7
| -0.01348 | -11.6
| -0.014822 | -11.5
| -0.016193 | -11.4
| -0.017594 | -11.3
| -0.019026 | -11.2
| -0.02049 | -11.1
| -0.021987 | -11
| -0.023517 | -10.9
| -0.025083 | -10.8
| -0.026686 | -10.7
| -0.028326 | -10.6
| -0.030005 | -10.5
| -0.031725 | -10.4
| -0.033487 | -10.3
| -0.035293 | -10.2
| -0.037143 | -10.1
| -0.039041 | -10
| -0.040987 | -9.9
| -0.042984 | -9.8
| -0.045034 | -9.7
| -0.047138 | -9.6
| -0.0493 | -9.5
| -0.05152 | -9.4
| -0.053803 | -9.3
| -0.05615 | -9.2
| -0.058564 | -9.1
| -0.061048 | -9
| -0.063605 | -8.9
| -0.066239 | -8.8
| -0.068953 | -8.7
| -0.071751 | -8.6
| -0.074636 | -8.5
| -0.077613 | -8.4
| -0.080686 | -8.3
| -0.08386 | -8.2
| -0.08714 | -8.1
| -0.090532 | -8
| -0.09404 | -7.9
| -0.097673 | -7.8
| -0.10143 | -7.7
| -0.10533 | -7.6
| -0.10938 | -7.5
| -0.11358 | -7.4
| -0.11793 | -7.3
| -0.12247 | -7.2
| -0.12718 | -7.1
| -0.13208 | -7
| -0.13719 | -6.9
| -0.14252 | -6.8
| -0.14808 | -6.7
| -0.15389 | -6.6
| -0.15996 | -6.5
| -0.16632 | -6.4
| -0.17298 | -6.3
| -0.17996 | -6.2
| -0.1873 | -6.1
| -0.19501 | -6
| -0.20313 | -5.9
| -0.2117 | -5.8
| -0.22073 | -5.7
| -0.23029 | -5.6
| -0.24041 | -5.5
| -0.25115 | -5.4
| -0.26256 | -5.3
| -0.27472 | -5.2
| -0.28768 | -5.1
| -0.30154 | -5
| -0.3164 | -4.9
| -0.33236 | -4.8
| -0.34956 | -4.7
| -0.36814 | -4.6
| -0.38828 | -4.5
| -0.41018 | -4.4
| -0.43409 | -4.3
| -0.46028 | -4.2
| -0.48911 | -4.1
| -0.521 | -4
| -0.55645 | -3.9
| -0.59611 | -3.8
| -0.64077 | -3.7
| -0.69143 | -3.6
| -0.74941 | -3.5
| -0.81639 | -3.4
| -0.89466 | -3.3
| -0.98733 | -3.2
| -1.0988 | -3.1
| -1.2354 | -3
| -1.4067 | -2.9
| -1.6279 | -2.8
| -1.9245 | -2.7
| -2.343 | -2.6
| -2.9778 | -2.5
| -4.0554 | -2.4
| -6.2868 | -2.3
| -13.6605 | -2.2
| 90.5663 | -2.1
| 10.6704 | -2
| 5.7186 | -1.9
| 3.9288 | -1.8
| 3.0052 | -1.7
| 2.4416 | -1.6
| 2.0618 | -1.5
| 1.7885 | -1.4
| 1.5824 | -1.3
| 1.4214 | -1.2
| 1.2922 | -1.1
| 1.1863 | -1
| 1.0978 | -0.9
| 1.0227 | -0.8
| 0.95832 | -0.7
| 0.90242 | -0.6
| 0.85346 | -0.5
| 0.81021 | -0.4
| 0.77173 | -0.3
| 0.73728 | -0.2
| 0.70624 | -0.1
| 0.67815 | 4.0146e-13
| 0.65259 | 0.1
| 0.62924 | 0.2
| 0.60783 | 0.3
| 0.58812 | 0.4
| 0.56992 | 0.5
| 0.55306 | 0.6
| 0.5374 | 0.7
| 0.52282 | 0.8
| 0.5092 | 0.9
| 0.49646 | 1
| 0.48451 | 1.1
| 0.47328 | 1.2
| 0.46271 | 1.3
| 0.45275 | 1.4
| 0.44333 | 1.5
| 0.43442 | 1.6
| 0.42598 | 1.7
| 0.41797 | 1.8
| 0.41036 | 1.9
| 0.40312 | 2
| 0.39622 | 2.1
| 0.38964 | 2.2
| 0.38336 | 2.3
| 0.37736 | 2.4
| 0.37162 | 2.5
| 0.36612 | 2.6
| 0.36085 | 2.7
| 0.3558 | 2.8
| 0.35095 | 2.9
| 0.34628 | 3
| 0.3418 | 3.1
| 0.33748 | 3.2
| 0.33333 | 3.3
| 0.32932 | 3.4
| 0.32546 | 3.5
| 0.32173 | 3.6
| 0.31813 | 3.7
| 0.31466 | 3.8
| 0.3113 | 3.9
| 0.30804 | 4
| 0.3049 | 4.1
| 0.30185 | 4.2
| 0.2989 | 4.3
| 0.29604 | 4.4
| 0.29326 | 4.5
| 0.29057 | 4.6
| 0.28796 | 4.7
| 0.28542 | 4.8
| 0.28295 | 4.9
| 0.28056 | 5
| 0.27823 | 5.1
| 0.27596 | 5.2
| 0.27376 | 5.3
| 0.27161 | 5.4
| 0.26952 | 5.5
| 0.26748 | 5.6
| 0.2655 | 5.7
| 0.26357 | 5.8
| 0.26168 | 5.9
| 0.25984 | 6
| 0.25805 | 6.1
| 0.2563 | 6.2
| 0.25459 | 6.3
| 0.25292 | 6.4
| 0.25129 | 6.5
| 0.2497 | 6.6
| 0.24814 | 6.7
| 0.24662 | 6.8
| 0.24513 | 6.9
| 0.24368 | 7
| 0.24225 | 7.1
| 0.24086 | 7.2
| 0.2395 | 7.3
| 0.23816 | 7.4
| 0.23686 | 7.5
| 0.23558 | 7.6
| 0.23432 | 7.7
| 0.23309 | 7.8
| 0.23189 | 7.9
| 0.23071 | 8
| 0.22955 | 8.1
| 0.22842 | 8.2
| 0.2273 | 8.3
| 0.22621 | 8.4
| 0.22514 | 8.5
| 0.22409 | 8.6
| 0.22306 | 8.7
| 0.22204 | 8.8
| 0.22105 | 8.9
| 0.22007 | 9
| 0.21911 | 9.1
| 0.21817 | 9.2
| 0.21725 | 9.3
| 0.21634 | 9.4
| 0.21544 | 9.5
| 0.21456 | 9.6
| 0.2137 | 9.7
| 0.21285 | 9.8
| 0.21202 | 9.9
| 0.21119 | 10
| 0.21039 | 10.1
| 0.20959 | 10.2
| 0.20881 | 10.3
| 0.20804 | 10.4
| 0.20728 | 10.5
| 0.20654 | 10.6
| 0.2058 | 10.7
| 0.20508 | 10.8
| 0.20437 | 10.9
| 0.20367 | 11
| 0.20298 | 11.1
| 0.2023 | 11.2
| 0.20163 | 11.3
| 0.20097 | 11.4
| 0.20032 | 11.5
| 0.19968 | 11.6
| 0.19905 | 11.7
| 0.19843 | 11.8
| 0.19781 | 11.9
| 0.19721 | 12
| 0.19661 | 12.1
| 0.19603 | 12.2
| 0.19545 | 12.3
| 0.19488 | 12.4
| 0.19431 | 12.5
| 0.19376 | 12.6
| 0.19321 | 12.7
| 0.19267 | 12.8
| 0.19213 | 12.9
| 0.19161 | 13
| 0.19109 | 13.1
| 0.19057 | 13.2
| 0.19007 | 13.3
| 0.18957 | 13.4
| 0.18907 | 13.5
| 0.18859 | 13.6
| 0.18811 | 13.7
| 0.18763 | 13.8
| 0.18716 | 13.9
| 0.1867 | 14
| 0.18624 | 14.1
| 0.18579 | 14.2
| 0.18534 | 14.3
| 0.1849 | 14.4
| 0.18447 | 14.5
| 0.18403 | 14.6
| 0.18361 | 14.7
| 0.18319 | 14.8
| 0.18277 | 14.9
| 0.18236 | 15
| 0.18196 | 15.1
| 0.18156 | 15.2
| 0.18116 | 15.3
| 0.18077 | 15.4
| 0.18038 | 15.5
| 0.18 | 15.6
| 0.17962 | 15.7
| 0.17924 | 15.8
| 0.17887 | 15.9
| 0.17851 | 16
| 0.17814 | 16.1
| 0.17779 | 16.2
| 0.17743 | 16.3
| 0.17708 | 16.4
| 0.17673 | 16.5
| 0.17639 | 16.6
| 0.17605 | 16.7
| 0.17572 | 16.8
| 0.17538 | 16.9
| 0.17505 | 17
| 0.17473 | 17.1
| 0.17441 | 17.2
| 0.17409 | 17.3
| 0.17377 | 17.4
| 0.17346 | 17.5
| 0.17315 | 17.6
| 0.17284 | 17.7
| 0.17254 | 17.8
| 0.17224 | 17.9
| 0.17195 | 18
| 0.17165 | 18.1
| 0.17136 | 18.2
| 0.17107 | 18.3
| 0.17079 | 18.4
| 0.1705 | 18.5
| 0.17022 | 18.6
| 0.16995 | 18.7
| 0.16967 | 18.8
| 0.1694 | 18.9
| 0.16913 | 19
| 0.16886 | 19.1
| 0.1686 | 19.2
| 0.16834 | 19.3
| 0.16808 | 19.4
| 0.16782 | 19.5
| 0.16757 | 19.6
| 0.16731 | 19.7
| 0.16706 | 19.8
| 0.16682 | 19.9
| 0.16657 | 20
| 0.16633 | 20.1
| 0.16609 | 20.2
| 0.16585 | 20.3
| 0.16561 | 20.4
| 0.16538 | 20.5
| 0.16514 | 20.6
| 0.16491 | 20.7
| 0.16468 | 20.8
| 0.16446 | 20.9
| 0.16423 | 21
| 0.16401 | 21.1
| 0.16379 | 21.2
| 0.16357 | 21.3
| 0.16335 | 21.4
| 0.16314 | 21.5
| 0.16292 | 21.6
| 0.16271 | 21.7
| 0.1625 | 21.8
| 0.16229 | 21.9
| 0.16209 | 22
| 0.16188 | 22.1
| 0.16168 | 22.2
| 0.16148 | 22.3
| 0.16128 | 22.4
| 0.16108 | 22.5
| 0.16088 | 22.6
| 0.16069 | 22.7
| 0.1605 | 22.8
| 0.1603 | 22.9
| 0.16011 | 23
| 0.15992 | 23.1
| 0.15974 | 23.2
| 0.15955 | 23.3
| 0.15937 | 23.4
| 0.15918 | 23.5
| 0.159 | 23.6
| 0.15882 | 23.7
| 0.15864 | 23.8
| 0.15847 | 23.9
| 0.15829 | 24
| 0.15812 | 24.1
| 0.15794 | 24.2
| 0.15777 | 24.3
| 0.1576 | 24.4
| 0.15743 | 24.5
| 0.15726 | 24.6
| 0.1571 | 24.7
| 0.15693 | 24.8
| 0.15677 | 24.9
| 0.1566 | 25
| 0.15644 | 25.1
| 0.15628 | 25.2
| 0.15612 | 25.3
| 0.15596 | 25.4
| 0.1558 | 25.5
| 0.15565 | 25.6
| 0.15549 | 25.7
| 0.15534 | 25.8
| 0.15519 | 25.9
| 0.15503 | 26
| 0.15488 | 26.1
| 0.15473 | 26.2
| 0.15459 | 26.3
| 0.15444 | 26.4
| 0.15429 | 26.5
| 0.15415 | 26.6
| 0.154 | 26.7
| 0.15386 | 26.8
| 0.15372 | 26.9
| 0.15357 | 27
| 0.15343 | 27.1
| 0.15329 | 27.2
| 0.15315 | 27.3
| 0.15302 | 27.4
| 0.15288 | 27.5
| 0.15274 | 27.6
| 0.15261 | 27.7
| 0.15248 | 27.8
| 0.15234 | 27.9
| 0.15221 | 28
| 0.15208 | 28.1
| 0.15195 | 28.2
| 0.15182 | 28.3
| 0.15169 | 28.4
| 0.15156 | 28.5
| 0.15143 | 28.6
| 0.15131 | 28.7
| 0.15118 | 28.8
| 0.15106 | 28.9
| 0.15093 | 29
| 0.15081 | 29.1
| 0.15069 | 29.2
| 0.15057 | 29.3
| 0.15045 | 29.4
| 0.15033 | 29.5
| 0.15021 | 29.6
| 0.15009 | 29.7
| 0.14997 | 29.8
| 0.14985 | 29.9
| 0.14974 | 30
| 0.14962 | 30.1
| 0.14951 | 30.2
| 0.14939 | 30.3
| 0.14928 | 30.4
| 0.14917 | 30.5
| 0.14906 | 30.6
| 0.14894 | 30.7
| 0.14883 | 30.8
| 0.14872 | 30.9
| 0.14861 | 31
| 0.14851 | 31.1
| 0.1484 | 31.2
| 0.14829 | 31.3
| 0.14818 | 31.4
| 0.14808 | 31.5
| 0.14797 | 31.6
| 0.14787 | 31.7
| 0.14776 | 31.8
| 0.14766 | 31.9
| 0.14756 | 32
| 0.14745 | 32.1
| 0.14735 | 32.2
| 0.14725 | 32.3
| 0.14715 | 32.4
| 0.14705 | 32.5
| 0.14695 | 32.6
| 0.14685 | 32.7
| 0.14675 | 32.8
| 0.14666 | 32.9
| 0.14656 | 33
| 0.14646 | 33.1
| 0.14637 | 33.2
| 0.14627 | 33.3
| 0.14617 | 33.4
| 0.14608 | 33.5
| 0.14599 | 33.6
| 0.14589 | 33.7
| 0.1458 | 33.8
| 0.14571 | 33.9
| 0.14562 | 34
| 0.14552 | 34.1
| 0.14543 | 34.2
| 0.14534 | 34.3
| 0.14525 | 34.4
| 0.14516 | 34.5
| 0.14507 | 34.6
| 0.14499 | 34.7
| 0.1449 | 34.8
| 0.14481 | 34.9
| 0.14472 | 35
| 0.14464 | 35.1
| 0.14455 | 35.2
| 0.14447 | 35.3
| 0.14438 | 35.4
| 0.1443 | 35.5
| 0.14421 | 35.6
| 0.14413 | 35.7
| 0.14404 | 35.8
| 0.14396 | 35.9
| 0.14388 | 36
| 0.1438 | 36.1
| 0.14371 | 36.2
| 0.14363 | 36.3
| 0.14355 | 36.4
| 0.14347 | 36.5
| 0.14339 | 36.6
| 0.14331 | 36.7
| 0.14323 | 36.8
| 0.14315 | 36.9
| 0.14308 | 37
| 0.143 | 37.1
| 0.14292 | 37.2
| 0.14284 | 37.3
| 0.14277 | 37.4
| 0.14269 | 37.5
| 0.14261 | 37.6
| 0.14254 | 37.7
| 0.14246 | 37.8
| 0.14239 | 37.9
| 0.14231 | 38
| 0.14224 | 38.1
| 0.14217 | 38.2
| 0.14209 | 38.3
| 0.14202 | 38.4
| 0.14195 | 38.5
| 0.14188 | 38.6
| 0.1418 | 38.7
| 0.14173 | 38.8
| 0.14166 | 38.9
| 0.14159 | 39
| 0.14152 | 39.1
| 0.14145 | 39.2
| 0.14138 | 39.3
| 0.14131 | 39.4
| 0.14124 | 39.5
| 0.14117 | 39.6
| 0.1411 | 39.7
| 0.14103 | 39.8
| 0.14097 | 39.9
| 0.1409 | 40
| 0.14083 | 40.1
| 0.14077 | 40.2
| 0.1407 | 40.3
| 0.14063 | 40.4
| 0.14057 | 40.5
| 0.1405 | 40.6
| 0.14044 | 40.7
| 0.14037 | 40.8
| 0.14031 | 40.9
| 0.14024 | 41
| 0.14018 | 41.1
| 0.14011 | 41.2
| 0.14005 | 41.3
| 0.13999 | 41.4
| 0.13992 | 41.5
| 0.13986 | 41.6
| 0.1398 | 41.7
| 0.13974 | 41.8
| 0.13967 | 41.9
| 0.13961 | 42
| 0.13955 | 42.1
| 0.13949 | 42.2
| 0.13943 | 42.3
| 0.13937 | 42.4
| 0.13931 | 42.5
| 0.13925 | 42.6
| 0.13919 | 42.7
| 0.13913 | 42.8
| 0.13907 | 42.9
| 0.13901 | 43
| 0.13895 | 43.1
| 0.1389 | 43.2
| 0.13884 | 43.3
| 0.13878 | 43.4
| 0.13872 | 43.5
| 0.13866 | 43.6
| 0.13861 | 43.7
| 0.13855 | 43.8
| 0.13849 | 43.9
| 0.13844 | 44
| 0.13838 | 44.1
| 0.13833 | 44.2
| 0.13827 | 44.3
| 0.13821 | 44.4
| 0.13816 | 44.5
| 0.1381 | 44.6
| 0.13805 | 44.7
| 0.138 | 44.8
| 0.13794 | 44.9
| 0.13789 | 45
| 0.13783 | 45.1
| 0.13778 | 45.2
| 0.13773 | 45.3
| 0.13767 | 45.4
| 0.13762 | 45.5
| 0.13757 | 45.6
| 0.13752 | 45.7
| 0.13746 | 45.8
| 0.13741 | 45.9
| 0.13736 | 46
| 0.13731 | 46.1
| 0.13726 | 46.2
| 0.13721 | 46.3
| 0.13716 | 46.4
| 0.1371 | 46.5
| 0.13705 | 46.6
| 0.137 | 46.7
| 0.13695 | 46.8
| 0.1369 | 46.9
| 0.13685 | 47
| 0.1368 | 47.1
| 0.13676 | 47.2
| 0.13671 | 47.3
| 0.13666 | 47.4
| 0.13661 | 47.5
| 0.13656 | 47.6
| 0.13651 | 47.7
| 0.13646 | 47.8
| 0.13642 | 47.9
| 0.13637 | 48
| 0.13632 | 48.1
| 0.13627 | 48.2
| 0.13623 | 48.3
| 0.13618 | 48.4
| 0.13613 | 48.5
| 0.13609 | 48.6
| 0.13604 | 48.7
| 0.13599 | 48.8
| 0.13595 | 48.9
| 0.1359 | 49
| 0.13586 | 49.1
| 0.13581 | 49.2
| 0.13577 | 49.3
| 0.13572 | 49.4
| 0.13568 | 49.5
| 0.13563 | 49.6
| 0.13559 | 49.7
| 0.13554 | 49.8
| 0.1355 | 49.9
| 0.13545 | 50
| 0.13541 | 50.1
| 0.13537 | 50.2
| 0.13532 | 50.3
| 0.13528 | 50.4
| 0.13524 | 50.5
| 0.13519 | 50.6
| 0.13515 | 50.7
| 0.13511 | 50.8
| 0.13506 | 50.9
| 0.13502 | 51
| 0.13498 | 51.1
| 0.13494 | 51.2
| 0.13489 | 51.3
| 0.13485 | 51.4
| 0.13481 | 51.5
| 0.13477 | 51.6
| 0.13473 | 51.7
| 0.13469 | 51.8
| 0.13465 | 51.9
| 0.13461 | 52
| 0.13456 | 52.1
| 0.13452 | 52.2
| 0.13448 | 52.3
| 0.13444 | 52.4
| 0.1344 | 52.5
| 0.13436 | 52.6
| 0.13432 | 52.7
| 0.13428 | 52.8
| 0.13424 | 52.9
| 0.1342 | 53
| 0.13417 | 53.1
| 0.13413 | 53.2
| 0.13409 | 53.3
| 0.13405 | 53.4
| 0.13401 | 53.5
| 0.13397 | 53.6
| 0.13393 | 53.7
| 0.13389 | 53.8
| 0.13386 | 53.9
| 0.13382 | 54
| 0.13378 | 54.1
| 0.13374 | 54.2
| 0.1337 | 54.3
| 0.13367 | 54.4
| 0.13363 | 54.5
| 0.13359 | 54.6
| 0.13356 | 54.7
| 0.13352 | 54.8
| 0.13348 | 54.9
| 0.13345 | 55
| 0.13341 | 55.1
| 0.13337 | 55.2
| 0.13334 | 55.3
| 0.1333 | 55.4
| 0.13326 | 55.5
| 0.13323 | 55.6
| 0.13319 | 55.7
| 0.13316 | 55.8
| 0.13312 | 55.9
| 0.13309 | 56
| 0.13305 | 56.1
| 0.13301 | 56.2
| 0.13298 | 56.3
| 0.13294 | 56.4
| 0.13291 | 56.5
| 0.13287 | 56.6
| 0.13284 | 56.7
| 0.13281 | 56.8
| 0.13277 | 56.9
| 0.13274 | 57
| 0.1327 | 57.1
| 0.13267 | 57.2
| 0.13263 | 57.3
| 0.1326 | 57.4
| 0.13257 | 57.5
| 0.13253 | 57.6
| 0.1325 | 57.7
| 0.13247 | 57.8
| 0.13243 | 57.9
| 0.1324 | 58
| 0.13237 | 58.1
| 0.13233 | 58.2
| 0.1323 | 58.3
| 0.13227 | 58.4
| 0.13224 | 58.5
| 0.1322 | 58.6
| 0.13217 | 58.7
| 0.13214 | 58.8
| 0.13211 | 58.9
| 0.13208 | 59
| 0.13204 | 59.1
| 0.13201 | 59.2
| 0.13198 | 59.3
| 0.13195 | 59.4
| 0.13192 | 59.5
| 0.13189 | 59.6
| 0.13185 | 59.7
| 0.13182 | 59.8
| 0.13179 | 59.9
| 0.13176 | 60

## rational cal 3
| Calibration data, 16-bit # | [10 empty] | Film | Position | Dose Measurement (cGy) | Scaled Dose (cGy) |  | Film Raw Reading (ImageJ) | film norm | calc film response | sum of squares
| Dose, cGy | Red | Green | Blue | [7 empty] | separate blank sheet |  |  | 0 | 0 | 43519.8124 | 0.66407 | 0.66758 | 1.2325e-05 | 0 | 0 | 21785.1125
| 300 | 22704.438 | 26601.2067 | 19960.2125 | [7 empty] | A | 0 | 45.3832 | 45.3832 | 0.45383 | 39156.755 | 0.59749 | 0.59265 | 2.3478e-05 | 5 | 45.3832 | 26695.787
| 200 | 26397.3885 | 30047.5799 | 21188.3132 | [7 empty] | B | 0 | 45.3832 | 45.3832 | 0.45383 | 39057.5935 | 0.59598 | 0.59265 | 1.1104e-05 | 10 | 75.6558 | 29431.6357
| 100 | 32547.5085 | 35131.9428 | 22942.7761 | [7 empty] | A | -2.5 |  | 75.6558 | 0.75656 | 36068.584 | 0.55037 | 0.5509 | 2.8166e-07 | 15 | 108.9881 | 32054.9261
| 0 | 44440.5987 | 41860.923 | 24778.5695 | [7 empty] | B | 2.5 |  | 108.9881 | 1.0899 | 33372.4325 | 0.50923 | 0.51087 | 2.6978e-06 | 20 | 136.4914 | 33966.2062
[12 empty] | A | -5 |  | 136.4914 | 1.3649 | 31447.0776 | 0.47985 | 0.48171 | 3.4495e-06 | 25 | 167.0991 | 35868.7089
[12 empty] | B | 5 |  | 167.0991 | 1.671 | 29534.7283 | 0.45067 | 0.45268 | 4.0301e-06 | 30 | 180.6057 | 36642.5889
| Calibration data, 16-bit #/65535 | [10 empty] | A | -7.5 |  | 180.6057 | 1.8061 | 28835.196 | 0.44 | 0.44087 | 7.6221e-07 | 35 | 208.5511 | 38131.6495
| Dose, Gy | Red | Green | Blue | [7 empty] | B | 7.5 |  | 208.5511 | 2.0855 | 27340.3198 | 0.41719 | 0.41815 | 9.2503e-07 | 40 | 216.2383 | 38516.9189
| 3 | 0.34645 | 0.40591 | 0.30457 | [7 empty] | A | -10 |  | 216.2383 | 2.1624 | 27013.1678 | 0.41219 | 0.41227 | 5.6209e-09 | 45 | 241.7777 | 39728.8972
| 2 | 0.4028 | 0.4585 | 0.32331 | [7 empty] | B | 10 |  | 241.7777 | 2.4178 | 25794.6459 | 0.3936 | 0.39378 | 3.0562e-08 | 50 | 243.6134 | 39812.2076
| 1 | 0.49664 | 0.53608 | 0.35008 |  |  |  |  | k=1.22476 |  |  | A | -12.5 |  | 243.6134 | 2.4361 | 25846.9674 | 0.3944 | 0.3925 | 3.5902e-06 | 55 | 266.1126 | 40795.0453
| 0 | 0.67812 | 0.63876 | 0.3781 | [7 empty] | B | 12.5 |  | 266.1126 | 2.6611 | 24846.9167 | 0.37914 | 0.37751 | 2.6639e-06
6.5344e-05
| Calibration curve | [11 empty] | initial values
|  | X(D) = A + B/(Dose -C) | [10 empty] | a | -0.040473
where X(D) is response expressed as 16-bit #/65535, D is dose in Gy and A, B and C are constants | [12 empty] | b | 2.7151
[13 empty] | c | -3.8345
|  | Red | Green | Blue
| A | 0.11252 | 0.02256 | 0.063139
| B | 1.1953 | 3.0228 | 3.0654
| C | -2.1132 | -4.903 | -9.7235
| X(D) | Dose(Gy)
| 0.069654 | -30
| 0.0695 | -29.9
| 0.069344 | -29.8
| 0.069188 | -29.7
| 0.06903 | -29.6
| 0.068871 | -29.5
| 0.068711 | -29.4
| 0.06855 | -29.3
| 0.068388 | -29.2
| 0.068224 | -29.1
| 0.06806 | -29
| 0.067894 | -28.9
| 0.067727 | -28.8
| 0.067558 | -28.7
| 0.067388 | -28.6
| 0.067217 | -28.5
| 0.067045 | -28.4
| 0.066871 | -28.3 | 0.11
| 0.066696 | -28.2
| 0.06652 | -28.1
| 0.066342 | -28
| 0.066163 | -27.9
| 0.065983 | -27.8
| 0.065801 | -27.7
| 0.065618 | -27.6
| 0.065433 | -27.5
| 0.065247 | -27.4
| 0.065059 | -27.3
| 0.06487 | -27.2
| 0.064679 | -27.1
| 0.064487 | -27 |  | X(D) | Dose(Gy)
| 0.064293 | -26.9 |  | 0.67815 | 4.0146e-13
| 0.064098 | -26.8 |  | 0.65259 | 0.1
| 0.063901 | -26.7 |  | 0.62924 | 0.2
| 0.063702 | -26.6 |  | 0.60783 | 0.3
| 0.063502 | -26.5 |  | 0.58812 | 0.4
| 0.0633 | -26.4 |  | 0.56992 | 0.5
| 0.063097 | -26.3 |  | 0.55306 | 0.6
| 0.062892 | -26.2 |  | 0.5374 | 0.7
| 0.062685 | -26.1 |  | 0.52282 | 0.8
| 0.062476 | -26 |  | 0.5092 | 0.9
| 0.062266 | -25.9 |  | 0.49646 | 1
| 0.062054 | -25.8 |  | 0.48451 | 1.1
| 0.06184 | -25.7 |  | 0.47328 | 1.2
| 0.061624 | -25.6 |  | 0.46271 | 1.3
| 0.061406 | -25.5 |  | 0.45275 | 1.4
| 0.061187 | -25.4 |  | 0.44333 | 1.5
| 0.060966 | -25.3 |  | 0.43442 | 1.6
| 0.060742 | -25.2 |  | 0.42598 | 1.7
| 0.060517 | -25.1 |  | 0.41797 | 1.8
| 0.06029 | -25 |  | 0.41036 | 1.9
| 0.060061 | -24.9 |  | 0.40312 | 2
| 0.059829 | -24.8 |  | 0.39622 | 2.1
| 0.059596 | -24.7 |  | 0.38964 | 2.2
| 0.059361 | -24.6 |  | 0.38336 | 2.3
| 0.059123 | -24.5 |  | 0.37736 | 2.4
| 0.058884 | -24.4 |  | 0.37162 | 2.5
| 0.058642 | -24.3 |  | 0.36612 | 2.6
| 0.058398 | -24.2 |  | 0.36085 | 2.7
| 0.058152 | -24.1 |  | 0.3558 | 2.8
| 0.057904 | -24 |  | 0.35095 | 2.9
| 0.057653 | -23.9 |  | 0.34628 | 3
| 0.0574 | -23.8 |  | 0.3418 | 3.1
| 0.057145 | -23.7 |  | 0.33748 | 3.2
| 0.056887 | -23.6 |  | 0.33333 | 3.3
| 0.056627 | -23.5 |  | 0.32932 | 3.4
| 0.056364 | -23.4 |  | 0.32546 | 3.5
| 0.056099 | -23.3 |  | 0.32173 | 3.6
| 0.055832 | -23.2 |  | 0.31813 | 3.7
| 0.055562 | -23.1 |  | 0.31466 | 3.8
| 0.055289 | -23 |  | 0.3113 | 3.9
| 0.055014 | -22.9 |  | 0.30804 | 4
| 0.054736 | -22.8 |  | 0.3049 | 4.1
| 0.054455 | -22.7 |  | 0.30185 | 4.2
| 0.054172 | -22.6 |  | 0.2989 | 4.3
| 0.053885 | -22.5 |  | 0.29604 | 4.4
| 0.053596 | -22.4 |  | 0.29326 | 4.5
| 0.053304 | -22.3 |  | 0.29057 | 4.6
| 0.05301 | -22.2 |  | 0.28796 | 4.7
| 0.052712 | -22.1 |  | 0.28542 | 4.8
| 0.052411 | -22 |  | 0.28295 | 4.9
| 0.052107 | -21.9 |  | 0.28056 | 5
| 0.051801 | -21.8 |  | 0.27823 | 5.1
| 0.051491 | -21.7 |  | 0.27596 | 5.2
| 0.051177 | -21.6 |  | 0.27376 | 5.3
| 0.050861 | -21.5 |  | 0.27161 | 5.4
| 0.050541 | -21.4 |  | 0.26952 | 5.5
| 0.050218 | -21.3 |  | 0.26748 | 5.6
| 0.049892 | -21.2 |  | 0.2655 | 5.7
| 0.049562 | -21.1 |  | 0.26357 | 5.8
| 0.049229 | -21 |  | 0.26168 | 5.9
| 0.048892 | -20.9 |  | 0.25984 | 6
| 0.048551 | -20.8 |  | 0.25805 | 6.1
| 0.048207 | -20.7 |  | 0.2563 | 6.2
| 0.047859 | -20.6 |  | 0.25459 | 6.3
| 0.047508 | -20.5 |  | 0.25292 | 6.4
| 0.047152 | -20.4 |  | 0.25129 | 6.5
| 0.046793 | -20.3 |  | 0.2497 | 6.6
| 0.04643 | -20.2 |  | 0.24814 | 6.7 |  | this equation, below, enables us to convert X(D) which is defined as 16 bit#/65535 , to dose; this can be used in the commissioning to ascertain what difference in response of the scanner across the scanner plate, or indeed what change in response over time, corresponds to which dose. The dose can then be assessed to see if the response of scanner is significant, if it needs to be corrected for during normal use or whether this difference can be subsumed into a considered allowed uncertainty budget.
| 0.046062 | -20.1 |  | 0.24662 | 6.8 |  |  |  |  | y = 547.95x4 - 1149.5x3 + 910.56x2 - 330.58x + 48.104
| 0.045691 | -20 |  | 0.24513 | 6.9 |  |  |  |  | y = 5688.1x6 - 17186x5 + 21549x4 - 14418x3 + 5477.3x2 - 1141.8x + 106.2
| 0.045315 | -19.9 |  | 0.24368 | 7
| 0.044935 | -19.8
| 0.044551 | -19.7
| 0.044162 | -19.6
| 0.043769 | -19.5
| 0.043371 | -19.4
| 0.042969 | -19.3
| 0.042562 | -19.2
| 0.04215 | -19.1
| 0.041733 | -19 | [5 empty] | change in X(D) |  |  | change in dose
| 0.041312 | -18.9 | [5 empty] | first X(D) | Second (XD) | first dose | second dose | change in dose
| 0.040885 | -18.8 | [5 empty] | 0.36011 | 0.38453 | 2.6739 | 2.2465 | 0.4274 | this one is the difference across the scan plate, using an already irradiated film to about 3Gy
| 0.040453 | -18.7 |  |  |  | 43900 | 43700 | 0.66987 | 0.66682 | 0.056687 | 0.055152 | 0.0015356 | this one is the difference measured over a day , it was the third day that I repeated these measurements, this time I used a non-irradiated film
| 0.040016 | -18.6 |  |  |  | 43850 | 43550 | 0.66911 | 0.66453 | 0.056173 | 0.054894 | 0.0012791 | this one is the variation in response or dose when the film repeatedly scanned, from the first batch of measurements on the third day scanning
| 0.039574 | -18.5 |  |  |  | 43460 | 44100 | 0.66316 | 0.67292 | 0.055097 | 0.059635 | -0.0045379 | this one is variation between middle and edge of scan plate when considering an unirradiated piece of film
| 0.039126 | -18.4 |  |  |  | 43900 | 43700 | 0.66987 | 0.66682 | 0.12973 | 0.13759 | -0.0078551 | this and the next two are the same as the previous 3 except they are using the order 6 polynomial ; the order of polynomial was increased because the 4th order one was appearing on the graph not to describe the low dose area very well
| 0.038672 | -18.3 |  |  |  | 43850 | 43550 | 0.66911 | 0.66453 | 0.13163 | 0.14394 | -0.012315
| 0.038213 | -18.2 |  |  |  | 43460 | 44100 | 0.66316 | 0.67292 | 0.14793 | 0.12264 | 0.025289
| 0.037748 | -18.1 |  |  |  | 43940 | 43520 | 0.67048 | 0.66407 | 0.12825 | 0.14526 | -0.017007 | first day looking at change over time
| 0.037278 | -18 |  |  |  | 26300 | 25900 | 0.40131 | 0.39521 | 2.0454 | 2.1343 | -0.088849 | second day over time
| 0.036801 | -17.9 |  |  |  | 46500 | 43250 | 0.70954 | 0.65995 | 0.12839 | 0.15773 | -0.029343
| 0.036319 | -17.8 |  |  |  | 45500 | 44000 | 0.69429 | 0.6714 | 0.10107 | 0.12609 | -0.025018
| 0.03583 | -17.7 |  |  |  | 46000 | 44000 | 0.70192 | 0.6714 | 0.10916 | 0.12609 | -0.01693
| 0.035335 | -17.6
| 0.034833 | -17.5
| 0.034325 | -17.4
| 0.03381 | -17.3
| 0.033288 | -17.2
| 0.03276 | -17.1
| 0.032224 | -17
| 0.031681 | -16.9
| 0.03113 | -16.8
| 0.030572 | -16.7
| 0.030007 | -16.6
| 0.029433 | -16.5
| 0.028852 | -16.4
| 0.028262 | -16.3
| 0.027664 | -16.2
| 0.027057 | -16.1
| 0.026442 | -16
| 0.025818 | -15.9
| 0.025184 | -15.8
| 0.024541 | -15.7
| 0.023889 | -15.6
| 0.023227 | -15.5
| 0.022555 | -15.4
| 0.021873 | -15.3
| 0.02118 | -15.2
| 0.020477 | -15.1
| 0.019763 | -15
| 0.019037 | -14.9
| 0.0183 | -14.8
| 0.017552 | -14.7
| 0.016791 | -14.6
| 0.016018 | -14.5
| 0.015233 | -14.4
| 0.014435 | -14.3
| 0.013623 | -14.2
| 0.012798 | -14.1
| 0.011959 | -14
| 0.011106 | -13.9
| 0.010239 | -13.8
| 0.0093558 | -13.7
| 0.0084578 | -13.6
| 0.0075439 | -13.5
| 0.0066138 | -13.4
| 0.0056672 | -13.3
| 0.0047034 | -13.2
| 0.0037221 | -13.1
| 0.0027228 | -13
| 0.0017049 | -12.9
| 0.00066804 | -12.8
| -0.00038845 | -12.7
| -0.0014651 | -12.6
| -0.0025625 | -12.5
| -0.0036812 | -12.4
| -0.0048218 | -12.3
| -0.0059851 | -12.2
| -0.0071717 | -12.1
| -0.0083823 | -12
| -0.0096176 | -11.9
| -0.010878 | -11.8
| -0.012166 | -11.7
| -0.01348 | -11.6
| -0.014822 | -11.5
| -0.016193 | -11.4
| -0.017594 | -11.3
| -0.019026 | -11.2
| -0.02049 | -11.1
| -0.021987 | -11
| -0.023517 | -10.9
| -0.025083 | -10.8
| -0.026686 | -10.7
| -0.028326 | -10.6
| -0.030005 | -10.5
| -0.031725 | -10.4
| -0.033487 | -10.3
| -0.035293 | -10.2
| -0.037143 | -10.1
| -0.039041 | -10
| -0.040987 | -9.9
| -0.042984 | -9.8
| -0.045034 | -9.7
| -0.047138 | -9.6
| -0.0493 | -9.5
| -0.05152 | -9.4
| -0.053803 | -9.3
| -0.05615 | -9.2
| -0.058564 | -9.1
| -0.061048 | -9
| -0.063605 | -8.9
| -0.066239 | -8.8
| -0.068953 | -8.7
| -0.071751 | -8.6
| -0.074636 | -8.5
| -0.077613 | -8.4
| -0.080686 | -8.3
| -0.08386 | -8.2
| -0.08714 | -8.1
| -0.090532 | -8
| -0.09404 | -7.9
| -0.097673 | -7.8
| -0.10143 | -7.7
| -0.10533 | -7.6
| -0.10938 | -7.5
| -0.11358 | -7.4
| -0.11793 | -7.3
| -0.12247 | -7.2
| -0.12718 | -7.1
| -0.13208 | -7
| -0.13719 | -6.9
| -0.14252 | -6.8
| -0.14808 | -6.7
| -0.15389 | -6.6
| -0.15996 | -6.5
| -0.16632 | -6.4
| -0.17298 | -6.3
| -0.17996 | -6.2
| -0.1873 | -6.1
| -0.19501 | -6
| -0.20313 | -5.9
| -0.2117 | -5.8
| -0.22073 | -5.7
| -0.23029 | -5.6
| -0.24041 | -5.5
| -0.25115 | -5.4
| -0.26256 | -5.3
| -0.27472 | -5.2
| -0.28768 | -5.1
| -0.30154 | -5
| -0.3164 | -4.9
| -0.33236 | -4.8
| -0.34956 | -4.7
| -0.36814 | -4.6
| -0.38828 | -4.5
| -0.41018 | -4.4
| -0.43409 | -4.3
| -0.46028 | -4.2
| -0.48911 | -4.1
| -0.521 | -4
| -0.55645 | -3.9
| -0.59611 | -3.8
| -0.64077 | -3.7
| -0.69143 | -3.6
| -0.74941 | -3.5
| -0.81639 | -3.4
| -0.89466 | -3.3
| -0.98733 | -3.2
| -1.0988 | -3.1
| -1.2354 | -3
| -1.4067 | -2.9
| -1.6279 | -2.8
| -1.9245 | -2.7
| -2.343 | -2.6
| -2.9778 | -2.5
| -4.0554 | -2.4
| -6.2868 | -2.3
| -13.6605 | -2.2
| 90.5663 | -2.1
| 10.6704 | -2
| 5.7186 | -1.9
| 3.9288 | -1.8
| 3.0052 | -1.7
| 2.4416 | -1.6
| 2.0618 | -1.5
| 1.7885 | -1.4
| 1.5824 | -1.3
| 1.4214 | -1.2
| 1.2922 | -1.1
| 1.1863 | -1
| 1.0978 | -0.9
| 1.0227 | -0.8
| 0.95832 | -0.7
| 0.90242 | -0.6
| 0.85346 | -0.5
| 0.81021 | -0.4
| 0.77173 | -0.3
| 0.73728 | -0.2
| 0.70624 | -0.1
| 0.67815 | 4.0146e-13
| 0.65259 | 0.1
| 0.62924 | 0.2
| 0.60783 | 0.3
| 0.58812 | 0.4
| 0.56992 | 0.5
| 0.55306 | 0.6
| 0.5374 | 0.7
| 0.52282 | 0.8
| 0.5092 | 0.9
| 0.49646 | 1
| 0.48451 | 1.1
| 0.47328 | 1.2
| 0.46271 | 1.3
| 0.45275 | 1.4
| 0.44333 | 1.5
| 0.43442 | 1.6
| 0.42598 | 1.7
| 0.41797 | 1.8
| 0.41036 | 1.9
| 0.40312 | 2
| 0.39622 | 2.1
| 0.38964 | 2.2
| 0.38336 | 2.3
| 0.37736 | 2.4
| 0.37162 | 2.5
| 0.36612 | 2.6
| 0.36085 | 2.7
| 0.3558 | 2.8
| 0.35095 | 2.9
| 0.34628 | 3
| 0.3418 | 3.1
| 0.33748 | 3.2
| 0.33333 | 3.3
| 0.32932 | 3.4
| 0.32546 | 3.5
| 0.32173 | 3.6
| 0.31813 | 3.7
| 0.31466 | 3.8
| 0.3113 | 3.9
| 0.30804 | 4
| 0.3049 | 4.1
| 0.30185 | 4.2
| 0.2989 | 4.3
| 0.29604 | 4.4
| 0.29326 | 4.5
| 0.29057 | 4.6
| 0.28796 | 4.7
| 0.28542 | 4.8
| 0.28295 | 4.9
| 0.28056 | 5
| 0.27823 | 5.1
| 0.27596 | 5.2
| 0.27376 | 5.3
| 0.27161 | 5.4
| 0.26952 | 5.5
| 0.26748 | 5.6
| 0.2655 | 5.7
| 0.26357 | 5.8
| 0.26168 | 5.9
| 0.25984 | 6
| 0.25805 | 6.1
| 0.2563 | 6.2
| 0.25459 | 6.3
| 0.25292 | 6.4
| 0.25129 | 6.5
| 0.2497 | 6.6
| 0.24814 | 6.7
| 0.24662 | 6.8
| 0.24513 | 6.9
| 0.24368 | 7
| 0.24225 | 7.1
| 0.24086 | 7.2
| 0.2395 | 7.3
| 0.23816 | 7.4
| 0.23686 | 7.5
| 0.23558 | 7.6
| 0.23432 | 7.7
| 0.23309 | 7.8
| 0.23189 | 7.9
| 0.23071 | 8
| 0.22955 | 8.1
| 0.22842 | 8.2
| 0.2273 | 8.3
| 0.22621 | 8.4
| 0.22514 | 8.5
| 0.22409 | 8.6
| 0.22306 | 8.7
| 0.22204 | 8.8
| 0.22105 | 8.9
| 0.22007 | 9
| 0.21911 | 9.1
| 0.21817 | 9.2
| 0.21725 | 9.3
| 0.21634 | 9.4
| 0.21544 | 9.5
| 0.21456 | 9.6
| 0.2137 | 9.7
| 0.21285 | 9.8
| 0.21202 | 9.9
| 0.21119 | 10
| 0.21039 | 10.1
| 0.20959 | 10.2
| 0.20881 | 10.3
| 0.20804 | 10.4
| 0.20728 | 10.5
| 0.20654 | 10.6
| 0.2058 | 10.7
| 0.20508 | 10.8
| 0.20437 | 10.9
| 0.20367 | 11
| 0.20298 | 11.1
| 0.2023 | 11.2
| 0.20163 | 11.3
| 0.20097 | 11.4
| 0.20032 | 11.5
| 0.19968 | 11.6
| 0.19905 | 11.7
| 0.19843 | 11.8
| 0.19781 | 11.9
| 0.19721 | 12
| 0.19661 | 12.1
| 0.19603 | 12.2
| 0.19545 | 12.3
| 0.19488 | 12.4
| 0.19431 | 12.5
| 0.19376 | 12.6
| 0.19321 | 12.7
| 0.19267 | 12.8
| 0.19213 | 12.9
| 0.19161 | 13
| 0.19109 | 13.1
| 0.19057 | 13.2
| 0.19007 | 13.3
| 0.18957 | 13.4
| 0.18907 | 13.5
| 0.18859 | 13.6
| 0.18811 | 13.7
| 0.18763 | 13.8
| 0.18716 | 13.9
| 0.1867 | 14
| 0.18624 | 14.1
| 0.18579 | 14.2
| 0.18534 | 14.3
| 0.1849 | 14.4
| 0.18447 | 14.5
| 0.18403 | 14.6
| 0.18361 | 14.7
| 0.18319 | 14.8
| 0.18277 | 14.9
| 0.18236 | 15
| 0.18196 | 15.1
| 0.18156 | 15.2
| 0.18116 | 15.3
| 0.18077 | 15.4
| 0.18038 | 15.5
| 0.18 | 15.6
| 0.17962 | 15.7
| 0.17924 | 15.8
| 0.17887 | 15.9
| 0.17851 | 16
| 0.17814 | 16.1
| 0.17779 | 16.2
| 0.17743 | 16.3
| 0.17708 | 16.4
| 0.17673 | 16.5
| 0.17639 | 16.6
| 0.17605 | 16.7
| 0.17572 | 16.8
| 0.17538 | 16.9
| 0.17505 | 17
| 0.17473 | 17.1
| 0.17441 | 17.2
| 0.17409 | 17.3
| 0.17377 | 17.4
| 0.17346 | 17.5
| 0.17315 | 17.6
| 0.17284 | 17.7
| 0.17254 | 17.8
| 0.17224 | 17.9
| 0.17195 | 18
| 0.17165 | 18.1
| 0.17136 | 18.2
| 0.17107 | 18.3
| 0.17079 | 18.4
| 0.1705 | 18.5
| 0.17022 | 18.6
| 0.16995 | 18.7
| 0.16967 | 18.8
| 0.1694 | 18.9
| 0.16913 | 19
| 0.16886 | 19.1
| 0.1686 | 19.2
| 0.16834 | 19.3
| 0.16808 | 19.4
| 0.16782 | 19.5
| 0.16757 | 19.6
| 0.16731 | 19.7
| 0.16706 | 19.8
| 0.16682 | 19.9
| 0.16657 | 20
| 0.16633 | 20.1
| 0.16609 | 20.2
| 0.16585 | 20.3
| 0.16561 | 20.4
| 0.16538 | 20.5
| 0.16514 | 20.6
| 0.16491 | 20.7
| 0.16468 | 20.8
| 0.16446 | 20.9
| 0.16423 | 21
| 0.16401 | 21.1
| 0.16379 | 21.2
| 0.16357 | 21.3
| 0.16335 | 21.4
| 0.16314 | 21.5
| 0.16292 | 21.6
| 0.16271 | 21.7
| 0.1625 | 21.8
| 0.16229 | 21.9
| 0.16209 | 22
| 0.16188 | 22.1
| 0.16168 | 22.2
| 0.16148 | 22.3
| 0.16128 | 22.4
| 0.16108 | 22.5
| 0.16088 | 22.6
| 0.16069 | 22.7
| 0.1605 | 22.8
| 0.1603 | 22.9
| 0.16011 | 23
| 0.15992 | 23.1
| 0.15974 | 23.2
| 0.15955 | 23.3
| 0.15937 | 23.4
| 0.15918 | 23.5
| 0.159 | 23.6
| 0.15882 | 23.7
| 0.15864 | 23.8
| 0.15847 | 23.9
| 0.15829 | 24
| 0.15812 | 24.1
| 0.15794 | 24.2
| 0.15777 | 24.3
| 0.1576 | 24.4
| 0.15743 | 24.5
| 0.15726 | 24.6
| 0.1571 | 24.7
| 0.15693 | 24.8
| 0.15677 | 24.9
| 0.1566 | 25
| 0.15644 | 25.1
| 0.15628 | 25.2
| 0.15612 | 25.3
| 0.15596 | 25.4
| 0.1558 | 25.5
| 0.15565 | 25.6
| 0.15549 | 25.7
| 0.15534 | 25.8
| 0.15519 | 25.9
| 0.15503 | 26
| 0.15488 | 26.1
| 0.15473 | 26.2
| 0.15459 | 26.3
| 0.15444 | 26.4
| 0.15429 | 26.5
| 0.15415 | 26.6
| 0.154 | 26.7
| 0.15386 | 26.8
| 0.15372 | 26.9
| 0.15357 | 27
| 0.15343 | 27.1
| 0.15329 | 27.2
| 0.15315 | 27.3
| 0.15302 | 27.4
| 0.15288 | 27.5
| 0.15274 | 27.6
| 0.15261 | 27.7
| 0.15248 | 27.8
| 0.15234 | 27.9
| 0.15221 | 28
| 0.15208 | 28.1
| 0.15195 | 28.2
| 0.15182 | 28.3
| 0.15169 | 28.4
| 0.15156 | 28.5
| 0.15143 | 28.6
| 0.15131 | 28.7
| 0.15118 | 28.8
| 0.15106 | 28.9
| 0.15093 | 29
| 0.15081 | 29.1
| 0.15069 | 29.2
| 0.15057 | 29.3
| 0.15045 | 29.4
| 0.15033 | 29.5
| 0.15021 | 29.6
| 0.15009 | 29.7
| 0.14997 | 29.8
| 0.14985 | 29.9
| 0.14974 | 30
| 0.14962 | 30.1
| 0.14951 | 30.2
| 0.14939 | 30.3
| 0.14928 | 30.4
| 0.14917 | 30.5
| 0.14906 | 30.6
| 0.14894 | 30.7
| 0.14883 | 30.8
| 0.14872 | 30.9
| 0.14861 | 31
| 0.14851 | 31.1
| 0.1484 | 31.2
| 0.14829 | 31.3
| 0.14818 | 31.4
| 0.14808 | 31.5
| 0.14797 | 31.6
| 0.14787 | 31.7
| 0.14776 | 31.8
| 0.14766 | 31.9
| 0.14756 | 32
| 0.14745 | 32.1
| 0.14735 | 32.2
| 0.14725 | 32.3
| 0.14715 | 32.4
| 0.14705 | 32.5
| 0.14695 | 32.6
| 0.14685 | 32.7
| 0.14675 | 32.8
| 0.14666 | 32.9
| 0.14656 | 33
| 0.14646 | 33.1
| 0.14637 | 33.2
| 0.14627 | 33.3
| 0.14617 | 33.4
| 0.14608 | 33.5
| 0.14599 | 33.6
| 0.14589 | 33.7
| 0.1458 | 33.8
| 0.14571 | 33.9
| 0.14562 | 34
| 0.14552 | 34.1
| 0.14543 | 34.2
| 0.14534 | 34.3
| 0.14525 | 34.4
| 0.14516 | 34.5
| 0.14507 | 34.6
| 0.14499 | 34.7
| 0.1449 | 34.8
| 0.14481 | 34.9
| 0.14472 | 35
| 0.14464 | 35.1
| 0.14455 | 35.2
| 0.14447 | 35.3
| 0.14438 | 35.4
| 0.1443 | 35.5
| 0.14421 | 35.6
| 0.14413 | 35.7
| 0.14404 | 35.8
| 0.14396 | 35.9
| 0.14388 | 36
| 0.1438 | 36.1
| 0.14371 | 36.2
| 0.14363 | 36.3
| 0.14355 | 36.4
| 0.14347 | 36.5
| 0.14339 | 36.6
| 0.14331 | 36.7
| 0.14323 | 36.8
| 0.14315 | 36.9
| 0.14308 | 37
| 0.143 | 37.1
| 0.14292 | 37.2
| 0.14284 | 37.3
| 0.14277 | 37.4
| 0.14269 | 37.5
| 0.14261 | 37.6
| 0.14254 | 37.7
| 0.14246 | 37.8
| 0.14239 | 37.9
| 0.14231 | 38
| 0.14224 | 38.1
| 0.14217 | 38.2
| 0.14209 | 38.3
| 0.14202 | 38.4
| 0.14195 | 38.5
| 0.14188 | 38.6
| 0.1418 | 38.7
| 0.14173 | 38.8
| 0.14166 | 38.9
| 0.14159 | 39
| 0.14152 | 39.1
| 0.14145 | 39.2
| 0.14138 | 39.3
| 0.14131 | 39.4
| 0.14124 | 39.5
| 0.14117 | 39.6
| 0.1411 | 39.7
| 0.14103 | 39.8
| 0.14097 | 39.9
| 0.1409 | 40
| 0.14083 | 40.1
| 0.14077 | 40.2
| 0.1407 | 40.3
| 0.14063 | 40.4
| 0.14057 | 40.5
| 0.1405 | 40.6
| 0.14044 | 40.7
| 0.14037 | 40.8
| 0.14031 | 40.9
| 0.14024 | 41
| 0.14018 | 41.1
| 0.14011 | 41.2
| 0.14005 | 41.3
| 0.13999 | 41.4
| 0.13992 | 41.5
| 0.13986 | 41.6
| 0.1398 | 41.7
| 0.13974 | 41.8
| 0.13967 | 41.9
| 0.13961 | 42
| 0.13955 | 42.1
| 0.13949 | 42.2
| 0.13943 | 42.3
| 0.13937 | 42.4
| 0.13931 | 42.5
| 0.13925 | 42.6
| 0.13919 | 42.7
| 0.13913 | 42.8
| 0.13907 | 42.9
| 0.13901 | 43
| 0.13895 | 43.1
| 0.1389 | 43.2
| 0.13884 | 43.3
| 0.13878 | 43.4
| 0.13872 | 43.5
| 0.13866 | 43.6
| 0.13861 | 43.7
| 0.13855 | 43.8
| 0.13849 | 43.9
| 0.13844 | 44
| 0.13838 | 44.1
| 0.13833 | 44.2
| 0.13827 | 44.3
| 0.13821 | 44.4
| 0.13816 | 44.5
| 0.1381 | 44.6
| 0.13805 | 44.7
| 0.138 | 44.8
| 0.13794 | 44.9
| 0.13789 | 45
| 0.13783 | 45.1
| 0.13778 | 45.2
| 0.13773 | 45.3
| 0.13767 | 45.4
| 0.13762 | 45.5
| 0.13757 | 45.6
| 0.13752 | 45.7
| 0.13746 | 45.8
| 0.13741 | 45.9
| 0.13736 | 46
| 0.13731 | 46.1
| 0.13726 | 46.2
| 0.13721 | 46.3
| 0.13716 | 46.4
| 0.1371 | 46.5
| 0.13705 | 46.6
| 0.137 | 46.7
| 0.13695 | 46.8
| 0.1369 | 46.9
| 0.13685 | 47
| 0.1368 | 47.1
| 0.13676 | 47.2
| 0.13671 | 47.3
| 0.13666 | 47.4
| 0.13661 | 47.5
| 0.13656 | 47.6
| 0.13651 | 47.7
| 0.13646 | 47.8
| 0.13642 | 47.9
| 0.13637 | 48
| 0.13632 | 48.1
| 0.13627 | 48.2
| 0.13623 | 48.3
| 0.13618 | 48.4
| 0.13613 | 48.5
| 0.13609 | 48.6
| 0.13604 | 48.7
| 0.13599 | 48.8
| 0.13595 | 48.9
| 0.1359 | 49
| 0.13586 | 49.1
| 0.13581 | 49.2
| 0.13577 | 49.3
| 0.13572 | 49.4
| 0.13568 | 49.5
| 0.13563 | 49.6
| 0.13559 | 49.7
| 0.13554 | 49.8
| 0.1355 | 49.9
| 0.13545 | 50
| 0.13541 | 50.1
| 0.13537 | 50.2
| 0.13532 | 50.3
| 0.13528 | 50.4
| 0.13524 | 50.5
| 0.13519 | 50.6
| 0.13515 | 50.7
| 0.13511 | 50.8
| 0.13506 | 50.9
| 0.13502 | 51
| 0.13498 | 51.1
| 0.13494 | 51.2
| 0.13489 | 51.3
| 0.13485 | 51.4
| 0.13481 | 51.5
| 0.13477 | 51.6
| 0.13473 | 51.7
| 0.13469 | 51.8
| 0.13465 | 51.9
| 0.13461 | 52
| 0.13456 | 52.1
| 0.13452 | 52.2
| 0.13448 | 52.3
| 0.13444 | 52.4
| 0.1344 | 52.5
| 0.13436 | 52.6
| 0.13432 | 52.7
| 0.13428 | 52.8
| 0.13424 | 52.9
| 0.1342 | 53
| 0.13417 | 53.1
| 0.13413 | 53.2
| 0.13409 | 53.3
| 0.13405 | 53.4
| 0.13401 | 53.5
| 0.13397 | 53.6
| 0.13393 | 53.7
| 0.13389 | 53.8
| 0.13386 | 53.9
| 0.13382 | 54
| 0.13378 | 54.1
| 0.13374 | 54.2
| 0.1337 | 54.3
| 0.13367 | 54.4
| 0.13363 | 54.5
| 0.13359 | 54.6
| 0.13356 | 54.7
| 0.13352 | 54.8
| 0.13348 | 54.9
| 0.13345 | 55
| 0.13341 | 55.1
| 0.13337 | 55.2
| 0.13334 | 55.3
| 0.1333 | 55.4
| 0.13326 | 55.5
| 0.13323 | 55.6
| 0.13319 | 55.7
| 0.13316 | 55.8
| 0.13312 | 55.9
| 0.13309 | 56
| 0.13305 | 56.1
| 0.13301 | 56.2
| 0.13298 | 56.3
| 0.13294 | 56.4
| 0.13291 | 56.5
| 0.13287 | 56.6
| 0.13284 | 56.7
| 0.13281 | 56.8
| 0.13277 | 56.9
| 0.13274 | 57
| 0.1327 | 57.1
| 0.13267 | 57.2
| 0.13263 | 57.3
| 0.1326 | 57.4
| 0.13257 | 57.5
| 0.13253 | 57.6
| 0.1325 | 57.7
| 0.13247 | 57.8
| 0.13243 | 57.9
| 0.1324 | 58
| 0.13237 | 58.1
| 0.13233 | 58.2
| 0.1323 | 58.3
| 0.13227 | 58.4
| 0.13224 | 58.5
| 0.1322 | 58.6
| 0.13217 | 58.7
| 0.13214 | 58.8
| 0.13211 | 58.9
| 0.13208 | 59
| 0.13204 | 59.1
| 0.13201 | 59.2
| 0.13198 | 59.3
| 0.13195 | 59.4
| 0.13192 | 59.5
| 0.13189 | 59.6
| 0.13185 | 59.7
| 0.13182 | 59.8
| 0.13179 | 59.9
| 0.13176 | 60

## rational cal 4 red
dose cGy | dose Gy | measured scanner value | measured scanner value normalised | calculated scanner value | sum of squares
1950 | 19.5 | 9853.722 | 0.15036 | 0.15297 | 6.837e-06
1500 | 15 | 12040.34 | 0.18372 | 0.18613 | 5.7841e-06
1350 | 13.5 | 13257.07 | 0.20229 | 0.20055 | 3.0442e-06
1200 | 12 | 14459.44 | 0.22064 | 0.21734 | 1.0893e-05
1050 | 10.5 | 15622.84 | 0.23839 | 0.23714 | 1.5546e-06
900 | 9 | 17246.94 | 0.26317 | 0.26085 | 5.3667e-06
700 | 7 | 19556.78 | 0.29842 | 0.30085 | 5.9088e-06
400 | 4 | 25538.42 | 0.38969 | 0.39022 | 2.7526e-07
200 | 2 | 31724.05 | 0.48408 | 0.48608 | 4.0015e-06
100 | 1 | 36454.16 | 0.55625 | 0.55395 | 5.2986e-06
50 | 0.5 | 38866.02 | 0.59306 | 0.59548 | 5.8531e-06
0 | 0 | 42281.41 | 0.64517 | 0.64369 | 2.2077e-06
5.7024e-05
Calibration equation
X(D) = A + B/(Dose -C)
where X(D) is response expressed as 16-bit #/65535, D is dose in Gy and A, B and C are constants
Initial and calculated constants
|  | Red | Green
| a | -0.0033024
| b | 4.0179
| c | -6.2101
| X(D) | Dose(Gy)
| 0.10766 | 30
| 0.11414 | 28
| 0.12144 | 26
| 0.12969 | 24
| 0.13912 | 22
| 0.14436 | 21
| 0.14999 | 20
| 0.15607 | 19
| 0.16266 | 18
| 0.16981 | 17
| 0.1776 | 16
| 0.18613 | 15
| 0.1955 | 14
| 0.20585 | 13
| 0.21734 | 12
| 0.23016 | 11
| 0.24456 | 10
| 0.26085 | 9
| 0.27944 | 8
| 0.30085 | 7
| 0.32576 | 6
| 0.35511 | 5
| 0.39022 | 4
| 0.43294 | 3
| 0.48608 | 2
| 0.55395 | 1
| 0.64369 | 4.0146e-13
Formatted for RIT
DOSE | SCANNER | LOCATION
3000 | 7055.3036 | 0
2800 | 7480.4258 | 0
2600 | 7958.3416 | 0
2400 | 8499.5365 | 0
2200 | 9117.4692 | 0
2100 | 9460.5001 | 0
2000 | 9829.7066 | 0
1900 | 10228.2034 | 0
1800 | 10659.6201 | 0
1700 | 11128.2118 | 0
1600 | 11638.9998 | 0
1500 | 12197.9525 | 0
1400 | 12812.2195 | 0
1300 | 13490.4391 | 0
1200 | 14243.147 | 0
1100 | 15083.328 | 0
1000 | 16027.1705 | 0
900 | 17095.1207 | 0
800 | 18313.3797 | 0
700 | 19716.0828 | 0
600 | 21348.5475 | 0
500 | 23272.2619 | 0
400 | 25572.8032 | 0
300 | 28372.9153 | 0
200 | 31855.144 | 0
100 | 36303.3073 | 0
4.0146e-11 | 42184.0353 | 0

## rational cal 5 green
dose cGy | dose Gy | measured scanner value | measured scanner value normalised | calculated scanner value | sum of squares
1950 | 19.5 | 9862.721 | 0.1505 | 0.15175 | 1.5845e-06
1500 | 15 | 11716.7 | 0.17879 | 0.18515 | 4.0564e-05
1350 | 13.5 | 13249.63 | 0.20218 | 0.19967 | 6.2893e-06
1200 | 12 | 14444.21 | 0.2204 | 0.21657 | 1.4723e-05
1050 | 10.5 | 15632.5 | 0.23854 | 0.23649 | 4.1828e-06
900 | 9 | 17256.61 | 0.26332 | 0.26033 | 8.9177e-06
700 | 7 | 19562.53 | 0.29851 | 0.30051 | 4.0296e-06
400 | 4 | 25528.14 | 0.38953 | 0.39016 | 3.8737e-07
200 | 2 | 31686.83 | 0.48351 | 0.4861 | 6.7152e-06
100 | 1 | 36422.06 | 0.55577 | 0.5539 | 3.4768e-06
50 | 0.5 | 38867.81 | 0.59308 | 0.59532 | 5.0114e-06
0 | 0 | 42283.39 | 0.6452 | 0.64336 | 3.3796e-06
9.9261e-05
Calibration equation
X(D) = A + B/(Dose -C)
where X(D) is response expressed as 16-bit #/65535, D is dose in Gy and A, B and C are constants
Initial and calculated constants
|  | Red | Green
| a | -0.0060334
| b | 4.0644
| c | -6.2587
| X(D) | Dose(Gy)
| 0.10606 | 30
| 0.11261 | 28
| 0.11996 | 26
| 0.12829 | 24
| 0.1378 | 22
| 0.14307 | 21
| 0.14875 | 20
| 0.15488 | 19
| 0.16151 | 18
| 0.16871 | 17
| 0.17657 | 16
| 0.18515 | 15
| 0.19459 | 14
| 0.20501 | 13
| 0.21657 | 12
| 0.22947 | 11
| 0.24395 | 10
| 0.26033 | 9
| 0.27901 | 8
| 0.30051 | 7
| 0.32552 | 6
| 0.35497 | 5
| 0.39016 | 4
| 0.43295 | 3
| 0.4861 | 2
| 0.5539 | 1
| 0.64336 | 4.0146e-13
Formatted for RIT
DOSE | SCANNER | LOCATION
3000 | 6950.7225 | 0
2800 | 7379.5839 | 0
2600 | 7861.623 | 0
2400 | 8407.3843 | 0
2200 | 9030.3978 | 0
2100 | 9376.1878 | 0
2000 | 9748.315 | 0
1900 | 10149.9074 | 0
1800 | 10584.6089 | 0
1700 | 11056.6901 | 0
1600 | 11571.1889 | 0
1500 | 12134.0912 | 0
1400 | 12752.5648 | 0
1300 | 13435.2663 | 0
1200 | 14192.7486 | 0
1100 | 15038.0105 | 0
1000 | 15987.2487 | 0
900 | 17060.9059 | 0
800 | 18285.1596 | 0
700 | 19694.0846 | 0
600 | 21332.8743 | 0
500 | 23262.7783 | 0
400 | 25568.9285 | 0
300 | 28373.2353 | 0
200 | 31856.6552 | 0
100 | 36299.8624 | 0
4.0146e-11 | 42162.9119 | 0

## unif profiles px val
centre across |  |  | near to scan array, across |  | opposite to scan array, across |  | centre, down |  | left side, down |  | right side, down |  | centre, down repeated |  | left side, down, repeated |  | right side , down, repeated
0 | 43717.8242 | 21817.1758 | 0 | 43242.6953 | 0 | 45511 | 0 | 43949 | 0 | 43265.0469 | 0 | 45790.5898 | 0 | 44048.9609 | 0 | 43328.2383 | 0 | 45893.4805
1 | 43660.1758 | 21874.8242 | 1 | 43235.6094 | 1 | 45398.2188 | 1 | 43955.2734 | 1 | 43268.2734 | 1 | 45776.5469 | 1 | 44011.4805 | 1 | 43321.6797 | 1 | 45858.3594
2 | 43717.6953 | 21817.3047 | 2 | 43307.1758 | 2 | 45373.7383 | 2 | 43948.4531 | 2 | 43226.1836 | 2 | 45768.4531 | 2 | 44005.9219 | 2 | 43315.1992 | 2 | 45851.1992
3 | 43674.3047 | 21860.6953 | 3 | 43329.6953 | 3 | 45315.6953 | 3 | 44004.3164 | 3 | 43255.1367 | 3 | 45740.7266 | 3 | 44022.8789 | 3 | 43334.6406 | 3 | 45866.8008
4 | 43622.5664 | 21912.4336 | 4 | 43259.1758 | 4 | 45232.3047 | 4 | 43955.2266 | 4 | 43280.3633 | 4 | 45733.8633 | 4 | 44082 | 4 | 43351.2383 | 4 | 45872.4805
5 | 43713.6523 | 21821.3477 | 5 | 43340 | 5 | 45290.2188 | 5 | 43921.7266 | 5 | 43291.5 | 5 | 45749.7734 | 5 | 44090.2812 | 5 | 43385.5586 | 5 | 45888.0781
6 | 43699.6094 | 21835.3906 | 6 | 43375.3047 | 6 | 45273.1289 | 6 | 43875.1367 | 6 | 43325.7734 | 6 | 45716.0469 | 6 | 44102.8789 | 6 | 43363.6406 | 6 | 45904.4414
7 | 43605 | 21930 | 7 | 43253.3906 | 7 | 45160.3047 | 7 | 43964.5 | 7 | 43329.2266 | 7 | 45688.4102 | 7 | 44142.6797 | 7 | 43402.5586 | 7 | 45936
8 | 43586.2188 | 21948.7812 | 8 | 43233.043 | 8 | 45121.2188 | 8 | 43819.9102 | 8 | 43354.4531 | 8 | 45667.5898 | 8 | 44069.8789 | 8 | 43373.7617 | 8 | 45897.6016
9 | 43641.8711 | 21893.1289 | 9 | 43276.7383 | 9 | 45071.4766 | 9 | 43888 | 9 | 43360.7266 | 9 | 45679.5469 | 9 | 44086.2812 | 9 | 43380.2383 | 9 | 45879.8008
10 | 43665.0859 | 21869.9141 | 10 | 43337.8711 | 10 | 45116.3906 | 10 | 43918.4531 | 10 | 43301.7734 | 10 | 45693.1836 | 10 | 44048.8789 | 10 | 43423.6406 | 10 | 45870.8789
11 | 43615.2617 | 21919.7383 | 11 | 43345.6094 | 11 | 45145.5664 | 11 | 44009.5 | 11 | 43341.5898 | 11 | 45686.4531 | 11 | 44017.4805 | 11 | 43472.8008 | 11 | 45822.2812
12 | 43660.1758 | 21874.8242 | 12 | 43324.3477 | 12 | 45087.9141 | 12 | 43947.1836 | 12 | 43313.8164 | 12 | 45638.6836 | 12 | 43993.3203 | 12 | 43458.4414 | 12 | 45802.6016
13 | 43601.4766 | 21933.5234 | 13 | 43278.0859 | 13 | 45006.7812 | 13 | 43907.8633 | 13 | 43338.4102 | 13 | 45652.1367 | 13 | 43971.6406 | 13 | 43490.0781 | 13 | 45805.3984
14 | 43639.043 | 21895.957 | 14 | 43317.2188 | 14 | 45073.5234 | 14 | 43963.4531 | 14 | 43280.8164 | 14 | 45619 | 14 | 43997.3984 | 14 | 43432.8398 | 14 | 45772.3203
15 | 43605.3906 | 21929.6094 | 15 | 43304.0859 | 15 | 45015.1758 | 15 | 44000.3164 | 15 | 43292 | 15 | 45602.8164 | 15 | 43922.7617 | 15 | 43373.1992 | 15 | 45745.6016
16 | 43531.043 | 22003.957 | 16 | 43284.043 | 16 | 44981.7812 | 16 | 43999.0898 | 16 | 43248.5 | 16 | 45597.0469 | 16 | 43847.6406 | 16 | 43344.8789 | 16 | 45720.5586
17 | 43474.3906 | 22060.6094 | 17 | 43338.043 | 17 | 44965.1758 | 17 | 44005.4102 | 17 | 43249.0469 | 17 | 45559.2266 | 17 | 43876.7617 | 17 | 43293.7188 | 17 | 45724
18 | 43457.0859 | 22077.9141 | 18 | 43290.7812 | 18 | 44943.4766 | 18 | 43963.8164 | 18 | 43248.3164 | 18 | 45612.2266 | 18 | 43926.0781 | 18 | 43244.3984 | 18 | 45723.2383
19 | 43458.7812 | 22076.2188 | 19 | 43236 | 19 | 44922.4766 | 19 | 43921.4102 | 19 | 43257.1836 | 19 | 45652.7734 | 19 | 43908.7617 | 19 | 43223.1602 | 19 | 45690.4414
20 | 43550 | 21985 | 20 | 43280.5664 | 20 | 44970 | 20 | 43890.5469 | 20 | 43271.3164 | 20 | 45710.5898 | 20 | 43878.5195 | 20 | 43161.3203 | 20 | 45709.6406
21 | 43505.6523 | 22029.3477 | 21 | 43262.4766 | 21 | 44938.4336 | 21 | 43865.5 | 21 | 43303.1367 | 21 | 45714.3164 | 21 | 43864.1602 | 21 | 43146.7188 | 21 | 45670.1992
22 | 43495.043 | 22039.957 | 22 | 43268.1289 | 22 | 44893.043 | 22 | 43819 | 22 | 43295.5469 | 22 | 45728.3633 | 22 | 43880.4805 | 22 | 43148.3203 | 22 | 45640.0781
23 | 43510.3477 | 22024.6523 | 23 | 43264.3906 | 23 | 44866.6523 | 23 | 43760.8164 | 23 | 43286 | 23 | 45710.5898 | 23 | 43916.3984 | 23 | 43132 | 23 | 45648.3203
24 | 43500.3477 | 22034.6523 | 24 | 43265.043 | 24 | 44849.8711 | 24 | 43754.6367 | 24 | 43285.9102 | 24 | 45640.0898 | 24 | 43894.8789 | 24 | 43148.8789 | 24 | 45625.0781
25 | 43470 | 22065 | 25 | 43239.8711 | 25 | 44806.8711 | 25 | 43775.3164 | 25 | 43265.2266 | 25 | 45573.9531 | 25 | 43917.5195 | 25 | 43186.2383 | 25 | 45627.5586
26 | 43414.043 | 22120.957 | 26 | 43204 | 26 | 44788.043 | 26 | 43767.0469 | 26 | 43287.5469 | 26 | 45547.7266 | 26 | 43913.1602 | 26 | 43223.9219 | 26 | 45618.1602
27 | 43431.957 | 22103.043 | 27 | 43219.6953 | 27 | 44796.3047 | 27 | 43805.2734 | 27 | 43260.6367 | 27 | 45506.9102 | 27 | 43860.8008 | 27 | 43237.8008 | 27 | 45565.7188
28 | 43478.6523 | 22056.3477 | 28 | 43278.8711 | 28 | 44818.3906 | 28 | 43824.1836 | 28 | 43273.9102 | 28 | 45482.6367 | 28 | 43813.0391 | 28 | 43274.7188 | 28 | 45550.1992
29 | 43437.9141 | 22097.0859 | 29 | 43284.3047 | 29 | 44780.4336 | 29 | 43813.7734 | 29 | 43284.2266 | 29 | 45492 | 29 | 43888.9219 | 29 | 43270.8789 | 29 | 45558.3594
30 | 43457.043 | 22077.957 | 30 | 43296.6094 | 30 | 44769.043 | 30 | 43856.9531 | 30 | 43324.7734 | 30 | 45470 | 30 | 43784.0781 | 30 | 43226.2383 | 30 | 45585.3594
31 | 43469.0859 | 22065.9141 | 31 | 43289.8711 | 31 | 44750.3906 | 31 | 43900.9102 | 31 | 43320.5898 | 31 | 45463.9102 | 31 | 43861.6406 | 31 | 43175.7188 | 31 | 45545.5586
32 | 43413.8242 | 22121.1758 | 32 | 43257.5664 | 32 | 44656.8711 | 32 | 43866.9102 | 32 | 43362.6836 | 32 | 45467.5898 | 32 | 43937.1211 | 32 | 43225.6797 | 32 | 45538.6797
33 | 43461.1758 | 22073.8242 | 33 | 43325.8242 | 33 | 44695.8711 | 33 | 43880.8164 | 33 | 43393 | 33 | 45453.3633 | 33 | 43989.1211 | 33 | 43197 | 33 | 45543.5586
34 | 43393.8711 | 22141.1289 | 34 | 43235 | 34 | 44598.2617 | 34 | 43858.3164 | 34 | 43342.1836 | 34 | 45470.4102 | 34 | 43923.1992 | 34 | 43270 | 34 | 45612.1211
35 | 43413.957 | 22121.043 | 35 | 43262.2188 | 35 | 44648.6953 | 35 | 43865.6836 | 35 | 43311.2266 | 35 | 45479.4531 | 35 | 43868.3203 | 35 | 43254.8789 | 35 | 45668.8008
36 | 43432.2617 | 22102.7383 | 36 | 43243.0859 | 36 | 44658.1289 | 36 | 43847.3633 | 36 | 43257.0898 | 36 | 45449 | 36 | 43887 | 36 | 43239.1602 | 36 | 45661.1992
37 | 43455 | 22080 | 37 | 43318.957 | 37 | 44704 | 37 | 43851.4531 | 37 | 43279.9102 | 37 | 45431.0469 | 37 | 43934.0781 | 37 | 43193.8008 | 37 | 45677.9609
38 | 43450.957 | 22084.043 | 38 | 43276.2188 | 38 | 44635.8711 | 38 | 43849.6367 | 38 | 43217.5469 | 38 | 45391 | 38 | 43922.5195 | 38 | 43168.2383 | 38 | 45642.8008
39 | 43515.1758 | 22019.8242 | 39 | 43334.8242 | 39 | 44660.043 | 39 | 43834.1836 | 39 | 43167.3633 | 39 | 45379.3633 | 39 | 43951.4414 | 39 | 43207.2812 | 39 | 45585.7188
40 | 43468.6094 | 22066.3906 | 40 | 43280.7383 | 40 | 44579.6523 | 40 | 43854.3633 | 40 | 43233.8633 | 40 | 45407.8164 | 40 | 43911.4414 | 40 | 43178.7188 | 40 | 45541.5586
41 | 43504.6094 | 22030.3906 | 41 | 43273.7383 | 41 | 44620.3047 | 41 | 43803.1367 | 41 | 43275.7266 | 41 | 45398.0898 | 41 | 43843.2383 | 41 | 43211.3594 | 41 | 45518.6797
42 | 43528.6094 | 22006.3906 | 42 | 43318.957 | 42 | 44625.8242 | 42 | 43779.2266 | 42 | 43305.0469 | 42 | 45429.6367 | 42 | 43825.9219 | 42 | 43235.8008 | 42 | 45474.6016
43 | 43525.3477 | 22009.6523 | 43 | 43350.9141 | 43 | 44619 | 43 | 43833.9531 | 43 | 43285.9102 | 43 | 45418.5898 | 43 | 43826.6016 | 43 | 43225.0391 | 43 | 45463.7188
44 | 43502.8711 | 22032.1289 | 44 | 43331.7812 | 44 | 44604.9141 | 44 | 43863.8164 | 44 | 43261.5898 | 44 | 45397.2266 | 44 | 43776.8008 | 44 | 43227.8789 | 44 | 45478.6797
45 | 43465.5664 | 22069.4336 | 45 | 43298.3906 | 45 | 44555.6953 | 45 | 43870.5898 | 45 | 43227 | 45 | 45380.7266 | 45 | 43723.6797 | 45 | 43232.4414 | 45 | 45423.1992
46 | 43455.8711 | 22079.1289 | 46 | 43294.7812 | 46 | 44538.9141 | 46 | 43873.9531 | 46 | 43269.5 | 46 | 45362 | 46 | 43735.5195 | 46 | 43224.5195 | 46 | 45377.3594
47 | 43471.5234 | 22063.4766 | 47 | 43307.8711 | 47 | 44561.4766 | 47 | 43832.5469 | 47 | 43239.7734 | 47 | 45345.5898 | 47 | 43735.9609 | 47 | 43227.9219 | 47 | 45406.2812
48 | 43443.8711 | 22091.1289 | 48 | 43272.4766 | 48 | 44561.3477 | 48 | 43809.3164 | 48 | 43236.5898 | 48 | 45353.4531 | 48 | 43712.4414 | 48 | 43218.6016 | 48 | 45411
49 | 43447 | 22088 | 49 | 43289.1289 | 49 | 44547.7812 | 49 | 43826.7266 | 49 | 43259.8164 | 49 | 45386.2266 | 49 | 43751 | 49 | 43206.7188 | 49 | 45405
50 | 43429.3047 | 22105.6953 | 50 | 43301.957 | 50 | 44506.0859 | 50 | 43864.5 | 50 | 43280.2266 | 50 | 45423.0898 | 50 | 43771.9219 | 50 | 43192.5195 | 50 | 45428.6016
51 | 43438.4766 | 22096.5234 | 51 | 43328.8711 | 51 | 44515.2188 | 51 | 43851.6367 | 51 | 43270.9531 | 51 | 45410.9102 | 51 | 43795.7617 | 51 | 43228.5195 | 51 | 45412.0781
52 | 43390.5234 | 22144.4766 | 52 | 43334.7383 | 52 | 44527.1289 | 52 | 43843.5898 | 52 | 43283 | 52 | 45400.5 | 52 | 43811.8398 | 52 | 43245.5586 | 52 | 45394.6016
53 | 43388.7812 | 22146.2188 | 53 | 43363.3906 | 53 | 44526.3477 | 53 | 43842.1367 | 53 | 43288.5469 | 53 | 45360.7266 | 53 | 43885.0391 | 53 | 43283.9609 | 53 | 45369.9609
54 | 43386.2188 | 22148.7812 | 54 | 43393.3906 | 54 | 44490.9141 | 54 | 43844.7734 | 54 | 43278.9531 | 54 | 45359.5898 | 54 | 43846.9609 | 54 | 43300.7617 | 54 | 45356.7188
55 | 43378.1289 | 22156.8711 | 55 | 43360.7383 | 55 | 44494.9141 | 55 | 43879.7734 | 55 | 43318.6367 | 55 | 45401 | 55 | 43838.6406 | 55 | 43260.1211 | 55 | 45354.3594
56 | 43375.1289 | 22159.8711 | 56 | 43317.4336 | 56 | 44502.6094 | 56 | 43918.4531 | 56 | 43345.0469 | 56 | 45344.1836 | 56 | 43810.4414 | 56 | 43210.9219 | 56 | 45366.3594
57 | 43379.2617 | 22155.7383 | 57 | 43304.4766 | 57 | 44488.5234 | 57 | 43826.6367 | 57 | 43358.3633 | 57 | 45320.2266 | 57 | 43814.7188 | 57 | 43166.0781 | 57 | 45380.7188
58 | 43427.1289 | 22107.8711 | 58 | 43382.3906 | 58 | 44517.6953 | 58 | 43856.0898 | 58 | 43404.9102 | 58 | 45363.7266 | 58 | 43844.8008 | 58 | 43183.1992 | 58 | 45347.2812
59 | 43432.957 | 22102.043 | 59 | 43416.2617 | 59 | 44569.2617 | 59 | 43865.3164 | 59 | 43373.0898 | 59 | 45369.4102 | 59 | 43838.0391 | 59 | 43129.2383 | 59 | 45331.4414
60 | 43299.3477 | 22235.6523 | 60 | 43380.5234 | 60 | 44490.1289 | 60 | 43795.7266 | 60 | 43331.8633 | 60 | 45354.0469 | 60 | 43823.6406 | 60 | 43133.3203 | 60 | 45319.6016
61 | 43432.2617 | 22102.7383 | 61 | 43359.7812 | 61 | 44532.3477 | 61 | 43796.3633 | 61 | 43233.2266 | 61 | 45346.7266 | 61 | 43797.1211 | 61 | 43195.9219 | 61 | 45316.8789
62 | 43399.7812 | 22135.2188 | 62 | 43324.4336 | 62 | 44488.1758 | 62 | 43826.8633 | 62 | 43248.1836 | 62 | 45388.7266 | 62 | 43830.7617 | 62 | 43223.2812 | 62 | 45333.8398
63 | 43397.2617 | 22137.7383 | 63 | 43427.5234 | 63 | 44461.2188 | 63 | 43801.6367 | 63 | 43114.7734 | 63 | 45391.2734 | 63 | 43791.1992 | 63 | 43226 | 63 | 45321.5195
64 | 43411.6953 | 22123.3047 | 64 | 43395.4336 | 64 | 44446.3906 | 64 | 43789.2734 | 64 | 43216.0898 | 64 | 45344.4102 | 64 | 43741.1992 | 64 | 43185.8008 | 64 | 45327.6406
65 | 43410.6953 | 22124.3047 | 65 | 43425.4336 | 65 | 44506.5664 | 65 | 43824.1836 | 65 | 43177.8633 | 65 | 45318.5 | 65 | 43773.8398 | 65 | 43158.8398 | 65 | 45344.5586
66 | 43425.3047 | 22109.6953 | 66 | 43430.3906 | 66 | 44495.7383 | 66 | 43876.9531 | 66 | 43174.6367 | 66 | 45245.4531 | 66 | 43775.6016 | 66 | 43147.6797 | 66 | 45339.3594
67 | 43427.9141 | 22107.0859 | 67 | 43417.3477 | 67 | 44455.1758 | 67 | 43835.1836 | 67 | 43188.4102 | 67 | 45177.6836 | 67 | 43794.6016 | 67 | 43187.4414 | 67 | 45346.6016
68 | 43438.6953 | 22096.3047 | 68 | 43406.8242 | 68 | 44414.1758 | 68 | 43869.8164 | 68 | 43179.2734 | 68 | 45146.8164 | 68 | 43792.1992 | 68 | 43165.4805 | 68 | 45317.8789
69 | 43440.9141 | 22094.0859 | 69 | 43415.6094 | 69 | 44417.957 | 69 | 43884.5898 | 69 | 43254.4102 | 69 | 45184.6836 | 69 | 43754.4805 | 69 | 43158.0781 | 69 | 45316.4805
70 | 43409.4336 | 22125.5664 | 70 | 43431.1758 | 70 | 44359.0859 | 70 | 43886 | 70 | 43363.4102 | 70 | 45218.5898 | 70 | 43745.0391 | 70 | 43199.3594 | 70 | 45332.1602
71 | 43391.6523 | 22143.3477 | 71 | 43444.3477 | 71 | 44341.3477 | 71 | 43877.4102 | 71 | 43445.0469 | 71 | 45246.0898 | 71 | 43788.1211 | 71 | 43239.5586 | 71 | 45308.0781
72 | 43396.8711 | 22138.1289 | 72 | 43464.1758 | 72 | 44360.6953 | 72 | 43892.0469 | 72 | 43464 | 72 | 45290.8633 | 72 | 43841.3594 | 72 | 43235.1211 | 72 | 45301.1992
73 | 43407.5234 | 22127.4766 | 73 | 43453.1758 | 73 | 44337.8242 | 73 | 43953 | 73 | 43494.6836 | 73 | 45311.8633 | 73 | 43838.6797 | 73 | 43258.4805 | 73 | 45303.0781
74 | 43396.2617 | 22138.7383 | 74 | 43466.8711 | 74 | 44338.8711 | 74 | 44000.8164 | 74 | 43517.9531 | 74 | 45308.0469 | 74 | 43824.0781 | 74 | 43269.3203 | 74 | 45309.3203
75 | 43400.957 | 22134.043 | 75 | 43530.2188 | 75 | 44292.3906 | 75 | 44046.5898 | 75 | 43510.6836 | 75 | 45282.2266 | 75 | 43803.9609 | 75 | 43241.2383 | 75 | 45320.3594
76 | 43427 | 22108 | 76 | 43540.8242 | 76 | 44275.0859 | 76 | 43985.2266 | 76 | 43478.1836 | 76 | 45283.9102 | 76 | 43802.0391 | 76 | 43243.0781 | 76 | 45322.1602
77 | 43412.8711 | 22122.1289 | 77 | 43571.8711 | 77 | 44357.9141 | 77 | 44003.1836 | 77 | 43421 | 77 | 45289.1367 | 77 | 43819.9609 | 77 | 43260.6016 | 77 | 45328.0781
78 | 43424.4336 | 22110.5664 | 78 | 43521.4766 | 78 | 44281.8711 | 78 | 44083.8633 | 78 | 43401.4531 | 78 | 45298.0898 | 78 | 43871.6406 | 78 | 43293.6406 | 78 | 45328.0391
79 | 43405.3906 | 22129.6094 | 79 | 43483.5234 | 79 | 44217.2188 | 79 | 43982.7734 | 79 | 43373.5 | 79 | 45292.6836 | 79 | 43818.5195 | 79 | 43302.0781 | 79 | 45304.4805
80 | 43311.5664 | 22223.4336 | 80 | 43513.0859 | 80 | 44204.6523 | 80 | 43998.4531 | 80 | 43370.5 | 80 | 45337.5 | 80 | 43807.0781 | 80 | 43287.7188 | 80 | 45261.3984
81 | 43360.043 | 22174.957 | 81 | 43526.1289 | 81 | 44257.7812 | 81 | 43851.1836 | 81 | 43364.5469 | 81 | 45287.2734 | 81 | 43822.8398 | 81 | 43268 | 81 | 45240.8398
82 | 43463.6094 | 22071.3906 | 82 | 43613.043 | 82 | 44358.5234 | 82 | 43933.4531 | 82 | 43370.3633 | 82 | 45295.2734 | 82 | 43783.8008 | 82 | 43215.9609 | 82 | 45149.5586
83 | 43448.6094 | 22086.3906 | 83 | 43630.6094 | 83 | 44385.043 | 83 | 43904.2734 | 83 | 43361.0469 | 83 | 45276.4531 | 83 | 43801.3984 | 83 | 43191.4414 | 83 | 45158.2383
84 | 43343.3906 | 22191.6094 | 84 | 43497.3047 | 84 | 44248.4766 | 84 | 43894.8633 | 84 | 43360.1836 | 84 | 45302.1836 | 84 | 43792.8398 | 84 | 43132.3203 | 84 | 45145.6406
85 | 43358.3047 | 22176.6953 | 85 | 43480.9141 | 85 | 44177.1289 | 85 | 43865.6836 | 85 | 43323.4531 | 85 | 45252.5 | 85 | 43749.2383 | 85 | 43163.3984 | 85 | 45177.3594
86 | 43338.8711 | 22196.1289 | 86 | 43500.5234 | 86 | 44194.043 | 86 | 43907.4531 | 86 | 43335.1367 | 86 | 45241.6367 | 86 | 43750.8398 | 86 | 43163.8789 | 86 | 45204.3594
87 | 43365.5234 | 22169.4766 | 87 | 43548.043 | 87 | 44197.3047 | 87 | 43912.6836 | 87 | 43316.7734 | 87 | 45175.3633 | 87 | 43804.3594 | 87 | 43121.6406 | 87 | 45236.5195
88 | 43356.043 | 22178.957 | 88 | 43549.3906 | 88 | 44199.6094 | 88 | 43873.3633 | 88 | 43302.8633 | 88 | 45194.1836 | 88 | 43854.8398 | 88 | 43150.6797 | 88 | 45265.5195
89 | 43314.6953 | 22220.3047 | 89 | 43497.6953 | 89 | 44103.4766 | 89 | 43844.0469 | 89 | 43267.5898 | 89 | 45189.6836 | 89 | 43801.0781 | 89 | 43165.8398 | 89 | 45270.5195
90 | 43357.7383 | 22177.2617 | 90 | 43552.8711 | 90 | 44151 | 90 | 43766.3164 | 90 | 43233.0898 | 90 | 45209.5898 | 90 | 43819.8789 | 90 | 43230.6406 | 90 | 45250.7617
91 | 43350.3477 | 22184.6523 | 91 | 43546 | 91 | 44146.2617 | 91 | 43713.5898 | 91 | 43249.9531 | 91 | 45146.3164 | 91 | 43829.3984 | 91 | 43306.2812 | 91 | 45252.0781
92 | 43340 | 22195 | 92 | 43497.7812 | 92 | 44064.1289 | 92 | 43729.5469 | 92 | 43254.4531 | 92 | 45137.2734 | 92 | 43835.9609 | 92 | 43347.2383 | 92 | 45220.5586
93 | 43331.3477 | 22203.6523 | 93 | 43475.5234 | 93 | 44041.0859 | 93 | 43723.6836 | 93 | 43286.3633 | 93 | 45134.7734 | 93 | 43813.1992 | 93 | 43389.4414 | 93 | 45211.9219
94 | 43350.7812 | 22184.2188 | 94 | 43466.4766 | 94 | 44044.0859 | 94 | 43713.5469 | 94 | 43285.2734 | 94 | 45114.8633 | 94 | 43829.7188 | 94 | 43417.5195 | 94 | 45238.1602
95 | 43396.0859 | 22138.9141 | 95 | 43485 | 95 | 44055.3906 | 95 | 43683.2266 | 95 | 43289.8164 | 95 | 45078.7734 | 95 | 43901.7617 | 95 | 43432.8789 | 95 | 45278.8008
96 | 43366.5234 | 22168.4766 | 96 | 43512.8711 | 96 | 44061.3047 | 96 | 43683.5898 | 96 | 43301.2266 | 96 | 45050.2266 | 96 | 43982.1602 | 96 | 43422.5195 | 96 | 45290.3984
97 | 43309.8242 | 22225.1758 | 97 | 43500.2188 | 97 | 44006.957 | 97 | 43700.8164 | 97 | 43297.9102 | 97 | 45039.8164 | 97 | 44009.4805 | 97 | 43383.7188 | 97 | 45263.7617
98 | 43299.1758 | 22235.8242 | 98 | 43470.7812 | 98 | 44027.8711 | 98 | 43740.5 | 98 | 43289.4102 | 98 | 45039.3633 | 98 | 43972.0781 | 98 | 43334.2812 | 98 | 45224.7188
99 | 43303.957 | 22231.043 | 99 | 43438.043 | 99 | 44032.9141 | 99 | 43789.3633 | 99 | 43294.2734 | 99 | 45042.0469 | 99 | 43949.9219 | 99 | 43302.4805 | 99 | 45269.8008
100 | 43300.6523 | 22234.3477 | 100 | 43449.957 | 100 | 44031.6953 | 100 | 43741.6367 | 100 | 43249.8633 | 100 | 45060.9102 | 100 | 43980.3203 | 100 | 43335.8008 | 100 | 45220.1602
101 | 43291.3477 | 22243.6523 | 101 | 43452.1758 | 101 | 44026.6094 | 101 | 43752 | 101 | 43275.7734 | 101 | 45077.6836 | 101 | 43919.1211 | 101 | 43348.0781 | 101 | 45159.4414
102 | 43239.957 | 22295.043 | 102 | 43435.2188 | 102 | 44000.957 | 102 | 43856.5898 | 102 | 43295.9531 | 102 | 45052.3164 | 102 | 43855.5586 | 102 | 43308.5586 | 102 | 45155.6406
103 | 43252.4336 | 22282.5664 | 103 | 43399.5234 | 103 | 43984.2617 | 103 | 43862.1367 | 103 | 43272.0898 | 103 | 45033.6836 | 103 | 43861.5586 | 103 | 43314.1992 | 103 | 45129.6016
104 | 43275.6523 | 22259.3477 | 104 | 43438.5664 | 104 | 44020.3047 | 104 | 43834.4531 | 104 | 43306.2734 | 104 | 45077.5 | 104 | 43862.8008 | 104 | 43315.9219 | 104 | 45141.1211
105 | 43244.4766 | 22290.5234 | 105 | 43412.5234 | 105 | 43959.3047 | 105 | 43840.0469 | 105 | 43294.3633 | 105 | 45090 | 105 | 43841.9609 | 105 | 43285.6406 | 105 | 45139.1992
106 | 43253.7383 | 22281.2617 | 106 | 43425.1758 | 106 | 43933.8242 | 106 | 43841.6836 | 106 | 43333.6367 | 106 | 45101.7266 | 106 | 43854.2812 | 106 | 43257.6797 | 106 | 45153.7617
107 | 43256.9141 | 22278.0859 | 107 | 43455.8711 | 107 | 44041.5234 | 107 | 43825.2266 | 107 | 43318.6836 | 107 | 45086.1367 | 107 | 43791.7188 | 107 | 43269.7188 | 107 | 45113.7617
108 | 43303.4336 | 22231.5664 | 108 | 43477.7812 | 108 | 44043.7383 | 108 | 43772.5898 | 108 | 43297.5 | 108 | 45066.4531 | 108 | 43859.7617 | 108 | 43242.7617 | 108 | 45114.8008
109 | 43208.2617 | 22326.7383 | 109 | 43375.7383 | 109 | 43977.1758 | 109 | 43712.5898 | 109 | 43290.5 | 109 | 45082.3164 | 109 | 43849.5195 | 109 | 43248.2812 | 109 | 45090.8789
110 | 43263.2188 | 22271.7812 | 110 | 43433.8242 | 110 | 43988.3906 | 110 | 43750.0898 | 110 | 43245.8164 | 110 | 45126.4102 | 110 | 43813.7188 | 110 | 43206.2383 | 110 | 45037.0781
111 | 43263.957 | 22271.043 | 111 | 43451.2188 | 111 | 43999.4766 | 111 | 43745.6836 | 111 | 43287.8633 | 111 | 45120 | 111 | 43787.4414 | 111 | 43211.3203 | 111 | 45023.0781
112 | 43302.4766 | 22232.5234 | 112 | 43517.8242 | 112 | 44045.8711 | 112 | 43674.2734 | 112 | 43306.2734 | 112 | 45087.2266 | 112 | 43716.1992 | 112 | 43177.4414 | 112 | 45020.8398
113 | 43307.4336 | 22227.5664 | 113 | 43517.1289 | 113 | 44051.4766 | 113 | 43723 | 113 | 43269.5 | 113 | 45015.3633 | 113 | 43676.7188 | 113 | 43195.6016 | 113 | 44998.0781
114 | 43275.2617 | 22259.7383 | 114 | 43490.6953 | 114 | 43987.6523 | 114 | 43721.1367 | 114 | 43233.9102 | 114 | 45004.9102 | 114 | 43639.0391 | 114 | 43207.8789 | 114 | 44992
115 | 43267 | 22268 | 115 | 43470.5664 | 115 | 43994.3906 | 115 | 43698.7266 | 115 | 43189.8633 | 115 | 45039.5469 | 115 | 43631.7188 | 115 | 43204.3203 | 115 | 45000.7617
116 | 43316.2617 | 22218.7383 | 116 | 43521.957 | 116 | 44036.9141 | 116 | 43697 | 116 | 43165.5469 | 116 | 45015.5898 | 116 | 43663.6797 | 116 | 43207.7188 | 116 | 45036.4414
117 | 43357.043 | 22177.957 | 117 | 43502.5664 | 117 | 44034.6953 | 117 | 43673.2734 | 117 | 43179.9102 | 117 | 45026.3633 | 117 | 43693.6797 | 117 | 43214.9219 | 117 | 45025.6406
118 | 43382.6523 | 22152.3477 | 118 | 43485.1758 | 118 | 44031.3906 | 118 | 43706.2734 | 118 | 43204.3164 | 118 | 45005.8633 | 118 | 43725 | 118 | 43214.3203 | 118 | 44988.5586
119 | 43392.3906 | 22142.6094 | 119 | 43497.1289 | 119 | 44027.4766 | 119 | 43804.7734 | 119 | 43138.2266 | 119 | 44965.4102 | 119 | 43694.1602 | 119 | 43245.6797 | 119 | 45022.9219
120 | 43350.9141 | 22184.0859 | 120 | 43509.5664 | 120 | 43990 | 120 | 43848.5469 | 120 | 43135.0469 | 120 | 44934.1367 | 120 | 43651.7617 | 120 | 43236.1211 | 120 | 45013.0391
121 | 43365.4336 | 22169.5664 | 121 | 43496.9141 | 121 | 43984.8711 | 121 | 43856.1367 | 121 | 43216.9102 | 121 | 44912.6367 | 121 | 43682.6797 | 121 | 43201.8008 | 121 | 45031.9609
122 | 43305.3047 | 22229.6953 | 122 | 43462.3477 | 122 | 43940.3906 | 122 | 43912.1836 | 122 | 43293.9102 | 122 | 44932.7266 | 122 | 43799.8789 | 122 | 43229.4805 | 122 | 45069.4414
123 | 43314.3047 | 22220.6953 | 123 | 43507.2617 | 123 | 44030.957 | 123 | 43813.4531 | 123 | 43351.0898 | 123 | 44938.9102 | 123 | 43762.1211 | 123 | 43236.3203 | 123 | 45036.1992
124 | 43400.3047 | 22134.6953 | 124 | 43570.3477 | 124 | 44047.7383 | 124 | 43880.4102 | 124 | 43372.1836 | 124 | 44965.0469 | 124 | 43759 | 124 | 43216.8398 | 124 | 45053.4805
125 | 43370.3047 | 22164.6953 | 125 | 43594.3906 | 125 | 44032.8711 | 125 | 43816 | 125 | 43355.1836 | 125 | 44949.2266 | 125 | 43800.6797 | 125 | 43236 | 125 | 45079.8398
126 | 43305.5234 | 22229.4766 | 126 | 43585.043 | 126 | 43982.7383 | 126 | 43828.4531 | 126 | 43350.2734 | 126 | 44949.3633 | 126 | 43812.6406 | 126 | 43246.9219 | 126 | 45080.5586
127 | 43356.8242 | 22178.1758 | 127 | 43573.043 | 127 | 44059.6953 | 127 | 43896.2734 | 127 | 43352.9531 | 127 | 44984.6836 | 127 | 43772.6016 | 127 | 43260.5586 | 127 | 45015.2383
128 | 43399.6094 | 22135.3906 | 128 | 43616.9141 | 128 | 44080.2188 | 128 | 43908.1836 | 128 | 43379.3633 | 128 | 44996.9102 | 128 | 43817.0391 | 128 | 43251.9219 | 128 | 44955.1211
129 | 43364.7812 | 22170.2188 | 129 | 43526.7383 | 129 | 44050.6953 | 129 | 43877.8633 | 129 | 43376.5469 | 129 | 44969.8164 | 129 | 43804.0391 | 129 | 43228.3594 | 129 | 44982.5586
130 | 43460.4766 | 22074.5234 | 130 | 43642.3906 | 130 | 44174.6094 | 130 | 43936.2266 | 130 | 43391.4531 | 130 | 44947.8164 | 130 | 43765 | 130 | 43247.2812 | 130 | 45003.4414
131 | 43376.957 | 22158.043 | 131 | 43628.4766 | 131 | 44063.3477 | 131 | 43845.9102 | 131 | 43450.9102 | 131 | 44954.8633 | 131 | 43726.3594 | 131 | 43226.2812 | 131 | 44969.5586
132 | 43403.6953 | 22131.3047 | 132 | 43623.2617 | 132 | 44089.5664 | 132 | 43782.1836 | 132 | 43402.1836 | 132 | 44977.3164 | 132 | 43711.6016 | 132 | 43243.1602 | 132 | 44978.1211
133 | 43408.043 | 22126.957 | 133 | 43642.0859 | 133 | 44048.3477 | 133 | 43777.5 | 133 | 43394.4102 | 133 | 44972.8164 | 133 | 43665.2812 | 133 | 43253.0781 | 133 | 44983.6016
134 | 43449.3047 | 22085.6953 | 134 | 43639.5664 | 134 | 44060.043 | 134 | 43756.0469 | 134 | 43357.3633 | 134 | 44957.7734 | 134 | 43642.6406 | 134 | 43206.0781 | 134 | 44939.2812
135 | 43478.8711 | 22056.1289 | 135 | 43627.8711 | 135 | 44088.5234 | 135 | 43729.9531 | 135 | 43300.2734 | 135 | 44952.0898 | 135 | 43652.6016 | 135 | 43218.8789 | 135 | 44902.8789
136 | 43422.0859 | 22112.9141 | 136 | 43642.2188 | 136 | 44062.3477 | 136 | 43809.4102 | 136 | 43276.3164 | 136 | 44928 | 136 | 43649.2812 | 136 | 43161.7188 | 136 | 44884.9609
137 | 43453 | 22082 | 137 | 43633.043 | 137 | 44080 | 137 | 43776.9102 | 137 | 43303.4102 | 137 | 44908.6836 | 137 | 43641.8398 | 137 | 43142.4805 | 137 | 44902.6016
138 | 43409.8242 | 22125.1758 | 138 | 43648.6094 | 138 | 44040.6094 | 138 | 43723.5469 | 138 | 43310.4531 | 138 | 44917.5469 | 138 | 43622.6406 | 138 | 43095.8789 | 138 | 44918.8398
139 | 43391.7383 | 22143.2617 | 139 | 43706.3477 | 139 | 44046.3477 | 139 | 43685.0469 | 139 | 43265.6367 | 139 | 44895.7266 | 139 | 43663.0781 | 139 | 43108.0391 | 139 | 44913.4414
140 | 43392.7812 | 22142.2188 | 140 | 43699.7383 | 140 | 44040.1289 | 140 | 43695.0898 | 140 | 43289.0898 | 140 | 44936.2734 | 140 | 43693.3594 | 140 | 43079.8008 | 140 | 44906.8398
141 | 43332.6094 | 22202.3906 | 141 | 43603 | 141 | 43986.3477 | 141 | 43716.0898 | 141 | 43276.8633 | 141 | 44969.6367 | 141 | 43666.1602 | 141 | 43063.6406 | 141 | 44916.5586
142 | 43387.6523 | 22147.3477 | 142 | 43700.7812 | 142 | 44062.3906 | 142 | 43682.8633 | 142 | 43297.4102 | 142 | 44958.6836 | 142 | 43751.9219 | 142 | 43144.6016 | 142 | 44914.1992
143 | 43426.8711 | 22108.1289 | 143 | 43734.9141 | 143 | 44071.043 | 143 | 43716.6836 | 143 | 43319.6836 | 143 | 44959.1836 | 143 | 43767.6406 | 143 | 43213.7188 | 143 | 44930.3984
144 | 43419.2188 | 22115.7812 | 144 | 43724.2188 | 144 | 44030.3477 | 144 | 43738.1836 | 144 | 43285.9102 | 144 | 44915.2266 | 144 | 43731.7188 | 144 | 43259.1211 | 144 | 44935.5195
145 | 43400.6953 | 22134.3047 | 145 | 43673.8242 | 145 | 43954.2188 | 145 | 43786 | 145 | 43232.1367 | 145 | 44874.0898 | 145 | 43864.6016 | 145 | 43264 | 145 | 44945.4805
146 | 43415.3047 | 22119.6953 | 146 | 43715.5234 | 146 | 43949.2188 | 146 | 43716.9102 | 146 | 43237 | 146 | 44859.9531 | 146 | 43820.3203 | 146 | 43303.3203 | 146 | 44950.6406
147 | 43344.6523 | 22190.3477 | 147 | 43649.3906 | 147 | 43917.6094 | 147 | 43648.0469 | 147 | 43202.2266 | 147 | 44819.0469 | 147 | 43865.6016 | 147 | 43307.1992 | 147 | 44919.2812
148 | 43396.957 | 22138.043 | 148 | 43709.3906 | 148 | 43957.5234 | 148 | 43672.3633 | 148 | 43225.0898 | 148 | 44789.2266 | 148 | 43793.9609 | 148 | 43286.7617 | 148 | 44920.1992
149 | 43320.3906 | 22214.6094 | 149 | 43631.7383 | 149 | 43890.3906 | 149 | 43694.4102 | 149 | 43307 | 149 | 44791.2266 | 149 | 43825.9219 | 149 | 43296.3594 | 149 | 44922.3203
150 | 43378.6523 | 22156.3477 | 150 | 43662.8242 | 150 | 43933.8242 | 150 | 43718.1367 | 150 | 43287.7266 | 150 | 44791.9531 | 150 | 43827.0391 | 150 | 43357.2383 | 150 | 44907.9219
151 | 43410.3477 | 22124.6523 | 151 | 43652.6523 | 151 | 43970.3047 | 151 | 43717.2734 | 151 | 43241.8633 | 151 | 44802.7734 | 151 | 43806.5586 | 151 | 43362.5195 | 151 | 44895.0781
152 | 43400.043 | 22134.957 | 152 | 43659.8711 | 152 | 43986.3047 | 152 | 43628.5898 | 152 | 43201.5469 | 152 | 44774.5 | 152 | 43748.9609 | 152 | 43330.3203 | 152 | 44842.1602
153 | 43371.1289 | 22163.8711 | 153 | 43649.2617 | 153 | 43899.5664 | 153 | 43689.2266 | 153 | 43194.4102 | 153 | 44775.4531 | 153 | 43787.3203 | 153 | 43314.2383 | 153 | 44858.5195
154 | 43296.1289 | 22238.8711 | 154 | 43698.3047 | 154 | 43930.1758 | 154 | 43711.0469 | 154 | 43200.1836 | 154 | 44829.5 | 154 | 43774.4805 | 154 | 43290.6016 | 154 | 44859.9219
155 | 43391.0859 | 22143.9141 | 155 | 43692.957 | 155 | 43939.8242 | 155 | 43765.3633 | 155 | 43189.3633 | 155 | 44799 | 155 | 43690.0781 | 155 | 43283.2812 | 155 | 44859.5586
156 | 43415.1758 | 22119.8242 | 156 | 43691.6523 | 156 | 43950.8242 | 156 | 43739.8164 | 156 | 43215.8633 | 156 | 44774.8633 | 156 | 43677.4414 | 156 | 43251.7188 | 156 | 44903.0391
157 | 43381.5664 | 22153.4336 | 157 | 43641.4336 | 157 | 43946.3047 | 157 | 43657.0898 | 157 | 43204.3633 | 157 | 44815.6367 | 157 | 43695.1211 | 157 | 43210.1992 | 157 | 44884.6016
158 | 43352.4336 | 22182.5664 | 158 | 43687.6523 | 158 | 43939.4336 | 158 | 43652.2266 | 158 | 43195.9102 | 158 | 44856.7266 | 158 | 43709.9219 | 158 | 43226.9219 | 158 | 44874.1992
159 | 43321.6094 | 22213.3906 | 159 | 43707 | 159 | 43959.9141 | 159 | 43660.5 | 159 | 43136.8633 | 159 | 44809.3633 | 159 | 43692.9609 | 159 | 43250.3594 | 159 | 44854.3203
160 | 43384.0859 | 22150.9141 | 160 | 43726 | 160 | 43977.6094 | 160 | 43689.2266 | 160 | 43168.7734 | 160 | 44793.3633 | 160 | 43695.9219 | 160 | 43208 | 160 | 44832.9609
161 | 43385 | 22150 | 161 | 43713.5234 | 161 | 43970.1758 | 161 | 43696.8164 | 161 | 43162.2266 | 161 | 44816.2734 | 161 | 43667.6406 | 161 | 43217.0391 | 161 | 44804.4414
162 | 43356.7812 | 22178.2188 | 162 | 43688.1289 | 162 | 43941.8711 | 162 | 43751 | 162 | 43190.8633 | 162 | 44776.2266 | 162 | 43641.1602 | 162 | 43243.5586 | 162 | 44787.9609
163 | 43372.6953 | 22162.3047 | 163 | 43732.4766 | 163 | 43981.9141 | 163 | 43753.6367 | 163 | 43215.0898 | 163 | 44745.4531 | 163 | 43635.9609 | 163 | 43278.3203 | 163 | 44757.8789
164 | 43395.0859 | 22139.9141 | 164 | 43770.3477 | 164 | 43960.4766 | 164 | 43684.4531 | 164 | 43242.0469 | 164 | 44775.7734 | 164 | 43633.3984 | 164 | 43261.4414 | 164 | 44754.7617
165 | 43422.4336 | 22112.5664 | 165 | 43794.6094 | 165 | 43997.6094 | 165 | 43754.2734 | 165 | 43234 | 165 | 44749.2266 | 165 | 43661.8398 | 165 | 43210.3984 | 165 | 44760.1602
166 | 43439.6523 | 22095.3477 | 166 | 43750.8711 | 166 | 44021.0859 | 166 | 43820.2734 | 166 | 43230.3633 | 166 | 44763.5469 | 166 | 43737.9609 | 166 | 43150.7188 | 166 | 44749.0391
167 | 43422.3477 | 22112.6523 | 167 | 43666.4766 | 167 | 43967.1758 | 167 | 43762.8164 | 167 | 43236.0469 | 167 | 44777.4102 | 167 | 43804.0781 | 167 | 43118.3203 | 167 | 44739.7188
168 | 43392.2188 | 22142.7812 | 168 | 43693.0859 | 168 | 43923.9141 | 168 | 43703.5469 | 168 | 43247.7266 | 168 | 44780.8164 | 168 | 43720 | 168 | 43116.2812 | 168 | 44739.1602
169 | 43401.6523 | 22133.3477 | 169 | 43760.4336 | 169 | 43903.043 | 169 | 43629.3164 | 169 | 43290.4102 | 169 | 44785.1367 | 169 | 43622.8008 | 169 | 43172.3203 | 169 | 44760.8398
170 | 43383.7812 | 22151.2188 | 170 | 43797.2188 | 170 | 43858.3047 | 170 | 43598.4102 | 170 | 43309.8164 | 170 | 44754.9531 | 170 | 43626.1992 | 170 | 43245.4805 | 170 | 44766.5586
171 | 43354.3047 | 22180.6953 | 171 | 43803.9141 | 171 | 43831.2617 | 171 | 43606.2266 | 171 | 43303.5898 | 171 | 44706.5 | 171 | 43672.8789 | 171 | 43226.9609 | 171 | 44711.9609
172 | 43319.4336 | 22215.5664 | 172 | 43650.2617 | 172 | 43819.3047 | 172 | 43566.5469 | 172 | 43280.5898 | 172 | 44698.9531 | 172 | 43653.5195 | 172 | 43187.3984 | 172 | 44776.7188
173 | 43391.043 | 22143.957 | 173 | 43698.9141 | 173 | 43871.8242 | 173 | 43485.8633 | 173 | 43208.9531 | 173 | 44688.2266 | 173 | 43666.0781 | 173 | 43195.9219 | 173 | 44771.1211
174 | 43366.9141 | 22168.0859 | 174 | 43681.6094 | 174 | 43845.0859 | 174 | 43500.3164 | 174 | 43157.4102 | 174 | 44681.6367 | 174 | 43616.1211 | 174 | 43158.7188 | 174 | 44739.5586
175 | 43377.5234 | 22157.4766 | 175 | 43713.2188 | 175 | 43846.7383 | 175 | 43638.5898 | 175 | 43137.5898 | 175 | 44659.3164 | 175 | 43645.0781 | 175 | 43158.1992 | 175 | 44733.7188
176 | 43450.6953 | 22084.3047 | 176 | 43729.9141 | 176 | 43895.957 | 176 | 43636.5 | 176 | 43127.7266 | 176 | 44635.7266 | 176 | 43638.9219 | 176 | 43125.7188 | 176 | 44782.0781
177 | 43323.3047 | 22211.6953 | 177 | 43664.957 | 177 | 43802.9141 | 177 | 43681.9531 | 177 | 43148.2266 | 177 | 44647.5469 | 177 | 43703.3984 | 177 | 43150.5586 | 177 | 44749.5195
178 | 43333.1758 | 22201.8242 | 178 | 43667.0859 | 178 | 43809.957 | 178 | 43628.0469 | 178 | 43218.6367 | 178 | 44660.0469 | 178 | 43664.7188 | 178 | 43080.9609 | 178 | 44731.8008
179 | 43184.7383 | 22350.2617 | 179 | 43725.2617 | 179 | 43766.1758 | 179 | 43608.2266 | 179 | 43291.5469 | 179 | 44689.0898 | 179 | 43617.8789 | 179 | 43142.4414 | 179 | 44737.7188
180 | 43352.6523 | 22182.3477 | 180 | 43761.2188 | 180 | 43847.5664 | 180 | 43606.9531 | 180 | 43329.0469 | 180 | 44700.5469 | 180 | 43624.3203 | 180 | 43097.2383 | 180 | 44711.2812
181 | 43367.7812 | 22167.2188 | 181 | 43770.5234 | 181 | 43861.043 | 181 | 43562.4102 | 181 | 43326.1836 | 181 | 44726.3633 | 181 | 43605.6016 | 181 | 43104.0781 | 181 | 44729.0781
182 | 43347.3906 | 22187.6094 | 182 | 43725.043 | 182 | 43817.6953 | 182 | 43525.9531 | 182 | 43232.6367 | 182 | 44794.7734 | 182 | 43642.5195 | 182 | 43111.3594 | 182 | 44740.5195
183 | 43323.8711 | 22211.1289 | 183 | 43708.1758 | 183 | 43808.957 | 183 | 43550.2734 | 183 | 43150.9102 | 183 | 44766.3164 | 183 | 43669.5586 | 183 | 43138.8789 | 183 | 44716.6797
184 | 43377.6953 | 22157.3047 | 184 | 43775.9141 | 184 | 43868.8242 | 184 | 43573.8633 | 184 | 43200.3633 | 184 | 44734.4531 | 184 | 43676.3984 | 184 | 43136.5195 | 184 | 44690.3984
185 | 43340.1289 | 22194.8711 | 185 | 43741.2188 | 185 | 43846.6953 | 185 | 43574.1367 | 185 | 43141.5469 | 185 | 44683.3164 | 185 | 43660.4414 | 185 | 43150.4805 | 185 | 44694.1211
186 | 43320.7812 | 22214.2188 | 186 | 43746.3477 | 186 | 43818.957 | 186 | 43580.2266 | 186 | 43158.3164 | 186 | 44626.3164 | 186 | 43660.2812 | 186 | 43154.4414 | 186 | 44667.8398
187 | 43312.2188 | 22222.7812 | 187 | 43746.1289 | 187 | 43809.6523 | 187 | 43596.7266 | 187 | 43188.4102 | 187 | 44633.5 | 187 | 43685.2383 | 187 | 43155.6406 | 187 | 44657.3984
188 | 43360.4766 | 22174.5234 | 188 | 43709.6094 | 188 | 43848.7812 | 188 | 43628.9531 | 188 | 43194.2734 | 188 | 44583.6367 | 188 | 43705.3984 | 188 | 43181.2812 | 188 | 44655.2812
189 | 43395.5664 | 22139.4336 | 189 | 43699.6094 | 189 | 43896.1289 | 189 | 43620.4102 | 189 | 43248.3633 | 189 | 44538.7266 | 189 | 43680.9609 | 189 | 43195.0391 | 189 | 44634.7617
190 | 43374.0859 | 22160.9141 | 190 | 43685.3906 | 190 | 43866.5664 | 190 | 43588.9102 | 190 | 43255.5898 | 190 | 44526.1836 | 190 | 43612.8789 | 190 | 43229.1211 | 190 | 44591.8398
191 | 43406.7383 | 22128.2617 | 191 | 43753.4766 | 191 | 43867.6523 | 191 | 43568.6367 | 191 | 43245.9531 | 191 | 44516.3164 | 191 | 43547.6797 | 191 | 43258.7617 | 191 | 44599.4805
192 | 43404.1289 | 22130.8711 | 192 | 43753.2617 | 192 | 43860.1289 | 192 | 43591.8164 | 192 | 43218.4102 | 192 | 44542.6836 | 192 | 43532.3984 | 192 | 43213.2383 | 192 | 44628.4805
193 | 43413.9141 | 22121.0859 | 193 | 43727.3477 | 193 | 43863.7383 | 193 | 43564.8164 | 193 | 43214.2734 | 193 | 44588.7266 | 193 | 43548.5195 | 193 | 43180.4414 | 193 | 44567.6797
194 | 43444.0859 | 22090.9141 | 194 | 43780.4336 | 194 | 43850.957 | 194 | 43560.6836 | 194 | 43218.5 | 194 | 44604.1367 | 194 | 43498.1602 | 194 | 43133.0781 | 194 | 44605.8008
195 | 43421.6523 | 22113.3477 | 195 | 43769.5664 | 195 | 43887.5664 | 195 | 43586.2734 | 195 | 43231.6836 | 195 | 44629.6367 | 195 | 43481.1211 | 195 | 43055.6016 | 195 | 44633.8008
196 | 43423.7383 | 22111.2617 | 196 | 43752.6523 | 196 | 43829.4336 | 196 | 43638.6367 | 196 | 43272 | 196 | 44631.0898 | 196 | 43445.3203 | 196 | 43051 | 196 | 44680.6016
197 | 43481.8242 | 22053.1758 | 197 | 43761.043 | 197 | 43829.0859 | 197 | 43640.0898 | 197 | 43315 | 197 | 44602.7734 | 197 | 43547.1992 | 197 | 43085.6016 | 197 | 44751.3594
198 | 43446.043 | 22088.957 | 198 | 43827.8711 | 198 | 43831.6523 | 198 | 43590.6836 | 198 | 43388.9531 | 198 | 44594.8633 | 198 | 43534.2812 | 198 | 43127.2383 | 198 | 44721.4805
199 | 43440.8242 | 22094.1758 | 199 | 43809.6953 | 199 | 43869.5234 | 199 | 43556.7266 | 199 | 43381.6367 | 199 | 44644.5469 | 199 | 43596.3984 | 199 | 43167.1602 | 199 | 44706.2383
200 | 43414.7383 | 22120.2617 | 200 | 43732.3477 | 200 | 43857.6953 | 200 | 43567.4531 | 200 | 43340.4531 | 200 | 44670.7734 | 200 | 43539.8789 | 200 | 43235.0781 | 200 | 44647.8008
201 | 43428.0859 | 22106.9141 | 201 | 43795.957 | 201 | 43841.3047 | 201 | 43498.2734 | 201 | 43323.6367 | 201 | 44651.5 | 201 | 43534.4805 | 201 | 43205.7617 | 201 | 44605.3594
202 | 43424.1289 | 22110.8711 | 202 | 43786.0859 | 202 | 43822.5664 | 202 | 43456.5469 | 202 | 43297.4102 | 202 | 44660.5898 | 202 | 43527.3594 | 202 | 43216.4414 | 202 | 44588.3203
203 | 43380.6953 | 22154.3047 | 203 | 43712.6094 | 203 | 43722.4766 | 203 | 43507.9531 | 203 | 43263.5898 | 203 | 44649.5469 | 203 | 43483.9609 | 203 | 43116.8789 | 203 | 44551.7188
204 | 43486.8242 | 22048.1758 | 204 | 43825.1758 | 204 | 43791.3047 | 204 | 43572.3633 | 204 | 43231.1836 | 204 | 44620.7734 | 204 | 43460.3984 | 204 | 43050.3984 | 204 | 44489.7617
205 | 43488.6523 | 22046.3477 | 205 | 43806.7383 | 205 | 43822.3477 | 205 | 43605.2266 | 205 | 43235.7734 | 205 | 44635.9102 | 205 | 43490.3203 | 205 | 43104.2383 | 205 | 44467
206 | 43487.957 | 22047.043 | 206 | 43803.1758 | 206 | 43837.6523 | 206 | 43624.0898 | 206 | 43232.6836 | 206 | 44637.6367 | 206 | 43513.8008 | 206 | 43096.6016 | 206 | 44461.8398
207 | 43463.3047 | 22071.6953 | 207 | 43768.1289 | 207 | 43826.6094 | 207 | 43632.3633 | 207 | 43232.8633 | 207 | 44589.5 | 207 | 43524.5195 | 207 | 43101.1602 | 207 | 44501.5586
208 | 43545.7812 | 21989.2188 | 208 | 43788.1758 | 208 | 43904.3906 | 208 | 43666.5 | 208 | 43236.0898 | 208 | 44595.8633 | 208 | 43539 | 208 | 43090.7188 | 208 | 44546.4414
209 | 43510.6094 | 22024.3906 | 209 | 43813.9141 | 209 | 43894.3906 | 209 | 43691.7266 | 209 | 43165.6836 | 209 | 44614.3164 | 209 | 43542.8398 | 209 | 43150.3203 | 209 | 44568.0391
210 | 43470.0859 | 22064.9141 | 210 | 43796.2188 | 210 | 43825.9141 | 210 | 43642.3164 | 210 | 43148.5 | 210 | 44555.1367 | 210 | 43548.8789 | 210 | 43166.2812 | 210 | 44594
211 | 43461.5234 | 22073.4766 | 211 | 43774.4766 | 211 | 43844.7812 | 211 | 43621 | 211 | 43173.1836 | 211 | 44542.2266 | 211 | 43540.8398 | 211 | 43182 | 211 | 44603.7617
212 | 43492.6094 | 22042.3906 | 212 | 43835.2188 | 212 | 43896.3477 | 212 | 43584.3164 | 212 | 43233.5 | 212 | 44556.4102 | 212 | 43535.8789 | 212 | 43173.3203 | 212 | 44576.9609
213 | 43513.3047 | 22021.6953 | 213 | 43842.6953 | 213 | 43898.043 | 213 | 43592.5 | 213 | 43215.8164 | 213 | 44629.9531 | 213 | 43528.7617 | 213 | 43183.8008 | 213 | 44543
214 | 43480.6953 | 22054.3047 | 214 | 43818.1289 | 214 | 43866.0859 | 214 | 43580.5 | 214 | 43223 | 214 | 44655.4531 | 214 | 43544.1992 | 214 | 43164.7617 | 214 | 44589.4805
215 | 43520.6094 | 22014.3906 | 215 | 43866.2617 | 215 | 43914.9141 | 215 | 43538.0898 | 215 | 43263.8633 | 215 | 44650.0898 | 215 | 43515.2812 | 215 | 43140.5586 | 215 | 44599.8398
216 | 43497.7383 | 22037.2617 | 216 | 43901.3477 | 216 | 43942.0859 | 216 | 43554.7734 | 216 | 43275.6367 | 216 | 44664.1367 | 216 | 43507.9609 | 216 | 43145.3984 | 216 | 44584.8789
217 | 43510.6094 | 22024.3906 | 217 | 43899.6523 | 217 | 43966.0859 | 217 | 43569.0469 | 217 | 43267.5898 | 217 | 44627 | 217 | 43564.6016 | 217 | 43175.5195 | 217 | 44592.2812
218 | 43531.7383 | 22003.2617 | 218 | 43864.8711 | 218 | 43926.043 | 218 | 43512.9102 | 218 | 43241.9102 | 218 | 44585 | 218 | 43577.6797 | 218 | 43237.9609 | 218 | 44591.3984
219 | 43513.9141 | 22021.0859 | 219 | 43835.2617 | 219 | 43904.3477 | 219 | 43490.0469 | 219 | 43231.7266 | 219 | 44589.9531 | 219 | 43590.0391 | 219 | 43289.0391 | 219 | 44589.7617
220 | 43555.1758 | 21979.8242 | 220 | 43850.0859 | 220 | 43940.5664 | 220 | 43469.3164 | 220 | 43196.5898 | 220 | 44559.1367 | 220 | 43538.0391 | 220 | 43262.5586 | 220 | 44588.7617
221 | 43558.6523 | 21976.3477 | 221 | 43844.2617 | 221 | 43964.6953 | 221 | 43480.4102 | 221 | 43201.0469 | 221 | 44516.2734 | 221 | 43497.0781 | 221 | 43219.4805 | 221 | 44569.6016
222 | 43508.0859 | 22026.9141 | 222 | 43824.0859 | 222 | 43862.4766 | 222 | 43429.9102 | 222 | 43195.0469 | 222 | 44475.0898 | 222 | 43467.3594 | 222 | 43226.9219 | 222 | 44544
223 | 43550.5664 | 21984.4336 | 223 | 43875.4336 | 223 | 43903.0859 | 223 | 43427.8164 | 223 | 43179.1367 | 223 | 44453.6836 | 223 | 43428.6406 | 223 | 43200.8398 | 223 | 44561.7188
224 | 43584.6523 | 21950.3477 | 224 | 43908.9141 | 224 | 43920.957 | 224 | 43434.9102 | 224 | 43218 | 224 | 44459.0469 | 224 | 43434.8398 | 224 | 43174.6406 | 224 | 44527.7188
225 | 43533.6953 | 22001.3047 | 225 | 43800.9141 | 225 | 43857.957 | 225 | 43439.4102 | 225 | 43158.8633 | 225 | 44489 | 225 | 43466.3203 | 225 | 43152.8398 | 225 | 44478.5586
226 | 43622.5234 | 21912.4766 | 226 | 43872.1289 | 226 | 43880.8711 | 226 | 43491.7266 | 226 | 43129.5 | 226 | 44473.7734 | 226 | 43510.1992 | 226 | 43130.8398 | 226 | 44481.5586
227 | 43618.7383 | 21916.2617 | 227 | 43811.8242 | 227 | 43908.5234 | 227 | 43457.5469 | 227 | 43177.2266 | 227 | 44457.0898 | 227 | 43559.5586 | 227 | 43165.3984 | 227 | 44512.9609
228 | 43555.6953 | 21979.3047 | 228 | 43784.7812 | 228 | 43819.5234 | 228 | 43439.7266 | 228 | 43217.1836 | 228 | 44479.6367 | 228 | 43593.1992 | 228 | 43153.2383 | 228 | 44573.9219
229 | 43555.3906 | 21979.6094 | 229 | 43862.3906 | 229 | 43828.4766 | 229 | 43441.1367 | 229 | 43273.6836 | 229 | 44528.7266 | 229 | 43608.6016 | 229 | 43152.8008 | 229 | 44588.1992
230 | 43505.6523 | 22029.3477 | 230 | 43824.5664 | 230 | 43850.6094 | 230 | 43405.1367 | 230 | 43255.9102 | 230 | 44530.5898 | 230 | 43617.5195 | 230 | 43099.9609 | 230 | 44604.3594
231 | 43547.6953 | 21987.3047 | 231 | 43874.3477 | 231 | 43873.3906 | 231 | 43400.2734 | 231 | 43285.9531 | 231 | 44489 | 231 | 43627.6797 | 231 | 43120.8789 | 231 | 44602.3984
232 | 43555.5234 | 21979.4766 | 232 | 43912.6094 | 232 | 43915.6094 | 232 | 43425.9531 | 232 | 43287.3164 | 232 | 44492.3633 | 232 | 43610.7188 | 232 | 43130.9609 | 232 | 44568.7188
233 | 43516.6094 | 22018.3906 | 233 | 43890.2188 | 233 | 43903.2188 | 233 | 43468.7266 | 233 | 43308.5469 | 233 | 44429.6836 | 233 | 43578.0391 | 233 | 43189.1211 | 233 | 44538.3203
234 | 43557.0859 | 21977.9141 | 234 | 43931 | 234 | 43921.3047 | 234 | 43513.6367 | 234 | 43242.0469 | 234 | 44431.5469 | 234 | 43530.1211 | 234 | 43176.0781 | 234 | 44535.8398
235 | 43487.7383 | 22047.2617 | 235 | 43861.4766 | 235 | 43859.6094 | 235 | 43542.9531 | 235 | 43240.8633 | 235 | 44443.8633 | 235 | 43539.8008 | 235 | 43196.8008 | 235 | 44503.2812
236 | 43484.5234 | 22050.4766 | 236 | 43889.8711 | 236 | 43838.6094 | 236 | 43520.6367 | 236 | 43195.4102 | 236 | 44444.9531 | 236 | 43542.7188 | 236 | 43232.2383 | 236 | 44470.9609
237 | 43518.5664 | 22016.4336 | 237 | 43934.3906 | 237 | 43850.6523 | 237 | 43525.8633 | 237 | 43193.7266 | 237 | 44429.8633 | 237 | 43538.3594 | 237 | 43232.0781 | 237 | 44448.5195
238 | 43547.8242 | 21987.1758 | 238 | 43941.8711 | 238 | 43868.1758 | 238 | 43502.6836 | 238 | 43185.5 | 238 | 44423.8164 | 238 | 43542.8008 | 238 | 43208.1211 | 238 | 44438.5195
239 | 43494.957 | 22040.043 | 239 | 43959.7383 | 239 | 43862.4766 | 239 | 43500.6367 | 239 | 43166.2734 | 239 | 44394.8164 | 239 | 43510.3203 | 239 | 43187.9609 | 239 | 44426.9219
240 | 43498 | 22037 | 240 | 43890.7383 | 240 | 43847.3047 | 240 | 43508.0898 | 240 | 43132.3633 | 240 | 44376.7266 | 240 | 43483.3203 | 240 | 43186.1602 | 240 | 44439.3203
241 | 43537.043 | 21997.957 | 241 | 43905.5664 | 241 | 43848.5234 | 241 | 43485.6367 | 241 | 43112.0898 | 241 | 44377.5469 | 241 | 43440.7617 | 241 | 43167.4805 | 241 | 44448.5586
242 | 43504.043 | 22030.957 | 242 | 43864.1289 | 242 | 43856.5234 | 242 | 43482.1367 | 242 | 43111.6836 | 242 | 44425.9102 | 242 | 43442.3594 | 242 | 43152.3594 | 242 | 44429.8789
243 | 43542.1758 | 21992.8242 | 243 | 43854.2188 | 243 | 43897.5664 | 243 | 43456.0898 | 243 | 43159.6836 | 243 | 44464.0469 | 243 | 43431.3984 | 243 | 43144.4414 | 243 | 44450.2383
244 | 43527.8242 | 22007.1758 | 244 | 43937 | 244 | 43936.1289 | 244 | 43428.8633 | 244 | 43214.9102 | 244 | 44470.8633 | 244 | 43380.3984 | 244 | 43131.6016 | 244 | 44482.6797
245 | 43557.5664 | 21977.4336 | 245 | 43867.9141 | 245 | 43942.8242 | 245 | 43472.1836 | 245 | 43223.6836 | 245 | 44496.4531 | 245 | 43386.8789 | 245 | 43130.2812 | 245 | 44447
246 | 43542.9141 | 21992.0859 | 246 | 43811.6953 | 246 | 43931.2188 | 246 | 43467.4531 | 246 | 43234.2734 | 246 | 44491.2266 | 246 | 43390.8398 | 246 | 43085.4805 | 246 | 44438.9609
247 | 43560.043 | 21974.957 | 247 | 43788.6523 | 247 | 43910.2188 | 247 | 43486.2266 | 247 | 43193.6836 | 247 | 44313.6367 | 247 | 43406.5195 | 247 | 43045.6797 | 247 | 44454.9609
248 | 43547.3047 | 21987.6953 | 248 | 43899.8711 | 248 | 43935.0859 | 248 | 43441.6836 | 248 | 43223.9102 | 248 | 44322.6367 | 248 | 43429.9219 | 248 | 43077.6406 | 248 | 44395.3984
249 | 43572.3477 | 21962.6523 | 249 | 43872.8711 | 249 | 43924.2617 | 249 | 43421.6836 | 249 | 43244.4102 | 249 | 44403.6367 | 249 | 43434.9609 | 249 | 43144.1602 | 249 | 44396.7188
250 | 43512.7383 | 22022.2617 | 250 | 43856.6094 | 250 | 43897.043 | 250 | 43450.2266 | 250 | 43266.3164 | 250 | 44419.5469 | 250 | 43433.9609 | 250 | 43196.0781 | 250 | 44395.4805
251 | 43571.9141 | 21963.0859 | 251 | 43845.8242 | 251 | 43942.6523 | 251 | 43496.9102 | 251 | 43249.0469 | 251 | 44393.1836 | 251 | 43403.3594 | 251 | 43191.4805 | 251 | 44400.6016
252 | 43497.1289 | 22037.8711 | 252 | 43795.2188 | 252 | 43960 | 252 | 43491.8164 | 252 | 43299.5898 | 252 | 44414.7734 | 252 | 43363.8398 | 252 | 43180.7617 | 252 | 44397.2383
253 | 43506.8711 | 22028.1289 | 253 | 43801.7812 | 253 | 43960.5664 | 253 | 43517.3633 | 253 | 43263.1836 | 253 | 44396.2266 | 253 | 43375.5586 | 253 | 43199.1602 | 253 | 44412.8789
254 | 43523.6094 | 22011.3906 | 254 | 43865.6094 | 254 | 43981.1289 | 254 | 43532.0469 | 254 | 43219.4531 | 254 | 44443.3633 | 254 | 43395.2812 | 254 | 43244.3984 | 254 | 44377.4805
255 | 43590.6523 | 21944.3477 | 255 | 43883.8242 | 255 | 44007.043 | 255 | 43545.8633 | 255 | 43289.1367 | 255 | 44439.4102 | 255 | 43417.2812 | 255 | 43209.4414 | 255 | 44350.3984
256 | 43531.3477 | 22003.6523 | 256 | 43885.8711 | 256 | 43958.1289 | 256 | 43520.2266 | 256 | 43325.1367 | 256 | 44448.6367 | 256 | 43438.5195 | 256 | 43168.1992 | 256 | 44367.6406
257 | 43551.2617 | 21983.7383 | 257 | 43838.4766 | 257 | 43958.5664 | 257 | 43473.7266 | 257 | 43357.1836 | 257 | 44439.8164 | 257 | 43468.2812 | 257 | 43168.6016 | 257 | 44392.7617
258 | 43583.7812 | 21951.2188 | 258 | 43890.6523 | 258 | 43993.4766 | 258 | 43470.6367 | 258 | 43382.6367 | 258 | 44354.9102 | 258 | 43449.5586 | 258 | 43175.1602 | 258 | 44424.8008
259 | 43560.5664 | 21974.4336 | 259 | 43871.043 | 259 | 43956.5234 | 259 | 43481 | 259 | 43354.5469 | 259 | 44297 | 259 | 43472.1602 | 259 | 43144.2383 | 259 | 44438.5195
260 | 43589.2617 | 21945.7383 | 260 | 43870.3906 | 260 | 43932.9141 | 260 | 43494.8164 | 260 | 43373.3633 | 260 | 44233.9531 | 260 | 43471.5586 | 260 | 43131 | 260 | 44445.7617
261 | 43617.5664 | 21917.4336 | 261 | 43899.9141 | 261 | 43951.8711 | 261 | 43523.1836 | 261 | 43320.4102 | 261 | 44228.4102 | 261 | 43475.6797 | 261 | 43112.6797 | 261 | 44424.9609
262 | 43619.9141 | 21915.0859 | 262 | 43900.4336 | 262 | 43955.5234 | 262 | 43501.7734 | 262 | 43284.9531 | 262 | 44280.3633 | 262 | 43447.8789 | 262 | 43081.0781 | 262 | 44310.5586
263 | 43590.5664 | 21944.4336 | 263 | 43860.1289 | 263 | 43935.957 | 263 | 43536.5898 | 263 | 43308.9102 | 263 | 44354.3164 | 263 | 43423.8008 | 263 | 43064.2812 | 263 | 44290.7188
264 | 43501.9141 | 22033.0859 | 264 | 43830.6953 | 264 | 43878.7383 | 264 | 43552.3164 | 264 | 43308.7734 | 264 | 44411.3164 | 264 | 43437.4414 | 264 | 43089.7617 | 264 | 44373.7188
265 | 43519.0859 | 22015.9141 | 265 | 43852.9141 | 265 | 43881.7383 | 265 | 43495.7734 | 265 | 43315.6367 | 265 | 44426.9102 | 265 | 43426.1992 | 265 | 43106.8008 | 265 | 44378.7188
266 | 43532.6953 | 22002.3047 | 266 | 43883.3047 | 266 | 43892.1758 | 266 | 43509.5469 | 266 | 43343.9531 | 266 | 44396.5469 | 266 | 43392.1602 | 266 | 43119.3594 | 266 | 44371.8008
267 | 43519.2617 | 22015.7383 | 267 | 43917.5664 | 267 | 43900.2188 | 267 | 43488.4102 | 267 | 43346.0898 | 267 | 44380.5898 | 267 | 43412.3594 | 267 | 43134.7188 | 267 | 44394.8008
268 | 43471.4766 | 22063.5234 | 268 | 43869.6094 | 268 | 43836.7383 | 268 | 43478.8633 | 268 | 43334.7734 | 268 | 44415.0898 | 268 | 43388 | 268 | 43122.5195 | 268 | 44399.0781
269 | 43480.4336 | 22054.5664 | 269 | 43844.2617 | 269 | 43879.043 | 269 | 43496.8164 | 269 | 43312.7734 | 269 | 44418.6836 | 269 | 43403.5195 | 269 | 43167.8398 | 269 | 44411.9609
270 | 43552.6953 | 21982.3047 | 270 | 43952.6523 | 270 | 43914.6094 | 270 | 43520.3164 | 270 | 43309.7734 | 270 | 44373.3633 | 270 | 43312.8008 | 270 | 43211.0781 | 270 | 44385.7617
271 | 43371.9141 | 22163.0859 | 271 | 43909.5234 | 271 | 43839.8242 | 271 | 43545.7266 | 271 | 43345.0898 | 271 | 44385.2734 | 271 | 43411.0781 | 271 | 43200.3203 | 271 | 44384.4414
272 | 43545.1758 | 21989.8242 | 272 | 43911.6094 | 272 | 43907.6094 | 272 | 43520.5469 | 272 | 43370.0898 | 272 | 44407.0898 | 272 | 43413.7188 | 272 | 43212.0391 | 272 | 44400.2383
273 | 43527.6953 | 22007.3047 | 273 | 43900.957 | 273 | 43888.8711 | 273 | 43479.5469 | 273 | 43419.7734 | 273 | 44408.6836 | 273 | 43420.1992 | 273 | 43220.5195 | 273 | 44334.3594
274 | 43540.9141 | 21994.0859 | 274 | 43913.7383 | 274 | 43905.7812 | 274 | 43505.4531 | 274 | 43376.2266 | 274 | 44390.6836 | 274 | 43432.4414 | 274 | 43143.6797 | 274 | 44267.7188
275 | 43551.957 | 21983.043 | 275 | 43896.957 | 275 | 43903.8711 | 275 | 43498.3164 | 275 | 43374.4531 | 275 | 44394.4102 | 275 | 43438.9219 | 275 | 43162.6016 | 275 | 44226.9609
276 | 43559.8242 | 21975.1758 | 276 | 43921.4766 | 276 | 43974.8242 | 276 | 43489.3164 | 276 | 43368.3633 | 276 | 44379.8164 | 276 | 43516.2383 | 276 | 43237.6797 | 276 | 44205.9219
277 | 43517.8711 | 22017.1289 | 277 | 43849.7383 | 277 | 43934.1758 | 277 | 43536.5 | 277 | 43415.2734 | 277 | 44363.1836 | 277 | 43499.3594 | 277 | 43266.3203 | 277 | 44233.9609
278 | 43585.7383 | 21949.2617 | 278 | 43912.6523 | 278 | 43967.2617 | 278 | 43597.0469 | 278 | 43390 | 278 | 44278.5 | 278 | 43485 | 278 | 43299.2812 | 278 | 44304.0781
279 | 43558.4336 | 21976.5664 | 279 | 43889.7812 | 279 | 43968.0859 | 279 | 43606.5898 | 279 | 43415.8633 | 279 | 44311.9102 | 279 | 43476 | 279 | 43348.9609 | 279 | 44346.7188
280 | 43558.2188 | 21976.7812 | 280 | 43897.6523 | 280 | 43996.5234 | 280 | 43647.6836 | 280 | 43470.7734 | 280 | 44309.4531 | 280 | 43458.4414 | 280 | 43326.2383 | 280 | 44357.5195
281 | 43609.6953 | 21925.3047 | 281 | 43939.7812 | 281 | 44049.6094 | 281 | 43670.4531 | 281 | 43450.0469 | 281 | 44283.2734 | 281 | 43445.8008 | 281 | 43300.1602 | 281 | 44341.7188
282 | 43581.8242 | 21953.1758 | 282 | 43953.8242 | 282 | 44028.4336 | 282 | 43658.7266 | 282 | 43505 | 282 | 44301.3164 | 282 | 43442.6406 | 282 | 43254.0391 | 282 | 44326.7617
283 | 43621.4766 | 21913.5234 | 283 | 43944.4766 | 283 | 44050.1758 | 283 | 43638.5 | 283 | 43543.1367 | 283 | 44342.5898 | 283 | 43478 | 283 | 43246.8789 | 283 | 44335.9219
284 | 43664.2617 | 21870.7383 | 284 | 43967.7812 | 284 | 44052.6523 | 284 | 43607.5 | 284 | 43476.3164 | 284 | 44367 | 284 | 43482.0391 | 284 | 43250.1992 | 284 | 44339.3984
285 | 43708.1289 | 21826.8711 | 285 | 43989.6953 | 285 | 44037.7812 | 285 | 43601.8164 | 285 | 43423.2266 | 285 | 44430.5898 | 285 | 43514.0781 | 285 | 43228.3203 | 285 | 44315.4414
286 | 43698.5664 | 21836.4336 | 286 | 43978.5664 | 286 | 44018.9141 | 286 | 43615.4102 | 286 | 43400.4102 | 286 | 44460.3164 | 286 | 43492.1992 | 286 | 43252.1211 | 286 | 44344.8398
287 | 43689.2188 | 21845.7812 | 287 | 44001 | 287 | 43999 | 287 | 43634.3633 | 287 | 43361.7734 | 287 | 44489.6836 | 287 | 43482.3984 | 287 | 43275.8398 | 287 | 44395.8789
288 | 43698.8711 | 21836.1289 | 288 | 44041.4766 | 288 | 44026.2617 | 288 | 43657.6367 | 288 | 43369.6836 | 288 | 44499.7266 | 288 | 43490.5195 | 288 | 43228.7617 | 288 | 44393
289 | 43726.7812 | 21808.2188 | 289 | 44061.3906 | 289 | 44025.0859 | 289 | 43654.7266 | 289 | 43431.8633 | 289 | 44512.8164 | 289 | 43451.2812 | 289 | 43228.1602 | 289 | 44365.7188
290 | 43723.3477 | 21811.6523 | 290 | 44064.1289 | 290 | 43976.3477 | 290 | 43662.0898 | 290 | 43495.8164 | 290 | 44533.9531 | 290 | 43438.9219 | 290 | 43223 | 290 | 44354.3594
291 | 43723.5234 | 21811.4766 | 291 | 44029.9141 | 291 | 43887.1758 | 291 | 43640.0469 | 291 | 43543.7734 | 291 | 44528.6836 | 291 | 43491.3203 | 291 | 43210.4805 | 291 | 44353.2383
292 | 43719.0859 | 21815.9141 | 292 | 44030.957 | 292 | 43927.1289 | 292 | 43553.5898 | 292 | 43528.4531 | 292 | 44486.5469 | 292 | 43482.8008 | 292 | 43210.3594 | 292 | 44324.6406
293 | 43725.7383 | 21809.2617 | 293 | 44042.7383 | 293 | 43987 | 293 | 43468.3633 | 293 | 43465.2734 | 293 | 44407.6836 | 293 | 43564.6797 | 293 | 43255.5195 | 293 | 44227.3203
294 | 43702.3477 | 21832.6523 | 294 | 44060.6523 | 294 | 43953.6094 | 294 | 43314.3164 | 294 | 43451.7266 | 294 | 44449.0469 | 294 | 43470.1992 | 294 | 43262.4414 | 294 | 44281.4805
295 | 43684.2617 | 21850.7383 | 295 | 43995.3477 | 295 | 43996.4766 | 295 | 43480.3633 | 295 | 43439.2266 | 295 | 44464.7734 | 295 | 43430.2383 | 295 | 43299.3203 | 295 | 44265.6016
296 | 43717.7383 | 21817.2617 | 296 | 44029.4766 | 296 | 43998.5664 | 296 | 43542.3164 | 296 | 43400.0898 | 296 | 44479.6836 | 296 | 43460.0781 | 296 | 43302.5586 | 296 | 44262.0781
297 | 43659.8711 | 21875.1289 | 297 | 43946 | 297 | 43919.5234 | 297 | 43493.7734 | 297 | 43359.7734 | 297 | 44463.0898 | 297 | 43457.9609 | 297 | 43366 | 297 | 44260.0391
298 | 43693.3477 | 21841.6523 | 298 | 43964.6094 | 298 | 43954.4336 | 298 | 43530.3164 | 298 | 43392.5 | 298 | 44453.5898 | 298 | 43490.1992 | 298 | 43355.2812 | 298 | 44288.2812
299 | 43657.6953 | 21877.3047 | 299 | 44013.3906 | 299 | 44044.7383 | 299 | 43480.6836 | 299 | 43405 | 299 | 44394 | 299 | 43521.2812 | 299 | 43333.9219 | 299 | 44352.8398
300 | 43638.1289 | 21896.8711 | 300 | 44039.4766 | 300 | 44038.6953 | 300 | 43489.8633 | 300 | 43420.5 | 300 | 44473.8164 | 300 | 43577.7188 | 300 | 43353.8398 | 300 | 44378.9609
301 | 43645 | 21890 | 301 | 44010.8242 | 301 | 44018.3477 | 301 | 43426.2734 | 301 | 43391.8164 | 301 | 44487.7266 | 301 | 43553.0391 | 301 | 43368.6406 | 301 | 44409.7617
302 | 43656.8242 | 21878.1758 | 302 | 44085.3906 | 302 | 44098.7812 | 302 | 43469.4102 | 302 | 43480.2734 | 302 | 44480.1836 | 302 | 43582.3594 | 302 | 43379.8008 | 302 | 44450.2812
303 | 43625.7383 | 21909.2617 | 303 | 44042.2188 | 303 | 44033.3047 | 303 | 43523.9531 | 303 | 43466.6836 | 303 | 44428.4102 | 303 | 43633.1211 | 303 | 43474.3203 | 303 | 44437.9219
304 | 43708.1758 | 21826.8242 | 304 | 44119.7812 | 304 | 44036.5664 | 304 | 43516.3164 | 304 | 43451.6367 | 304 | 44403.4102 | 304 | 43636.2383 | 304 | 43479.4414 | 304 | 44439.6797
305 | 43662.3047 | 21872.6953 | 305 | 44064.3477 | 305 | 44042.5234 | 305 | 43532.6367 | 305 | 43461.8633 | 305 | 44325.4102 | 305 | 43604.6797 | 305 | 43418.1211 | 305 | 44457.8008
306 | 43614 | 21921 | 306 | 44003.957 | 306 | 44003.7383 | 306 | 43571.4531 | 306 | 43428.0898 | 306 | 44335.2734 | 306 | 43573.8789 | 306 | 43390.8398 | 306 | 44466.1602
307 | 43607.7383 | 21927.2617 | 307 | 44001.3906 | 307 | 43969.043 | 307 | 43581.3633 | 307 | 43441.6836 | 307 | 44278.1836 | 307 | 43561.2812 | 307 | 43347.3203 | 307 | 44433.1992
308 | 43609.1758 | 21925.8242 | 308 | 44014.1758 | 308 | 43966.3477 | 308 | 43606.5469 | 308 | 43465.5469 | 308 | 44300 | 308 | 43557.0781 | 308 | 43310.0781 | 308 | 44371.3984
309 | 43604.1289 | 21930.8711 | 309 | 44013.2617 | 309 | 43973.1289 | 309 | 43601.6836 | 309 | 43505.5898 | 309 | 44296.7266 | 309 | 43578.6406 | 309 | 43340.2812 | 309 | 44388.4414
310 | 43645.2188 | 21889.7812 | 310 | 44074.7383 | 310 | 44026.1289 | 310 | 43613.7266 | 310 | 43511.8164 | 310 | 44322.6367 | 310 | 43607.3984 | 310 | 43416.9609 | 310 | 44409.4805
311 | 43676.5664 | 21858.4336 | 311 | 44127.1289 | 311 | 44059.8242 | 311 | 43627.8164 | 311 | 43517.9531 | 311 | 44335.4102 | 311 | 43638.3594 | 311 | 43489.5586 | 311 | 44434.0781
312 | 43667.2188 | 21867.7812 | 312 | 44114.9141 | 312 | 44059.6523 | 312 | 43613.4102 | 312 | 43519.6836 | 312 | 44336 | 312 | 43658.4805 | 312 | 43523.2812 | 312 | 44428.2383
313 | 43654.2617 | 21880.7383 | 313 | 44077.9141 | 313 | 44031.8711 | 313 | 43591.0898 | 313 | 43543.7266 | 313 | 44319.8633 | 313 | 43595.5195 | 313 | 43469.6406 | 313 | 44416.4805
314 | 43648.4336 | 21886.5664 | 314 | 44044.8711 | 314 | 44051.4766 | 314 | 43623.7266 | 314 | 43556.1836 | 314 | 44369.2266 | 314 | 43540.0781 | 314 | 43439.3594 | 314 | 44407.3594
315 | 43686.043 | 21848.957 | 315 | 44098.4336 | 315 | 44061.9141 | 315 | 43639.5469 | 315 | 43570.0898 | 315 | 44397.7266 | 315 | 43451.3984 | 315 | 43387.4805 | 315 | 44405.8789
316 | 43698.9141 | 21836.0859 | 316 | 44081.6953 | 316 | 44058 | 316 | 43614.3633 | 316 | 43588.2266 | 316 | 44403.4531 | 316 | 43289.1211 | 316 | 43369.3984 | 316 | 44433.7188
317 | 43669.4336 | 21865.5664 | 317 | 44037.4336 | 317 | 44022.5234 | 317 | 43661.1367 | 317 | 43627 | 317 | 44399.1367 | 317 | 43455.6016 | 317 | 43310.6406 | 317 | 44457.7617
318 | 43705.957 | 21829.043 | 318 | 44097.2617 | 318 | 44020.8242 | 318 | 43678.4102 | 318 | 43614.2734 | 318 | 44385.7266 | 318 | 43472.0391 | 318 | 43280.2383 | 318 | 44366.2383
319 | 43700.3477 | 21834.6523 | 319 | 44048 | 319 | 44035.5664 | 319 | 43599.0469 | 319 | 43584.6367 | 319 | 44400.9531 | 319 | 43467.7188 | 319 | 43301.3984 | 319 | 44364.2812
320 | 43698.5664 | 21836.4336 | 320 | 43984.7383 | 320 | 43956.3477 | 320 | 43654.8633 | 320 | 43562.7266 | 320 | 44417.2266 | 320 | 43436.7617 | 320 | 43346.7617 | 320 | 44290
321 | 43772.6094 | 21762.3906 | 321 | 44048.0859 | 321 | 44079.6094 | 321 | 43567 | 321 | 43531.6836 | 321 | 44419.1367 | 321 | 43450.7188 | 321 | 43345.2812 | 321 | 44256.9219
322 | 43745.6953 | 21789.3047 | 322 | 44037.7383 | 322 | 44075.2188 | 322 | 43572.2734 | 322 | 43513.0898 | 322 | 44430.3633 | 322 | 43457.6016 | 322 | 43388.7617 | 322 | 44242.5586
323 | 43655.4336 | 21879.5664 | 323 | 43983.6523 | 323 | 43965.957 | 323 | 43556.0898 | 323 | 43514.8633 | 323 | 44428.7734 | 323 | 43422.3594 | 323 | 43407.5586 | 323 | 44258.6406
324 | 43650.043 | 21884.957 | 324 | 44027.8711 | 324 | 43987.6094 | 324 | 43492.8633 | 324 | 43524.3633 | 324 | 44429.2266 | 324 | 43407.1602 | 324 | 43396.8008 | 324 | 44249.5586
325 | 43733.5664 | 21801.4336 | 325 | 44093.1289 | 325 | 44078.9141 | 325 | 43447.4102 | 325 | 43570.0898 | 325 | 44394.4531 | 325 | 43459.2812 | 325 | 43380.8008 | 325 | 44278.6406
326 | 43685.7812 | 21849.2188 | 326 | 44038.3906 | 326 | 44006.3906 | 326 | 43451.1836 | 326 | 43502.0469 | 326 | 44416.0469 | 326 | 43481.8398 | 326 | 43393.5586 | 326 | 44294.7188
327 | 43789.8711 | 21745.1289 | 327 | 44064.1289 | 327 | 44066.0859 | 327 | 43474.6836 | 327 | 43502.9531 | 327 | 44433.4102 | 327 | 43471.1992 | 327 | 43373.8398 | 327 | 44315.8008
328 | 43734.6523 | 21800.3477 | 328 | 44015.957 | 328 | 44056.7812 | 328 | 43509.8164 | 328 | 43502.7734 | 328 | 44390.5 | 328 | 43496.0781 | 328 | 43368 | 328 | 44295.0391
329 | 43740.6953 | 21794.3047 | 329 | 44018.2188 | 329 | 44032.3477 | 329 | 43507.6367 | 329 | 43467.3164 | 329 | 44410.1367 | 329 | 43525.0781 | 329 | 43392 | 329 | 44295.5195
330 | 43730.1758 | 21804.8242 | 330 | 44005.3906 | 330 | 44017.8711 | 330 | 43522.2266 | 330 | 43426.0898 | 330 | 44401.2734 | 330 | 43554.8789 | 330 | 43413.2383 | 330 | 44320.8008
331 | 43724 | 21811 | 331 | 44058.3047 | 331 | 44030.2617 | 331 | 43520.5 | 331 | 43386.7266 | 331 | 44369.6367 | 331 | 43572.2812 | 331 | 43427.4805 | 331 | 44347.6797
332 | 43779.6953 | 21755.3047 | 332 | 44059.6953 | 332 | 44017.2188 | 332 | 43476.3633 | 332 | 43369.2266 | 332 | 44291.2266 | 332 | 43568.6016 | 332 | 43448 | 332 | 44324.3594
333 | 43806.4766 | 21728.5234 | 333 | 44086.7812 | 333 | 44023.6094 | 333 | 43496.4102 | 333 | 43414.9531 | 333 | 44255.5898 | 333 | 43582.8789 | 333 | 43459.8789 | 333 | 44309.0391
334 | 43755 | 21780 | 334 | 44085.4336 | 334 | 44050.6953 | 334 | 43533.7266 | 334 | 43494.3164 | 334 | 44224.2734 | 334 | 43590.1211 | 334 | 43492.3984 | 334 | 44340.9219
335 | 43738.3047 | 21796.6953 | 335 | 44093.9141 | 335 | 44042.3047 | 335 | 43544.6367 | 335 | 43557.5469 | 335 | 44222.5469 | 335 | 43559.9219 | 335 | 43526.6406 | 335 | 44345.8008
336 | 43737.1289 | 21797.8711 | 336 | 44108.957 | 336 | 44012.3047 | 336 | 43581.1367 | 336 | 43575.0898 | 336 | 44230.0469 | 336 | 43618.0391 | 336 | 43550.9609 | 336 | 44343.6016
337 | 43731.3047 | 21803.6953 | 337 | 44067.957 | 337 | 43976.6953 | 337 | 43592.1836 | 337 | 43596.2734 | 337 | 44225.7266 | 337 | 43632.7617 | 337 | 43560.9609 | 337 | 44354.0781
338 | 43725.8711 | 21809.1289 | 338 | 44076.6523 | 338 | 43975.5234 | 338 | 43589.6836 | 338 | 43571.2734 | 338 | 44211 | 338 | 43560.0391 | 338 | 43556.2812 | 338 | 44353.8398
339 | 43710.8711 | 21824.1289 | 339 | 44039.2617 | 339 | 44009.1289 | 339 | 43597.2734 | 339 | 43594.1836 | 339 | 44200.9531 | 339 | 43591.8008 | 339 | 43537.8398 | 339 | 44355.5195
340 | 43726.5664 | 21808.4336 | 340 | 44087.3906 | 340 | 44026.0859 | 340 | 43640.4531 | 340 | 43603.7266 | 340 | 44202.0469 | 340 | 43599.4805 | 340 | 43520.1211 | 340 | 44367
341 | 43665.3477 | 21869.6523 | 341 | 44068.3047 | 341 | 43997.957 | 341 | 43587.4531 | 341 | 43611.1836 | 341 | 44215.5469 | 341 | 43612.3984 | 341 | 43470.4414 | 341 | 44363.2812
342 | 43648.3477 | 21886.6523 | 342 | 44029 | 342 | 43965.9141 | 342 | 43611.1367 | 342 | 43573.9102 | 342 | 44195.8164 | 342 | 43597.4805 | 342 | 43482.4414 | 342 | 44381.4414
343 | 43729.0859 | 21805.9141 | 343 | 44038.3047 | 343 | 43988.9141 | 343 | 43643.5469 | 343 | 43586.4102 | 343 | 44207 | 343 | 43544.1211 | 343 | 43509.5586 | 343 | 44335.0781
344 | 43755.4766 | 21779.5234 | 344 | 44104.5664 | 344 | 43998.1758 | 344 | 43597.5 | 344 | 43603.1836 | 344 | 44232.7734 | 344 | 43545.0781 | 344 | 43482.9219 | 344 | 44365.3984
345 | 43714.1758 | 21820.8242 | 345 | 44091.0859 | 345 | 43993.1758 | 345 | 43612.5898 | 345 | 43612.2266 | 345 | 44311.3633 | 345 | 43520.8398 | 345 | 43497.7617 | 345 | 44340.2812
346 | 43685.5234 | 21849.4766 | 346 | 44068.7383 | 346 | 43959.1289 | 346 | 43632.8164 | 346 | 43634.0469 | 346 | 44405.8164 | 346 | 43486.9609 | 346 | 43487.1602 | 346 | 44300.3984
347 | 43740.3906 | 21794.6094 | 347 | 44158.3047 | 347 | 44003.2188 | 347 | 43687.1836 | 347 | 43620.2734 | 347 | 44451.6836 | 347 | 43419.6016 | 347 | 43487.6406 | 347 | 44231.5586
348 | 43737.3906 | 21797.6094 | 348 | 44120.2617 | 348 | 43967.6523 | 348 | 43712.3164 | 348 | 43602.7734 | 348 | 44435.1836 | 348 | 43431.6016 | 348 | 43462.8398 | 348 | 44196.6016
349 | 43740.4766 | 21794.5234 | 349 | 44107.1289 | 349 | 43994.957 | 349 | 43752.4102 | 349 | 43695.3633 | 349 | 44519.0469 | 349 | 43436.5195 | 349 | 43443.0781 | 349 | 44186.5195
350 | 43720 | 21815 | 350 | 44152.5664 | 350 | 44008.8242 | 350 | 43754.2266 | 350 | 43677.8633 | 350 | 44492.7734 | 350 | 43462.4805 | 350 | 43421.5586 | 350 | 44179.3984
351 | 43715.2617 | 21819.7383 | 351 | 44120.5234 | 351 | 44035.8711 | 351 | 43741.2266 | 351 | 43698.9102 | 351 | 44490.8633 | 351 | 43454.5195 | 351 | 43408.4805 | 351 | 44144.8789
352 | 43675.2617 | 21859.7383 | 352 | 44076.5234 | 352 | 43950.8711 | 352 | 43700.0469 | 352 | 43666.6836 | 352 | 44443.0469 | 352 | 43475.6797 | 352 | 43384 | 352 | 44134.3594
353 | 43650.2617 | 21884.7383 | 353 | 44108.2617 | 353 | 43954.8242 | 353 | 43674.8164 | 353 | 43707.1836 | 353 | 44410.7734 | 353 | 43473.3984 | 353 | 43378.6406 | 353 | 44142.5195
354 | 43640.5234 | 21894.4766 | 354 | 44120.1289 | 354 | 43985.957 | 354 | 43617.0898 | 354 | 43694.3164 | 354 | 44422.0898 | 354 | 43445.2383 | 354 | 43409.4805 | 354 | 44141.7617
355 | 43647.8711 | 21887.1289 | 355 | 44075.2617 | 355 | 44009.043 | 355 | 43574.4102 | 355 | 43660.7266 | 355 | 44391.6367 | 355 | 43440.5195 | 355 | 43452.6406 | 355 | 44127.8008
356 | 43692.2188 | 21842.7812 | 356 | 44105.957 | 356 | 44092.7383 | 356 | 43613.5 | 356 | 43653.0469 | 356 | 44372.5469 | 356 | 43448.8398 | 356 | 43505.7617 | 356 | 44131.4805
357 | 43689.8242 | 21845.1758 | 357 | 44113.6523 | 357 | 44090.957 | 357 | 43612.1836 | 357 | 43625.2266 | 357 | 44414.9531 | 357 | 43524 | 357 | 43494.5195 | 357 | 44151.8789
358 | 43683.4336 | 21851.5664 | 358 | 44084.3047 | 358 | 44058.6094 | 358 | 43577.8164 | 358 | 43611.0898 | 358 | 44417.5898 | 358 | 43547.4414 | 358 | 43514.5195 | 358 | 44172.1602
359 | 43685.8711 | 21849.1289 | 359 | 44082.8711 | 359 | 44046.4336 | 359 | 43586.3633 | 359 | 43570.8164 | 359 | 44387.5898 | 359 | 43551.9219 | 359 | 43498.0781 | 359 | 44198.5586
360 | 43668 | 21867 | 360 | 44088.8242 | 360 | 44040.3477 | 360 | 43560.6836 | 360 | 43571.4531 | 360 | 44388.5 | 360 | 43586.8789 | 360 | 43518.5195 | 360 | 44282.2383
361 | 43657.7383 | 21877.2617 | 361 | 44093.6094 | 361 | 44026.4766 | 361 | 43561.4531 | 361 | 43594.5 | 361 | 44387.0469 | 361 | 43567.4805 | 361 | 43561.3984 | 361 | 44361.3203
362 | 43702.1758 | 21832.8242 | 362 | 44114.957 | 362 | 44049.2617 | 362 | 43612.4102 | 362 | 43612.1836 | 362 | 44350.6836 | 362 | 43628.8398 | 362 | 43589.6797 | 362 | 44387.1602
363 | 43645.043 | 21889.957 | 363 | 44078.3477 | 363 | 43988.5664 | 363 | 43645.9531 | 363 | 43592.3164 | 363 | 44355.8164 | 363 | 43586.2812 | 363 | 43580.0391 | 363 | 44413.4805
364 | 43627.043 | 21907.957 | 364 | 44125.3906 | 364 | 44003.8242 | 364 | 43667.8633 | 364 | 43557.1367 | 364 | 44381.3633 | 364 | 43609.8398 | 364 | 43585.4805 | 364 | 44471.3594
365 | 43615.3047 | 21919.6953 | 365 | 44129 | 365 | 43989.957 | 365 | 43680.4531 | 365 | 43551.5469 | 365 | 44341.5898 | 365 | 43599.3594 | 365 | 43622.6016 | 365 | 44448.9219
366 | 43622.5664 | 21912.4336 | 366 | 44101.4336 | 366 | 43959.0859 | 366 | 43719.7734 | 366 | 43620.4102 | 366 | 44276.5 | 366 | 43621.1211 | 366 | 43577.6797 | 366 | 44456.6016
367 | 43676.957 | 21858.043 | 367 | 44162.5234 | 367 | 43995.7812 | 367 | 43693.3633 | 367 | 43633.1367 | 367 | 44200 | 367 | 43581.9219 | 367 | 43634.6406 | 367 | 44403.3984
368 | 43563.6094 | 21971.3906 | 368 | 44054.6094 | 368 | 43921.3906 | 368 | 43653.2266 | 368 | 43701.3633 | 368 | 44153.1836 | 368 | 43625 | 368 | 43588.1211 | 368 | 44379.5586
369 | 43568.7383 | 21966.2617 | 369 | 44108.0859 | 369 | 43941.1758 | 369 | 43738.6367 | 369 | 43705 | 369 | 44167.0898 | 369 | 43690.6406 | 369 | 43580.7617 | 369 | 44363.9219
370 | 43576.3047 | 21958.6953 | 370 | 44147.9141 | 370 | 43952.6094 | 370 | 43695.0898 | 370 | 43703.2734 | 370 | 44175.6836 | 370 | 43700.9219 | 370 | 43657.3984 | 370 | 44385.3203
371 | 43603.7812 | 21931.2188 | 371 | 44240.7383 | 371 | 43952 | 371 | 43691.6836 | 371 | 43694.3164 | 371 | 44227.7266 | 371 | 43707.9609 | 371 | 43640.5586 | 371 | 44336.3203
372 | 43657.5234 | 21877.4766 | 372 | 44195.5664 | 372 | 43934.6523 | 372 | 43642.7734 | 372 | 43665.1367 | 372 | 44323.6836 | 372 | 43701.1602 | 372 | 43679.6016 | 372 | 44342.3203
373 | 43660.5664 | 21874.4336 | 373 | 44153.6094 | 373 | 43922.3477 | 373 | 43696.5 | 373 | 43655.0898 | 373 | 44360.1367 | 373 | 43696.2383 | 373 | 43623.9609 | 373 | 44350.7188
374 | 43638.5234 | 21896.4766 | 374 | 44148.6094 | 374 | 43901.9141 | 374 | 43697.2734 | 374 | 43716.6836 | 374 | 44349.6367 | 374 | 43653.0781 | 374 | 43659.3203 | 374 | 44345.1602
375 | 43646.5664 | 21888.4336 | 375 | 44149.5664 | 375 | 43866.0859 | 375 | 43689.8633 | 375 | 43694.7266 | 375 | 44341.2734 | 375 | 43620.7188 | 375 | 43612.1602 | 375 | 44335.8789
376 | 43670.7812 | 21864.2188 | 376 | 44157.7383 | 376 | 43841.2617 | 376 | 43701.9531 | 376 | 43730 | 376 | 44334.9102 | 376 | 43586.1992 | 376 | 43609.5195 | 376 | 44335.4805
377 | 43668.6523 | 21866.3477 | 377 | 44155.3906 | 377 | 43862.6953 | 377 | 43703.8164 | 377 | 43744.8164 | 377 | 44334.9531 | 377 | 43567.2383 | 377 | 43605.8398 | 377 | 44281.6016
378 | 43651.6953 | 21883.3047 | 378 | 44129.7812 | 378 | 43906.2188 | 378 | 43667.9531 | 378 | 43703.4102 | 378 | 44375.6836 | 378 | 43563.7617 | 378 | 43598.0391 | 378 | 44278.5195
379 | 43693.1758 | 21841.8242 | 379 | 44173.8242 | 379 | 43892.043 | 379 | 43681.0469 | 379 | 43756.4531 | 379 | 44415.2734 | 379 | 43567.9219 | 379 | 43525.1602 | 379 | 44303.9609
380 | 43655.4766 | 21879.5234 | 380 | 44188.5664 | 380 | 43858.957 | 380 | 43716.7266 | 380 | 43782.5 | 380 | 44390.1367 | 380 | 43557.3203 | 380 | 43502.4805 | 380 | 44285.3594
381 | 43652.6094 | 21882.3906 | 381 | 44199.2188 | 381 | 43853.3906 | 381 | 43741.9102 | 381 | 43770.6367 | 381 | 44426.1367 | 381 | 43519.8789 | 381 | 43522.6406 | 381 | 44240.5586
382 | 43677.6094 | 21857.3906 | 382 | 44209.1289 | 382 | 43855.1289 | 382 | 43778.5 | 382 | 43723.1367 | 382 | 44392.6836 | 382 | 43540.3203 | 382 | 43558.1992 | 382 | 44157.4805
383 | 43671.2188 | 21863.7812 | 383 | 44197.6523 | 383 | 43892.043 | 383 | 43700.8633 | 383 | 43739.9531 | 383 | 44377.2266 | 383 | 43564.7188 | 383 | 43567.1992 | 383 | 44086
384 | 43700.7812 | 21834.2188 | 384 | 44236.2188 | 384 | 43906.9141 | 384 | 43657.6836 | 384 | 43759.7734 | 384 | 44355.5898 | 384 | 43561.2383 | 384 | 43547.0391 | 384 | 44143.4805
385 | 43628.6953 | 21906.3047 | 385 | 44211.2188 | 385 | 43894.6523 | 385 | 43650.5469 | 385 | 43702 | 385 | 44300.4531 | 385 | 43571.0391 | 385 | 43523.2812 | 385 | 44152.6797
386 | 43558.4336 | 21976.5664 | 386 | 44090.6094 | 386 | 43902.3047 | 386 | 43646.9531 | 386 | 43642.5 | 386 | 44295.7734 | 386 | 43627.5586 | 386 | 43526.5195 | 386 | 44215.1992
387 | 43647.5664 | 21887.4336 | 387 | 44141.9141 | 387 | 43867.4336 | 387 | 43622.6367 | 387 | 43608.5898 | 387 | 44272.3164 | 387 | 43644.0391 | 387 | 43569.4805 | 387 | 44275.1992
388 | 43667.5664 | 21867.4336 | 388 | 44188.7812 | 388 | 43868.1289 | 388 | 43641.5 | 388 | 43649.2734 | 388 | 44301.8633 | 388 | 43645.9219 | 388 | 43581.1211 | 388 | 44323.3594
389 | 43668.4766 | 21866.5234 | 389 | 44226.9141 | 389 | 43870 | 389 | 43598.9531 | 389 | 43735.4531 | 389 | 44335.5 | 389 | 43654.6406 | 389 | 43648.5195 | 389 | 44328.9219
390 | 43591.3047 | 21943.6953 | 390 | 44190.6953 | 390 | 43931.8711 | 390 | 43633.0898 | 390 | 43804.7266 | 390 | 44336.3164 | 390 | 43650.1992 | 390 | 43661.5586 | 390 | 44290.8008
391 | 43623.1289 | 21911.8711 | 391 | 44203.6094 | 391 | 44018.9141 | 391 | 43650.9102 | 391 | 43833.7734 | 391 | 44325.6367 | 391 | 43656.8008 | 391 | 43666.1602 | 391 | 44276.5586
392 | 43633.1758 | 21901.8242 | 392 | 44243.2188 | 392 | 43979.043 | 392 | 43611.6367 | 392 | 43818.5 | 392 | 44337.0469 | 392 | 43685.6406 | 392 | 43659.9219 | 392 | 44275.6797
393 | 43644.1289 | 21890.8711 | 393 | 44297.957 | 393 | 43903.1758 | 393 | 43663.9102 | 393 | 43832.2734 | 393 | 44305.5469 | 393 | 43649.8008 | 393 | 43651.6016 | 393 | 44307
394 | 43650.4766 | 21884.5234 | 394 | 44234 | 394 | 43947.3906 | 394 | 43703.1367 | 394 | 43802.3164 | 394 | 44303.0469 | 394 | 43598.1211 | 394 | 43647.2383 | 394 | 44351.8008
395 | 43687.3477 | 21847.6523 | 395 | 44315.8242 | 395 | 43865.4766 | 395 | 43639.2266 | 395 | 43740.1836 | 395 | 44303.9102 | 395 | 43621.9609 | 395 | 43685.6406 | 395 | 44329
396 | 43660.4766 | 21874.5234 | 396 | 44216.7383 | 396 | 43830.2188 | 396 | 43663.6836 | 396 | 43697.5 | 396 | 44297.7266 | 396 | 43645.5586 | 396 | 43692.4414 | 396 | 44385.8789
397 | 43686.6953 | 21848.3047 | 397 | 44313.8242 | 397 | 44025 | 397 | 43587.1367 | 397 | 43679.4102 | 397 | 44303.0469 | 397 | 43688.3984 | 397 | 43672.5586 | 397 | 44349.7188
398 | 43731.5664 | 21803.4336 | 398 | 44334.6523 | 398 | 44091.6094 | 398 | 43560.4102 | 398 | 43714.3164 | 398 | 44296.4531 | 398 | 43660.2383 | 398 | 43712.0391 | 398 | 44322.0391
399 | 43693.6953 | 21841.3047 | 399 | 44315.5664 | 399 | 43928.9141 | 399 | 43486.3164 | 399 | 43738.3164 | 399 | 44239.5 | 399 | 43642.3984 | 399 | 43694.9219 | 399 | 44303.2383
400 | 43679.8242 | 21855.1758 | 400 | 44328.5234 | 400 | 43961.8242 | 400 | 43453.8633 | 400 | 43758.9102 | 400 | 44284 | 400 | 43666.7188 | 400 | 43743.5195 | 400 | 44254.5195
401 | 43687.2617 | 21847.7383 | 401 | 44324.8242 | 401 | 43975.5234 | 401 | 43469.7734 | 401 | 43732.1367 | 401 | 44271.7734 | 401 | 43688 | 401 | 43787.6406 | 401 | 44242.7617
402 | 43685.3047 | 21849.6953 | 402 | 44318.2617 | 402 | 44053.043 | 402 | 43453.1836 | 402 | 43723.0898 | 402 | 44274.2266 | 402 | 43680.1211 | 402 | 43786.3203 | 402 | 44236.6797
403 | 43691.4766 | 21843.5234 | 403 | 44380 | 403 | 44068.5664 | 403 | 43497.3633 | 403 | 43700.3164 | 403 | 44242.1836 | 403 | 43632.9219 | 403 | 43733.6406 | 403 | 44230.3203
404 | 43714.2188 | 21820.7812 | 404 | 44379.1758 | 404 | 44075.0859 | 404 | 43511.1836 | 404 | 43673.1367 | 404 | 44280.1367 | 404 | 43631.4805 | 404 | 43719.7617 | 404 | 44276.2383
405 | 43703.6523 | 21831.3477 | 405 | 44360.7383 | 405 | 44100.4766 | 405 | 43509.5 | 405 | 43661.7734 | 405 | 44257.5 | 405 | 43647.5195 | 405 | 43697.6406 | 405 | 44302
406 | 43714.2617 | 21820.7383 | 406 | 44395.957 | 406 | 44107.043 | 406 | 43584.2734 | 406 | 43694.0898 | 406 | 44258.2266 | 406 | 43678.1992 | 406 | 43665.3203 | 406 | 44284.4414
407 | 43766.6953 | 21768.3047 | 407 | 44422.1758 | 407 | 44112 | 407 | 43581.7266 | 407 | 43724.5 | 407 | 44263.1836 | 407 | 43732.3203 | 407 | 43627.5586 | 407 | 44261.6016
408 | 43776.3477 | 21758.6523 | 408 | 44405.6094 | 408 | 44069.2188 | 408 | 43643.9102 | 408 | 43749.2734 | 408 | 44241.9102 | 408 | 43656.9609 | 408 | 43618.3984 | 408 | 44246.3203
409 | 43813.8711 | 21721.1289 | 409 | 44489.3047 | 409 | 44123.0859 | 409 | 43704.7266 | 409 | 43774.6367 | 409 | 44222.8164 | 409 | 43620.6406 | 409 | 43648.8398 | 409 | 44233.8398
410 | 43824.5664 | 21710.4336 | 410 | 44473.4766 | 410 | 44052.8242 | 410 | 43712.2266 | 410 | 43790.7266 | 410 | 44223.9531 | 410 | 43669 | 410 | 43723.2383 | 410 | 44248.2812
411 | 43863.0859 | 21671.9141 | 411 | 44456.1758 | 411 | 44054 | 411 | 43813.0898 | 411 | 43800.9531 | 411 | 44241.0898 | 411 | 43647.2383 | 411 | 43775 | 411 | 44234.5195
412 | 43871.8711 | 21663.1289 | 412 | 44477.3047 | 412 | 44045.4336 | 412 | 43779.6367 | 412 | 43803.5469 | 412 | 44252.0898 | 412 | 43657.0391 | 412 | 43780.0781 | 412 | 44242.8789
413 | 43843.3477 | 21691.6523 | 413 | 44478.8242 | 413 | 44036.8711 | 413 | 43779.9102 | 413 | 43804.0898 | 413 | 44276.7734 | 413 | 43642.7617 | 413 | 43803 | 413 | 44250.9219
414 | 43903.8242 | 21631.1758 | 414 | 44502.8242 | 414 | 44024.7812 | 414 | 43722.5469 | 414 | 43815.8164 | 414 | 44277.3164 | 414 | 43676.7617 | 414 | 43807.9219 | 414 | 44229.8008
415 | 43925.4766 | 21609.5234 | 415 | 44452.6094 | 415 | 44031.5664 | 415 | 43753.6367 | 415 | 43851.3633 | 415 | 44278.3633 | 415 | 43657.4414 | 415 | 43780.5195 | 415 | 44255.7617
416 | 43923.5234 | 21611.4766 | 416 | 44449.7812 | 416 | 44000.6094 | 416 | 43724 | 416 | 43871.5898 | 416 | 44278.9102 | 416 | 43710.3984 | 416 | 43748.1602 | 416 | 44249.1602
417 | 43890.2617 | 21644.7383 | 417 | 44400.7812 | 417 | 43975.6094 | 417 | 43697.2266 | 417 | 43871.7266 | 417 | 44272.2266 | 417 | 43688.7617 | 417 | 43717.5586 | 417 | 44238.1992
418 | 43906.7383 | 21628.2617 | 418 | 44481.5664 | 418 | 43986.957 | 418 | 43751.2266 | 418 | 43852.9531 | 418 | 44312.6367 | 418 | 43602.8008 | 418 | 43685.1992 | 418 | 44193.5586
419 | 43877.3047 | 21657.6953 | 419 | 44467.5234 | 419 | 43977.8242 | 419 | 43752.3633 | 419 | 43809.5469 | 419 | 44322.7266 | 419 | 43627.2383 | 419 | 43700.8008 | 419 | 44233.6016
420 | 43825.2188 | 21709.7812 | 420 | 44422.2617 | 420 | 43969.5664 | 420 | 43746.9531 | 420 | 43845.5 | 420 | 44379.7734 | 420 | 43552.0391 | 420 | 43736.1992 | 420 | 44225.3594
421 | 43848.2188 | 21686.7812 | 421 | 44490.9141 | 421 | 43997.3047 | 421 | 43752.8164 | 421 | 43831.4531 | 421 | 44455.0898 | 421 | 43533.7617 | 421 | 43756.0391 | 421 | 44213.4805
422 | 43769.8242 | 21765.1758 | 422 | 44439 | 422 | 43924.4336 | 422 | 43730.5469 | 422 | 43813.6367 | 422 | 44436.7734 | 422 | 43511.3984 | 422 | 43709.7188 | 422 | 44210.5195
423 | 43831.043 | 21703.957 | 423 | 44503.0859 | 423 | 43993.4766 | 423 | 43752.5898 | 423 | 43858.7266 | 423 | 44409.4102 | 423 | 43455.8789 | 423 | 43706.1602 | 423 | 44208.6016
424 | 43833.7383 | 21701.2617 | 424 | 44492.3906 | 424 | 43964.5664 | 424 | 43717.1836 | 424 | 43925.5898 | 424 | 44411.5898 | 424 | 43473.6016 | 424 | 43693.4805 | 424 | 44189.6797
425 | 43871.6953 | 21663.3047 | 425 | 44568.043 | 425 | 43992.5664 | 425 | 43701.7266 | 425 | 43950.6836 | 425 | 44394.0469 | 425 | 43491.7617 | 425 | 43656.3203 | 425 | 44150.8008
426 | 43785.6094 | 21749.3906 | 426 | 44534.4766 | 426 | 43929.9141 | 426 | 43656.2734 | 426 | 43982.6836 | 426 | 44411.7266 | 426 | 43523.3594 | 426 | 43635.1211 | 426 | 44162.5195
427 | 43797.5234 | 21737.4766 | 427 | 44585.6953 | 427 | 43969.4766 | 427 | 43629.3164 | 427 | 43964.9102 | 427 | 44375.7734 | 427 | 43539.0781 | 427 | 43662.0781 | 427 | 44196.0781
428 | 43752.2188 | 21782.7812 | 428 | 44540.3047 | 428 | 43927.7812 | 428 | 43596.7734 | 428 | 43932.2734 | 428 | 44318.7266 | 428 | 43590.7188 | 428 | 43707.3594 | 428 | 44234.0391
429 | 43802.1289 | 21732.8711 | 429 | 44601.3047 | 429 | 43948.9141 | 429 | 43634.1836 | 429 | 43922.1836 | 429 | 44299.8164 | 429 | 43599.8398 | 429 | 43731.2383 | 429 | 44239.6797
430 | 43834.7812 | 21700.2188 | 430 | 44640.043 | 430 | 43965.4766 | 430 | 43634.9102 | 430 | 43866.6367 | 430 | 44277.2734 | 430 | 43629.3594 | 430 | 43751.1211 | 430 | 44210.1602
431 | 43843.4766 | 21691.5234 | 431 | 44660.6094 | 431 | 43988.3906 | 431 | 43655.8164 | 431 | 43830.4531 | 431 | 44244.5898 | 431 | 43701.8398 | 431 | 43766.9609 | 431 | 44242.8398
432 | 43813.3047 | 21721.6953 | 432 | 44592.3477 | 432 | 43925.6953 | 432 | 43684.5 | 432 | 43841.9102 | 432 | 44254.1836 | 432 | 43728.7188 | 432 | 43767.8398 | 432 | 44241.2812
433 | 43747.3477 | 21787.6523 | 433 | 44550.043 | 433 | 43851.1758 | 433 | 43682.0469 | 433 | 43833.2266 | 433 | 44300.5469 | 433 | 43736.1992 | 433 | 43775.7188 | 433 | 44276.9219
434 | 43723.4336 | 21811.5664 | 434 | 44520.9141 | 434 | 43861.4336 | 434 | 43650.3164 | 434 | 43835.5898 | 434 | 44336.4102 | 434 | 43779.8789 | 434 | 43813.4805 | 434 | 44292.0781
435 | 43815.957 | 21719.043 | 435 | 44637.6523 | 435 | 43976.4336 | 435 | 43659.6836 | 435 | 43870.8164 | 435 | 44355.1367 | 435 | 43766.2383 | 435 | 43835.0781 | 435 | 44354.0781
436 | 43809.2188 | 21725.7812 | 436 | 44629.957 | 436 | 43961.3906 | 436 | 43661.4102 | 436 | 43904.6836 | 436 | 44338.7734 | 436 | 43799.7617 | 436 | 43872.6797 | 436 | 44413.3984
437 | 43843.6094 | 21691.3906 | 437 | 44673.1758 | 437 | 43988.5234 | 437 | 43567.0898 | 437 | 43974.8633 | 437 | 44290.2266 | 437 | 43738.4805 | 437 | 43857.3203 | 437 | 44414
438 | 43831.6523 | 21703.3477 | 438 | 44708.6094 | 438 | 43991.7383 | 438 | 43613.3164 | 438 | 44033.2734 | 438 | 44278.5469 | 438 | 43705.6016 | 438 | 43842.4805 | 438 | 44362.8398
439 | 43802.6523 | 21732.3477 | 439 | 44712.7812 | 439 | 43923.3477 | 439 | 43601.0469 | 439 | 44083.9102 | 439 | 44256.9531 | 439 | 43703.6797 | 439 | 43861 | 439 | 44384.4805
440 | 43852.8711 | 21682.1289 | 440 | 44710.4766 | 440 | 43909.6094 | 440 | 43685.6367 | 440 | 44150.9531 | 440 | 44230.5469 | 440 | 43732.9609 | 440 | 43821.4805 | 440 | 44375.4805
441 | 43920.8242 | 21614.1758 | 441 | 44794.8242 | 441 | 43944 | 441 | 43685.0469 | 441 | 44137.5898 | 441 | 44235.9102 | 441 | 43701.4414 | 441 | 43799.3203 | 441 | 44367.6406
442 | 43847.4766 | 21687.5234 | 442 | 44682.3477 | 442 | 43843.957 | 442 | 43705.6367 | 442 | 44078.0898 | 442 | 44250.9531 | 442 | 43704.3594 | 442 | 43846.0781 | 442 | 44346.6797
443 | 43843.2188 | 21691.7812 | 443 | 44691.3477 | 443 | 43839 | 443 | 43726.9102 | 443 | 44065.2266 | 443 | 44279.9102 | 443 | 43685.6797 | 443 | 43795.7617 | 443 | 44310.8398
444 | 43862.2617 | 21672.7383 | 444 | 44751.1758 | 444 | 43850.1289 | 444 | 43723.6836 | 444 | 44079.7734 | 444 | 44248.6367 | 444 | 43744.1211 | 444 | 43896.8789 | 444 | 44260.4414
445 | 43869.8711 | 21665.1289 | 445 | 44753.6523 | 445 | 43877.8711 | 445 | 43780.4531 | 445 | 44036.8164 | 445 | 44239.1836 | 445 | 43723.2812 | 445 | 43905.3984 | 445 | 44231.6797
446 | 43890.7812 | 21644.2188 | 446 | 44793.7383 | 446 | 43880.4766 | 446 | 43782.3633 | 446 | 44050.5469 | 446 | 44100.9531 | 446 | 43706.2383 | 446 | 43935.7617 | 446 | 44208.5586
447 | 43902.4336 | 21632.5664 | 447 | 44841.8242 | 447 | 43855.043 | 447 | 43731.2266 | 447 | 43995.6836 | 447 | 44160.1836 | 447 | 43678.6797 | 447 | 43990.8398 | 447 | 44229.3594
448 | 43902.1758 | 21632.8242 | 448 | 44802.2188 | 448 | 43866.3047 | 448 | 43667.0898 | 448 | 44008.4531 | 448 | 44146.6367 | 448 | 43665.1992 | 448 | 43985.8789 | 448 | 44270.2812
449 | 43946.5664 | 21588.4336 | 449 | 44768.4336 | 449 | 43844.2188 | 449 | 43670.7266 | 449 | 43989.5469 | 449 | 44106.6367 | 449 | 43616 | 449 | 43952.2383 | 449 | 44305.9219
450 | 43927.1289 | 21607.8711 | 450 | 44739.5664 | 450 | 43795.1289 | 450 | 43646.4531 | 450 | 43946.9531 | 450 | 44143.1836 | 450 | 43560.9609 | 450 | 43950.4805 | 450 | 44301.3984
451 | 43976.4336 | 21558.5664 | 451 | 44793.2188 | 451 | 43885.043 | 451 | 43632.2734 | 451 | 43897.7734 | 451 | 44141.1367 | 451 | 43604.2812 | 451 | 43885.6406 | 451 | 44303
452 | 43948.3047 | 21586.6953 | 452 | 44804 | 452 | 43851.8711 | 452 | 43627.9102 | 452 | 43872.3633 | 452 | 44142.7734 | 452 | 43612.2812 | 452 | 43853.5195 | 452 | 44280.5586
453 | 43927.957 | 21607.043 | 453 | 44782.4336 | 453 | 43827.8242 | 453 | 43642.5469 | 453 | 43887 | 453 | 44176.6836 | 453 | 43630.3594 | 453 | 43843.5195 | 453 | 44246
454 | 43906.6094 | 21628.3906 | 454 | 44805.9141 | 454 | 43834.6523 | 454 | 43638.1367 | 454 | 43920.0898 | 454 | 44184.8164 | 454 | 43669.7188 | 454 | 43827.6797 | 454 | 44231.5195
455 | 43875.3906 | 21659.6094 | 455 | 44810.5664 | 455 | 43809.4766 | 455 | 43618.9531 | 455 | 43922.1367 | 455 | 44193.4102 | 455 | 43673.2812 | 455 | 43833.1211 | 455 | 44200.2812
456 | 43873.6094 | 21661.3906 | 456 | 44791.5234 | 456 | 43818.9141 | 456 | 43618.8164 | 456 | 43875.9531 | 456 | 44180.8164 | 456 | 43651.3203 | 456 | 43861.1602 | 456 | 44213.8789
457 | 43895.3477 | 21639.6523 | 457 | 44817.1289 | 457 | 43836.1289 | 457 | 43620.0469 | 457 | 43907.8164 | 457 | 44204.3164 | 457 | 43659 | 457 | 43889 | 457 | 44212.2383
458 | 43955.5234 | 21579.4766 | 458 | 44864.0859 | 458 | 43866.8242 | 458 | 43613.8633 | 458 | 43938.2734 | 458 | 44238.2266 | 458 | 43678.9219 | 458 | 43956.1992 | 458 | 44230.1602
459 | 43912.9141 | 21622.0859 | 459 | 44819 | 459 | 43881.8711 | 459 | 43608.9531 | 459 | 43919.5469 | 459 | 44266.0469 | 459 | 43572.0781 | 459 | 43977.4805 | 459 | 44185.6406
460 | 43917.1289 | 21617.8711 | 460 | 44790.9141 | 460 | 43859.3477 | 460 | 43620.0469 | 460 | 43891.1836 | 460 | 44232 | 460 | 43592.3594 | 460 | 44012.9609 | 460 | 44197.9219
461 | 43910.4336 | 21624.5664 | 461 | 44755.4336 | 461 | 43818.9141 | 461 | 43655.7266 | 461 | 43924.9531 | 461 | 44169.5469 | 461 | 43601.2812 | 461 | 44072.9609 | 461 | 44048.0391
462 | 43892.3047 | 21642.6953 | 462 | 44805.2188 | 462 | 43846 | 462 | 43656.5 | 462 | 43915.9531 | 462 | 44132.5 | 462 | 43604.8398 | 462 | 44060.4805 | 462 | 44106.4805
463 | 43982.7812 | 21552.2188 | 463 | 44914.3906 | 463 | 43951.3477 | 463 | 43648.5898 | 463 | 43903.6836 | 463 | 44066.9531 | 463 | 43643.0781 | 463 | 44046.0781 | 463 | 44107.6797
464 | 43975.4766 | 21559.5234 | 464 | 44877.2188 | 464 | 43913.6094 | 464 | 43614.5898 | 464 | 43884.5469 | 464 | 44052.7734 | 464 | 43707.9609 | 464 | 44021.3594 | 464 | 44055.6406
465 | 43961.957 | 21573.043 | 465 | 44821.7812 | 465 | 43909.4336 | 465 | 43667 | 465 | 43873.6836 | 465 | 44060.5 | 465 | 43676.8398 | 465 | 44030.8789 | 465 | 44068.0781
466 | 43929.6094 | 21605.3906 | 466 | 44809.7383 | 466 | 43839.5664 | 466 | 43668.8633 | 466 | 43870.3164 | 466 | 44093 | 466 | 43665.1211 | 466 | 44037.6016 | 466 | 44091.4805
467 | 43961.043 | 21573.957 | 467 | 44877.3047 | 467 | 43933.1289 | 467 | 43657.2266 | 467 | 43838 | 467 | 44121.5469 | 467 | 43689.3203 | 467 | 44039.9219 | 467 | 44097.4414
468 | 43952.1289 | 21582.8711 | 468 | 44876.5664 | 468 | 43873.1289 | 468 | 43686.5898 | 468 | 43802.1836 | 468 | 44136.8633 | 468 | 43715.0391 | 468 | 44022.1211 | 468 | 44104.2812
469 | 44001.4336 | 21533.5664 | 469 | 44935.7383 | 469 | 43940.8711 | 469 | 43706.0469 | 469 | 43835.6836 | 469 | 44123.5898 | 469 | 43694.0391 | 469 | 43973.6797 | 469 | 44102.7188
470 | 43995.2617 | 21539.7383 | 470 | 44980.3047 | 470 | 43933.6953 | 470 | 43724.8164 | 470 | 43819.2734 | 470 | 44141.0898 | 470 | 43675.0781 | 470 | 43925.9219 | 470 | 44133.4805
471 | 44009.3906 | 21525.6094 | 471 | 44938.5234 | 471 | 43889.5234 | 471 | 43753.5 | 471 | 43899.9531 | 471 | 44178.9531 | 471 | 43676.8789 | 471 | 43901.2812 | 471 | 44126.9219
472 | 44022.5664 | 21512.4336 | 472 | 44986.1289 | 472 | 43902.6523 | 472 | 43736.9531 | 472 | 43948.5898 | 472 | 44170 | 472 | 43640.5195 | 472 | 43891.1992 | 472 | 44144.7617
473 | 44088.6094 | 21446.3906 | 473 | 45035.4336 | 473 | 43951.6094 | 473 | 43727.5898 | 473 | 43993.7266 | 473 | 44179.5 | 473 | 43627 | 473 | 43874.9219 | 473 | 44182.9219
474 | 44064.5664 | 21470.4336 | 474 | 45013 | 474 | 43899.3477 | 474 | 43770.5469 | 474 | 43954.6836 | 474 | 44244.5898 | 474 | 43603.1992 | 474 | 43884.8789 | 474 | 44192.3203
475 | 44031.1289 | 21503.8711 | 475 | 44958.4336 | 475 | 43855.4766 | 475 | 43758.1367 | 475 | 43962.1367 | 475 | 44243 | 475 | 43624.3203 | 475 | 43894.3203 | 475 | 44175.0391
476 | 44059.9141 | 21475.0859 | 476 | 44967.3906 | 476 | 43885.5664 | 476 | 43730.6836 | 476 | 43931.9102 | 476 | 44220.6836 | 476 | 43596.3984 | 476 | 43911.6016 | 476 | 44144.1992
477 | 44103.2188 | 21431.7812 | 477 | 45020.4336 | 477 | 43915 | 477 | 43705.1367 | 477 | 43918.5469 | 477 | 44190.1367 | 477 | 43547.9219 | 477 | 43894.5586 | 477 | 44080.7617
478 | 44104.8242 | 21430.1758 | 478 | 45008.957 | 478 | 43874 | 478 | 43660 | 478 | 43930.5469 | 478 | 44183.5469 | 478 | 43547.1211 | 478 | 43923.2812 | 478 | 44017.8789
479 | 44102.1758 | 21432.8242 | 479 | 45024.0859 | 479 | 43874.043 | 479 | 43660.9102 | 479 | 43963.6367 | 479 | 44160.0469 | 479 | 43576.4805 | 479 | 43933.9219 | 479 | 44005.7617
480 | 44092.8711 | 21442.1289 | 480 | 45052.3047 | 480 | 43886.4336 | 480 | 43671.3633 | 480 | 43959.5469 | 480 | 44151.4102 | 480 | 43584.3203 | 480 | 43923.1992 | 480 | 44025.4414
481 | 44093.8711 | 21441.1289 | 481 | 45041.4336 | 481 | 43902.7812 | 481 | 43690.5898 | 481 | 43978.0898 | 481 | 44148.5898 | 481 | 43630.0781 | 481 | 43920.6797 | 481 | 44029.5195
482 | 44169.2617 | 21365.7383 | 482 | 45112.2188 | 482 | 43907.5234 | 482 | 43704 | 482 | 44025.3164 | 482 | 44143 | 482 | 43610.7188 | 482 | 43933 | 482 | 44068.6797
483 | 44145.957 | 21389.043 | 483 | 45110.2188 | 483 | 43913.3047 | 483 | 43708.2266 | 483 | 44090.1367 | 483 | 44166.3164 | 483 | 43635.0781 | 483 | 43891.4414 | 483 | 44080.1602
484 | 44110.3906 | 21424.6094 | 484 | 45059.7812 | 484 | 43864.3047 | 484 | 43667.5469 | 484 | 44102.7734 | 484 | 44176.4531 | 484 | 43596.7617 | 484 | 43899.1602 | 484 | 44067.3203
485 | 44154.1289 | 21380.8711 | 485 | 45176.5234 | 485 | 43948.8242 | 485 | 43610.6836 | 485 | 44100.0898 | 485 | 44209.4102 | 485 | 43559.6016 | 485 | 43842.2812 | 485 | 44104.8008
486 | 44138.3047 | 21396.6953 | 486 | 45144.0859 | 486 | 43599.8242 | 486 | 43683.9531 | 486 | 44126.9531 | 486 | 44224.1836 | 486 | 43578.3984 | 486 | 43857.0781 | 486 | 44123.5586
487 | 44152.6094 | 21382.3906 | 487 | 45092.0859 | 487 | 43814.7812 | 487 | 43718.7734 | 487 | 44102.1367 | 487 | 44231.1367 | 487 | 43654.9609 | 487 | 43859.1211 | 487 | 44130
488 | 44165.7383 | 21369.2617 | 488 | 45123.4336 | 488 | 43897.7383 | 488 | 43792.0898 | 488 | 44078.9102 | 488 | 44212.2266 | 488 | 43631.6797 | 488 | 43808.1211 | 488 | 44132.3203
489 | 44171.6094 | 21363.3906 | 489 | 45146.4336 | 489 | 43935 | 489 | 43781.3164 | 489 | 44067.3164 | 489 | 44179.7734 | 489 | 43666.1211 | 489 | 43785.7617 | 489 | 44188.3594
490 | 44182.1289 | 21352.8711 | 490 | 45153.1289 | 490 | 43932.5664 | 490 | 43775.1836 | 490 | 44114.1836 | 490 | 44198.5898 | 490 | 43673.0391 | 490 | 43818.6406 | 490 | 44219.6016
491 | 44154.5234 | 21380.4766 | 491 | 45113.4766 | 491 | 43874.8242 | 491 | 43770.9102 | 491 | 44136.1836 | 491 | 44197.6367 | 491 | 43698.8398 | 491 | 43784.4805 | 491 | 44196.2383
492 | 44193.1758 | 21341.8242 | 492 | 45120.6094 | 492 | 43859.5664 | 492 | 43761.3164 | 492 | 44146.1836 | 492 | 44223.8633 | 492 | 43722 | 492 | 43866.6406 | 492 | 44156.3203
493 | 44168.4336 | 21366.5664 | 493 | 45161.2617 | 493 | 43885.6523 | 493 | 43726.7734 | 493 | 44158.1367 | 493 | 44214.8633 | 493 | 43719.1602 | 493 | 43949.1211 | 493 | 44148.9219
494 | 44224.957 | 21310.043 | 494 | 45257.2188 | 494 | 43932.043 | 494 | 43717.6367 | 494 | 44104.9102 | 494 | 44171.8633 | 494 | 43712.6406 | 494 | 43960.3984 | 494 | 44140.1211
495 | 44168.3906 | 21366.6094 | 495 | 45209.7383 | 495 | 43881.8711 | 495 | 43731.8633 | 495 | 44057.1367 | 495 | 44172.1367 | 495 | 43705.1602 | 495 | 43944.5195 | 495 | 44116.8789
496 | 44189.6953 | 21345.3047 | 496 | 45217.6094 | 496 | 43925 | 496 | 43787.0898 | 496 | 44072.2266 | 496 | 44191.7266 | 496 | 43720.4414 | 496 | 43946.9609 | 496 | 44107.3594
497 | 44256.7383 | 21278.2617 | 497 | 45296.5234 | 497 | 43970.3906 | 497 | 43788.9531 | 497 | 44061.3164 | 497 | 44185 | 497 | 43724.8398 | 497 | 43901.7617 | 497 | 44107.2812
498 | 44265.3047 | 21269.6953 | 498 | 45331.5664 | 498 | 44001.8242 | 498 | 43824.0898 | 498 | 44033.9102 | 498 | 44206.9531 | 498 | 43703.1211 | 498 | 43892.9609 | 498 | 44118.6797
499 | 44297.9141 | 21237.0859 | 499 | 45373.5664 | 499 | 43994 | 499 | 43852.7266 | 499 | 44026.8633 | 499 | 44222.4102 | 499 | 43674.3594 | 499 | 43898.8398 | 499 | 44127.7188
500 | 44317.957 | 21217.043 | 500 | 45420.3047 | 500 | 43984.043 | 500 | 43890.2266 | 500 | 44039.5469 | 500 | 44245.6367 | 500 | 43653.4414 | 500 | 43908.8789 | 500 | 44147
501 | 44333.6953 | 21201.3047 | 501 | 45421.2617 | 501 | 44005.2188 | 501 | 43905.4531 | 501 | 44055.1836 | 501 | 44282.7266 | 501 | 43634.8008 | 501 | 43938.1211 | 501 | 44159.6016
502 | 44354.7383 | 21180.2617 | 502 | 45428.8242 | 502 | 44016.0859 | 502 | 43913.9102 | 502 | 44034.8633 | 502 | 44250.6836 | 502 | 43648.3984 | 502 | 43971.2383 | 502 | 44154.1992
503 | 44349.5234 | 21185.4766 | 503 | 45450.7812 | 503 | 43996.3906 | 503 | 43898 | 503 | 44032.3164 | 503 | 44258.0469 | 503 | 43662.9609 | 503 | 44019.8008 | 503 | 44171.1211
504 | 44348.9141 | 21186.0859 | 504 | 45464.9141 | 504 | 44017.5664 | 504 | 43893.5898 | 504 | 44027.2266 | 504 | 44264.5469 | 504 | 43687.7188 | 504 | 44064.5586 | 504 | 44138.1602
505 | 44365.2617 | 21169.7383 | 505 | 45498.0859 | 505 | 44037 | 505 | 43894.6836 | 505 | 44018.1836 | 505 | 44274.4531 | 505 | 43678.7188 | 505 | 44082.6797 | 505 | 44141.9219
506 | 44376.5664 | 21158.4336 | 506 | 45531.8242 | 506 | 44055.1758 | 506 | 43896.0469 | 506 | 44005.6367 | 506 | 44274.1367 | 506 | 43673.9609 | 506 | 44088.0391 | 506 | 44175.0781
507 | 44345.0859 | 21189.9141 | 507 | 45476.8711 | 507 | 44024.5234 | 507 | 43845.2734 | 507 | 44024.8633 | 507 | 44270.4102 | 507 | 43622.8398 | 507 | 44101.4805 | 507 | 44151.9609
508 | 44357.2617 | 21177.7383 | 508 | 45497.043 | 508 | 43987.6094 | 508 | 43847.1367 | 508 | 44077.5469 | 508 | 44263.7734 | 508 | 43677.6406 | 508 | 44087.7617 | 508 | 44143.3984
509 | 44391.4766 | 21143.5234 | 509 | 45577.5234 | 509 | 44025.8242 | 509 | 43791.1367 | 509 | 44132.3164 | 509 | 44234.5898 | 509 | 43726.4414 | 509 | 44088.2812 | 509 | 44129.0391
510 | 44321.6523 | 21213.3477 | 510 | 45496.1289 | 510 | 43955.6523 | 510 | 43768.4531 | 510 | 44145.1836 | 510 | 44262.3633 | 510 | 43757.3984 | 510 | 44075.6016 | 510 | 44123.7617
511 | 44390.4336 | 21144.5664 | 511 | 45611.1289 | 511 | 44032.5664 | 511 | 43731.5 | 511 | 44123.2734 | 511 | 44277.3633 | 511 | 43749.6406 | 511 | 44067.6797 | 511 | 44180.9219
512 | 44408.8711 | 21126.1289 | 512 | 45617.043 | 512 | 44056.5234 | 512 | 43766.4102 | 512 | 44078.0898 | 512 | 44247.1367 | 512 | 43741.7188 | 512 | 44113.8008 | 512 | 44172.8789
513 | 44425.6523 | 21109.3477 | 513 | 45646.7812 | 513 | 44077.8242 | 513 | 43732.8164 | 513 | 44067.8633 | 513 | 44241.7266 | 513 | 43737.9219 | 513 | 44136.8398 | 513 | 44184.1992
514 | 44420.1758 | 21114.8242 | 514 | 45643.6523 | 514 | 44082 | 514 | 43746.9531 | 514 | 44042.4102 | 514 | 44201.7266 | 514 | 43714.8008 | 514 | 44136.3203 | 514 | 44201.0391
515 | 44464.8242 | 21070.1758 | 515 | 45638.5234 | 515 | 44092.7383 | 515 | 43742.7734 | 515 | 44021.9102 | 515 | 44174.0898 | 515 | 43689.6797 | 515 | 44087.9219 | 515 | 44194.5586
516 | 44476.5234 | 21058.4766 | 516 | 45637.6953 | 516 | 44058.9141 | 516 | 43729.8164 | 516 | 44003.3633 | 516 | 44160.1836 | 516 | 43684.2812 | 516 | 44070.3984 | 516 | 44241.7188
517 | 44430.5664 | 21104.4336 | 517 | 45648.9141 | 517 | 44022.5664 | 517 | 43718.4531 | 517 | 44010.8164 | 517 | 44166.0898 | 517 | 43705.4414 | 517 | 44078.6406 | 517 | 44246.3203
518 | 44431.0859 | 21103.9141 | 518 | 45669.1758 | 518 | 43992.957 | 518 | 43718.1367 | 518 | 43987.5 | 518 | 44120.0898 | 518 | 43747.4414 | 518 | 44052.6016 | 518 | 44254.8789
519 | 44399.7812 | 21135.2188 | 519 | 45691.3906 | 519 | 43967.0859 | 519 | 43729 | 519 | 43991.0469 | 519 | 44086.5898 | 519 | 43781.6016 | 519 | 44051.8398 | 519 | 44255
520 | 44424.043 | 21110.957 | 520 | 45791.6523 | 520 | 44007.1758 | 520 | 43729.9531 | 520 | 44000.8164 | 520 | 44076.7734 | 520 | 43810.9219 | 520 | 44034.6406 | 520 | 44235.6406
521 | 44476.3906 | 21058.6094 | 521 | 45773.3906 | 521 | 44014.8711 | 521 | 43676.1367 | 521 | 44001.5469 | 521 | 44083.5469 | 521 | 43853.9219 | 521 | 44051.3203 | 521 | 44237.9219
522 | 44468.043 | 21066.957 | 522 | 45750.8711 | 522 | 43981.2188 | 522 | 43619.5 | 522 | 44027.5898 | 522 | 43991.1836 | 522 | 43891.3203 | 522 | 44046.0391 | 522 | 44241.6016
523 | 44437.8711 | 21097.1289 | 523 | 45811.0859 | 523 | 43978.8242 | 523 | 43614.2734 | 523 | 44055.9531 | 523 | 44034.4531 | 523 | 43907.0391 | 523 | 44049.5195 | 523 | 44224.7617
524 | 44434.4766 | 21100.5234 | 524 | 45786.8242 | 524 | 43963.0859 | 524 | 43638 | 524 | 44048.7734 | 524 | 44091.2734 | 524 | 43895.8008 | 524 | 44049.8789 | 524 | 44195.1992
525 | 44450.4766 | 21084.5234 | 525 | 45801.9141 | 525 | 43967.1758 | 525 | 43665.9102 | 525 | 44098.3164 | 525 | 44144.4102 | 525 | 43892.7188 | 525 | 44060 | 525 | 44186.3594
526 | 44441.2617 | 21093.7383 | 526 | 45809.5664 | 526 | 43957.5234 | 526 | 43695.2266 | 526 | 44095.8164 | 526 | 44146.8164 | 526 | 43867.6016 | 526 | 44048.8398 | 526 | 44205.3594
527 | 44437.9141 | 21097.0859 | 527 | 45824.3047 | 527 | 43938.4766 | 527 | 43744.7266 | 527 | 44099.6367 | 527 | 44146.0898 | 527 | 43879.1602 | 527 | 44039.8008 | 527 | 44200.5195
528 | 44463.2617 | 21071.7383 | 528 | 45849.3906 | 528 | 43965.5664 | 528 | 43774.5 | 528 | 44130.5898 | 528 | 44152 | 528 | 43868.5195 | 528 | 44055.7188 | 528 | 44206.6797
529 | 44494.3906 | 21040.6094 | 529 | 45869.3047 | 529 | 43987.957 | 529 | 43803.2734 | 529 | 44135.3164 | 529 | 44154.2266 | 529 | 43814.2383 | 529 | 44081.8008 | 529 | 44180.9219
530 | 44518.5664 | 21016.4336 | 530 | 45876.0859 | 530 | 43979.4336 | 530 | 43799.8164 | 530 | 44169.9102 | 530 | 44136.7734 | 530 | 43817.5586 | 530 | 44122.5195 | 530 | 44144.8789
531 | 44551.3477 | 20983.6523 | 531 | 45905.6094 | 531 | 43995.0859 | 531 | 43749.1836 | 531 | 44171.1367 | 531 | 44120.2266 | 531 | 43761.6016 | 531 | 44149.3984 | 531 | 44148.1211
532 | 44599.1758 | 20935.8242 | 532 | 45940.043 | 532 | 44031 | 532 | 43766.6836 | 532 | 44148.6367 | 532 | 44114.5898 | 532 | 43732.6406 | 532 | 44147.5586 | 532 | 44130.8008
533 | 44511.1289 | 21023.8711 | 533 | 45917.3906 | 533 | 43987.4766 | 533 | 43727.6367 | 533 | 44151.4102 | 533 | 44111.1836 | 533 | 43717.7617 | 533 | 44108.8789 | 533 | 44061.7617
534 | 44513.043 | 21021.957 | 534 | 45908.2617 | 534 | 43983.4766 | 534 | 43715.0898 | 534 | 44140.3164 | 534 | 44101.0898 | 534 | 43717.1992 | 534 | 44098.2812 | 534 | 44054
535 | 44484.4766 | 21050.5234 | 535 | 45901.957 | 535 | 43989.1289 | 535 | 43729.1836 | 535 | 44141.7734 | 535 | 44071.8633 | 535 | 43671.7617 | 535 | 44051 | 535 | 44049.7188
536 | 44456 | 21079 | 536 | 45888.4336 | 536 | 43936.8711 | 536 | 43707.8164 | 536 | 44117.5898 | 536 | 44101.3164 | 536 | 43698.8008 | 536 | 44036.8789 | 536 | 44029.4414
537 | 44501.8242 | 21033.1758 | 537 | 45968.8242 | 537 | 43987.4336 | 537 | 43683.6367 | 537 | 44059.9102 | 537 | 44089.2734 | 537 | 43708.5586 | 537 | 44003.8398 | 537 | 43985.9609
538 | 44546 | 20989 | 538 | 46038.1758 | 538 | 44035.8711 | 538 | 43636.8633 | 538 | 44043.7734 | 538 | 44061.9531 | 538 | 43672.7617 | 538 | 44003.7617 | 538 | 44008.1992
539 | 44498.6094 | 21036.3906 | 539 | 46041.0859 | 539 | 43989.2188 | 539 | 43649.5469 | 539 | 44004.6836 | 539 | 44086.3164 | 539 | 43653.4414 | 539 | 43963.4414 | 539 | 44039.1602
540 | 44531.3047 | 21003.6953 | 540 | 46045.8711 | 540 | 44001.5664 | 540 | 43643.2266 | 540 | 44017.0469 | 540 | 44062.5469 | 540 | 43674.1211 | 540 | 43937.9609 | 540 | 44097.3203
541 | 44570.043 | 20964.957 | 541 | 46052.6523 | 541 | 43989 | 541 | 43652.0898 | 541 | 44024.6836 | 541 | 44039.1367 | 541 | 43704.7188 | 541 | 43985.6406 | 541 | 44122.9609
542 | 44575.2617 | 20959.7383 | 542 | 46070.3906 | 542 | 43947.3477 | 542 | 43740.9102 | 542 | 44066.7266 | 542 | 44094.3633 | 542 | 43695 | 542 | 44000.2812 | 542 | 44114.4805
543 | 44717.1758 | 20817.8242 | 543 | 46169.1758 | 543 | 44070.957 | 543 | 43763.7266 | 543 | 44110.7734 | 543 | 44115.1367 | 543 | 43682.3984 | 543 | 44043.3984 | 543 | 44108.5195
544 | 44646.9141 | 20888.0859 | 544 | 46137.8711 | 544 | 44023.043 | 544 | 43773.3633 | 544 | 44143.0469 | 544 | 44110.3164 | 544 | 43647.1602 | 544 | 44041.8398 | 544 | 44100.2812
545 | 44721.5234 | 20813.4766 | 545 | 46192.7383 | 545 | 44043.3477 | 545 | 43781.1836 | 545 | 44188.5898 | 545 | 44123.5898 | 545 | 43627.3203 | 545 | 44033.2383 | 545 | 44098.8008
546 | 44711.8711 | 20823.1289 | 546 | 46242.4766 | 546 | 44075.6523 | 546 | 43803.9102 | 546 | 44212.5898 | 546 | 44136.2266 | 546 | 43652.8789 | 546 | 44104.3203 | 546 | 44098.3594
547 | 44745.3477 | 20789.6523 | 547 | 46296.3906 | 547 | 44110.2188 | 547 | 43836.1836 | 547 | 44262.9531 | 547 | 44078.1367 | 547 | 43687.1211 | 547 | 44124 | 547 | 44091.2812
548 | 44732.5234 | 20802.4766 | 548 | 46290.4336 | 548 | 44106.4766 | 548 | 43863.3164 | 548 | 44287.5898 | 548 | 44093.5 | 548 | 43711.5195 | 548 | 44120.8008 | 548 | 44096.0781
549 | 44704.0859 | 20830.9141 | 549 | 46264.7812 | 549 | 44048.3906 | 549 | 43872.3633 | 549 | 44253.8164 | 549 | 44096.5 | 549 | 43709.0781 | 549 | 44133.0391 | 549 | 44073.4805
550 | 44743.3047 | 20791.6953 | 550 | 46314.8711 | 550 | 44088.7383 | 550 | 43869.9531 | 550 | 44270.9531 | 550 | 44084.8164 | 550 | 43752.8789 | 550 | 44122.7617 | 550 | 44044.3594
551 | 44792.9141 | 20742.0859 | 551 | 46411.957 | 551 | 44163.9141 | 551 | 43892.5 | 551 | 44314.5469 | 551 | 44076.5469 | 551 | 43776.8789 | 551 | 44146.9219 | 551 | 44056.3984
552 | 44796.3047 | 20738.6953 | 552 | 46402.9141 | 552 | 44158.2188 | 552 | 43900.9531 | 552 | 44396.5 | 552 | 44082.4102 | 552 | 43792.6797 | 552 | 44172.0781 | 552 | 44039.9609
[7 empty] | 553 | 43907.3164 | 553 | 44412.3164 | 553 | 44082.0898 | 553 | 43760.5195 | 553 | 44176.2812 | 553 | 44005.8008
[7 empty] | 554 | 43946.2266 | 554 | 44410.7266 | 554 | 44091 | 554 | 43728.0781 | 554 | 44175 | 554 | 44038.4805
[7 empty] | 555 | 43972 | 555 | 44431.1367 | 555 | 44083.7734 | 555 | 43740.2383 | 555 | 44168.3203 | 555 | 44051.2383
[7 empty] | 556 | 43962.0898 | 556 | 44416.2266 | 556 | 44114.3633 | 556 | 43711.0781 | 556 | 44157.4805 | 556 | 44029.5586
[7 empty] | 557 | 43940.9102 | 557 | 44401 | 557 | 44133.7734 | 557 | 43699.9609 | 557 | 44137.4414 | 557 | 44039.2812
[7 empty] | 558 | 43928.9531 | 558 | 44407.0898 | 558 | 44135 | 558 | 43673.9609 | 558 | 44106.7188 | 558 | 44052.7188
[7 empty] | 559 | 43956.8633 | 559 | 44392.7734 | 559 | 44183.3164 | 559 | 43672.8789 | 559 | 44133.5586 | 559 | 44065.6016
[7 empty] | 560 | 43966.7734 | 560 | 44372.1367 | 560 | 44228.4102 | 560 | 43642.0781 | 560 | 44097.2812 | 560 | 44098.0391
[7 empty] | 561 | 43941.1836 | 561 | 44341 | 561 | 44237.6367 | 561 | 43615.5195 | 561 | 44068.8008 | 561 | 44076.2383
[7 empty] | 562 | 43924.7266 | 562 | 44337.9531 | 562 | 44218.9531 | 562 | 43641.3984 | 562 | 44049.2812 | 562 | 44058.6016
[7 empty] | 563 | 43899.6367 | 563 | 44311.5469 | 563 | 44175.8633 | 563 | 43644.6797 | 563 | 44109.9219 | 563 | 44049.9609
[7 empty] | 564 | 43832.1367 | 564 | 44319.1367 | 564 | 44185.3164 | 564 | 43716.1992 | 564 | 44136.8789 | 564 | 44035.1602
[7 empty] | 565 | 43837.0898 | 565 | 44315.6367 | 565 | 44168 | 565 | 43739.1211 | 565 | 44164.8008 | 565 | 44032.3594
[7 empty] | 566 | 43835.3164 | 566 | 44337.5469 | 566 | 44149.4102 | 566 | 43752 | 566 | 44193.0781 | 566 | 44043.3984
[7 empty] | 567 | 43809.5 | 567 | 44287.8164 | 567 | 44137.4531 | 567 | 43756.4805 | 567 | 44228.1992 | 567 | 44047.0781
[7 empty] | 568 | 43737.2266 | 568 | 44309.5 | 568 | 44185.6367 | 568 | 43789.0781 | 568 | 44264.1211 | 568 | 44053
[7 empty] | 569 | 43678.5898 | 569 | 44303.1367 | 569 | 44188.4531 | 569 | 43848.7617 | 569 | 44318.6016 | 569 | 44058.8789
[7 empty] | 570 | 43720.3164 | 570 | 44284.8633 | 570 | 44185.3164 | 570 | 43867.6797 | 570 | 44325.8789 | 570 | 44064.8398
[7 empty] | 571 | 43768.7734 | 571 | 44278.5 | 571 | 44212.7734 | 571 | 43864.5586 | 571 | 44329.9219 | 571 | 44070.3203
[7 empty] | 572 | 43780.7266 | 572 | 44293.9102 | 572 | 44228.1836 | 572 | 43874.4805 | 572 | 44350.6016 | 572 | 44096.1211
[7 empty] | 573 | 43788.6836 | 573 | 44395.2734 | 573 | 44186.4102 | 573 | 43872.7188 | 573 | 44388.4414 | 573 | 44115.8789
[7 empty] | 574 | 43822.0898 | 574 | 44369.8164 | 574 | 44080.5898 | 574 | 43909.8789 | 574 | 44426.0391 | 574 | 44140.5586
[7 empty] | 575 | 43856.3633 | 575 | 44343.0898 | 575 | 44027.5469 | 575 | 43908.1992 | 575 | 44434.3203 | 575 | 44182.1211
[7 empty] | 576 | 43897.7266 | 576 | 44399.1367 | 576 | 44012.2266 | 576 | 43928.6016 | 576 | 44441.1992 | 576 | 44197.7188
[7 empty] | 577 | 43907.3164 | 577 | 44422.9102 | 577 | 43973.9531 | 577 | 43939.4414 | 577 | 44437.6016 | 577 | 44192.9609
[7 empty] | 578 | 43899.5 | 578 | 44454.8164 | 578 | 43987.7734 | 578 | 43947.7617 | 578 | 44440.5195 | 578 | 44155.8008
[7 empty] | 579 | 43881.4102 | 579 | 44472.4531 | 579 | 43978.6836 | 579 | 43915.8789 | 579 | 44435.7188 | 579 | 44144.9609
[7 empty] | 580 | 43891.6367 | 580 | 44466.3164 | 580 | 43985.5469 | 580 | 43881.3203 | 580 | 44462.5586 | 580 | 44104.4414
[7 empty] | 581 | 43879.3164 | 581 | 44442.2266 | 581 | 43979.1367 | 581 | 43893.2383 | 581 | 44421.7617 | 581 | 44069.8398
[7 empty] | 582 | 43806.8633 | 582 | 44391.0469 | 582 | 43976.7266 | 582 | 43913.5195 | 582 | 44399.9219 | 582 | 44077.9609
[7 empty] | 583 | 43787.1836 | 583 | 44397.2734 | 583 | 44000.4102 | 583 | 43921.7617 | 583 | 44370.2812 | 583 | 44099
[7 empty] | 584 | 43780.3633 | 584 | 44383.6836 | 584 | 44025.5898 | 584 | 43868.6016 | 584 | 44342.1992 | 584 | 44091.2812
[7 empty] | 585 | 43716.2266 | 585 | 44395.1367 | 585 | 44027.5 | 585 | 43882.1211 | 585 | 44327.4805 | 585 | 44131.6797
[7 empty] | 586 | 43770.5 | 586 | 44405.0469 | 586 | 44028.6367 | 586 | 43792.0781 | 586 | 44325.0781 | 586 | 44169.3203
[7 empty] | 587 | 43792.5469 | 587 | 44386.6836 | 587 | 44025.6836 | 587 | 43779.6797 | 587 | 44309 | 587 | 44208.9219
[7 empty] | 588 | 43836.8164 | 588 | 44405.0898 | 588 | 44031.0898 | 588 | 43813.4414 | 588 | 44295.7617 | 588 | 44124.4414
[7 empty] | 589 | 43821.5 | 589 | 44358.7266 | 589 | 44013.1836 | 589 | 43775.3594 | 589 | 44288.6797 | 589 | 44049.9609
[7 empty] | 590 | 43829.9531 | 590 | 44389.3633 | 590 | 44053.6367 | 590 | 43710.4805 | 590 | 44241 | 590 | 43992.1602
[7 empty] | 591 | 43847.1836 | 591 | 44389.4531 | 591 | 44072.6367 | 591 | 43658.2383 | 591 | 44268.7188 | 591 | 43956.7617
[7 empty] | 592 | 43874.7266 | 592 | 44424.8164 | 592 | 44057.7266 | 592 | 43684.8789 | 592 | 44265.9219 | 592 | 43918.9219
[7 empty] | 593 | 43872.7266 | 593 | 44438.2266 | 593 | 44007.0898 | 593 | 43750.6406 | 593 | 44294.8008 | 593 | 43931.8789
[7 empty] | 594 | 43855.6367 | 594 | 44420.3164 | 594 | 44023.9102 | 594 | 43762.2383 | 594 | 44385.5195 | 594 | 43923.6406
[7 empty] | 595 | 43832.6836 | 595 | 44389.3633 | 595 | 44014.6836 | 595 | 43766.1602 | 595 | 44389.1602 | 595 | 43921.9219
[7 empty] | 596 | 43820.5469 | 596 | 44436.6367 | 596 | 44016.0469 | 596 | 43803.2383 | 596 | 44361.1992 | 596 | 43933.2383
[7 empty] | 597 | 43814.5 | 597 | 44452.5898 | 597 | 44033.6367 | 597 | 43852.8398 | 597 | 44381.6016 | 597 | 43940.6797
[7 empty] | 598 | 43744.1367 | 598 | 44473.0469 | 598 | 44032.4531 | 598 | 43899.9609 | 598 | 44391.2383 | 598 | 43970.8398
[7 empty] | 599 | 43715.8633 | 599 | 44467.4102 | 599 | 44069.2734 | 599 | 43841.3594 | 599 | 44448.5195 | 599 | 43973.6016
[7 empty] | 600 | 43743.1836 | 600 | 44428.2734 | 600 | 44081.7266 | 600 | 43876.1602 | 600 | 44500.3984 | 600 | 43999
[7 empty] | 601 | 43801.5 | 601 | 44446.5898 | 601 | 44062.6836 | 601 | 43872.0781 | 601 | 44505.3594 | 601 | 43992.4414
[7 empty] | 602 | 43867.2734 | 602 | 44470.0469 | 602 | 44049.6836 | 602 | 43872.0391 | 602 | 44505.0391 | 602 | 43982.1992
[7 empty] | 603 | 43916 | 603 | 44427.6836 | 603 | 44048.6367 | 603 | 43857.6797 | 603 | 44442.9219 | 603 | 44004.8789
[7 empty] | 604 | 43944.6836 | 604 | 44383 | 604 | 43991.8164 | 604 | 43790.3203 | 604 | 44368.2383 | 604 | 43973.8398
[7 empty] | 605 | 43916.8164 | 605 | 44388.7734 | 605 | 43960 | 605 | 43740.8008 | 605 | 44373.7188 | 605 | 44016.4414
[7 empty] | 606 | 43898.3164 | 606 | 44385.1367 | 606 | 43954.3633 | 606 | 43728.6406 | 606 | 44407.4414 | 606 | 44025.9219
[7 empty] | 607 | 43926.9102 | 607 | 44387.5898 | 607 | 43949.2734 | 607 | 43671.2383 | 607 | 44408.6797 | 607 | 44019.7617
[7 empty] | 608 | 43823.8633 | 608 | 44416 | 608 | 43969.6836 | 608 | 43730.0781 | 608 | 44378.9609 | 608 | 43979.8398
[7 empty] | 609 | 43839.8164 | 609 | 44436.2734 | 609 | 44052.4102 | 609 | 43795.7188 | 609 | 44406.8008 | 609 | 43995.0391
[7 empty] | 610 | 43870.1367 | 610 | 44429 | 610 | 44025.9102 | 610 | 43818.4414 | 610 | 44382.9219 | 610 | 43985.4414
[7 empty] | 611 | 43837 | 611 | 44456.1367 | 611 | 44031.6367 | 611 | 43802.6797 | 611 | 44365.3203 | 611 | 43979.1992
[7 empty] | 612 | 43811.9531 | 612 | 44540.2266 | 612 | 43976.6836 | 612 | 43834.6016 | 612 | 44394.1211 | 612 | 43993.0781
[7 empty] | 613 | 43816.0469 | 613 | 44521.0898 | 613 | 43966.3164 | 613 | 43851.1211 | 613 | 44415.5195 | 613 | 44001.4414
[7 empty] | 614 | 43788.0469 | 614 | 44546.6367 | 614 | 43952.2266 | 614 | 43868.1602 | 614 | 44413.0781 | 614 | 44021.3984
[7 empty] | 615 | 43788.7266 | 615 | 44587.7734 | 615 | 43968.4102 | 615 | 43854.6797 | 615 | 44420.2812 | 615 | 44022.0391
[7 empty] | 616 | 43776.5 | 616 | 44573.2266 | 616 | 43949.1367 | 616 | 43848.1992 | 616 | 44413.2383 | 616 | 43987.2812
[7 empty] | 617 | 43762 | 617 | 44569.0898 | 617 | 43970.2266 | 617 | 43845.5586 | 617 | 44423.1992 | 617 | 43990.2383
[7 empty] | 618 | 43763.5898 | 618 | 44594.2734 | 618 | 43977.6836 | 618 | 43843.4805 | 618 | 44457.1211 | 618 | 43979.6406
[7 empty] | 619 | 43798.7266 | 619 | 44589.4531 | 619 | 44005.1367 | 619 | 43818.8008 | 619 | 44500.2812 | 619 | 43938.5586
[7 empty] | 620 | 43813.7266 | 620 | 44622.7734 | 620 | 44006.4531 | 620 | 43767.3984 | 620 | 44479.8008 | 620 | 43916.1602
[7 empty] | 621 | 43833.9102 | 621 | 44642.1836 | 621 | 43994.9102 | 621 | 43715.2383 | 621 | 44473.4805 | 621 | 43922.0781
[7 empty] | 622 | 43842.6367 | 622 | 44570.1367 | 622 | 44028 | 622 | 43750.9219 | 622 | 44496.4805 | 622 | 43911.9219
[7 empty] | 623 | 43883.2266 | 623 | 44511.1836 | 623 | 44015.6367 | 623 | 43799.9609 | 623 | 44487.2812 | 623 | 43915.4414
[7 empty] | 624 | 43855.2266 | 624 | 44473.2734 | 624 | 44003.6367 | 624 | 43839.2383 | 624 | 44441.9609 | 624 | 43964.2383
[7 empty] | 625 | 43882.3633 | 625 | 44450.7734 | 625 | 44010.0469 | 625 | 43878.3594 | 625 | 44442.3594 | 625 | 43968.3594
[7 empty] | 626 | 43870.9102 | 626 | 44456.2734 | 626 | 44004.7266 | 626 | 43898.4805 | 626 | 44423.0781 | 626 | 43981.0781
[7 empty] | 627 | 43868.5898 | 627 | 44455.2266 | 627 | 44010.4531 | 627 | 43896.6016 | 627 | 44391.8789 | 627 | 43936.3984
[7 empty] | 628 | 43897.3633 | 628 | 44503.8633 | 628 | 44028.0469 | 628 | 43903.6016 | 628 | 44445.5195 | 628 | 43901.0391
[7 empty] | 629 | 43930.1836 | 629 | 44525.0898 | 629 | 44044.0469 | 629 | 43889.1211 | 629 | 44465.4414 | 629 | 43956.4805
[7 empty] | 630 | 43934.8164 | 630 | 44564.6367 | 630 | 44036.4102 | 630 | 43784.4414 | 630 | 44492.9219 | 630 | 43954.3594
[7 empty] | 631 | 43926.1836 | 631 | 44593.8164 | 631 | 44020.6836 | 631 | 43850.3984 | 631 | 44528.5586 | 631 | 43917.5586
[7 empty] | 632 | 43910.4531 | 632 | 44615.8633 | 632 | 43992.5469 | 632 | 43831 | 632 | 44548.8789 | 632 | 43948.7617
[7 empty] | 633 | 43964.3633 | 633 | 44614.5898 | 633 | 43997.2266 | 633 | 43812.8008 | 633 | 44553.0781 | 633 | 43953.7188
[7 empty] | 634 | 43951.0469 | 634 | 44608.6367 | 634 | 43939.5469 | 634 | 43792.7188 | 634 | 44562.8008 | 634 | 43921.9219
[7 empty] | 635 | 43947 | 635 | 44631.5 | 635 | 43958.5898 | 635 | 43757.8008 | 635 | 44550.3984 | 635 | 43947.9609
[7 empty] | 636 | 43917.7266 | 636 | 44673.8633 | 636 | 43979.6367 | 636 | 43748.6406 | 636 | 44555.9609 | 636 | 43940.2383
[7 empty] | 637 | 43977.9102 | 637 | 44646.7266 | 637 | 43994.8164 | 637 | 43773.5195 | 637 | 44565.5195 | 637 | 43971.2812
[7 empty] | 638 | 44010.7266 | 638 | 44635.5 | 638 | 44032.9531 | 638 | 43773.1602 | 638 | 44596.0781 | 638 | 43969.3594
[7 empty] | 639 | 44029.0469 | 639 | 44613.5469 | 639 | 44059.3164 | 639 | 43767.7617 | 639 | 44616.5586 | 639 | 43958.5195
[7 empty] | 640 | 44042.4102 | 640 | 44644.7734 | 640 | 44081.0898 | 640 | 43784.6797 | 640 | 44633.1992 | 640 | 43950.1602
[7 empty] | 641 | 44019.5 | 641 | 44696.5469 | 641 | 44094.4102 | 641 | 43804.8789 | 641 | 44666.6797 | 641 | 43941.7617
[7 empty] | 642 | 44014.3164 | 642 | 44731.4531 | 642 | 44042.1836 | 642 | 43816.0391 | 642 | 44669.5586 | 642 | 43931.2812
[7 empty] | 643 | 44034.5898 | 643 | 44730.7266 | 643 | 44035.6836 | 643 | 43839.6016 | 643 | 44627.7617 | 643 | 43955.6016
[7 empty] | 644 | 44033.7734 | 644 | 44734.3164 | 644 | 44028.7734 | 644 | 43845.6406 | 644 | 44554.8789 | 644 | 43978.3594
[7 empty] | 645 | 44021.1836 | 645 | 44793.2266 | 645 | 44023.5469 | 645 | 43869.5586 | 645 | 44530.0781 | 645 | 43982.3984
[7 empty] | 646 | 44019.1836 | 646 | 44837.9531 | 646 | 44040.7734 | 646 | 43861.8789 | 646 | 44527.2383 | 646 | 43941.4805
[7 empty] | 647 | 43998 | 647 | 44841.1836 | 647 | 44078 | 647 | 43854.7188 | 647 | 44537.9219 | 647 | 43915.8789
[7 empty] | 648 | 44017.5469 | 648 | 44822.4102 | 648 | 44121.2734 | 648 | 43838.6797 | 648 | 44539.8398 | 648 | 43939.7188
[7 empty] | 649 | 44039.7266 | 649 | 44836.9531 | 649 | 44127 | 649 | 43878.8008 | 649 | 44558.7617 | 649 | 43887.1602
[7 empty] | 650 | 44078.0898 | 650 | 44845.9531 | 650 | 44118.8164 | 650 | 43880.8008 | 650 | 44572.8008 | 650 | 43881.1602
[7 empty] | 651 | 44104 | 651 | 44811.0469 | 651 | 44116 | 651 | 43895.9219 | 651 | 44604.5586 | 651 | 43914.0391
[7 empty] | 652 | 44142.2734 | 652 | 44805.5 | 652 | 44135.0469 | 652 | 43949.8789 | 652 | 44619.0781 | 652 | 43948.7188
[7 empty] | 653 | 44125.6836 | 653 | 44818.0469 | 653 | 44147.7266 | 653 | 43943.0391 | 653 | 44612.8008 | 653 | 43977.3594
[7 empty] | 654 | 44142.0898 | 654 | 44852.5898 | 654 | 44139.6836 | 654 | 43908.9219 | 654 | 44628.5195 | 654 | 43987.9609
[13 empty] | 655 | 43962.9609 | 655 | 44637.6016 | 655 | 44015.2383
[13 empty] | 656 | 43955.5586 | 656 | 44644.6406 | 656 | 43997.8789
[13 empty] | 657 | 43924.3594 | 657 | 44681.2812 | 657 | 43998.4805
[13 empty] | 658 | 43916.0391 | 658 | 44654.5586 | 658 | 44006.1992
[13 empty] | 659 | 43939.2383 | 659 | 44604.9219 | 659 | 43986.6797
[13 empty] | 660 | 44003.8789 | 660 | 44617 | 660 | 43998.6406
[13 empty] | 661 | 44026.7617 | 661 | 44657.3594 | 661 | 43990.4805
[13 empty] | 662 | 44049.8008 | 662 | 44663.8008 | 662 | 43997.1602
[13 empty] | 663 | 44029.1211 | 663 | 44714.6016 | 663 | 44030.0781
[13 empty] | 664 | 43989.5586 | 664 | 44735.1992 | 664 | 44042.1992
[13 empty] | 665 | 44012.4805 | 665 | 44755.8008 | 665 | 44045.6406
[13 empty] | 666 | 44025.2383 | 666 | 44797.7617 | 666 | 44042.8398
[13 empty] | 667 | 43999.5586 | 667 | 44813.1992 | 667 | 44049.7617
[13 empty] | 668 | 44005.1211 | 668 | 44840.0391 | 668 | 44076.8789
[13 empty] | 669 | 44013.2383 | 669 | 44840.5586 | 669 | 44088.7617
[13 empty] | 670 | 44026.7188 | 670 | 44863.4414 | 670 | 44087.5195
[13 empty] | 671 | 44039.7188 | 671 | 44885.6406 | 671 | 44089.1992
[13 empty] | 672 | 44068.6016 | 672 | 44870.4414 | 672 | 44058.8789
[13 empty] | 673 | 44085 | 673 | 44839.6797 | 673 | 44056.6406
[13 empty] | 674 | 44123.4805 | 674 | 44857.1992 | 674 | 43991.6797
[13 empty] | 675 | 44082.8008 | 675 | 44867.6016 | 675 | 43947.5195
[13 empty] | 676 | 44137.2383 | 676 | 44856.9219 | 676 | 43926.3203
[13 empty] | 677 | 44096 | 677 | 44878.7617 | 677 | 43928.7188
[13 empty] | 678 | 44089.3203 | 678 | 44891.2383 | 678 | 43946.5586
[13 empty] | 679 | 44025.1992 | 679 | 44897.0391 | 679 | 43904.5195
[13 empty] | 680 | 44028.5195 | 680 | 44916.7188 | 680 | 43931.6797
[13 empty] | 681 | 43991.7617 | 681 | 44860.9609 | 681 | 43950.5195
[13 empty] | 682 | 44024.3203 | 682 | 44859.9609 | 682 | 43975.6797
[13 empty] | 683 | 44066.9219 | 683 | 44915.8398 | 683 | 44025.3594
[13 empty] | 684 | 44066.5195 | 684 | 44902.8398 | 684 | 44039.1602
[13 empty] | 685 | 44045.6797 | 685 | 44913.0391 | 685 | 44061.1992
[13 empty] | 686 | 44071.5586 | 686 | 44921.7188 | 686 | 43997.5586
[13 empty] | 687 | 44071.6406 | 687 | 44957.5195 | 687 | 43987.4414
[13 empty] | 688 | 44077.1602 | 688 | 44959.5195 | 688 | 44013.0391
[13 empty] | 689 | 44055.9609 | 689 | 45006.9219 | 689 | 44002.3984

## unif profiles px val 2
y=ax^2+bx+c
a | 5e-07
b | -0.0498
c | 1168.4
centre across |  |  | near to scan array, across |  | opposite to scan array, across |  | centre, down |  | left side, down |  | right side, down |  | centre, down repeated |  | left side, down, repeated |  | right side , down, repeated
0 | -53.1236 | 319.8992 | 0 | -50.1209 | 0 | -62.4222 | 0 | 43949 | 0 | 43265.0469 | 0 | 45790.5898 | 0 | 44048.9609 | 0 | 43328.2383 | 0 | 45893.4805
1 | -52.7713 | 318.2877 | 1 | -50.0744 | 1 | -61.9322 | 1 | 43955.2734 | 1 | 43268.2734 | 1 | 45776.5469 | 1 | 44011.4805 | 1 | 43321.6797 | 1 | 45858.3594
2 | -53.1228 | 319.8956 | 2 | -50.5416 | 2 | -61.8241 | 2 | 43948.4531 | 2 | 43226.1836 | 2 | 45768.4531 | 2 | 44005.9219 | 2 | 43315.1992 | 2 | 45851.1992
3 | -52.8579 | 318.6824 | 3 | -50.6876 | 3 | -61.5655 | 3 | 44004.3164 | 3 | 43255.1367 | 3 | 45740.7266 | 3 | 44022.8789 | 3 | 43334.6406 | 3 | 45866.8008
4 | -52.5397 | 317.2382 | 4 | -50.2288 | 4 | -61.1881 | 4 | 43955.2266 | 4 | 43280.3633 | 4 | 45733.8633 | 4 | 44082 | 4 | 43351.2383 | 4 | 45872.4805
5 | -53.0982 | 319.7825 | 5 | -50.7542 | 5 | -61.4509 | 5 | 43921.7266 | 5 | 43291.5 | 5 | 45749.7734 | 5 | 44090.2812 | 5 | 43385.5586 | 5 | 45888.0781
6 | -53.0126 | 319.3897 | 6 | -50.9816 | 6 | -61.3737 | 6 | 43875.1367 | 6 | 43325.7734 | 6 | 45716.0469 | 6 | 44102.8789 | 6 | 43363.6406 | 6 | 45904.4414
7 | -52.431 | 316.7485 | 7 | -50.191 | 7 | -60.8566 | 7 | 43964.5 | 7 | 43329.2266 | 7 | 45688.4102 | 7 | 44142.6797 | 7 | 43402.5586 | 7 | 45936
8 | -52.3145 | 316.2252 | 8 | -50.0575 | 8 | -60.6745 | 8 | 43819.9102 | 8 | 43354.4531 | 8 | 45667.5898 | 8 | 44069.8789 | 8 | 43373.7617 | 8 | 45897.6016
9 | -52.6587 | 317.7767 | 9 | -50.3435 | 9 | -60.4405 | 9 | 43888 | 9 | 43360.7266 | 9 | 45679.5469 | 9 | 44086.2812 | 9 | 43380.2383 | 9 | 45879.8008
10 | -52.8014 | 318.4249 | 10 | -50.7404 | 10 | -60.6519 | 10 | 43918.4531 | 10 | 43301.7734 | 10 | 45693.1836 | 10 | 44048.8789 | 10 | 43423.6406 | 10 | 45870.8789
11 | -52.4945 | 317.0345 | 11 | -50.7904 | 11 | -60.7881 | 11 | 44009.5 | 11 | 43341.5898 | 11 | 45686.4531 | 11 | 44017.4805 | 11 | 43472.8008 | 11 | 45822.2812
12 | -52.7713 | 318.2877 | 12 | -50.653 | 12 | -60.5181 | 12 | 43947.1836 | 12 | 43313.8164 | 12 | 45638.6836 | 12 | 43993.3203 | 12 | 43458.4414 | 12 | 45802.6016
13 | -52.4092 | 316.6503 | 13 | -50.3523 | 13 | -60.1325 | 13 | 43907.8633 | 13 | 43338.4102 | 13 | 45652.1367 | 13 | 43971.6406 | 13 | 43490.0781 | 13 | 45805.3984
14 | -52.6413 | 317.6978 | 14 | -50.6068 | 14 | -60.4502 | 14 | 43963.4531 | 14 | 43280.8164 | 14 | 45619 | 14 | 43997.3984 | 14 | 43432.8398 | 14 | 45772.3203
15 | -52.4334 | 316.7593 | 15 | -50.5216 | 15 | -60.1727 | 15 | 44000.3164 | 15 | 43292 | 15 | 45602.8164 | 15 | 43922.7617 | 15 | 43373.1992 | 15 | 45745.6016
16 | -51.9701 | 314.69 | 16 | -50.3912 | 16 | -60.0124 | 16 | 43999.0898 | 16 | 43248.5 | 16 | 45597.0469 | 16 | 43847.6406 | 16 | 43344.8789 | 16 | 45720.5586
17 | -51.6133 | 313.1169 | 17 | -50.7416 | 17 | -59.9322 | 17 | 44005.4102 | 17 | 43249.0469 | 17 | 45559.2266 | 17 | 43876.7617 | 17 | 43293.7188 | 17 | 45724
18 | -51.5037 | 312.637 | 18 | -50.435 | 18 | -59.8271 | 18 | 43963.8164 | 18 | 43248.3164 | 18 | 45612.2266 | 18 | 43926.0781 | 18 | 43244.3984 | 18 | 45723.2383
19 | -51.5145 | 312.684 | 19 | -50.077 | 19 | -59.7249 | 19 | 43921.4102 | 19 | 43257.1836 | 19 | 45652.7734 | 19 | 43908.7617 | 19 | 43223.1602 | 19 | 45690.4414
20 | -52.0887 | 315.2171 | 20 | -50.3685 | 20 | -59.9555 | 20 | 43890.5469 | 20 | 43271.3164 | 20 | 45710.5898 | 20 | 43878.5195 | 20 | 43161.3203 | 20 | 45709.6406
21 | -51.8106 | 313.9846 | 21 | -50.2504 | 21 | -59.8026 | 21 | 43865.5 | 21 | 43303.1367 | 21 | 45714.3164 | 21 | 43864.1602 | 21 | 43146.7188 | 21 | 45670.1992
22 | -51.7438 | 313.69 | 22 | -50.2873 | 22 | -59.5809 | 22 | 43819 | 22 | 43295.5469 | 22 | 45728.3633 | 22 | 43880.4805 | 22 | 43148.3203 | 22 | 45640.0781
23 | -51.8401 | 314.115 | 23 | -50.2629 | 23 | -59.451 | 23 | 43760.8164 | 23 | 43286 | 23 | 45710.5898 | 23 | 43916.3984 | 23 | 43132 | 23 | 45648.3203
24 | -51.7772 | 313.8373 | 24 | -50.2672 | 24 | -59.3681 | 24 | 43754.6367 | 24 | 43285.9102 | 24 | 45640.0898 | 24 | 43894.8789 | 24 | 43148.8789 | 24 | 45625.0781
25 | -51.5856 | 312.9951 | 25 | -50.1024 | 25 | -59.1543 | 25 | 43775.3164 | 25 | 43265.2266 | 25 | 45573.9531 | 25 | 43917.5195 | 25 | 43186.2383 | 25 | 45627.5586
26 | -51.2298 | 311.4447 | 26 | -49.8664 | 26 | -59.0601 | 26 | 43767.0469 | 26 | 43287.5469 | 26 | 45547.7266 | 26 | 43913.1602 | 26 | 43223.9219 | 26 | 45618.1602
27 | -51.344 | 311.9407 | 27 | -49.9698 | 27 | -59.1015 | 27 | 43805.2734 | 27 | 43260.6367 | 27 | 45506.9102 | 27 | 43860.8008 | 27 | 43237.8008 | 27 | 45565.7188
28 | -51.6403 | 313.2351 | 28 | -50.3574 | 28 | -59.2118 | 28 | 43824.1836 | 28 | 43273.9102 | 28 | 45482.6367 | 28 | 43813.0391 | 28 | 43274.7188 | 28 | 45550.1992
29 | -51.3819 | 312.1057 | 29 | -50.3929 | 29 | -59.022 | 29 | 43813.7734 | 29 | 43284.2266 | 29 | 45492 | 29 | 43888.9219 | 29 | 43270.8789 | 29 | 45558.3594
30 | -51.5034 | 312.6358 | 30 | -50.473 | 30 | -58.9647 | 30 | 43856.9531 | 30 | 43324.7734 | 30 | 45470 | 30 | 43784.0781 | 30 | 43226.2383 | 30 | 45585.3594
31 | -51.5798 | 312.9698 | 31 | -50.4291 | 31 | -58.8707 | 31 | 43900.9102 | 31 | 43320.5898 | 31 | 45463.9102 | 31 | 43861.6406 | 31 | 43175.7188 | 31 | 45545.5586
32 | -51.2284 | 311.4387 | 32 | -50.2183 | 32 | -58.3941 | 32 | 43866.9102 | 32 | 43362.6836 | 32 | 45467.5898 | 32 | 43937.1211 | 32 | 43225.6797 | 32 | 45538.6797
33 | -51.5297 | 312.7504 | 33 | -50.6625 | 33 | -58.5939 | 33 | 43880.8164 | 33 | 43393 | 33 | 45453.3633 | 33 | 43989.1211 | 33 | 43197 | 33 | 45543.5586
34 | -51.1008 | 310.8866 | 34 | -50.0704 | 34 | -58.091 | 34 | 43858.3164 | 34 | 43342.1836 | 34 | 45470.4102 | 34 | 43923.1992 | 34 | 43270 | 34 | 45612.1211
35 | -51.2292 | 311.4423 | 35 | -50.2487 | 35 | -58.352 | 35 | 43865.6836 | 35 | 43311.2266 | 35 | 45479.4531 | 35 | 43868.3203 | 35 | 43254.8789 | 35 | 45668.8008
36 | -51.346 | 311.9492 | 36 | -50.1234 | 36 | -58.4006 | 36 | 43847.3633 | 36 | 43257.0898 | 36 | 45449 | 36 | 43887 | 36 | 43239.1602 | 36 | 45661.1992
37 | -51.4905 | 312.5792 | 37 | -50.618 | 37 | -58.6354 | 37 | 43851.4531 | 37 | 43279.9102 | 37 | 45431.0469 | 37 | 43934.0781 | 37 | 43193.8008 | 37 | 45677.9609
38 | -51.4648 | 312.4671 | 38 | -50.3401 | 38 | -58.2859 | 38 | 43849.6367 | 38 | 43217.5469 | 38 | 45391 | 38 | 43922.5195 | 38 | 43168.2383 | 38 | 45642.8008
39 | -51.8705 | 314.2491 | 39 | -50.7208 | 39 | -58.4104 | 39 | 43834.1836 | 39 | 43167.3633 | 39 | 45379.3633 | 39 | 43951.4414 | 39 | 43207.2812 | 39 | 45585.7188
40 | -51.5767 | 312.9565 | 40 | -50.3696 | 40 | -57.994 | 40 | 43854.3633 | 40 | 43233.8633 | 40 | 45407.8164 | 40 | 43911.4414 | 40 | 43178.7188 | 40 | 45541.5586
41 | -51.804 | 313.9556 | 41 | -50.324 | 41 | -58.2054 | 41 | 43803.1367 | 41 | 43275.7266 | 41 | 45398.0898 | 41 | 43843.2383 | 41 | 43211.3594 | 41 | 45518.6797
42 | -51.9548 | 314.6224 | 42 | -50.618 | 42 | -58.234 | 42 | 43779.2266 | 42 | 43305.0469 | 42 | 45429.6367 | 42 | 43825.9219 | 42 | 43235.8008 | 42 | 45474.6016
43 | -51.9344 | 314.5317 | 43 | -50.8246 | 43 | -58.1986 | 43 | 43833.9531 | 43 | 43285.9102 | 43 | 45418.5898 | 43 | 43826.6016 | 43 | 43225.0391 | 43 | 45463.7188
44 | -51.7931 | 313.9073 | 44 | -50.7011 | 44 | -58.1255 | 44 | 43863.8164 | 44 | 43261.5898 | 44 | 45397.2266 | 44 | 43776.8008 | 44 | 43227.8789 | 44 | 45478.6797
45 | -51.5575 | 312.8722 | 45 | -50.4845 | 45 | -57.8686 | 45 | 43870.5898 | 45 | 43227 | 45 | 45380.7266 | 45 | 43723.6797 | 45 | 43232.4414 | 45 | 45423.1992
46 | -51.496 | 312.6033 | 46 | -50.4611 | 46 | -57.7805 | 46 | 43873.9531 | 46 | 43269.5 | 46 | 45362 | 46 | 43735.5195 | 46 | 43224.5195 | 46 | 45377.3594
47 | -51.5952 | 313.0374 | 47 | -50.5461 | 47 | -57.8989 | 47 | 43832.5469 | 47 | 43239.7734 | 47 | 45345.5898 | 47 | 43735.9609 | 47 | 43227.9219 | 47 | 45406.2812
48 | -51.4198 | 312.2708 | 48 | -50.3157 | 48 | -57.8983 | 48 | 43809.3164 | 48 | 43236.5898 | 48 | 45353.4531 | 48 | 43712.4414 | 48 | 43218.6016 | 48 | 45411
49 | -51.4397 | 312.3575 | 49 | -50.4243 | 49 | -57.8271 | 49 | 43826.7266 | 49 | 43259.8164 | 49 | 45386.2266 | 49 | 43751 | 49 | 43206.7188 | 49 | 45405
50 | -51.3271 | 311.8673 | 50 | -50.5077 | 50 | -57.6072 | 50 | 43864.5 | 50 | 43280.2266 | 50 | 45423.0898 | 50 | 43771.9219 | 50 | 43192.5195 | 50 | 45428.6016
51 | -51.3855 | 312.1213 | 51 | -50.6822 | 51 | -57.6555 | 51 | 43851.6367 | 51 | 43270.9531 | 51 | 45410.9102 | 51 | 43795.7617 | 51 | 43228.5195 | 51 | 45412.0781
52 | -51.0793 | 310.794 | 52 | -50.7202 | 52 | -57.7184 | 52 | 43843.5898 | 52 | 43283 | 52 | 45400.5 | 52 | 43811.8398 | 52 | 43245.5586 | 52 | 45394.6016
53 | -51.0681 | 310.7458 | 53 | -50.905 | 53 | -57.7143 | 53 | 43842.1367 | 53 | 43288.5469 | 53 | 45360.7266 | 53 | 43885.0391 | 53 | 43283.9609 | 53 | 45369.9609
54 | -51.0517 | 310.6749 | 54 | -51.0977 | 54 | -57.5268 | 54 | 43844.7734 | 54 | 43278.9531 | 54 | 45359.5898 | 54 | 43846.9609 | 54 | 43300.7617 | 54 | 45356.7188
55 | -50.9998 | 310.4513 | 55 | -50.888 | 55 | -57.548 | 55 | 43879.7734 | 55 | 43318.6367 | 55 | 45401 | 55 | 43838.6406 | 55 | 43260.1211 | 55 | 45354.3594
56 | -50.9805 | 310.3684 | 56 | -50.6082 | 56 | -57.5888 | 56 | 43918.4531 | 56 | 43345.0469 | 56 | 45344.1836 | 56 | 43810.4414 | 56 | 43210.9219 | 56 | 45366.3594
57 | -51.0071 | 310.4826 | 57 | -50.5241 | 57 | -57.5141 | 57 | 43826.6367 | 57 | 43358.3633 | 57 | 45320.2266 | 57 | 43814.7188 | 57 | 43166.0781 | 57 | 45380.7188
58 | -51.3133 | 311.807 | 58 | -51.0271 | 58 | -57.6686 | 58 | 43856.0898 | 58 | 43404.9102 | 58 | 45363.7266 | 58 | 43844.8008 | 58 | 43183.1992 | 58 | 45347.2812
59 | -51.3504 | 311.9684 | 59 | -51.2439 | 59 | -57.9397 | 59 | 43865.3164 | 59 | 43373.0898 | 59 | 45369.4102 | 59 | 43838.0391 | 59 | 43129.2383 | 59 | 45331.4414
60 | -50.4908 | 308.2766 | 60 | -51.0152 | 60 | -57.5226 | 60 | 43795.7266 | 60 | 43331.8633 | 60 | 45354.0469 | 60 | 43823.6406 | 60 | 43133.3203 | 60 | 45319.6016
61 | -51.346 | 311.9492 | 61 | -50.8818 | 61 | -57.7459 | 61 | 43796.3633 | 61 | 43233.2266 | 61 | 45346.7266 | 61 | 43797.1211 | 61 | 43195.9219 | 61 | 45316.8789
62 | -51.1386 | 311.0501 | 62 | -50.6535 | 62 | -57.5123 | 62 | 43826.8633 | 62 | 43248.1836 | 62 | 45388.7266 | 62 | 43830.7617 | 62 | 43223.2812 | 62 | 45333.8398
63 | -51.1225 | 310.9804 | 63 | -51.3158 | 63 | -57.3687 | 63 | 43801.6367 | 63 | 43114.7734 | 63 | 45391.2734 | 63 | 43791.1992 | 63 | 43226 | 63 | 45321.5195
64 | -51.2148 | 311.3797 | 64 | -51.1108 | 64 | -57.2894 | 64 | 43789.2734 | 64 | 43216.0898 | 64 | 45344.4102 | 64 | 43741.1992 | 64 | 43185.8008 | 64 | 45327.6406
65 | -51.2084 | 311.3521 | 65 | -51.3025 | 65 | -57.6098 | 65 | 43824.1836 | 65 | 43177.8633 | 65 | 45318.5 | 65 | 43773.8398 | 65 | 43158.8398 | 65 | 45344.5586
66 | -51.3016 | 311.7565 | 66 | -51.334 | 66 | -57.5524 | 66 | 43876.9531 | 66 | 43174.6367 | 66 | 45245.4531 | 66 | 43775.6016 | 66 | 43147.6797 | 66 | 45339.3594
67 | -51.3183 | 311.8287 | 67 | -51.2509 | 67 | -57.3364 | 67 | 43835.1836 | 67 | 43188.4102 | 67 | 45177.6836 | 67 | 43794.6016 | 67 | 43187.4414 | 67 | 45346.6016
68 | -51.3869 | 312.1274 | 68 | -51.1837 | 68 | -57.1164 | 68 | 43869.8164 | 68 | 43179.2734 | 68 | 45146.8164 | 68 | 43792.1992 | 68 | 43165.4805 | 68 | 45317.8789
69 | -51.401 | 312.1888 | 69 | -51.2398 | 69 | -57.1368 | 69 | 43884.5898 | 69 | 43254.4102 | 69 | 45184.6836 | 69 | 43754.4805 | 69 | 43158.0781 | 69 | 45316.4805
70 | -51.2003 | 311.3171 | 70 | -51.339 | 70 | -56.8182 | 70 | 43886 | 70 | 43363.4102 | 70 | 45218.5898 | 70 | 43745.0391 | 70 | 43199.3594 | 70 | 45332.1602
71 | -51.0865 | 310.8252 | 71 | -51.4228 | 71 | -56.7216 | 71 | 43877.4102 | 71 | 43445.0469 | 71 | 45246.0898 | 71 | 43788.1211 | 71 | 43239.5586 | 71 | 45308.0781
72 | -51.12 | 310.9696 | 72 | -51.5487 | 72 | -56.827 | 72 | 43892.0469 | 72 | 43464 | 72 | 45290.8633 | 72 | 43841.3594 | 72 | 43235.1211 | 72 | 45301.1992
73 | -51.1881 | 311.2643 | 73 | -51.4789 | 73 | -56.7023 | 73 | 43953 | 73 | 43494.6836 | 73 | 45311.8633 | 73 | 43838.6797 | 73 | 43258.4805 | 73 | 45303.0781
74 | -51.1161 | 310.9527 | 74 | -51.5657 | 74 | -56.708 | 74 | 44000.8164 | 74 | 43517.9531 | 74 | 45308.0469 | 74 | 43824.0781 | 74 | 43269.3203 | 74 | 45309.3203
75 | -51.1461 | 311.0826 | 75 | -51.9649 | 75 | -56.4531 | 75 | 44046.5898 | 75 | 43510.6836 | 75 | 45282.2266 | 75 | 43803.9609 | 75 | 43241.2383 | 75 | 45320.3594
76 | -51.3124 | 311.8034 | 76 | -52.0314 | 76 | -56.3577 | 76 | 43985.2266 | 76 | 43478.1836 | 76 | 45283.9102 | 76 | 43802.0391 | 76 | 43243.0781 | 76 | 45322.1602
77 | -51.2223 | 311.4123 | 77 | -52.2252 | 77 | -56.8119 | 77 | 44003.1836 | 77 | 43421 | 77 | 45289.1367 | 77 | 43819.9609 | 77 | 43260.6016 | 77 | 45328.0781
78 | -51.2961 | 311.7324 | 78 | -51.9101 | 78 | -56.3951 | 78 | 44083.8633 | 78 | 43401.4531 | 78 | 45298.0898 | 78 | 43871.6406 | 78 | 43293.6406 | 78 | 45328.0391
79 | -51.1745 | 311.2053 | 79 | -51.6711 | 79 | -56.0363 | 79 | 43982.7734 | 79 | 43373.5 | 79 | 45292.6836 | 79 | 43818.5195 | 79 | 43302.0781 | 79 | 45304.4805
80 | -50.5701 | 308.6135 | 80 | -51.8574 | 80 | -55.966 | 80 | 43998.4531 | 80 | 43370.5 | 80 | 45337.5 | 80 | 43807.0781 | 80 | 43287.7188 | 80 | 45261.3984
81 | -50.8835 | 309.9515 | 81 | -51.9393 | 81 | -56.2619 | 81 | 43851.1836 | 81 | 43364.5469 | 81 | 45287.2734 | 81 | 43822.8398 | 81 | 43268 | 81 | 45240.8398
82 | -51.5451 | 312.8179 | 82 | -52.4808 | 82 | -56.8152 | 82 | 43933.4531 | 82 | 43370.3633 | 82 | 45295.2734 | 82 | 43783.8008 | 82 | 43215.9609 | 82 | 45149.5586
83 | -51.4499 | 312.4021 | 83 | -52.5893 | 83 | -56.9591 | 83 | 43904.2734 | 83 | 43361.0469 | 83 | 45276.4531 | 83 | 43801.3984 | 83 | 43191.4414 | 83 | 45158.2383
84 | -50.7761 | 309.4916 | 84 | -51.758 | 84 | -56.2103 | 84 | 43894.8633 | 84 | 43360.1836 | 84 | 45302.1836 | 84 | 43792.8398 | 84 | 43132.3203 | 84 | 45145.6406
85 | -50.8723 | 309.9035 | 85 | -51.6546 | 85 | -55.8117 | 85 | 43865.6836 | 85 | 43323.4531 | 85 | 45252.5 | 85 | 43749.2383 | 85 | 43163.3984 | 85 | 45177.3594
86 | -50.7469 | 309.3668 | 86 | -51.7783 | 86 | -55.9066 | 86 | 43907.4531 | 86 | 43335.1367 | 86 | 45241.6367 | 86 | 43750.8398 | 86 | 43163.8789 | 86 | 45204.3594
87 | -50.9188 | 310.1029 | 87 | -52.0765 | 87 | -55.9249 | 87 | 43912.6836 | 87 | 43316.7734 | 87 | 45175.3633 | 87 | 43804.3594 | 87 | 43121.6406 | 87 | 45236.5195
88 | -50.8577 | 309.841 | 88 | -52.0849 | 88 | -55.9378 | 88 | 43873.3633 | 88 | 43302.8633 | 88 | 45194.1836 | 88 | 43854.8398 | 88 | 43150.6797 | 88 | 45265.5195
89 | -50.5904 | 308.6998 | 89 | -51.7605 | 89 | -55.3948 | 89 | 43844.0469 | 89 | 43267.5898 | 89 | 45189.6836 | 89 | 43801.0781 | 89 | 43165.8398 | 89 | 45270.5195
90 | -50.8686 | 309.8878 | 90 | -52.1067 | 90 | -55.6644 | 90 | 43766.3164 | 90 | 43233.0898 | 90 | 45209.5898 | 90 | 43819.8789 | 90 | 43230.6406 | 90 | 45250.7617
91 | -50.821 | 309.6837 | 91 | -52.0637 | 91 | -55.6376 | 91 | 43713.5898 | 91 | 43249.9531 | 91 | 45146.3164 | 91 | 43829.3984 | 91 | 43306.2812 | 91 | 45252.0781
92 | -50.7542 | 309.398 | 92 | -51.761 | 92 | -55.1699 | 92 | 43729.5469 | 92 | 43254.4531 | 92 | 45137.2734 | 92 | 43835.9609 | 92 | 43347.2383 | 92 | 45220.5586
93 | -50.6983 | 309.1592 | 93 | -51.6205 | 93 | -55.0375 | 93 | 43723.6836 | 93 | 43286.3633 | 93 | 45134.7734 | 93 | 43813.1992 | 93 | 43389.4414 | 93 | 45211.9219
94 | -50.8238 | 309.6957 | 94 | -51.5632 | 94 | -55.0547 | 94 | 43713.5469 | 94 | 43285.2734 | 94 | 45114.8633 | 94 | 43829.7188 | 94 | 43417.5195 | 94 | 45238.1602
95 | -51.1149 | 310.9478 | 95 | -51.6804 | 95 | -55.1197 | 95 | 43683.2266 | 95 | 43289.8164 | 95 | 45078.7734 | 95 | 43901.7617 | 95 | 43432.8789 | 95 | 45278.8008
96 | -50.9252 | 310.1305 | 96 | -51.856 | 96 | -55.1537 | 96 | 43683.5898 | 96 | 43301.2266 | 96 | 45050.2266 | 96 | 43982.1602 | 96 | 43422.5195 | 96 | 45290.3984
97 | -50.5588 | 308.5655 | 97 | -51.7764 | 97 | -54.8403 | 97 | 43700.8164 | 97 | 43297.9102 | 97 | 45039.8164 | 97 | 44009.4805 | 97 | 43383.7188 | 97 | 45263.7617
98 | -50.4896 | 308.2719 | 98 | -51.5905 | 98 | -54.9613 | 98 | 43740.5 | 98 | 43289.4102 | 98 | 45039.3633 | 98 | 43972.0781 | 98 | 43334.2812 | 98 | 45224.7188
99 | -50.5207 | 308.4037 | 99 | -51.3828 | 99 | -54.9904 | 99 | 43789.3633 | 99 | 43294.2734 | 99 | 45042.0469 | 99 | 43949.9219 | 99 | 43302.4805 | 99 | 45269.8008
100 | -50.4992 | 308.3126 | 100 | -51.4585 | 100 | -54.9833 | 100 | 43741.6367 | 100 | 43249.8633 | 100 | 45060.9102 | 100 | 43980.3203 | 100 | 43335.8008 | 100 | 45220.1602
101 | -50.4387 | 308.0561 | 101 | -51.4726 | 101 | -54.954 | 101 | 43752 | 101 | 43275.7734 | 101 | 45077.6836 | 101 | 43919.1211 | 101 | 43348.0781 | 101 | 45159.4414
102 | -50.1029 | 306.6413 | 102 | -51.3648 | 102 | -54.8056 | 102 | 43856.5898 | 102 | 43295.9531 | 102 | 45052.3164 | 102 | 43855.5586 | 102 | 43308.5586 | 102 | 45155.6406
103 | -50.1847 | 306.9846 | 103 | -51.1369 | 103 | -54.7086 | 103 | 43862.1367 | 103 | 43272.0898 | 103 | 45033.6836 | 103 | 43861.5586 | 103 | 43314.1992 | 103 | 45129.6016
104 | -50.3364 | 307.6238 | 104 | -51.3861 | 104 | -54.9176 | 104 | 43834.4531 | 104 | 43306.2734 | 104 | 45077.5 | 104 | 43862.8008 | 104 | 43315.9219 | 104 | 45141.1211
105 | -50.1326 | 306.7657 | 105 | -51.2201 | 105 | -54.5631 | 105 | 43840.0469 | 105 | 43294.3633 | 105 | 45090 | 105 | 43841.9609 | 105 | 43285.6406 | 105 | 45139.1992
106 | -50.1932 | 307.0205 | 106 | -51.3008 | 106 | -54.414 | 106 | 43841.6836 | 106 | 43333.6367 | 106 | 45101.7266 | 106 | 43854.2812 | 106 | 43257.6797 | 106 | 45153.7617
107 | -50.214 | 307.1079 | 107 | -51.496 | 107 | -55.04 | 107 | 43825.2266 | 107 | 43318.6836 | 107 | 45086.1367 | 107 | 43791.7188 | 107 | 43269.7188 | 107 | 45113.7617
108 | -50.5173 | 308.3893 | 108 | -51.6348 | 108 | -55.0527 | 108 | 43772.5898 | 108 | 43297.5 | 108 | 45066.4531 | 108 | 43859.7617 | 108 | 43242.7617 | 108 | 45114.8008
109 | -49.8945 | 305.7701 | 109 | -50.9844 | 109 | -54.6674 | 109 | 43712.5898 | 109 | 43290.5 | 109 | 45082.3164 | 109 | 43849.5195 | 109 | 43248.2812 | 109 | 45090.8789
110 | -50.2552 | 307.2814 | 110 | -51.3559 | 110 | -54.7326 | 110 | 43750.0898 | 110 | 43245.8164 | 110 | 45126.4102 | 110 | 43813.7188 | 110 | 43206.2383 | 110 | 45037.0781
111 | -50.2601 | 307.3017 | 111 | -51.4665 | 111 | -54.797 | 111 | 43745.6836 | 111 | 43287.8633 | 111 | 45120 | 111 | 43787.4414 | 111 | 43211.3203 | 111 | 45023.0781
112 | -50.5111 | 308.3629 | 112 | -51.8871 | 112 | -55.065 | 112 | 43674.2734 | 112 | 43306.2734 | 112 | 45087.2266 | 112 | 43716.1992 | 112 | 43177.4414 | 112 | 45020.8398
113 | -50.5433 | 308.4995 | 113 | -51.8828 | 113 | -55.0972 | 113 | 43723 | 113 | 43269.5 | 113 | 45015.3633 | 113 | 43676.7188 | 113 | 43195.6016 | 113 | 44998.0781
114 | -50.3339 | 307.613 | 114 | -51.7163 | 114 | -54.7283 | 114 | 43721.1367 | 114 | 43233.9102 | 114 | 45004.9102 | 114 | 43639.0391 | 114 | 43207.8789 | 114 | 44992
115 | -50.28 | 307.3855 | 115 | -51.5891 | 115 | -54.7674 | 115 | 43698.7266 | 115 | 43189.8633 | 115 | 45039.5469 | 115 | 43631.7188 | 115 | 43204.3203 | 115 | 45000.7617
116 | -50.6006 | 308.743 | 116 | -51.9131 | 116 | -55.0134 | 116 | 43697 | 116 | 43165.5469 | 116 | 45015.5898 | 116 | 43663.6797 | 116 | 43207.7188 | 116 | 45036.4414
117 | -50.8642 | 309.8686 | 117 | -51.7912 | 117 | -55.0006 | 117 | 43673.2734 | 117 | 43179.9102 | 117 | 45026.3633 | 117 | 43693.6797 | 117 | 43214.9219 | 117 | 45025.6406
118 | -51.0288 | 310.5763 | 118 | -51.6815 | 118 | -54.9816 | 118 | 43706.2734 | 118 | 43204.3164 | 118 | 45005.8633 | 118 | 43725 | 118 | 43214.3203 | 118 | 44988.5586
119 | -51.0913 | 310.8456 | 119 | -51.7569 | 119 | -54.959 | 119 | 43804.7734 | 119 | 43138.2266 | 119 | 44965.4102 | 119 | 43694.1602 | 119 | 43245.6797 | 119 | 45022.9219
120 | -50.8246 | 309.6994 | 120 | -51.8352 | 120 | -54.7419 | 120 | 43848.5469 | 120 | 43135.0469 | 120 | 44934.1367 | 120 | 43651.7617 | 120 | 43236.1211 | 120 | 45013.0391
121 | -50.9182 | 310.1004 | 121 | -51.7556 | 121 | -54.7121 | 121 | 43856.1367 | 121 | 43216.9102 | 121 | 44912.6367 | 121 | 43682.6797 | 121 | 43201.8008 | 121 | 45031.9609
122 | -50.5295 | 308.4409 | 122 | -51.5371 | 122 | -54.4525 | 122 | 43912.1836 | 122 | 43293.9102 | 122 | 44932.7266 | 122 | 43799.8789 | 122 | 43229.4805 | 122 | 45069.4414
123 | -50.5879 | 308.689 | 123 | -51.8207 | 123 | -54.9791 | 123 | 43813.4531 | 123 | 43351.0898 | 123 | 44938.9102 | 123 | 43762.1211 | 123 | 43236.3203 | 123 | 45036.1992
124 | -51.1419 | 311.0645 | 124 | -52.2157 | 124 | -55.0757 | 124 | 43880.4102 | 124 | 43372.1836 | 124 | 44965.0469 | 124 | 43759 | 124 | 43216.8398 | 124 | 45053.4805
125 | -50.9495 | 310.235 | 125 | -52.3652 | 125 | -54.9901 | 125 | 43816 | 125 | 43355.1836 | 125 | 44949.2266 | 125 | 43800.6797 | 125 | 43236 | 125 | 45079.8398
126 | -50.5309 | 308.4469 | 126 | -52.3072 | 126 | -54.6997 | 126 | 43828.4531 | 126 | 43350.2734 | 126 | 44949.3633 | 126 | 43812.6406 | 126 | 43246.9219 | 126 | 45080.5586
127 | -50.8627 | 309.8626 | 127 | -52.2325 | 127 | -55.1445 | 127 | 43896.2734 | 127 | 43352.9531 | 127 | 44984.6836 | 127 | 43772.6016 | 127 | 43260.5586 | 127 | 45015.2383
128 | -51.1375 | 311.0453 | 128 | -52.5047 | 128 | -55.2621 | 128 | 43908.1836 | 128 | 43379.3633 | 128 | 44996.9102 | 128 | 43817.0391 | 128 | 43251.9219 | 128 | 44955.1211
129 | -50.914 | 310.0824 | 129 | -51.9431 | 129 | -55.0927 | 129 | 43877.8633 | 129 | 43376.5469 | 129 | 44969.8164 | 129 | 43804.0391 | 129 | 43228.3594 | 129 | 44982.5586
130 | -51.5252 | 312.731 | 130 | -52.6619 | 130 | -55.7975 | 130 | 43936.2266 | 130 | 43391.4531 | 130 | 44947.8164 | 130 | 43765 | 130 | 43247.2812 | 130 | 45003.4414
131 | -50.9923 | 310.4189 | 131 | -52.5761 | 131 | -55.1654 | 131 | 43845.9102 | 131 | 43450.9102 | 131 | 44954.8633 | 131 | 43726.3594 | 131 | 43226.2812 | 131 | 44969.5586
132 | -51.1636 | 311.1584 | 132 | -52.544 | 132 | -55.3155 | 132 | 43782.1836 | 132 | 43402.1836 | 132 | 44977.3164 | 132 | 43711.6016 | 132 | 43243.1602 | 132 | 44978.1211
133 | -51.1914 | 311.2787 | 133 | -52.66 | 133 | -55.0792 | 133 | 43777.5 | 133 | 43394.4102 | 133 | 44972.8164 | 133 | 43665.2812 | 133 | 43253.0781 | 133 | 44983.6016
134 | -51.4543 | 312.4213 | 134 | -52.6445 | 134 | -55.1464 | 134 | 43756.0469 | 134 | 43357.3633 | 134 | 44957.7734 | 134 | 43642.6406 | 134 | 43206.0781 | 134 | 44939.2812
135 | -51.6417 | 313.2412 | 135 | -52.5724 | 135 | -55.3095 | 135 | 43729.9531 | 135 | 43300.2734 | 135 | 44952.0898 | 135 | 43652.6016 | 135 | 43218.8789 | 135 | 44902.8789
136 | -51.2811 | 311.6674 | 136 | -52.6609 | 136 | -55.1597 | 136 | 43809.4102 | 136 | 43276.3164 | 136 | 44928 | 136 | 43649.2812 | 136 | 43161.7188 | 136 | 44884.9609
137 | -51.4778 | 312.5238 | 137 | -52.6043 | 137 | -55.2608 | 137 | 43776.9102 | 137 | 43303.4102 | 137 | 44908.6836 | 137 | 43641.8398 | 137 | 43142.4805 | 137 | 44902.6016
138 | -51.2028 | 311.3279 | 138 | -52.7002 | 138 | -55.0347 | 138 | 43723.5469 | 138 | 43310.4531 | 138 | 44917.5469 | 138 | 43622.6406 | 138 | 43095.8789 | 138 | 44918.8398
139 | -51.0871 | 310.8276 | 139 | -53.0537 | 139 | -55.0677 | 139 | 43685.0469 | 139 | 43265.6367 | 139 | 44895.7266 | 139 | 43663.0781 | 139 | 43108.0391 | 139 | 44913.4414
140 | -51.0938 | 310.8564 | 140 | -53.0134 | 140 | -55.0319 | 140 | 43695.0898 | 140 | 43289.0898 | 140 | 44936.2734 | 140 | 43693.3594 | 140 | 43079.8008 | 140 | 44906.8398
141 | -50.7064 | 309.194 | 141 | -52.4186 | 141 | -54.7207 | 141 | 43716.0898 | 141 | 43276.8633 | 141 | 44969.6367 | 141 | 43666.1602 | 141 | 43063.6406 | 141 | 44916.5586
142 | -51.0609 | 310.7146 | 142 | -53.0198 | 142 | -55.1599 | 142 | 43682.8633 | 142 | 43297.4102 | 142 | 44958.6836 | 142 | 43751.9219 | 142 | 43144.6016 | 142 | 44914.1992
143 | -51.3116 | 311.7999 | 143 | -53.2274 | 143 | -55.2095 | 143 | 43716.6836 | 143 | 43319.6836 | 143 | 44959.1836 | 143 | 43767.6406 | 143 | 43213.7188 | 143 | 44930.3984
144 | -51.2628 | 311.588 | 144 | -53.1624 | 144 | -54.9756 | 144 | 43738.1836 | 144 | 43285.9102 | 144 | 44915.2266 | 144 | 43731.7188 | 144 | 43259.1211 | 144 | 44935.5195
145 | -51.1444 | 311.0753 | 145 | -52.855 | 145 | -54.5334 | 145 | 43786 | 145 | 43232.1367 | 145 | 44874.0898 | 145 | 43864.6016 | 145 | 43264 | 145 | 44945.4805
146 | -51.2378 | 311.4796 | 146 | -53.1096 | 146 | -54.5042 | 146 | 43716.9102 | 146 | 43237 | 146 | 44859.9531 | 146 | 43820.3203 | 146 | 43303.3203 | 146 | 44950.6406
147 | -50.7842 | 309.5265 | 147 | -52.705 | 147 | -54.3187 | 147 | 43648.0469 | 147 | 43202.2266 | 147 | 44819.0469 | 147 | 43865.6016 | 147 | 43307.1992 | 147 | 44919.2812
148 | -51.1205 | 310.9719 | 148 | -53.0722 | 148 | -54.5527 | 148 | 43672.3633 | 148 | 43225.0898 | 148 | 44789.2266 | 148 | 43793.9609 | 148 | 43286.7617 | 148 | 44920.1992
149 | -50.6273 | 308.8569 | 149 | -52.5963 | 149 | -54.1583 | 149 | 43694.4102 | 149 | 43307 | 149 | 44791.2266 | 149 | 43825.9219 | 149 | 43296.3594 | 149 | 44922.3203
150 | -51.0031 | 310.4658 | 150 | -52.7875 | 150 | -54.414 | 150 | 43718.1367 | 150 | 43287.7266 | 150 | 44791.9531 | 150 | 43827.0391 | 150 | 43357.2383 | 150 | 44907.9219
151 | -51.2062 | 311.3424 | 151 | -52.7251 | 151 | -54.6273 | 151 | 43717.2734 | 151 | 43241.8633 | 151 | 44802.7734 | 151 | 43806.5586 | 151 | 43362.5195 | 151 | 44895.0781
152 | -51.1403 | 311.0573 | 152 | -52.7694 | 152 | -54.7205 | 152 | 43628.5898 | 152 | 43201.5469 | 152 | 44774.5 | 152 | 43748.9609 | 152 | 43330.3203 | 152 | 44842.1602
153 | -50.9548 | 310.2578 | 153 | -52.7042 | 153 | -54.2124 | 153 | 43689.2266 | 153 | 43194.4102 | 153 | 44775.4531 | 153 | 43787.3203 | 153 | 43314.2383 | 153 | 44858.5195
154 | -50.4698 | 308.1879 | 154 | -53.0047 | 154 | -54.3926 | 154 | 43711.0469 | 154 | 43200.1836 | 154 | 44829.5 | 154 | 43774.4805 | 154 | 43290.6016 | 154 | 44859.9219
155 | -51.0829 | 310.8095 | 155 | -52.972 | 155 | -54.4492 | 155 | 43765.3633 | 155 | 43189.3633 | 155 | 44799 | 155 | 43690.0781 | 155 | 43283.2812 | 155 | 44859.5586
156 | -51.237 | 311.4761 | 156 | -52.964 | 156 | -54.5136 | 156 | 43739.8164 | 156 | 43215.8633 | 156 | 44774.8633 | 156 | 43677.4414 | 156 | 43251.7188 | 156 | 44903.0391
157 | -51.0219 | 310.5463 | 157 | -52.656 | 157 | -54.4871 | 157 | 43657.0898 | 157 | 43204.3633 | 157 | 44815.6367 | 157 | 43695.1211 | 157 | 43210.1992 | 157 | 44884.6016
158 | -50.8344 | 309.7413 | 158 | -52.9396 | 158 | -54.4469 | 158 | 43652.2266 | 158 | 43195.9102 | 158 | 44856.7266 | 158 | 43709.9219 | 158 | 43226.9219 | 158 | 44874.1992
159 | -50.6352 | 308.8905 | 159 | -53.0577 | 159 | -54.5667 | 159 | 43660.5 | 159 | 43136.8633 | 159 | 44809.3633 | 159 | 43692.9609 | 159 | 43250.3594 | 159 | 44854.3203
160 | -51.038 | 310.616 | 160 | -53.1733 | 160 | -54.6699 | 160 | 43689.2266 | 160 | 43168.7734 | 160 | 44793.3633 | 160 | 43695.9219 | 160 | 43208 | 160 | 44832.9609
161 | -51.0439 | 310.6413 | 161 | -53.0974 | 161 | -54.6266 | 161 | 43696.8164 | 161 | 43162.2266 | 161 | 44816.2734 | 161 | 43667.6406 | 161 | 43217.0391 | 161 | 44804.4414
162 | -50.8625 | 309.8614 | 162 | -52.9425 | 162 | -54.4612 | 162 | 43751 | 162 | 43190.8633 | 162 | 44776.2266 | 162 | 43641.1602 | 162 | 43243.5586 | 162 | 44787.9609
163 | -50.9649 | 310.3011 | 163 | -53.2126 | 163 | -54.6949 | 163 | 43753.6367 | 163 | 43215.0898 | 163 | 44745.4531 | 163 | 43635.9609 | 163 | 43278.3203 | 163 | 44757.8789
164 | -51.1085 | 310.9202 | 164 | -53.4416 | 164 | -54.57 | 164 | 43684.4531 | 164 | 43242.0469 | 164 | 44775.7734 | 164 | 43633.3984 | 164 | 43261.4414 | 164 | 44754.7617
165 | -51.2833 | 311.677 | 165 | -53.5876 | 165 | -54.7861 | 165 | 43754.2734 | 165 | 43234 | 165 | 44749.2266 | 165 | 43661.8398 | 165 | 43210.3984 | 165 | 44760.1602
166 | -51.393 | 312.1539 | 166 | -53.324 | 166 | -54.9221 | 166 | 43820.2734 | 166 | 43230.3633 | 166 | 44763.5469 | 166 | 43737.9609 | 166 | 43150.7188 | 166 | 44749.0391
167 | -51.2828 | 311.6746 | 167 | -52.8099 | 167 | -54.6091 | 167 | 43762.8164 | 167 | 43236.0469 | 167 | 44777.4102 | 167 | 43804.0781 | 167 | 43118.3203 | 167 | 44739.7188
168 | -51.0902 | 310.8409 | 168 | -52.9728 | 168 | -54.3558 | 168 | 43703.5469 | 168 | 43247.7266 | 168 | 44780.8164 | 168 | 43720 | 168 | 43116.2812 | 168 | 44739.1602
169 | -51.1506 | 311.1018 | 169 | -53.3818 | 169 | -54.2329 | 169 | 43629.3164 | 169 | 43290.4102 | 169 | 44785.1367 | 169 | 43622.8008 | 169 | 43172.3203 | 169 | 44760.8398
170 | -51.0361 | 310.6076 | 170 | -53.6033 | 170 | -53.9681 | 170 | 43598.4102 | 170 | 43309.8164 | 170 | 44754.9531 | 170 | 43626.1992 | 170 | 43245.4805 | 170 | 44766.5586
171 | -50.8465 | 309.793 | 171 | -53.6435 | 171 | -53.8071 | 171 | 43606.2266 | 171 | 43303.5898 | 171 | 44706.5 | 171 | 43672.8789 | 171 | 43226.9609 | 171 | 44711.9609
172 | -50.6211 | 308.8305 | 172 | -52.7104 | 172 | -53.7356 | 172 | 43566.5469 | 172 | 43280.5898 | 172 | 44698.9531 | 172 | 43653.5195 | 172 | 43187.3984 | 172 | 44776.7188
173 | -51.0826 | 310.8084 | 173 | -53.0084 | 173 | -54.0484 | 173 | 43485.8633 | 173 | 43208.9531 | 173 | 44688.2266 | 173 | 43666.0781 | 173 | 43195.9219 | 173 | 44771.1211
174 | -50.9277 | 310.1413 | 174 | -52.9026 | 174 | -53.8895 | 174 | 43500.3164 | 174 | 43157.4102 | 174 | 44681.6367 | 174 | 43616.1211 | 174 | 43158.7188 | 174 | 44739.5586
175 | -50.9959 | 310.4346 | 175 | -53.0955 | 175 | -53.8993 | 175 | 43638.5898 | 175 | 43137.5898 | 175 | 44659.3164 | 175 | 43645.0781 | 175 | 43158.1992 | 175 | 44733.7188
176 | -51.4632 | 312.4599 | 176 | -53.197 | 176 | -54.1911 | 176 | 43636.5 | 176 | 43127.7266 | 176 | 44635.7266 | 176 | 43638.9219 | 176 | 43125.7188 | 176 | 44782.0781
177 | -50.6462 | 308.9373 | 177 | -52.8006 | 177 | -53.6375 | 177 | 43681.9531 | 177 | 43148.2266 | 177 | 44647.5469 | 177 | 43703.3984 | 177 | 43150.5586 | 177 | 44749.5195
178 | -50.7101 | 309.2097 | 178 | -52.8137 | 178 | -53.6797 | 178 | 43628.0469 | 178 | 43218.6367 | 178 | 44660.0469 | 178 | 43664.7188 | 178 | 43080.9609 | 178 | 44731.8008
179 | -49.7392 | 305.1241 | 179 | -53.1688 | 179 | -53.4165 | 179 | 43608.2266 | 179 | 43291.5469 | 179 | 44689.0898 | 179 | 43617.8789 | 179 | 43142.4414 | 179 | 44737.7188
180 | -50.8359 | 309.7474 | 180 | -53.3866 | 180 | -53.9043 | 180 | 43606.9531 | 180 | 43329.0469 | 180 | 44700.5469 | 180 | 43624.3203 | 180 | 43097.2383 | 180 | 44711.2812
181 | -50.9333 | 310.1653 | 181 | -53.4427 | 181 | -53.9844 | 181 | 43562.4102 | 181 | 43326.1836 | 181 | 44726.3633 | 181 | 43605.6016 | 181 | 43104.0781 | 181 | 44729.0781
182 | -50.8019 | 309.6021 | 182 | -53.1674 | 182 | -53.726 | 182 | 43525.9531 | 182 | 43232.6367 | 182 | 44794.7734 | 182 | 43642.5195 | 182 | 43111.3594 | 182 | 44740.5195
183 | -50.6499 | 308.9529 | 183 | -53.0648 | 183 | -53.6737 | 183 | 43550.2734 | 183 | 43150.9102 | 183 | 44766.3164 | 183 | 43669.5586 | 183 | 43138.8789 | 183 | 44716.6797
184 | -50.997 | 310.4393 | 184 | -53.4752 | 184 | -54.0306 | 184 | 43573.8633 | 184 | 43200.3633 | 184 | 44734.4531 | 184 | 43676.3984 | 184 | 43136.5195 | 184 | 44690.3984
185 | -50.755 | 309.4016 | 185 | -53.2656 | 185 | -53.8991 | 185 | 43574.1367 | 185 | 43141.5469 | 185 | 44683.3164 | 185 | 43660.4414 | 185 | 43150.4805 | 185 | 44694.1211
186 | -50.6299 | 308.8677 | 186 | -53.2966 | 186 | -53.7336 | 186 | 43580.2266 | 186 | 43158.3164 | 186 | 44626.3164 | 186 | 43660.2812 | 186 | 43154.4414 | 186 | 44667.8398
187 | -50.5743 | 308.6315 | 187 | -53.2953 | 187 | -53.6779 | 187 | 43596.7266 | 187 | 43188.4102 | 187 | 44633.5 | 187 | 43685.2383 | 187 | 43155.6406 | 187 | 44657.3984
188 | -50.8863 | 309.9635 | 188 | -53.0736 | 188 | -53.9115 | 188 | 43628.9531 | 188 | 43194.2734 | 188 | 44583.6367 | 188 | 43705.3984 | 188 | 43181.2812 | 188 | 44655.2812
189 | -51.1116 | 310.9335 | 189 | -53.0126 | 189 | -54.1922 | 189 | 43620.4102 | 189 | 43248.3633 | 189 | 44538.7266 | 189 | 43680.9609 | 189 | 43195.0391 | 189 | 44634.7617
190 | -50.9738 | 310.3395 | 190 | -52.9258 | 190 | -54.0172 | 190 | 43588.9102 | 190 | 43255.5898 | 190 | 44526.1836 | 190 | 43612.8789 | 190 | 43229.1211 | 190 | 44591.8398
191 | -51.1831 | 311.2425 | 191 | -53.3398 | 191 | -54.0236 | 191 | 43568.6367 | 191 | 43245.9531 | 191 | 44516.3164 | 191 | 43547.6797 | 191 | 43258.7617 | 191 | 44599.4805
192 | -51.1664 | 311.1703 | 192 | -53.3385 | 192 | -53.979 | 192 | 43591.8164 | 192 | 43218.4102 | 192 | 44542.6836 | 192 | 43532.3984 | 192 | 43213.2383 | 192 | 44628.4805
193 | -51.229 | 311.4411 | 193 | -53.1814 | 193 | -54.0004 | 193 | 43564.8164 | 193 | 43214.2734 | 193 | 44588.7266 | 193 | 43548.5195 | 193 | 43180.4414 | 193 | 44567.6797
194 | -51.4212 | 312.2767 | 194 | -53.5024 | 194 | -53.9244 | 194 | 43560.6836 | 194 | 43218.5 | 194 | 44604.1367 | 194 | 43498.1602 | 194 | 43133.0781 | 194 | 44605.8008
195 | -51.2783 | 311.6554 | 195 | -53.4369 | 195 | -54.1416 | 195 | 43586.2734 | 195 | 43231.6836 | 195 | 44629.6367 | 195 | 43481.1211 | 195 | 43055.6016 | 195 | 44633.8008
196 | -51.2916 | 311.7131 | 196 | -53.3348 | 196 | -53.7962 | 196 | 43638.6367 | 196 | 43272 | 196 | 44631.0898 | 196 | 43445.3203 | 196 | 43051 | 196 | 44680.6016
197 | -51.6603 | 313.3231 | 197 | -53.3855 | 197 | -53.7941 | 197 | 43640.0898 | 197 | 43315 | 197 | 44602.7734 | 197 | 43547.1992 | 197 | 43085.6016 | 197 | 44751.3594
198 | -51.4336 | 312.331 | 198 | -53.7868 | 198 | -53.8094 | 198 | 43590.6836 | 198 | 43388.9531 | 198 | 44594.8633 | 198 | 43534.2812 | 198 | 43127.2383 | 198 | 44721.4805
199 | -51.4004 | 312.1863 | 199 | -53.6781 | 199 | -54.0347 | 199 | 43556.7266 | 199 | 43381.6367 | 199 | 44644.5469 | 199 | 43596.3984 | 199 | 43167.1602 | 199 | 44706.2383
200 | -51.2342 | 311.464 | 200 | -53.2118 | 200 | -53.9645 | 200 | 43567.4531 | 200 | 43340.4531 | 200 | 44670.7734 | 200 | 43539.8789 | 200 | 43235.0781 | 200 | 44647.8008
201 | -51.3194 | 311.8335 | 201 | -53.5957 | 201 | -53.867 | 201 | 43498.2734 | 201 | 43323.6367 | 201 | 44651.5 | 201 | 43534.4805 | 201 | 43205.7617 | 201 | 44605.3594
202 | -51.2941 | 311.7239 | 202 | -53.5364 | 202 | -53.7551 | 202 | 43456.5469 | 202 | 43297.4102 | 202 | 44660.5898 | 202 | 43527.3594 | 202 | 43216.4414 | 202 | 44588.3203
203 | -51.0163 | 310.5222 | 203 | -53.0918 | 203 | -53.1519 | 203 | 43507.9531 | 203 | 43263.5898 | 203 | 44649.5469 | 203 | 43483.9609 | 203 | 43116.8789 | 203 | 44551.7188
204 | -51.6919 | 313.4619 | 204 | -53.7707 | 204 | -53.5678 | 204 | 43572.3633 | 204 | 43231.1836 | 204 | 44620.7734 | 204 | 43460.3984 | 204 | 43050.3984 | 204 | 44489.7617
205 | -51.7034 | 313.5126 | 205 | -53.6604 | 205 | -53.7538 | 205 | 43605.2266 | 205 | 43235.7734 | 205 | 44635.9102 | 205 | 43490.3203 | 205 | 43104.2383 | 205 | 44467
206 | -51.6991 | 313.4933 | 206 | -53.639 | 206 | -53.8452 | 206 | 43624.0898 | 206 | 43232.6836 | 206 | 44637.6367 | 206 | 43513.8008 | 206 | 43096.6016 | 206 | 44461.8398
207 | -51.5431 | 312.8094 | 207 | -53.4283 | 207 | -53.7793 | 207 | 43632.3633 | 207 | 43232.8633 | 207 | 44589.5 | 207 | 43524.5195 | 207 | 43101.1602 | 207 | 44501.5586
208 | -52.0624 | 315.0998 | 208 | -53.549 | 208 | -54.2409 | 208 | 43666.5 | 208 | 43236.0898 | 208 | 44595.8633 | 208 | 43539 | 208 | 43090.7188 | 208 | 44546.4414
209 | -51.8418 | 314.1222 | 209 | -53.7034 | 209 | -54.1819 | 209 | 43691.7266 | 209 | 43165.6836 | 209 | 44614.3164 | 209 | 43542.8398 | 209 | 43150.3203 | 209 | 44568.0391
210 | -51.5861 | 312.9975 | 210 | -53.5973 | 210 | -53.7751 | 210 | 43642.3164 | 210 | 43148.5 | 210 | 44555.1367 | 210 | 43548.8789 | 210 | 43166.2812 | 210 | 44594
211 | -51.5319 | 312.7601 | 211 | -53.4665 | 211 | -53.8877 | 211 | 43621 | 211 | 43173.1836 | 211 | 44542.2266 | 211 | 43540.8398 | 211 | 43182 | 211 | 44603.7617
212 | -51.7284 | 313.6224 | 212 | -53.8307 | 212 | -54.1934 | 212 | 43584.3164 | 212 | 43233.5 | 212 | 44556.4102 | 212 | 43535.8789 | 212 | 43173.3203 | 212 | 44576.9609
213 | -51.8587 | 314.1971 | 213 | -53.8753 | 213 | -54.2035 | 213 | 43592.5 | 213 | 43215.8164 | 213 | 44629.9531 | 213 | 43528.7617 | 213 | 43183.8008 | 213 | 44543
214 | -51.6532 | 313.2918 | 214 | -53.7286 | 214 | -54.0143 | 214 | 43580.5 | 214 | 43223 | 214 | 44655.4531 | 214 | 43544.1992 | 214 | 43164.7617 | 214 | 44589.4805
215 | -51.9046 | 314.4 | 215 | -54.0154 | 215 | -54.3029 | 215 | 43538.0898 | 215 | 43263.8633 | 215 | 44650.0898 | 215 | 43515.2812 | 215 | 43140.5586 | 215 | 44599.8398
216 | -51.7607 | 313.7648 | 216 | -54.223 | 216 | -54.4624 | 216 | 43554.7734 | 216 | 43275.6367 | 216 | 44664.1367 | 216 | 43507.9609 | 216 | 43145.3984 | 216 | 44584.8789
217 | -51.8418 | 314.1222 | 217 | -54.2129 | 217 | -54.6027 | 217 | 43569.0469 | 217 | 43267.5898 | 217 | 44627 | 217 | 43564.6016 | 217 | 43175.5195 | 217 | 44592.2812
218 | -51.9744 | 314.7093 | 218 | -54.0071 | 218 | -54.3683 | 218 | 43512.9102 | 218 | 43241.9102 | 218 | 44585 | 218 | 43577.6797 | 218 | 43237.9609 | 218 | 44591.3984
219 | -51.8626 | 314.214 | 219 | -53.8309 | 219 | -54.2406 | 219 | 43490.0469 | 219 | 43231.7266 | 219 | 44589.9531 | 219 | 43590.0391 | 219 | 43289.0391 | 219 | 44589.7617
220 | -52.1211 | 315.3611 | 220 | -53.9193 | 220 | -54.4535 | 220 | 43469.3164 | 220 | 43196.5898 | 220 | 44559.1367 | 220 | 43538.0391 | 220 | 43262.5586 | 220 | 44588.7617
221 | -52.1428 | 315.4578 | 221 | -53.8846 | 221 | -54.5946 | 221 | 43480.4102 | 221 | 43201.0469 | 221 | 44516.2734 | 221 | 43497.0781 | 221 | 43219.4805 | 221 | 44569.6016
222 | -51.8259 | 314.0522 | 222 | -53.7642 | 222 | -53.9929 | 222 | 43429.9102 | 222 | 43195.0469 | 222 | 44475.0898 | 222 | 43467.3594 | 222 | 43226.9219 | 222 | 44544
223 | -52.0923 | 315.2329 | 223 | -54.0698 | 223 | -54.2332 | 223 | 43427.8164 | 223 | 43179.1367 | 223 | 44453.6836 | 223 | 43428.6406 | 223 | 43200.8398 | 223 | 44561.7188
224 | -52.3047 | 316.1816 | 224 | -54.2676 | 224 | -54.3384 | 224 | 43434.9102 | 224 | 43218 | 224 | 44459.0469 | 224 | 43434.8398 | 224 | 43174.6406 | 224 | 44527.7188
225 | -51.9867 | 314.7637 | 225 | -53.6255 | 225 | -53.9661 | 225 | 43439.4102 | 225 | 43158.8633 | 225 | 44489 | 225 | 43466.3203 | 225 | 43152.8398 | 225 | 44478.5586
226 | -52.5394 | 317.237 | 226 | -54.0502 | 226 | -54.102 | 226 | 43491.7266 | 226 | 43129.5 | 226 | 44473.7734 | 226 | 43510.1992 | 226 | 43130.8398 | 226 | 44481.5586
227 | -52.516 | 317.1314 | 227 | -53.6909 | 227 | -54.2653 | 227 | 43457.5469 | 227 | 43177.2266 | 227 | 44457.0898 | 227 | 43559.5586 | 227 | 43165.3984 | 227 | 44512.9609
228 | -52.1243 | 315.3755 | 228 | -53.5286 | 228 | -53.737 | 228 | 43439.7266 | 228 | 43217.1836 | 228 | 44479.6367 | 228 | 43593.1992 | 228 | 43153.2383 | 228 | 44573.9219
229 | -52.1224 | 315.3671 | 229 | -53.9924 | 229 | -53.7905 | 229 | 43441.1367 | 229 | 43273.6836 | 229 | 44528.7266 | 229 | 43608.6016 | 229 | 43152.8008 | 229 | 44588.1992
230 | -51.8106 | 313.9846 | 230 | -53.7671 | 230 | -53.9224 | 230 | 43405.1367 | 230 | 43255.9102 | 230 | 44530.5898 | 230 | 43617.5195 | 230 | 43099.9609 | 230 | 44604.3594
231 | -52.0743 | 315.153 | 231 | -54.0633 | 231 | -54.0577 | 231 | 43400.2734 | 231 | 43285.9531 | 231 | 44489 | 231 | 43627.6797 | 231 | 43120.8789 | 231 | 44602.3984
232 | -52.1233 | 315.3708 | 232 | -54.2893 | 232 | -54.307 | 232 | 43425.9531 | 232 | 43287.3164 | 232 | 44492.3633 | 232 | 43610.7188 | 232 | 43130.9609 | 232 | 44568.7188
233 | -51.8795 | 314.2889 | 233 | -54.1572 | 233 | -54.234 | 233 | 43468.7266 | 233 | 43308.5469 | 233 | 44429.6836 | 233 | 43578.0391 | 233 | 43189.1211 | 233 | 44538.3203
234 | -52.133 | 315.4142 | 234 | -54.3974 | 234 | -54.3405 | 234 | 43513.6367 | 234 | 43242.0469 | 234 | 44431.5469 | 234 | 43530.1211 | 234 | 43176.0781 | 234 | 44535.8398
235 | -51.6977 | 313.4872 | 235 | -53.987 | 235 | -53.9759 | 235 | 43542.9531 | 235 | 43240.8633 | 235 | 44443.8633 | 235 | 43539.8008 | 235 | 43196.8008 | 235 | 44503.2812
236 | -51.6774 | 313.398 | 236 | -54.1552 | 236 | -53.8509 | 236 | 43520.6367 | 236 | 43195.4102 | 236 | 44444.9531 | 236 | 43542.7188 | 236 | 43232.2383 | 236 | 44470.9609
237 | -51.8918 | 314.3433 | 237 | -54.4173 | 237 | -53.9226 | 237 | 43525.8633 | 237 | 43193.7266 | 237 | 44429.8633 | 237 | 43538.3594 | 237 | 43232.0781 | 237 | 44448.5195
238 | -52.0751 | 315.1566 | 238 | -54.4612 | 238 | -54.0267 | 238 | 43502.6836 | 238 | 43185.5 | 238 | 44423.8164 | 238 | 43542.8008 | 238 | 43208.1211 | 238 | 44438.5195
239 | -51.7432 | 313.6876 | 239 | -54.5657 | 239 | -53.9929 | 239 | 43500.6367 | 239 | 43166.2734 | 239 | 44394.8164 | 239 | 43510.3203 | 239 | 43187.9609 | 239 | 44426.9219
240 | -51.7624 | 313.7721 | 240 | -54.1603 | 240 | -53.9027 | 240 | 43508.0898 | 240 | 43132.3633 | 240 | 44376.7266 | 240 | 43483.3203 | 240 | 43186.1602 | 240 | 44439.3203
241 | -52.0077 | 314.8568 | 241 | -54.2478 | 241 | -53.91 | 241 | 43485.6367 | 241 | 43112.0898 | 241 | 44377.5469 | 241 | 43440.7617 | 241 | 43167.4805 | 241 | 44448.5586
242 | -51.8005 | 313.9399 | 242 | -54.0027 | 242 | -53.9575 | 242 | 43482.1367 | 242 | 43111.6836 | 242 | 44425.9102 | 242 | 43442.3594 | 242 | 43152.3594 | 242 | 44429.8789
243 | -52.0398 | 314.9995 | 243 | -53.9438 | 243 | -54.2006 | 243 | 43456.0898 | 243 | 43159.6836 | 243 | 44464.0469 | 243 | 43431.3984 | 243 | 43144.4414 | 243 | 44450.2383
244 | -51.9499 | 314.6005 | 244 | -54.4326 | 244 | -54.4275 | 244 | 43428.8633 | 244 | 43214.9102 | 244 | 44470.8633 | 244 | 43380.3984 | 244 | 43131.6016 | 244 | 44482.6797
245 | -52.136 | 315.4276 | 245 | -54.0252 | 245 | -54.4667 | 245 | 43472.1836 | 245 | 43223.6836 | 245 | 44496.4531 | 245 | 43386.8789 | 245 | 43130.2812 | 245 | 44447
246 | -52.0444 | 315.02 | 246 | -53.6901 | 246 | -54.3987 | 246 | 43467.4531 | 246 | 43234.2734 | 246 | 44491.2266 | 246 | 43390.8398 | 246 | 43085.4805 | 246 | 44438.9609
247 | -52.1515 | 315.4965 | 247 | -53.5518 | 247 | -54.2752 | 247 | 43486.2266 | 247 | 43193.6836 | 247 | 44313.6367 | 247 | 43406.5195 | 247 | 43045.6797 | 247 | 44454.9609
248 | -52.0719 | 315.1421 | 248 | -54.2142 | 248 | -54.4214 | 248 | 43441.6836 | 248 | 43223.9102 | 248 | 44322.6367 | 248 | 43429.9219 | 248 | 43077.6406 | 248 | 44395.3984
249 | -52.2282 | 315.839 | 249 | -54.0546 | 249 | -54.3578 | 249 | 43421.6836 | 249 | 43244.4102 | 249 | 44403.6367 | 249 | 43434.9609 | 249 | 43144.1602 | 249 | 44396.7188
250 | -51.8552 | 314.1814 | 250 | -53.9581 | 250 | -54.1975 | 250 | 43450.2266 | 250 | 43266.3164 | 250 | 44419.5469 | 250 | 43433.9609 | 250 | 43196.0781 | 250 | 44395.4805
251 | -52.2255 | 315.8269 | 251 | -53.8939 | 251 | -54.4657 | 251 | 43496.9102 | 251 | 43249.0469 | 251 | 44393.1836 | 251 | 43403.3594 | 251 | 43191.4805 | 251 | 44400.6016
252 | -51.7569 | 313.7479 | 252 | -53.5913 | 252 | -54.5672 | 252 | 43491.8164 | 252 | 43299.5898 | 252 | 44414.7734 | 252 | 43363.8398 | 252 | 43180.7617 | 252 | 44397.2383
253 | -51.8183 | 314.0184 | 253 | -53.6307 | 253 | -54.5705 | 253 | 43517.3633 | 253 | 43263.1836 | 253 | 44396.2266 | 253 | 43375.5586 | 253 | 43199.1602 | 253 | 44412.8789
254 | -51.9235 | 314.4834 | 254 | -54.0115 | 254 | -54.6904 | 254 | 43532.0469 | 254 | 43219.4531 | 254 | 44443.3633 | 254 | 43395.2812 | 254 | 43244.3984 | 254 | 44377.4805
255 | -52.342 | 316.3487 | 255 | -54.1194 | 255 | -54.8408 | 255 | 43545.8633 | 255 | 43289.1367 | 255 | 44439.4102 | 255 | 43417.2812 | 255 | 43209.4414 | 255 | 44350.3984
256 | -51.972 | 314.6985 | 256 | -54.1315 | 256 | -54.5563 | 256 | 43520.2266 | 256 | 43325.1367 | 256 | 44448.6367 | 256 | 43438.5195 | 256 | 43168.1992 | 256 | 44367.6406
257 | -52.0966 | 315.2522 | 257 | -53.8501 | 257 | -54.5588 | 257 | 43473.7266 | 257 | 43357.1836 | 257 | 44439.8164 | 257 | 43468.2812 | 257 | 43168.6016 | 257 | 44392.7617
258 | -52.2993 | 316.1573 | 258 | -54.1598 | 258 | -54.7621 | 258 | 43470.6367 | 258 | 43382.6367 | 258 | 44354.9102 | 258 | 43449.5586 | 258 | 43175.1602 | 258 | 44424.8008
259 | -52.1547 | 315.5111 | 259 | -54.0437 | 259 | -54.5469 | 259 | 43481 | 259 | 43354.5469 | 259 | 44297 | 259 | 43472.1602 | 259 | 43144.2383 | 259 | 44438.5195
260 | -52.3334 | 316.3099 | 260 | -54.0399 | 260 | -54.4087 | 260 | 43494.8164 | 260 | 43373.3633 | 260 | 44233.9531 | 260 | 43471.5586 | 260 | 43131 | 260 | 44445.7617
261 | -52.5088 | 317.0988 | 261 | -54.2145 | 261 | -54.5197 | 261 | 43523.1836 | 261 | 43320.4102 | 261 | 44228.4102 | 261 | 43475.6797 | 261 | 43112.6797 | 261 | 44424.9609
262 | -52.5233 | 317.1642 | 262 | -54.2176 | 262 | -54.541 | 262 | 43501.7734 | 262 | 43284.9531 | 262 | 44280.3633 | 262 | 43447.8789 | 262 | 43081.0781 | 262 | 44310.5586
263 | -52.3415 | 316.3463 | 263 | -53.979 | 263 | -54.4265 | 263 | 43536.5898 | 263 | 43308.9102 | 263 | 44354.3164 | 263 | 43423.8008 | 263 | 43064.2812 | 263 | 44290.7188
264 | -51.7871 | 313.8808 | 264 | -53.8037 | 264 | -54.0893 | 264 | 43552.3164 | 264 | 43308.7734 | 264 | 44411.3164 | 264 | 43437.4414 | 264 | 43089.7617 | 264 | 44373.7188
265 | -51.8951 | 314.3577 | 265 | -53.9361 | 265 | -54.1071 | 265 | 43495.7734 | 265 | 43315.6367 | 265 | 44426.9102 | 265 | 43426.1992 | 265 | 43106.8008 | 265 | 44378.7188
266 | -51.9804 | 314.7359 | 266 | -54.1164 | 266 | -54.1688 | 266 | 43509.5469 | 266 | 43343.9531 | 266 | 44396.5469 | 266 | 43392.1602 | 266 | 43119.3594 | 266 | 44371.8008
267 | -51.8962 | 314.3626 | 267 | -54.3185 | 267 | -54.2163 | 267 | 43488.4102 | 267 | 43346.0898 | 267 | 44380.5898 | 267 | 43412.3594 | 267 | 43134.7188 | 267 | 44394.8008
268 | -51.5949 | 313.0361 | 268 | -54.0352 | 268 | -53.8398 | 268 | 43478.8633 | 268 | 43334.7734 | 268 | 44415.0898 | 268 | 43388 | 268 | 43122.5195 | 268 | 44399.0781
269 | -51.6515 | 313.2845 | 269 | -53.8846 | 269 | -54.0911 | 269 | 43496.8164 | 269 | 43312.7734 | 269 | 44418.6836 | 269 | 43403.5195 | 269 | 43167.8398 | 269 | 44411.9609
270 | -52.1056 | 315.2921 | 270 | -54.5243 | 270 | -54.3011 | 270 | 43520.3164 | 270 | 43309.7734 | 270 | 44373.3633 | 270 | 43312.8008 | 270 | 43211.0781 | 270 | 44385.7617
271 | -50.9599 | 310.2795 | 271 | -54.2711 | 271 | -53.8582 | 271 | 43545.7266 | 271 | 43345.0898 | 271 | 44385.2734 | 271 | 43411.0781 | 271 | 43200.3203 | 271 | 44384.4414
272 | -52.0586 | 315.0829 | 272 | -54.2834 | 272 | -54.2599 | 272 | 43520.5469 | 272 | 43370.0898 | 272 | 44407.0898 | 272 | 43413.7188 | 272 | 43212.0391 | 272 | 44400.2383
273 | -51.9491 | 314.597 | 273 | -54.2206 | 273 | -54.1493 | 273 | 43479.5469 | 273 | 43419.7734 | 273 | 44408.6836 | 273 | 43420.1992 | 273 | 43220.5195 | 273 | 44334.3594
274 | -52.0319 | 314.9644 | 274 | -54.296 | 274 | -54.2491 | 274 | 43505.4531 | 274 | 43376.2266 | 274 | 44390.6836 | 274 | 43432.4414 | 274 | 43143.6797 | 274 | 44267.7188
275 | -52.101 | 315.2715 | 275 | -54.197 | 275 | -54.2378 | 275 | 43498.3164 | 275 | 43374.4531 | 275 | 44394.4102 | 275 | 43438.9219 | 275 | 43162.6016 | 275 | 44226.9609
276 | -52.1501 | 315.4904 | 276 | -54.3415 | 276 | -54.6537 | 276 | 43489.3164 | 276 | 43368.3633 | 276 | 44379.8164 | 276 | 43516.2383 | 276 | 43237.6797 | 276 | 44205.9219
277 | -51.8874 | 314.324 | 277 | -53.9172 | 277 | -54.4161 | 277 | 43536.5 | 277 | 43415.2734 | 277 | 44363.1836 | 277 | 43499.3594 | 277 | 43266.3203 | 277 | 44233.9609
278 | -52.3115 | 316.2118 | 278 | -54.2896 | 278 | -54.6096 | 278 | 43597.0469 | 278 | 43390 | 278 | 44278.5 | 278 | 43485 | 278 | 43299.2812 | 278 | 44304.0781
279 | -52.1414 | 315.4517 | 279 | -54.1547 | 279 | -54.6144 | 279 | 43606.5898 | 279 | 43415.8633 | 279 | 44311.9102 | 279 | 43476 | 279 | 43348.9609 | 279 | 44346.7188
280 | -52.1401 | 315.4458 | 280 | -54.2011 | 280 | -54.7798 | 280 | 43647.6836 | 280 | 43470.7734 | 280 | 44309.4531 | 280 | 43458.4414 | 280 | 43326.2383 | 280 | 44357.5195
281 | -52.4601 | 316.8793 | 281 | -54.4489 | 281 | -55.0865 | 281 | 43670.4531 | 281 | 43450.0469 | 281 | 44283.2734 | 281 | 43445.8008 | 281 | 43300.1602 | 281 | 44341.7188
282 | -52.2871 | 316.1028 | 282 | -54.5311 | 282 | -54.9645 | 282 | 43658.7266 | 282 | 43505 | 282 | 44301.3164 | 282 | 43442.6406 | 282 | 43254.0391 | 282 | 44326.7617
283 | -52.5329 | 317.2078 | 283 | -54.4764 | 283 | -55.0898 | 283 | 43638.5 | 283 | 43543.1367 | 283 | 44342.5898 | 283 | 43478 | 283 | 43246.8789 | 283 | 44335.9219
284 | -52.7964 | 318.4018 | 284 | -54.6126 | 284 | -55.104 | 284 | 43607.5 | 284 | 43476.3164 | 284 | 44367 | 284 | 43482.0391 | 284 | 43250.1992 | 284 | 44339.3984
285 | -53.0646 | 319.628 | 285 | -54.7402 | 285 | -55.0184 | 285 | 43601.8164 | 285 | 43423.2266 | 285 | 44430.5898 | 285 | 43514.0781 | 285 | 43228.3203 | 285 | 44315.4414
286 | -53.0063 | 319.3605 | 286 | -54.6755 | 286 | -54.9095 | 286 | 43615.4102 | 286 | 43400.4102 | 286 | 44460.3164 | 286 | 43492.1992 | 286 | 43252.1211 | 286 | 44344.8398
287 | -52.9492 | 319.0992 | 287 | -54.8058 | 287 | -54.7942 | 287 | 43634.3633 | 287 | 43361.7734 | 287 | 44489.6836 | 287 | 43482.3984 | 287 | 43275.8398 | 287 | 44395.8789
288 | -53.0081 | 319.369 | 288 | -55.0397 | 288 | -54.952 | 288 | 43657.6367 | 288 | 43369.6836 | 288 | 44499.7266 | 288 | 43490.5195 | 288 | 43228.7617 | 288 | 44393
289 | -53.178 | 320.1499 | 289 | -55.1542 | 289 | -54.9452 | 289 | 43654.7266 | 289 | 43431.8633 | 289 | 44512.8164 | 289 | 43451.2812 | 289 | 43228.1602 | 289 | 44365.7188
290 | -53.1571 | 320.0538 | 290 | -55.1699 | 290 | -54.6625 | 290 | 43662.0898 | 290 | 43495.8164 | 290 | 44533.9531 | 290 | 43438.9219 | 290 | 43223 | 290 | 44354.3594
291 | -53.1582 | 320.0587 | 291 | -54.9731 | 291 | -54.1393 | 291 | 43640.0469 | 291 | 43543.7734 | 291 | 44528.6836 | 291 | 43491.3203 | 291 | 43210.4805 | 291 | 44353.2383
292 | -53.1312 | 319.9345 | 292 | -54.9791 | 292 | -54.3747 | 292 | 43553.5898 | 292 | 43528.4531 | 292 | 44486.5469 | 292 | 43482.8008 | 292 | 43210.3594 | 292 | 44324.6406
293 | -53.1717 | 320.1207 | 293 | -55.047 | 293 | -54.7245 | 293 | 43468.3633 | 293 | 43465.2734 | 293 | 44407.6836 | 293 | 43564.6797 | 293 | 43255.5195 | 293 | 44227.3203
294 | -53.0293 | 319.4663 | 294 | -55.1499 | 294 | -54.5299 | 294 | 43314.3164 | 294 | 43451.7266 | 294 | 44449.0469 | 294 | 43470.1992 | 294 | 43262.4414 | 294 | 44281.4805
295 | -52.9189 | 318.9606 | 295 | -54.773 | 295 | -54.7796 | 295 | 43480.3633 | 295 | 43439.2266 | 295 | 44464.7734 | 295 | 43430.2383 | 295 | 43299.3203 | 295 | 44265.6016
296 | -53.123 | 319.8968 | 296 | -54.9705 | 296 | -54.7917 | 296 | 43542.3164 | 296 | 43400.0898 | 296 | 44479.6836 | 296 | 43460.0781 | 296 | 43302.5586 | 296 | 44262.0781
297 | -52.7694 | 318.2792 | 297 | -54.4853 | 297 | -54.33 | 297 | 43493.7734 | 297 | 43359.7734 | 297 | 44463.0898 | 297 | 43457.9609 | 297 | 43366 | 297 | 44260.0391
298 | -52.9744 | 319.2146 | 298 | -54.5941 | 298 | -54.5347 | 298 | 43530.3164 | 298 | 43392.5 | 298 | 44453.5898 | 298 | 43490.1992 | 298 | 43355.2812 | 298 | 44288.2812
299 | -52.756 | 318.2185 | 299 | -54.8776 | 299 | -55.0585 | 299 | 43480.6836 | 299 | 43405 | 299 | 44394 | 299 | 43521.2812 | 299 | 43333.9219 | 299 | 44352.8398
300 | -52.6357 | 317.6723 | 300 | -55.0282 | 300 | -55.0237 | 300 | 43489.8633 | 300 | 43420.5 | 300 | 44473.8164 | 300 | 43577.7188 | 300 | 43353.8398 | 300 | 44378.9609
301 | -52.678 | 317.8641 | 301 | -54.8627 | 301 | -54.9062 | 301 | 43426.2734 | 301 | 43391.8164 | 301 | 44487.7266 | 301 | 43553.0391 | 301 | 43368.6406 | 301 | 44409.7617
302 | -52.7507 | 318.1941 | 302 | -55.2916 | 302 | -55.3681 | 302 | 43469.4102 | 302 | 43480.2734 | 302 | 44480.1836 | 302 | 43582.3594 | 302 | 43379.8008 | 302 | 44450.2812
303 | -52.5592 | 317.3266 | 303 | -55.044 | 303 | -54.9926 | 303 | 43523.9531 | 303 | 43466.6836 | 303 | 44428.4102 | 303 | 43633.1211 | 303 | 43474.3203 | 303 | 44437.9219
304 | -53.0648 | 319.6293 | 304 | -55.4876 | 304 | -55.0114 | 304 | 43516.3164 | 304 | 43451.6367 | 304 | 44403.4102 | 304 | 43636.2383 | 304 | 43479.4414 | 304 | 44439.6797
305 | -52.7843 | 318.3472 | 305 | -55.1711 | 305 | -55.0457 | 305 | 43532.6367 | 305 | 43461.8633 | 305 | 44325.4102 | 305 | 43604.6797 | 305 | 43418.1211 | 305 | 44457.8008
306 | -52.4867 | 316.9993 | 306 | -54.8229 | 306 | -54.8217 | 306 | 43571.4531 | 306 | 43428.0898 | 306 | 44335.2734 | 306 | 43573.8789 | 306 | 43390.8398 | 306 | 44466.1602
307 | -52.4479 | 316.8248 | 307 | -54.8081 | 307 | -54.62 | 307 | 43581.3633 | 307 | 43441.6836 | 307 | 44278.1836 | 307 | 43561.2812 | 307 | 43347.3203 | 307 | 44433.1992
308 | -52.4568 | 316.8648 | 308 | -54.8821 | 308 | -54.6043 | 308 | 43606.5469 | 308 | 43465.5469 | 308 | 44300 | 308 | 43557.0781 | 308 | 43310.0781 | 308 | 44371.3984
309 | -52.4256 | 316.7242 | 309 | -54.8768 | 309 | -54.6438 | 309 | 43601.6836 | 309 | 43505.5898 | 309 | 44296.7266 | 309 | 43578.6406 | 309 | 43340.2812 | 309 | 44388.4414
310 | -52.6793 | 317.8702 | 310 | -55.2307 | 310 | -54.9512 | 310 | 43613.7266 | 310 | 43511.8164 | 310 | 44322.6367 | 310 | 43607.3984 | 310 | 43416.9609 | 310 | 44409.4805
311 | -52.8718 | 318.7456 | 311 | -55.5293 | 311 | -55.1452 | 311 | 43627.8164 | 311 | 43517.9531 | 311 | 44335.4102 | 311 | 43638.3594 | 311 | 43489.5586 | 311 | 44434.0781
312 | -52.8145 | 318.4844 | 312 | -55.4599 | 312 | -55.1442 | 312 | 43613.4102 | 312 | 43519.6836 | 312 | 44336 | 312 | 43658.4805 | 312 | 43523.2812 | 312 | 44428.2383
313 | -52.735 | 318.1226 | 313 | -55.2489 | 313 | -54.9843 | 313 | 43591.0898 | 313 | 43543.7266 | 313 | 44319.8633 | 313 | 43595.5195 | 313 | 43469.6406 | 313 | 44416.4805
314 | -52.6991 | 317.9599 | 314 | -55.0592 | 314 | -55.0972 | 314 | 43623.7266 | 314 | 43556.1836 | 314 | 44369.2266 | 314 | 43540.0781 | 314 | 43439.3594 | 314 | 44407.3594
315 | -52.9298 | 319.0104 | 315 | -55.3661 | 315 | -55.1572 | 315 | 43639.5469 | 315 | 43570.0898 | 315 | 44397.7266 | 315 | 43451.3984 | 315 | 43387.4805 | 315 | 44405.8789
316 | -53.0084 | 319.3702 | 316 | -55.2705 | 316 | -55.1347 | 316 | 43614.3633 | 316 | 43588.2266 | 316 | 44403.4531 | 316 | 43289.1211 | 316 | 43369.3984 | 316 | 44433.7188
317 | -52.8281 | 318.5463 | 317 | -55.0164 | 317 | -54.9304 | 317 | 43661.1367 | 317 | 43627 | 317 | 44399.1367 | 317 | 43455.6016 | 317 | 43310.6406 | 317 | 44457.7617
318 | -53.0513 | 319.5672 | 318 | -55.3594 | 318 | -54.9206 | 318 | 43678.4102 | 318 | 43614.2734 | 318 | 44385.7266 | 318 | 43472.0391 | 318 | 43280.2383 | 318 | 44366.2383
319 | -53.0171 | 319.4103 | 319 | -55.0772 | 319 | -55.0057 | 319 | 43599.0469 | 319 | 43584.6367 | 319 | 44400.9531 | 319 | 43467.7188 | 319 | 43301.3984 | 319 | 44364.2812
320 | -53.0063 | 319.3605 | 320 | -54.7114 | 320 | -54.5459 | 320 | 43654.8633 | 320 | 43562.7266 | 320 | 44417.2266 | 320 | 43436.7617 | 320 | 43346.7617 | 320 | 44290
321 | -53.4553 | 321.4338 | 321 | -55.0777 | 321 | -55.2586 | 321 | 43567 | 321 | 43531.6836 | 321 | 44419.1367 | 321 | 43450.7188 | 321 | 43345.2812 | 321 | 44256.9219
322 | -53.2927 | 320.6795 | 322 | -55.0182 | 322 | -55.2334 | 322 | 43572.2734 | 322 | 43513.0898 | 322 | 44430.3633 | 322 | 43457.6016 | 322 | 43388.7617 | 322 | 44242.5586
323 | -52.7422 | 318.1553 | 323 | -54.705 | 323 | -54.602 | 323 | 43556.0898 | 323 | 43514.8633 | 323 | 44428.7734 | 323 | 43422.3594 | 323 | 43407.5586 | 323 | 44258.6406
324 | -52.709 | 318.0048 | 324 | -54.9613 | 324 | -54.7281 | 324 | 43492.8633 | 324 | 43524.3633 | 324 | 44429.2266 | 324 | 43407.1602 | 324 | 43396.8008 | 324 | 44249.5586
325 | -53.2192 | 320.3399 | 325 | -55.3358 | 325 | -55.2546 | 325 | 43447.4102 | 325 | 43570.0898 | 325 | 44394.4531 | 325 | 43459.2812 | 325 | 43380.8008 | 325 | 44278.6406
326 | -52.9282 | 319.0031 | 326 | -55.0219 | 326 | -54.837 | 326 | 43451.1836 | 326 | 43502.0469 | 326 | 44416.0469 | 326 | 43481.8398 | 326 | 43393.5586 | 326 | 44294.7188
327 | -53.5592 | 321.9179 | 327 | -55.1699 | 327 | -55.1811 | 327 | 43474.6836 | 327 | 43502.9531 | 327 | 44433.4102 | 327 | 43471.1992 | 327 | 43373.8398 | 327 | 44315.8008
328 | -53.2258 | 320.3703 | 328 | -54.8924 | 328 | -55.1277 | 328 | 43509.8164 | 328 | 43502.7734 | 328 | 44390.5 | 328 | 43496.0781 | 328 | 43368 | 328 | 44295.0391
329 | -53.2624 | 320.5395 | 329 | -54.9055 | 329 | -54.9871 | 329 | 43507.6367 | 329 | 43467.3164 | 329 | 44410.1367 | 329 | 43525.0781 | 329 | 43392 | 329 | 44295.5195
330 | -53.1986 | 320.2449 | 330 | -54.8313 | 330 | -54.9035 | 330 | 43522.2266 | 330 | 43426.0898 | 330 | 44401.2734 | 330 | 43554.8789 | 330 | 43413.2383 | 330 | 44320.8008
331 | -53.1611 | 320.0721 | 331 | -55.1365 | 331 | -54.9751 | 331 | 43520.5 | 331 | 43386.7266 | 331 | 44369.6367 | 331 | 43572.2812 | 331 | 43427.4805 | 331 | 44347.6797
332 | -53.498 | 321.6325 | 332 | -55.1445 | 332 | -54.8997 | 332 | 43476.3633 | 332 | 43369.2266 | 332 | 44291.2266 | 332 | 43568.6016 | 332 | 43448 | 332 | 44324.3594
333 | -53.6588 | 322.3839 | 333 | -55.2996 | 333 | -54.9367 | 333 | 43496.4102 | 333 | 43414.9531 | 333 | 44255.5898 | 333 | 43582.8789 | 333 | 43459.8789 | 333 | 44309.0391
334 | -53.349 | 320.9402 | 334 | -55.2919 | 334 | -55.0927 | 334 | 43533.7266 | 334 | 43494.3164 | 334 | 44224.2734 | 334 | 43590.1211 | 334 | 43492.3984 | 334 | 44340.9219
335 | -53.2479 | 320.4725 | 335 | -55.3403 | 335 | -55.0445 | 335 | 43544.6367 | 335 | 43557.5469 | 335 | 44222.5469 | 335 | 43559.9219 | 335 | 43526.6406 | 335 | 44345.8008
336 | -53.2408 | 320.4396 | 336 | -55.426 | 336 | -54.8713 | 336 | 43581.1367 | 336 | 43575.0898 | 336 | 44230.0469 | 336 | 43618.0391 | 336 | 43550.9609 | 336 | 44343.6016
337 | -53.2055 | 320.2765 | 337 | -55.1918 | 337 | -54.6646 | 337 | 43592.1836 | 337 | 43596.2734 | 337 | 44225.7266 | 337 | 43632.7617 | 337 | 43560.9609 | 337 | 44354.0781
338 | -53.1725 | 320.1244 | 338 | -55.2416 | 338 | -54.6577 | 338 | 43589.6836 | 338 | 43571.2734 | 338 | 44211 | 338 | 43560.0391 | 338 | 43556.2812 | 338 | 44353.8398
339 | -53.0813 | 319.7047 | 339 | -55.0269 | 339 | -54.8529 | 339 | 43597.2734 | 339 | 43594.1836 | 339 | 44200.9531 | 339 | 43591.8008 | 339 | 43537.8398 | 339 | 44355.5195
340 | -53.1767 | 320.1439 | 340 | -55.303 | 340 | -54.951 | 340 | 43640.4531 | 340 | 43603.7266 | 340 | 44202.0469 | 340 | 43599.4805 | 340 | 43520.1211 | 340 | 44367
341 | -52.803 | 318.4322 | 341 | -55.1938 | 341 | -54.7881 | 341 | 43587.4531 | 341 | 43611.1836 | 341 | 44215.5469 | 341 | 43612.3984 | 341 | 43470.4414 | 341 | 44363.2812
342 | -52.6986 | 317.9575 | 342 | -54.9678 | 342 | -54.6017 | 342 | 43611.1367 | 342 | 43573.9102 | 342 | 44195.8164 | 342 | 43597.4805 | 342 | 43482.4414 | 342 | 44381.4414
343 | -53.192 | 320.2144 | 343 | -55.0214 | 343 | -54.7356 | 343 | 43643.5469 | 343 | 43586.4102 | 343 | 44207 | 343 | 43544.1211 | 343 | 43509.5586 | 343 | 44335.0781
344 | -53.3519 | 320.9536 | 344 | -55.401 | 344 | -54.7894 | 344 | 43597.5 | 344 | 43603.1836 | 344 | 44232.7734 | 344 | 43545.0781 | 344 | 43482.9219 | 344 | 44365.3984
345 | -53.1014 | 319.7971 | 345 | -55.3242 | 345 | -54.7604 | 345 | 43612.5898 | 345 | 43612.2266 | 345 | 44311.3633 | 345 | 43520.8398 | 345 | 43497.7617 | 345 | 44340.2812
346 | -52.9266 | 318.9959 | 346 | -55.1963 | 346 | -54.5621 | 346 | 43632.8164 | 346 | 43634.0469 | 346 | 44405.8164 | 346 | 43486.9609 | 346 | 43487.1602 | 346 | 44300.3984
347 | -53.2606 | 320.531 | 347 | -55.7056 | 347 | -54.8187 | 347 | 43687.1836 | 347 | 43620.2734 | 347 | 44451.6836 | 347 | 43419.6016 | 347 | 43487.6406 | 347 | 44231.5586
348 | -53.2424 | 320.4469 | 348 | -55.4903 | 348 | -54.6119 | 348 | 43712.3164 | 348 | 43602.7734 | 348 | 44435.1836 | 348 | 43431.6016 | 348 | 43462.8398 | 348 | 44196.6016
349 | -53.2611 | 320.5334 | 349 | -55.4156 | 349 | -54.7707 | 349 | 43752.4102 | 349 | 43695.3633 | 349 | 44519.0469 | 349 | 43436.5195 | 349 | 43443.0781 | 349 | 44186.5195
350 | -53.1368 | 319.9601 | 350 | -55.6732 | 350 | -54.8511 | 350 | 43754.2266 | 350 | 43677.8633 | 350 | 44492.7734 | 350 | 43462.4805 | 350 | 43421.5586 | 350 | 44179.3984
351 | -53.108 | 319.8275 | 351 | -55.4918 | 351 | -55.0074 | 351 | 43741.2266 | 351 | 43698.9102 | 351 | 44490.8633 | 351 | 43454.5195 | 351 | 43408.4805 | 351 | 44144.8789
352 | -52.8638 | 318.7091 | 352 | -55.2409 | 352 | -54.5138 | 352 | 43700.0469 | 352 | 43666.6836 | 352 | 44443.0469 | 352 | 43475.6797 | 352 | 43384 | 352 | 44134.3594
353 | -52.7104 | 318.0109 | 353 | -55.4221 | 353 | -54.537 | 353 | 43674.8164 | 353 | 43707.1836 | 353 | 44410.7734 | 353 | 43473.3984 | 353 | 43378.6406 | 353 | 44142.5195
354 | -52.6504 | 317.7391 | 354 | -55.4895 | 354 | -54.7185 | 354 | 43617.0898 | 354 | 43694.3164 | 354 | 44422.0898 | 354 | 43445.2383 | 354 | 43409.4805 | 354 | 44141.7617
355 | -52.6957 | 317.9442 | 355 | -55.2337 | 355 | -54.8524 | 355 | 43574.4102 | 355 | 43660.7266 | 355 | 44391.6367 | 355 | 43440.5195 | 355 | 43452.6406 | 355 | 44127.8008
356 | -52.9675 | 319.183 | 356 | -55.4089 | 356 | -55.3336 | 356 | 43613.5 | 356 | 43653.0469 | 356 | 44372.5469 | 356 | 43448.8398 | 356 | 43505.7617 | 356 | 44131.4805
357 | -52.9529 | 319.1161 | 357 | -55.4527 | 357 | -55.3234 | 357 | 43612.1836 | 357 | 43625.2266 | 357 | 44414.9531 | 357 | 43524 | 357 | 43494.5195 | 357 | 44151.8789
358 | -52.9138 | 318.9375 | 358 | -55.2854 | 358 | -55.1382 | 358 | 43577.8164 | 358 | 43611.0898 | 358 | 44417.5898 | 358 | 43547.4414 | 358 | 43514.5195 | 358 | 44172.1602
359 | -52.9287 | 319.0056 | 359 | -55.2772 | 359 | -55.0682 | 359 | 43586.3633 | 359 | 43570.8164 | 359 | 44387.5898 | 359 | 43551.9219 | 359 | 43498.0781 | 359 | 44198.5586
360 | -52.8193 | 318.5062 | 360 | -55.3112 | 360 | -55.0332 | 360 | 43560.6836 | 360 | 43571.4531 | 360 | 44388.5 | 360 | 43586.8789 | 360 | 43518.5195 | 360 | 44282.2383
361 | -52.7563 | 318.2197 | 361 | -55.3386 | 361 | -54.9532 | 361 | 43561.4531 | 361 | 43594.5 | 361 | 44387.0469 | 361 | 43567.4805 | 361 | 43561.3984 | 361 | 44361.3203
362 | -53.0283 | 319.4615 | 362 | -55.4601 | 362 | -55.0845 | 362 | 43612.4102 | 362 | 43612.1836 | 362 | 44350.6836 | 362 | 43628.8398 | 362 | 43589.6797 | 362 | 44387.1602
363 | -52.6783 | 317.8652 | 363 | -55.2513 | 363 | -54.7336 | 363 | 43645.9531 | 363 | 43592.3164 | 363 | 44355.8164 | 363 | 43586.2812 | 363 | 43580.0391 | 363 | 44413.4805
364 | -52.5673 | 317.363 | 364 | -55.5194 | 364 | -54.8222 | 364 | 43667.8633 | 364 | 43557.1367 | 364 | 44381.3633 | 364 | 43609.8398 | 364 | 43585.4805 | 364 | 44471.3594
365 | -52.4948 | 317.0357 | 365 | -55.5399 | 365 | -54.7417 | 365 | 43680.4531 | 365 | 43551.5469 | 365 | 44341.5898 | 365 | 43599.3594 | 365 | 43622.6016 | 365 | 44448.9219
366 | -52.5397 | 317.2382 | 366 | -55.3832 | 366 | -54.5619 | 366 | 43719.7734 | 366 | 43620.4102 | 366 | 44276.5 | 366 | 43621.1211 | 366 | 43577.6797 | 366 | 44456.6016
367 | -52.8742 | 318.7565 | 367 | -55.7294 | 367 | -54.7755 | 367 | 43693.3633 | 367 | 43633.1367 | 367 | 44200 | 367 | 43581.9219 | 367 | 43634.6406 | 367 | 44403.3984
368 | -52.1737 | 315.5957 | 368 | -55.1152 | 368 | -54.341 | 368 | 43653.2266 | 368 | 43701.3633 | 368 | 44153.1836 | 368 | 43625 | 368 | 43588.1211 | 368 | 44379.5586
369 | -52.2057 | 315.7385 | 369 | -55.4211 | 369 | -54.4571 | 369 | 43738.6367 | 369 | 43705 | 369 | 44167.0898 | 369 | 43690.6406 | 369 | 43580.7617 | 369 | 44363.9219
370 | -52.2528 | 315.9491 | 370 | -55.647 | 370 | -54.524 | 370 | 43695.0898 | 370 | 43703.2734 | 370 | 44175.6836 | 370 | 43700.9219 | 370 | 43657.3984 | 370 | 44385.3203
371 | -52.4234 | 316.7145 | 371 | -56.1673 | 371 | -54.5204 | 371 | 43691.6836 | 371 | 43694.3164 | 371 | 44227.7266 | 371 | 43707.9609 | 371 | 43640.5586 | 371 | 44336.3203
372 | -52.755 | 318.2137 | 372 | -55.9152 | 372 | -54.4188 | 372 | 43642.7734 | 372 | 43665.1367 | 372 | 44323.6836 | 372 | 43701.1602 | 372 | 43679.6016 | 372 | 44342.3203
373 | -52.7737 | 318.2986 | 373 | -55.6791 | 373 | -54.3466 | 373 | 43696.5 | 373 | 43655.0898 | 373 | 44360.1367 | 373 | 43696.2383 | 373 | 43623.9609 | 373 | 44350.7188
374 | -52.6381 | 317.6833 | 374 | -55.6509 | 374 | -54.2263 | 374 | 43697.2734 | 374 | 43716.6836 | 374 | 44349.6367 | 374 | 43653.0781 | 374 | 43659.3203 | 374 | 44345.1602
375 | -52.6876 | 317.9078 | 375 | -55.6563 | 375 | -54.0143 | 375 | 43689.8633 | 375 | 43694.7266 | 375 | 44341.2734 | 375 | 43620.7188 | 375 | 43612.1602 | 375 | 44335.8789
376 | -52.8363 | 318.5839 | 376 | -55.7024 | 376 | -53.8667 | 376 | 43701.9531 | 376 | 43730 | 376 | 44334.9102 | 376 | 43586.1992 | 376 | 43609.5195 | 376 | 44335.4805
377 | -52.8233 | 318.5245 | 377 | -55.6892 | 377 | -53.9942 | 377 | 43703.8164 | 377 | 43744.8164 | 377 | 44334.9531 | 377 | 43567.2383 | 377 | 43605.8398 | 377 | 44281.6016
378 | -52.7192 | 318.0509 | 378 | -55.5443 | 378 | -54.2517 | 378 | 43667.9531 | 378 | 43703.4102 | 378 | 44375.6836 | 378 | 43563.7617 | 378 | 43598.0391 | 378 | 44278.5195
379 | -52.9733 | 319.2098 | 379 | -55.7931 | 379 | -54.168 | 379 | 43681.0469 | 379 | 43756.4531 | 379 | 44415.2734 | 379 | 43567.9219 | 379 | 43525.1602 | 379 | 44303.9609
380 | -52.7424 | 318.1565 | 380 | -55.8759 | 380 | -53.972 | 380 | 43716.7266 | 380 | 43782.5 | 380 | 44390.1367 | 380 | 43557.3203 | 380 | 43502.4805 | 380 | 44285.3594
381 | -52.7248 | 318.0765 | 381 | -55.9356 | 381 | -53.9389 | 381 | 43741.9102 | 381 | 43770.6367 | 381 | 44426.1367 | 381 | 43519.8789 | 381 | 43522.6406 | 381 | 44240.5586
382 | -52.8782 | 318.7747 | 382 | -55.9911 | 382 | -53.9493 | 382 | 43778.5 | 382 | 43723.1367 | 382 | 44392.6836 | 382 | 43540.3203 | 382 | 43558.1992 | 382 | 44157.4805
383 | -52.839 | 318.5962 | 383 | -55.9269 | 383 | -54.168 | 383 | 43700.8633 | 383 | 43739.9531 | 383 | 44377.2266 | 383 | 43564.7188 | 383 | 43567.1992 | 383 | 44086
384 | -53.0198 | 319.4225 | 384 | -56.1422 | 384 | -54.2558 | 384 | 43657.6836 | 384 | 43759.7734 | 384 | 44355.5898 | 384 | 43561.2383 | 384 | 43547.0391 | 384 | 44143.4805
385 | -52.5775 | 317.4091 | 385 | -56.0028 | 385 | -54.1834 | 385 | 43650.5469 | 385 | 43702 | 385 | 44300.4531 | 385 | 43571.0391 | 385 | 43523.2812 | 385 | 44152.6797
386 | -52.1414 | 315.4517 | 386 | -55.3214 | 386 | -54.2286 | 386 | 43646.9531 | 386 | 43642.5 | 386 | 44295.7734 | 386 | 43627.5586 | 386 | 43526.5195 | 386 | 44215.1992
387 | -52.6938 | 317.9357 | 387 | -55.613 | 387 | -54.0223 | 387 | 43622.6367 | 387 | 43608.5898 | 387 | 44272.3164 | 387 | 43644.0391 | 387 | 43569.4805 | 387 | 44275.1992
388 | -52.8166 | 318.4941 | 388 | -55.8771 | 388 | -54.0265 | 388 | 43641.5 | 388 | 43649.2734 | 388 | 44301.8633 | 388 | 43645.9219 | 388 | 43581.1211 | 388 | 44323.3594
389 | -52.8222 | 318.5196 | 389 | -56.0904 | 389 | -54.0375 | 389 | 43598.9531 | 389 | 43735.4531 | 389 | 44335.5 | 389 | 43654.6406 | 389 | 43648.5195 | 389 | 44328.9219
390 | -52.3461 | 316.3669 | 390 | -55.8879 | 390 | -54.4025 | 390 | 43633.0898 | 390 | 43804.7266 | 390 | 44336.3164 | 390 | 43650.1992 | 390 | 43661.5586 | 390 | 44290.8008
391 | -52.5431 | 317.2539 | 391 | -55.9602 | 391 | -54.9095 | 391 | 43650.9102 | 391 | 43833.7734 | 391 | 44325.6367 | 391 | 43656.8008 | 391 | 43666.1602 | 391 | 44276.5586
392 | -52.6051 | 317.5341 | 392 | -56.1811 | 392 | -54.6782 | 392 | 43611.6367 | 392 | 43818.5 | 392 | 44337.0469 | 392 | 43685.6406 | 392 | 43659.9219 | 392 | 44275.6797
393 | -52.6726 | 317.8397 | 393 | -56.4838 | 393 | -54.2337 | 393 | 43663.9102 | 393 | 43832.2734 | 393 | 44305.5469 | 393 | 43649.8008 | 393 | 43651.6016 | 393 | 44307
394 | -52.7117 | 318.0169 | 394 | -56.1298 | 394 | -54.4935 | 394 | 43703.1367 | 394 | 43802.3164 | 394 | 44303.0469 | 394 | 43598.1211 | 394 | 43647.2383 | 394 | 44351.8008
395 | -52.9377 | 319.0469 | 395 | -56.5819 | 395 | -54.0107 | 395 | 43639.2266 | 395 | 43740.1836 | 395 | 44303.9102 | 395 | 43621.9609 | 395 | 43685.6406 | 395 | 44329
396 | -52.7731 | 318.2961 | 396 | -56.0336 | 396 | -53.8009 | 396 | 43663.6836 | 396 | 43697.5 | 396 | 44297.7266 | 396 | 43645.5586 | 396 | 43692.4414 | 396 | 44385.8789
397 | -52.9338 | 319.0286 | 397 | -56.5709 | 397 | -54.9447 | 397 | 43587.1367 | 397 | 43679.4102 | 397 | 44303.0469 | 397 | 43688.3984 | 397 | 43672.5586 | 397 | 44349.7188
398 | -53.2071 | 320.2839 | 398 | -56.685 | 398 | -55.3271 | 398 | 43560.4102 | 398 | 43714.3164 | 398 | 44296.4531 | 398 | 43660.2383 | 398 | 43712.0391 | 398 | 44322.0391
399 | -52.9765 | 319.2243 | 399 | -56.5805 | 399 | -54.3852 | 399 | 43486.3164 | 399 | 43738.3164 | 399 | 44239.5 | 399 | 43642.3984 | 399 | 43694.9219 | 399 | 44303.2383
400 | -52.8917 | 318.8366 | 400 | -56.6515 | 400 | -54.5779 | 400 | 43453.8633 | 400 | 43758.9102 | 400 | 44284 | 400 | 43666.7188 | 400 | 43743.5195 | 400 | 44254.5195
401 | -52.9372 | 319.0445 | 401 | -56.6312 | 401 | -54.6577 | 401 | 43469.7734 | 401 | 43732.1367 | 401 | 44271.7734 | 401 | 43688 | 401 | 43787.6406 | 401 | 44242.7617
402 | -52.9253 | 318.9898 | 402 | -56.5953 | 402 | -55.1062 | 402 | 43453.1836 | 402 | 43723.0898 | 402 | 44274.2266 | 402 | 43680.1211 | 402 | 43786.3203 | 402 | 44236.6797
403 | -52.963 | 319.1623 | 403 | -56.9318 | 403 | -55.1953 | 403 | 43497.3633 | 403 | 43700.3164 | 403 | 44242.1836 | 403 | 43632.9219 | 403 | 43733.6406 | 403 | 44230.3203
404 | -53.1016 | 319.7983 | 404 | -56.9273 | 404 | -55.2327 | 404 | 43511.1836 | 404 | 43673.1367 | 404 | 44280.1367 | 404 | 43631.4805 | 404 | 43719.7617 | 404 | 44276.2383
405 | -53.0373 | 319.5028 | 405 | -56.8272 | 405 | -55.3777 | 405 | 43509.5 | 405 | 43661.7734 | 405 | 44257.5 | 405 | 43647.5195 | 405 | 43697.6406 | 405 | 44302
406 | -53.1019 | 319.7995 | 406 | -57.0182 | 406 | -55.4151 | 406 | 43584.2734 | 406 | 43694.0898 | 406 | 44258.2266 | 406 | 43678.1992 | 406 | 43665.3203 | 406 | 44284.4414
407 | -53.4196 | 321.268 | 407 | -57.1595 | 407 | -55.4433 | 407 | 43581.7266 | 407 | 43724.5 | 407 | 44263.1836 | 407 | 43732.3203 | 407 | 43627.5586 | 407 | 44261.6016
408 | -53.4778 | 321.5386 | 408 | -57.0703 | 408 | -55.1991 | 408 | 43643.9102 | 408 | 43749.2734 | 408 | 44241.9102 | 408 | 43656.9609 | 408 | 43618.3984 | 408 | 44246.3203
409 | -53.7031 | 322.5915 | 409 | -57.5183 | 409 | -55.5063 | 409 | 43704.7266 | 409 | 43774.6367 | 409 | 44222.8164 | 409 | 43620.6406 | 409 | 43648.8398 | 409 | 44233.8398
410 | -53.7671 | 322.8919 | 410 | -57.4341 | 410 | -55.105 | 410 | 43712.2266 | 410 | 43790.7266 | 410 | 44223.9531 | 410 | 43669 | 410 | 43723.2383 | 410 | 44248.2812
411 | -53.9965 | 323.9746 | 411 | -57.3418 | 411 | -55.1117 | 411 | 43813.0898 | 411 | 43800.9531 | 411 | 44241.0898 | 411 | 43647.2383 | 411 | 43775 | 411 | 44234.5195
412 | -54.0486 | 324.2218 | 412 | -57.4545 | 412 | -55.0625 | 412 | 43779.6367 | 412 | 43803.5469 | 412 | 44252.0898 | 412 | 43657.0391 | 412 | 43780.0781 | 412 | 44242.8789
413 | -53.8791 | 323.4196 | 413 | -57.4625 | 413 | -55.0132 | 413 | 43779.9102 | 413 | 43804.0898 | 413 | 44276.7734 | 413 | 43642.7617 | 413 | 43803 | 413 | 44250.9219
414 | -54.2376 | 325.1213 | 414 | -57.59 | 414 | -54.9434 | 414 | 43722.5469 | 414 | 43815.8164 | 414 | 44277.3164 | 414 | 43676.7617 | 414 | 43807.9219 | 414 | 44229.8008
415 | -54.365 | 325.7315 | 415 | -57.3227 | 415 | -54.9826 | 415 | 43753.6367 | 415 | 43851.3633 | 415 | 44278.3633 | 415 | 43657.4414 | 415 | 43780.5195 | 415 | 44255.7617
416 | -54.3535 | 325.6764 | 416 | -57.3076 | 416 | -54.8035 | 416 | 43724 | 416 | 43871.5898 | 416 | 44278.9102 | 416 | 43710.3984 | 416 | 43748.1602 | 416 | 44249.1602
417 | -54.1575 | 324.7394 | 417 | -57.0442 | 417 | -54.6582 | 417 | 43697.2266 | 417 | 43871.7266 | 417 | 44272.2266 | 417 | 43688.7617 | 417 | 43717.5586 | 417 | 44238.1992
418 | -54.2547 | 325.2034 | 418 | -57.4771 | 418 | -54.7243 | 418 | 43751.2266 | 418 | 43852.9531 | 418 | 44312.6367 | 418 | 43602.8008 | 418 | 43685.1992 | 418 | 44193.5586
419 | -54.0808 | 324.3747 | 419 | -57.4023 | 419 | -54.6711 | 419 | 43752.3633 | 419 | 43809.5469 | 419 | 44322.7266 | 419 | 43627.2383 | 419 | 43700.8008 | 419 | 44233.6016
420 | -53.771 | 322.9102 | 420 | -57.16 | 420 | -54.623 | 420 | 43746.9531 | 420 | 43845.5 | 420 | 44379.7734 | 420 | 43552.0391 | 420 | 43736.1992 | 420 | 44225.3594
421 | -53.9081 | 323.5565 | 421 | -57.5268 | 421 | -54.7844 | 421 | 43752.8164 | 421 | 43831.4531 | 421 | 44455.0898 | 421 | 43533.7617 | 421 | 43756.0391 | 421 | 44213.4805
422 | -53.4385 | 321.3557 | 422 | -57.2498 | 422 | -54.3589 | 422 | 43730.5469 | 422 | 43813.6367 | 422 | 44436.7734 | 422 | 43511.3984 | 422 | 43709.7188 | 422 | 44210.5195
423 | -53.8058 | 323.0738 | 423 | -57.5914 | 423 | -54.7621 | 423 | 43752.5898 | 423 | 43858.7266 | 423 | 44409.4102 | 423 | 43455.8789 | 423 | 43706.1602 | 423 | 44208.6016
424 | -53.8219 | 323.1495 | 424 | -57.5346 | 424 | -54.5939 | 424 | 43717.1836 | 424 | 43925.5898 | 424 | 44411.5898 | 424 | 43473.6016 | 424 | 43693.4805 | 424 | 44189.6797
425 | -54.0476 | 324.2168 | 425 | -57.9333 | 425 | -54.7569 | 425 | 43701.7266 | 425 | 43950.6836 | 425 | 44394.0469 | 425 | 43491.7617 | 425 | 43656.3203 | 425 | 44150.8008
426 | -53.5336 | 321.7983 | 426 | -57.7571 | 426 | -54.391 | 426 | 43656.2734 | 426 | 43982.6836 | 426 | 44411.7266 | 426 | 43523.3594 | 426 | 43635.1211 | 426 | 44162.5195
427 | -53.6051 | 322.1326 | 427 | -58.0255 | 427 | -54.6225 | 427 | 43629.3164 | 427 | 43964.9102 | 427 | 44375.7734 | 427 | 43539.0781 | 427 | 43662.0781 | 427 | 44196.0781
428 | -53.3322 | 320.8623 | 428 | -57.7878 | 428 | -54.3785 | 428 | 43596.7734 | 428 | 43932.2734 | 428 | 44318.7266 | 428 | 43590.7188 | 428 | 43707.3594 | 428 | 44234.0391
429 | -53.6328 | 322.2619 | 429 | -58.1068 | 429 | -54.5024 | 429 | 43634.1836 | 429 | 43922.1836 | 429 | 44299.8164 | 429 | 43599.8398 | 429 | 43731.2383 | 429 | 44239.6797
430 | -53.8281 | 323.1789 | 430 | -58.3074 | 430 | -54.5992 | 430 | 43634.9102 | 430 | 43866.6367 | 430 | 44277.2734 | 430 | 43629.3594 | 430 | 43751.1211 | 430 | 44210.1602
431 | -53.8799 | 323.4232 | 431 | -58.4133 | 431 | -54.7326 | 431 | 43655.8164 | 431 | 43830.4531 | 431 | 44244.5898 | 431 | 43701.8398 | 431 | 43766.9609 | 431 | 44242.8398
432 | -53.6997 | 322.5756 | 432 | -58.0602 | 432 | -54.3663 | 432 | 43684.5 | 432 | 43841.9102 | 432 | 44254.1836 | 432 | 43728.7188 | 432 | 43767.8398 | 432 | 44241.2812
433 | -53.3027 | 320.7258 | 433 | -57.839 | 433 | -53.9257 | 433 | 43682.0469 | 433 | 43833.2266 | 433 | 44300.5469 | 433 | 43736.1992 | 433 | 43775.7188 | 433 | 44276.9219
434 | -53.1577 | 320.0562 | 434 | -57.6856 | 434 | -53.9867 | 434 | 43650.3164 | 434 | 43835.5898 | 434 | 44336.4102 | 434 | 43779.8789 | 434 | 43813.4805 | 434 | 44292.0781
435 | -53.7156 | 322.6501 | 435 | -58.2951 | 435 | -54.663 | 435 | 43659.6836 | 435 | 43870.8164 | 435 | 44355.1367 | 435 | 43766.2383 | 435 | 43835.0781 | 435 | 44354.0781
436 | -53.6753 | 322.4609 | 436 | -58.2553 | 436 | -54.5753 | 436 | 43661.4102 | 436 | 43904.6836 | 436 | 44338.7734 | 436 | 43799.7617 | 436 | 43872.6797 | 436 | 44413.3984
437 | -53.8807 | 323.427 | 437 | -58.4778 | 437 | -54.7334 | 437 | 43567.0898 | 437 | 43974.8633 | 437 | 44290.2266 | 437 | 43738.4805 | 437 | 43857.3203 | 437 | 44414
438 | -53.8094 | 323.0909 | 438 | -58.6589 | 438 | -54.752 | 438 | 43613.3164 | 438 | 44033.2734 | 438 | 44278.5469 | 438 | 43705.6016 | 438 | 43842.4805 | 438 | 44362.8398
439 | -53.6359 | 322.2766 | 439 | -58.6801 | 439 | -54.3525 | 439 | 43601.0469 | 439 | 44083.9102 | 439 | 44256.9531 | 439 | 43703.6797 | 439 | 43861 | 439 | 44384.4805
440 | -53.9358 | 323.6873 | 440 | -58.6684 | 440 | -54.2716 | 440 | 43685.6367 | 440 | 44150.9531 | 440 | 44230.5469 | 440 | 43732.9609 | 440 | 43821.4805 | 440 | 44375.4805
441 | -54.3376 | 325.6003 | 441 | -59.0941 | 441 | -54.4736 | 441 | 43685.0469 | 441 | 44137.5898 | 441 | 44235.9102 | 441 | 43701.4414 | 441 | 43799.3203 | 441 | 44367.6406
442 | -53.9037 | 323.5357 | 442 | -58.5248 | 442 | -53.8828 | 442 | 43705.6367 | 442 | 44078.0898 | 442 | 44250.9531 | 442 | 43704.3594 | 442 | 43846.0781 | 442 | 44346.6797
443 | -53.8784 | 323.416 | 443 | -58.5708 | 443 | -53.8532 | 443 | 43726.9102 | 443 | 44065.2266 | 443 | 44279.9102 | 443 | 43685.6797 | 443 | 43795.7617 | 443 | 44310.8398
444 | -53.9916 | 323.9514 | 444 | -58.8747 | 444 | -53.9195 | 444 | 43723.6836 | 444 | 44079.7734 | 444 | 44248.6367 | 444 | 43744.1211 | 444 | 43896.8789 | 444 | 44260.4414
445 | -54.0368 | 324.1655 | 445 | -58.8872 | 445 | -54.0842 | 445 | 43780.4531 | 445 | 44036.8164 | 445 | 44239.1836 | 445 | 43723.2812 | 445 | 43905.3984 | 445 | 44231.6797
446 | -54.1606 | 324.754 | 446 | -59.0887 | 446 | -54.0996 | 446 | 43782.3633 | 446 | 44050.5469 | 446 | 44100.9531 | 446 | 43706.2383 | 446 | 43935.7617 | 446 | 44208.5586
447 | -54.2294 | 325.0822 | 447 | -59.3282 | 447 | -53.9487 | 447 | 43731.2266 | 447 | 43995.6836 | 447 | 44160.1836 | 447 | 43678.6797 | 447 | 43990.8398 | 447 | 44229.3594
448 | -54.2278 | 325.0749 | 448 | -59.1311 | 448 | -54.0156 | 448 | 43667.0898 | 448 | 44008.4531 | 448 | 44146.6367 | 448 | 43665.1992 | 448 | 43985.8789 | 448 | 44270.2812
449 | -54.4887 | 326.3262 | 449 | -58.9617 | 449 | -53.8843 | 449 | 43670.7266 | 449 | 43989.5469 | 449 | 44106.6367 | 449 | 43616 | 449 | 43952.2383 | 449 | 44305.9219
450 | -54.3747 | 325.7781 | 450 | -58.816 | 450 | -53.5908 | 450 | 43646.4531 | 450 | 43946.9531 | 450 | 44143.1836 | 450 | 43560.9609 | 450 | 43950.4805 | 450 | 44301.3984
451 | -54.663 | 327.1693 | 451 | -59.0861 | 451 | -54.1266 | 451 | 43632.2734 | 451 | 43897.7734 | 451 | 44141.1367 | 451 | 43604.2812 | 451 | 43885.6406 | 451 | 44303
452 | -54.4988 | 326.3753 | 452 | -59.14 | 452 | -53.9299 | 452 | 43627.9102 | 452 | 43872.3633 | 452 | 44142.7734 | 452 | 43612.2812 | 452 | 43853.5195 | 452 | 44280.5586
453 | -54.3796 | 325.8014 | 453 | -59.032 | 453 | -53.7866 | 453 | 43642.5469 | 453 | 43887 | 453 | 44176.6836 | 453 | 43630.3594 | 453 | 43843.5195 | 453 | 44246
454 | -54.254 | 325.1998 | 454 | -59.1496 | 454 | -53.8273 | 454 | 43638.1367 | 454 | 43920.0898 | 454 | 44184.8164 | 454 | 43669.7188 | 454 | 43827.6797 | 454 | 44231.5195
455 | -54.0695 | 324.3208 | 455 | -59.1728 | 455 | -53.6768 | 455 | 43618.9531 | 455 | 43922.1367 | 455 | 44193.4102 | 455 | 43673.2812 | 455 | 43833.1211 | 455 | 44200.2812
456 | -54.0589 | 324.2707 | 456 | -59.0776 | 456 | -53.7333 | 456 | 43618.8164 | 456 | 43875.9531 | 456 | 44180.8164 | 456 | 43651.3203 | 456 | 43861.1602 | 456 | 44213.8789
457 | -54.1875 | 324.8826 | 457 | -59.2055 | 457 | -53.8361 | 457 | 43620.0469 | 457 | 43907.8164 | 457 | 44204.3164 | 457 | 43659 | 457 | 43889 | 457 | 44212.2383
458 | -54.541 | 326.579 | 458 | -59.4384 | 458 | -54.0187 | 458 | 43613.8633 | 458 | 43938.2734 | 458 | 44238.2266 | 458 | 43678.9219 | 458 | 43956.1992 | 458 | 44230.1602
459 | -54.2911 | 325.3774 | 459 | -59.2148 | 459 | -54.1079 | 459 | 43608.9531 | 459 | 43919.5469 | 459 | 44266.0469 | 459 | 43572.0781 | 459 | 43977.4805 | 459 | 44185.6406
460 | -54.3159 | 325.4962 | 460 | -59.0745 | 460 | -53.9743 | 460 | 43620.0469 | 460 | 43891.1836 | 460 | 44232 | 460 | 43592.3594 | 460 | 44012.9609 | 460 | 44197.9219
461 | -54.2765 | 325.3075 | 461 | -58.8962 | 461 | -53.7333 | 461 | 43655.7266 | 461 | 43924.9531 | 461 | 44169.5469 | 461 | 43601.2812 | 461 | 44072.9609 | 461 | 44048.0391
462 | -54.1696 | 324.7969 | 462 | -59.1461 | 462 | -53.8949 | 462 | 43656.5 | 462 | 43915.9531 | 462 | 44132.5 | 462 | 43604.8398 | 462 | 44060.4805 | 462 | 44106.4805
463 | -54.7 | 327.3486 | 463 | -59.6854 | 463 | -54.5166 | 463 | 43648.5898 | 463 | 43903.6836 | 463 | 44066.9531 | 463 | 43643.0781 | 463 | 44046.0781 | 463 | 44107.6797
464 | -54.6575 | 327.1423 | 464 | -59.5031 | 464 | -54.2952 | 464 | 43614.5898 | 464 | 43884.5469 | 464 | 44052.7734 | 464 | 43707.9609 | 464 | 44021.3594 | 464 | 44055.6406
465 | -54.5786 | 326.7606 | 465 | -59.2287 | 465 | -54.2706 | 465 | 43667 | 465 | 43873.6836 | 465 | 44060.5 | 465 | 43676.8398 | 465 | 44030.8789 | 465 | 44068.0781
466 | -54.3893 | 325.848 | 466 | -59.1686 | 466 | -53.8566 | 466 | 43668.8633 | 466 | 43870.3164 | 466 | 44093 | 466 | 43665.1211 | 466 | 44037.6016 | 466 | 44091.4805
467 | -54.5733 | 326.7348 | 467 | -59.5035 | 467 | -54.4099 | 467 | 43657.2266 | 467 | 43838 | 467 | 44121.5469 | 467 | 43689.3203 | 467 | 44039.9219 | 467 | 44097.4414
468 | -54.5212 | 326.4832 | 468 | -59.4999 | 468 | -54.0561 | 468 | 43686.5898 | 468 | 43802.1836 | 468 | 44136.8633 | 468 | 43715.0391 | 468 | 44022.1211 | 468 | 44104.2812
469 | -54.8083 | 327.8756 | 469 | -59.7895 | 469 | -54.4553 | 469 | 43706.0469 | 469 | 43835.6836 | 469 | 44123.5898 | 469 | 43694.0391 | 469 | 43973.6797 | 469 | 44102.7188
470 | -54.7725 | 327.7012 | 470 | -60.0053 | 470 | -54.4132 | 470 | 43724.8164 | 470 | 43819.2734 | 470 | 44141.0898 | 470 | 43675.0781 | 470 | 43925.9219 | 470 | 44133.4805
471 | -54.8544 | 328.1006 | 471 | -59.803 | 471 | -54.1531 | 471 | 43753.5 | 471 | 43899.9531 | 471 | 44178.9531 | 471 | 43676.8789 | 471 | 43901.2812 | 471 | 44126.9219
472 | -54.9306 | 328.4732 | 472 | -60.0333 | 472 | -54.2306 | 472 | 43736.9531 | 472 | 43948.5898 | 472 | 44170 | 472 | 43640.5195 | 472 | 43891.1992 | 472 | 44144.7617
473 | -55.31 | 330.3436 | 473 | -60.2695 | 473 | -54.5182 | 473 | 43727.5898 | 473 | 43993.7266 | 473 | 44179.5 | 473 | 43627 | 473 | 43874.9219 | 473 | 44182.9219
474 | -55.1724 | 329.6622 | 474 | -60.1623 | 474 | -54.2112 | 474 | 43770.5469 | 474 | 43954.6836 | 474 | 44244.5898 | 474 | 43603.1992 | 474 | 43884.8789 | 474 | 44192.3203
475 | -54.9801 | 328.7155 | 475 | -59.8996 | 475 | -53.9513 | 475 | 43758.1367 | 475 | 43962.1367 | 475 | 44243 | 475 | 43624.3203 | 475 | 43894.3203 | 475 | 44175.0391
476 | -55.1457 | 329.5304 | 476 | -59.9429 | 476 | -54.1297 | 476 | 43730.6836 | 476 | 43931.9102 | 476 | 44220.6836 | 476 | 43596.3984 | 476 | 43911.6016 | 476 | 44144.1992
477 | -55.3933 | 330.7579 | 477 | -60.1979 | 477 | -54.3034 | 477 | 43705.1367 | 477 | 43918.5469 | 477 | 44190.1367 | 477 | 43547.9219 | 477 | 43894.5586 | 477 | 44080.7617
478 | -55.4025 | 330.8035 | 478 | -60.143 | 478 | -54.0613 | 478 | 43660 | 478 | 43930.5469 | 478 | 44183.5469 | 478 | 43547.1211 | 478 | 43923.2812 | 478 | 44017.8789
479 | -55.3874 | 330.7283 | 479 | -60.2153 | 479 | -54.0615 | 479 | 43660.9102 | 479 | 43963.6367 | 479 | 44160.0469 | 479 | 43576.4805 | 479 | 43933.9219 | 479 | 44005.7617
480 | -55.3343 | 330.4644 | 480 | -60.3497 | 480 | -54.1349 | 480 | 43671.3633 | 480 | 43959.5469 | 480 | 44151.4102 | 480 | 43584.3203 | 480 | 43923.1992 | 480 | 44025.4414
481 | -55.34 | 330.4928 | 481 | -60.298 | 481 | -54.2314 | 481 | 43690.5898 | 481 | 43978.0898 | 481 | 44148.5898 | 481 | 43630.0781 | 481 | 43920.6797 | 481 | 44029.5195
482 | -55.7674 | 332.6336 | 482 | -60.6324 | 482 | -54.2594 | 482 | 43704 | 482 | 44025.3164 | 482 | 44143 | 482 | 43610.7188 | 482 | 43933 | 482 | 44068.6797
483 | -55.6359 | 331.9712 | 483 | -60.623 | 483 | -54.2934 | 483 | 43708.2266 | 483 | 44090.1367 | 483 | 44166.3164 | 483 | 43635.0781 | 483 | 43891.4414 | 483 | 44080.1602
484 | -55.4342 | 330.9614 | 484 | -60.3852 | 484 | -54.0038 | 484 | 43667.5469 | 484 | 44102.7734 | 484 | 44176.4531 | 484 | 43596.7617 | 484 | 43899.1602 | 484 | 44067.3203
485 | -55.6821 | 332.2034 | 485 | -60.9317 | 485 | -54.5019 | 485 | 43610.6836 | 485 | 44100.0898 | 485 | 44209.4102 | 485 | 43559.6016 | 485 | 43842.2812 | 485 | 44104.8008
486 | -55.5926 | 331.7539 | 486 | -60.7812 | 486 | -52.3989 | 486 | 43683.9531 | 486 | 44126.9531 | 486 | 44224.1836 | 486 | 43578.3984 | 486 | 43857.0781 | 486 | 44123.5586
487 | -55.6735 | 332.1603 | 487 | -60.5378 | 487 | -53.7086 | 487 | 43718.7734 | 487 | 44102.1367 | 487 | 44231.1367 | 487 | 43654.9609 | 487 | 43859.1211 | 487 | 44130
488 | -55.7475 | 332.5334 | 488 | -60.6849 | 488 | -54.2017 | 488 | 43792.0898 | 488 | 44078.9102 | 488 | 44212.2266 | 488 | 43631.6797 | 488 | 43808.1211 | 488 | 44132.3203
489 | -55.7806 | 332.7004 | 489 | -60.7922 | 489 | -54.4209 | 489 | 43781.3164 | 489 | 44067.3164 | 489 | 44179.7734 | 489 | 43666.1211 | 489 | 43785.7617 | 489 | 44188.3594
490 | -55.8398 | 332.9996 | 490 | -60.8233 | 490 | -54.4066 | 490 | 43775.1836 | 490 | 44114.1836 | 490 | 44198.5898 | 490 | 43673.0391 | 490 | 43818.6406 | 490 | 44219.6016
491 | -55.6843 | 332.2147 | 491 | -60.6382 | 491 | -54.0661 | 491 | 43770.9102 | 491 | 44136.1836 | 491 | 44197.6367 | 491 | 43698.8398 | 491 | 43784.4805 | 491 | 44196.2383
492 | -55.9018 | 333.3139 | 492 | -60.6717 | 492 | -53.9756 | 492 | 43761.3164 | 492 | 44146.1836 | 492 | 44223.8633 | 492 | 43722 | 492 | 43866.6406 | 492 | 44156.3203
493 | -55.7627 | 332.6101 | 493 | -60.8611 | 493 | -54.1302 | 493 | 43726.7734 | 493 | 44158.1367 | 493 | 44214.8633 | 493 | 43719.1602 | 493 | 43949.1211 | 493 | 44148.9219
494 | -56.0794 | 334.2188 | 494 | -61.3016 | 494 | -54.4035 | 494 | 43717.6367 | 494 | 44104.9102 | 494 | 44171.8633 | 494 | 43712.6406 | 494 | 43960.3984 | 494 | 44140.1211
495 | -55.7625 | 332.6089 | 495 | -61.0847 | 495 | -54.1079 | 495 | 43731.8633 | 495 | 44057.1367 | 495 | 44172.1367 | 495 | 43705.1602 | 495 | 43944.5195 | 495 | 44116.8789
496 | -55.8822 | 333.2148 | 496 | -61.1208 | 496 | -54.3622 | 496 | 43787.0898 | 496 | 44072.2266 | 496 | 44191.7266 | 496 | 43720.4414 | 496 | 43946.9609 | 496 | 44107.3594
497 | -56.2561 | 335.1248 | 497 | -61.4793 | 497 | -54.6278 | 497 | 43788.9531 | 497 | 44061.3164 | 497 | 44185 | 497 | 43724.8398 | 497 | 43901.7617 | 497 | 44107.2812
498 | -56.3036 | 335.3691 | 498 | -61.6366 | 498 | -54.8106 | 498 | 43824.0898 | 498 | 44033.9102 | 498 | 44206.9531 | 498 | 43703.1211 | 498 | 43892.9609 | 498 | 44118.6797
499 | -56.4835 | 336.3 | 499 | -61.8233 | 499 | -54.7652 | 499 | 43852.7266 | 499 | 44026.8633 | 499 | 44222.4102 | 499 | 43674.3594 | 499 | 43898.8398 | 499 | 44127.7188
500 | -56.5936 | 336.8727 | 500 | -62.0291 | 500 | -54.7073 | 500 | 43890.2266 | 500 | 44039.5469 | 500 | 44245.6367 | 500 | 43653.4414 | 500 | 43908.8789 | 500 | 44147
501 | -56.6798 | 337.3227 | 501 | -62.0333 | 501 | -54.8303 | 501 | 43905.4531 | 501 | 44055.1836 | 501 | 44282.7266 | 501 | 43634.8008 | 501 | 43938.1211 | 501 | 44159.6016
502 | -56.7946 | 337.9247 | 502 | -62.0664 | 502 | -54.8932 | 502 | 43913.9102 | 502 | 44034.8633 | 502 | 44250.6836 | 502 | 43648.3984 | 502 | 43971.2383 | 502 | 44154.1992
503 | -56.7662 | 337.7755 | 503 | -62.1621 | 503 | -54.7791 | 503 | 43898 | 503 | 44032.3164 | 503 | 44258.0469 | 503 | 43662.9609 | 503 | 44019.8008 | 503 | 44171.1211
504 | -56.7628 | 337.758 | 504 | -62.2235 | 504 | -54.9017 | 504 | 43893.5898 | 504 | 44027.2266 | 504 | 44264.5469 | 504 | 43687.7188 | 504 | 44064.5586 | 504 | 44138.1602
505 | -56.8518 | 338.2259 | 505 | -62.3668 | 505 | -55.0139 | 505 | 43894.6836 | 505 | 44018.1836 | 505 | 44274.4531 | 505 | 43678.7188 | 505 | 44082.6797 | 505 | 44141.9219
506 | -56.9132 | 338.5497 | 506 | -62.5113 | 506 | -55.1185 | 506 | 43896.0469 | 506 | 44005.6367 | 506 | 44274.1367 | 506 | 43673.9609 | 506 | 44088.0391 | 506 | 44175.0781
507 | -56.742 | 337.6485 | 507 | -62.2753 | 507 | -54.9419 | 507 | 43845.2734 | 507 | 44024.8633 | 507 | 44270.4102 | 507 | 43622.8398 | 507 | 44101.4805 | 507 | 44151.9609
508 | -56.8083 | 337.9969 | 508 | -62.3623 | 508 | -54.7281 | 508 | 43847.1367 | 508 | 44077.5469 | 508 | 44263.7734 | 508 | 43677.6406 | 508 | 44087.7617 | 508 | 44143.3984
509 | -56.9939 | 338.9768 | 509 | -62.7053 | 509 | -54.9494 | 509 | 43791.1367 | 509 | 44132.3164 | 509 | 44234.5898 | 509 | 43726.4414 | 509 | 44088.2812 | 509 | 44129.0391
510 | -56.6139 | 336.9783 | 510 | -62.3583 | 510 | -54.5418 | 510 | 43768.4531 | 510 | 44145.1836 | 510 | 44262.3633 | 510 | 43757.3984 | 510 | 44075.6016 | 510 | 44123.7617
511 | -56.9883 | 338.9469 | 511 | -62.8467 | 511 | -54.9884 | 511 | 43731.5 | 511 | 44123.2734 | 511 | 44277.3633 | 511 | 43749.6406 | 511 | 44067.6797 | 511 | 44180.9219
512 | -57.0879 | 339.4754 | 512 | -62.8714 | 512 | -55.1262 | 512 | 43766.4102 | 512 | 44078.0898 | 512 | 44247.1367 | 512 | 43741.7188 | 512 | 44113.8008 | 512 | 44172.8789
513 | -57.1782 | 339.9568 | 513 | -62.9954 | 513 | -55.2484 | 513 | 43732.8164 | 513 | 44067.8633 | 513 | 44241.7266 | 513 | 43737.9219 | 513 | 44136.8398 | 513 | 44184.1992
514 | -57.1487 | 339.7997 | 514 | -62.9824 | 514 | -55.2722 | 514 | 43746.9531 | 514 | 44042.4102 | 514 | 44201.7266 | 514 | 43714.8008 | 514 | 44136.3203 | 514 | 44201.0391
515 | -57.3879 | 341.0814 | 515 | -62.9611 | 515 | -55.3336 | 515 | 43742.7734 | 515 | 44021.9102 | 515 | 44174.0898 | 515 | 43689.6797 | 515 | 44087.9219 | 515 | 44194.5586
516 | -57.4503 | 341.4176 | 516 | -62.9576 | 516 | -55.14 | 516 | 43729.8164 | 516 | 44003.3633 | 516 | 44160.1836 | 516 | 43684.2812 | 516 | 44070.3984 | 516 | 44241.7188
517 | -57.2046 | 340.0978 | 517 | -63.0042 | 517 | -54.9306 | 517 | 43718.4531 | 517 | 44010.8164 | 517 | 44166.0898 | 517 | 43705.4414 | 517 | 44078.6406 | 517 | 44246.3203
518 | -57.2074 | 340.1127 | 518 | -63.0881 | 518 | -54.7591 | 518 | 43718.1367 | 518 | 43987.5 | 518 | 44120.0898 | 518 | 43747.4414 | 518 | 44052.6016 | 518 | 44254.8789
519 | -57.0388 | 339.2148 | 519 | -63.1797 | 519 | -54.6086 | 519 | 43729 | 519 | 43991.0469 | 519 | 44086.5898 | 519 | 43781.6016 | 519 | 44051.8398 | 519 | 44255
520 | -57.1695 | 339.9106 | 520 | -63.5866 | 520 | -54.8416 | 520 | 43729.9531 | 520 | 44000.8164 | 520 | 44076.7734 | 520 | 43810.9219 | 520 | 44034.6406 | 520 | 44235.6406
521 | -57.4496 | 341.4138 | 521 | -63.5132 | 521 | -54.8861 | 521 | 43676.1367 | 521 | 44001.5469 | 521 | 44083.5469 | 521 | 43853.9219 | 521 | 44051.3203 | 521 | 44237.9219
522 | -57.4051 | 341.1739 | 522 | -63.4223 | 522 | -54.6909 | 522 | 43619.5 | 522 | 44027.5898 | 522 | 43991.1836 | 522 | 43891.3203 | 522 | 44046.0391 | 522 | 44241.6016
523 | -57.2438 | 340.3074 | 523 | -63.6643 | 523 | -54.677 | 523 | 43614.2734 | 523 | 44055.9531 | 523 | 44034.4531 | 523 | 43907.0391 | 523 | 44049.5195 | 523 | 44224.7617
524 | -57.2256 | 340.21 | 524 | -63.5672 | 524 | -54.5852 | 524 | 43638 | 524 | 44048.7734 | 524 | 44091.2734 | 524 | 43895.8008 | 524 | 44049.8789 | 524 | 44195.1992
525 | -57.3113 | 340.6693 | 525 | -63.6277 | 525 | -54.6091 | 525 | 43665.9102 | 525 | 44098.3164 | 525 | 44144.4102 | 525 | 43892.7188 | 525 | 44060 | 525 | 44186.3594
526 | -57.262 | 340.4047 | 526 | -63.6582 | 526 | -54.5527 | 526 | 43695.2266 | 526 | 44095.8164 | 526 | 44146.8164 | 526 | 43867.6016 | 526 | 44048.8398 | 526 | 44205.3594
527 | -57.244 | 340.3086 | 527 | -63.7169 | 527 | -54.4413 | 527 | 43744.7266 | 527 | 44099.6367 | 527 | 44146.0898 | 527 | 43879.1602 | 527 | 44039.8008 | 527 | 44200.5195
528 | -57.3796 | 341.0365 | 528 | -63.8163 | 528 | -54.5997 | 528 | 43774.5 | 528 | 44130.5898 | 528 | 44152 | 528 | 43868.5195 | 528 | 44055.7188 | 528 | 44206.6797
529 | -57.5453 | 341.9313 | 529 | -63.8948 | 529 | -54.7301 | 529 | 43803.2734 | 529 | 44135.3164 | 529 | 44154.2266 | 529 | 43814.2383 | 529 | 44081.8008 | 529 | 44180.9219
530 | -57.6732 | 342.6268 | 530 | -63.9214 | 530 | -54.6805 | 530 | 43799.8164 | 530 | 44169.9102 | 530 | 44136.7734 | 530 | 43817.5586 | 530 | 44122.5195 | 530 | 44144.8789
531 | -57.8458 | 343.5709 | 531 | -64.0369 | 531 | -54.7715 | 531 | 43749.1836 | 531 | 44171.1367 | 531 | 44120.2266 | 531 | 43761.6016 | 531 | 44149.3984 | 531 | 44148.1211
532 | -58.0957 | 344.9503 | 532 | -64.1704 | 532 | -54.9793 | 532 | 43766.6836 | 532 | 44148.6367 | 532 | 44114.5898 | 532 | 43732.6406 | 532 | 44147.5586 | 532 | 44130.8008
533 | -57.6339 | 342.4128 | 533 | -64.0827 | 533 | -54.7273 | 533 | 43727.6367 | 533 | 44151.4102 | 533 | 44111.1836 | 533 | 43717.7617 | 533 | 44108.8789 | 533 | 44061.7617
534 | -57.644 | 342.4679 | 534 | -64.0472 | 534 | -54.704 | 534 | 43715.0898 | 534 | 44140.3164 | 534 | 44101.0898 | 534 | 43717.1992 | 534 | 44098.2812 | 534 | 44054
535 | -57.4926 | 341.6462 | 535 | -64.0226 | 535 | -54.7369 | 535 | 43729.1836 | 535 | 44141.7734 | 535 | 44071.8633 | 535 | 43671.7617 | 535 | 44051 | 535 | 44049.7188
536 | -57.3408 | 340.8279 | 536 | -63.9698 | 536 | -54.4319 | 536 | 43707.8164 | 536 | 44117.5898 | 536 | 44101.3164 | 536 | 43698.8008 | 536 | 44036.8789 | 536 | 44029.4414
537 | -57.5847 | 342.1451 | 537 | -64.281 | 537 | -54.727 | 537 | 43683.6367 | 537 | 44059.9102 | 537 | 44089.2734 | 537 | 43708.5586 | 537 | 44003.8398 | 537 | 43985.9609
538 | -57.8177 | 343.4169 | 538 | -64.5443 | 538 | -55.0074 | 538 | 43636.8633 | 538 | 44043.7734 | 538 | 44061.9531 | 538 | 43672.7617 | 538 | 44003.7617 | 538 | 44008.1992
539 | -57.5676 | 342.0526 | 539 | -64.5553 | 539 | -54.7374 | 539 | 43649.5469 | 539 | 44004.6836 | 539 | 44086.3164 | 539 | 43653.4414 | 539 | 43963.4414 | 539 | 44039.1602
540 | -57.7404 | 342.9936 | 540 | -64.5733 | 540 | -54.8091 | 540 | 43643.2266 | 540 | 44017.0469 | 540 | 44062.5469 | 540 | 43674.1211 | 540 | 43937.9609 | 540 | 44097.3203
541 | -57.9438 | 344.1099 | 541 | -64.5987 | 541 | -54.7361 | 541 | 43652.0898 | 541 | 44024.6836 | 541 | 44039.1367 | 541 | 43704.7188 | 541 | 43985.6406 | 541 | 44122.9609
542 | -57.9711 | 344.2603 | 542 | -64.665 | 542 | -54.4932 | 542 | 43740.9102 | 542 | 44066.7266 | 542 | 44094.3633 | 542 | 43695 | 542 | 44000.2812 | 542 | 44114.4805
543 | -58.7024 | 348.3633 | 543 | -65.0286 | 543 | -55.209 | 543 | 43763.7266 | 543 | 44110.7734 | 543 | 44115.1367 | 543 | 43682.3984 | 543 | 44043.3984 | 543 | 44108.5195
544 | -58.3429 | 346.3294 | 544 | -64.9144 | 544 | -54.9334 | 544 | 43773.3633 | 544 | 44143.0469 | 544 | 44110.3164 | 544 | 43647.1602 | 544 | 44041.8398 | 544 | 44100.2812
545 | -58.7245 | 348.4893 | 545 | -65.1138 | 545 | -55.0505 | 545 | 43781.1836 | 545 | 44188.5898 | 545 | 44123.5898 | 545 | 43627.3203 | 545 | 44033.2383 | 545 | 44098.8008
546 | -58.6755 | 348.2095 | 546 | -65.292 | 546 | -55.2359 | 546 | 43803.9102 | 546 | 44212.5898 | 546 | 44136.2266 | 546 | 43652.8789 | 546 | 44104.3203 | 546 | 44098.3594
547 | -58.8452 | 349.1801 | 547 | -65.4824 | 547 | -55.4332 | 547 | 43836.1836 | 547 | 44262.9531 | 547 | 44078.1367 | 547 | 43687.1211 | 547 | 44124 | 547 | 44091.2812
548 | -58.7803 | 348.8082 | 548 | -65.4615 | 548 | -55.4119 | 548 | 43863.3164 | 548 | 44287.5898 | 548 | 44093.5 | 548 | 43711.5195 | 548 | 44120.8008 | 548 | 44096.0781
549 | -58.6358 | 347.984 | 549 | -65.3711 | 549 | -55.0795 | 549 | 43872.3633 | 549 | 44253.8164 | 549 | 44096.5 | 549 | 43709.0781 | 549 | 44133.0391 | 549 | 44073.4805
550 | -58.8349 | 349.1209 | 550 | -65.5469 | 550 | -55.3107 | 550 | 43869.9531 | 550 | 44270.9531 | 550 | 44084.8164 | 550 | 43752.8789 | 550 | 44122.7617 | 550 | 44044.3594
551 | -59.0845 | 350.5612 | 551 | -65.8806 | 551 | -55.7373 | 551 | 43892.5 | 551 | 44314.5469 | 551 | 44076.5469 | 551 | 43776.8789 | 551 | 44146.9219 | 551 | 44056.3984
552 | -59.1015 | 350.6597 | 552 | -65.8499 | 552 | -55.7052 | 552 | 43900.9531 | 552 | 44396.5 | 552 | 44082.4102 | 552 | 43792.6797 | 552 | 44172.0781 | 552 | 44039.9609
[7 empty] | 553 | 43907.3164 | 553 | 44412.3164 | 553 | 44082.0898 | 553 | 43760.5195 | 553 | 44176.2812 | 553 | 44005.8008
[7 empty] | 554 | 43946.2266 | 554 | 44410.7266 | 554 | 44091 | 554 | 43728.0781 | 554 | 44175 | 554 | 44038.4805
[7 empty] | 555 | 43972 | 555 | 44431.1367 | 555 | 44083.7734 | 555 | 43740.2383 | 555 | 44168.3203 | 555 | 44051.2383
[7 empty] | 556 | 43962.0898 | 556 | 44416.2266 | 556 | 44114.3633 | 556 | 43711.0781 | 556 | 44157.4805 | 556 | 44029.5586
[7 empty] | 557 | 43940.9102 | 557 | 44401 | 557 | 44133.7734 | 557 | 43699.9609 | 557 | 44137.4414 | 557 | 44039.2812
[7 empty] | 558 | 43928.9531 | 558 | 44407.0898 | 558 | 44135 | 558 | 43673.9609 | 558 | 44106.7188 | 558 | 44052.7188
[7 empty] | 559 | 43956.8633 | 559 | 44392.7734 | 559 | 44183.3164 | 559 | 43672.8789 | 559 | 44133.5586 | 559 | 44065.6016
[7 empty] | 560 | 43966.7734 | 560 | 44372.1367 | 560 | 44228.4102 | 560 | 43642.0781 | 560 | 44097.2812 | 560 | 44098.0391
[7 empty] | 561 | 43941.1836 | 561 | 44341 | 561 | 44237.6367 | 561 | 43615.5195 | 561 | 44068.8008 | 561 | 44076.2383
[7 empty] | 562 | 43924.7266 | 562 | 44337.9531 | 562 | 44218.9531 | 562 | 43641.3984 | 562 | 44049.2812 | 562 | 44058.6016
[7 empty] | 563 | 43899.6367 | 563 | 44311.5469 | 563 | 44175.8633 | 563 | 43644.6797 | 563 | 44109.9219 | 563 | 44049.9609
[7 empty] | 564 | 43832.1367 | 564 | 44319.1367 | 564 | 44185.3164 | 564 | 43716.1992 | 564 | 44136.8789 | 564 | 44035.1602
[7 empty] | 565 | 43837.0898 | 565 | 44315.6367 | 565 | 44168 | 565 | 43739.1211 | 565 | 44164.8008 | 565 | 44032.3594
[7 empty] | 566 | 43835.3164 | 566 | 44337.5469 | 566 | 44149.4102 | 566 | 43752 | 566 | 44193.0781 | 566 | 44043.3984
[7 empty] | 567 | 43809.5 | 567 | 44287.8164 | 567 | 44137.4531 | 567 | 43756.4805 | 567 | 44228.1992 | 567 | 44047.0781
[7 empty] | 568 | 43737.2266 | 568 | 44309.5 | 568 | 44185.6367 | 568 | 43789.0781 | 568 | 44264.1211 | 568 | 44053
[7 empty] | 569 | 43678.5898 | 569 | 44303.1367 | 569 | 44188.4531 | 569 | 43848.7617 | 569 | 44318.6016 | 569 | 44058.8789
[7 empty] | 570 | 43720.3164 | 570 | 44284.8633 | 570 | 44185.3164 | 570 | 43867.6797 | 570 | 44325.8789 | 570 | 44064.8398
[7 empty] | 571 | 43768.7734 | 571 | 44278.5 | 571 | 44212.7734 | 571 | 43864.5586 | 571 | 44329.9219 | 571 | 44070.3203
[7 empty] | 572 | 43780.7266 | 572 | 44293.9102 | 572 | 44228.1836 | 572 | 43874.4805 | 572 | 44350.6016 | 572 | 44096.1211
[7 empty] | 573 | 43788.6836 | 573 | 44395.2734 | 573 | 44186.4102 | 573 | 43872.7188 | 573 | 44388.4414 | 573 | 44115.8789
[7 empty] | 574 | 43822.0898 | 574 | 44369.8164 | 574 | 44080.5898 | 574 | 43909.8789 | 574 | 44426.0391 | 574 | 44140.5586
[7 empty] | 575 | 43856.3633 | 575 | 44343.0898 | 575 | 44027.5469 | 575 | 43908.1992 | 575 | 44434.3203 | 575 | 44182.1211
[7 empty] | 576 | 43897.7266 | 576 | 44399.1367 | 576 | 44012.2266 | 576 | 43928.6016 | 576 | 44441.1992 | 576 | 44197.7188
[7 empty] | 577 | 43907.3164 | 577 | 44422.9102 | 577 | 43973.9531 | 577 | 43939.4414 | 577 | 44437.6016 | 577 | 44192.9609
[7 empty] | 578 | 43899.5 | 578 | 44454.8164 | 578 | 43987.7734 | 578 | 43947.7617 | 578 | 44440.5195 | 578 | 44155.8008
[7 empty] | 579 | 43881.4102 | 579 | 44472.4531 | 579 | 43978.6836 | 579 | 43915.8789 | 579 | 44435.7188 | 579 | 44144.9609
[7 empty] | 580 | 43891.6367 | 580 | 44466.3164 | 580 | 43985.5469 | 580 | 43881.3203 | 580 | 44462.5586 | 580 | 44104.4414
[7 empty] | 581 | 43879.3164 | 581 | 44442.2266 | 581 | 43979.1367 | 581 | 43893.2383 | 581 | 44421.7617 | 581 | 44069.8398
[7 empty] | 582 | 43806.8633 | 582 | 44391.0469 | 582 | 43976.7266 | 582 | 43913.5195 | 582 | 44399.9219 | 582 | 44077.9609
[7 empty] | 583 | 43787.1836 | 583 | 44397.2734 | 583 | 44000.4102 | 583 | 43921.7617 | 583 | 44370.2812 | 583 | 44099
[7 empty] | 584 | 43780.3633 | 584 | 44383.6836 | 584 | 44025.5898 | 584 | 43868.6016 | 584 | 44342.1992 | 584 | 44091.2812
[7 empty] | 585 | 43716.2266 | 585 | 44395.1367 | 585 | 44027.5 | 585 | 43882.1211 | 585 | 44327.4805 | 585 | 44131.6797
[7 empty] | 586 | 43770.5 | 586 | 44405.0469 | 586 | 44028.6367 | 586 | 43792.0781 | 586 | 44325.0781 | 586 | 44169.3203
[7 empty] | 587 | 43792.5469 | 587 | 44386.6836 | 587 | 44025.6836 | 587 | 43779.6797 | 587 | 44309 | 587 | 44208.9219
[7 empty] | 588 | 43836.8164 | 588 | 44405.0898 | 588 | 44031.0898 | 588 | 43813.4414 | 588 | 44295.7617 | 588 | 44124.4414
[7 empty] | 589 | 43821.5 | 589 | 44358.7266 | 589 | 44013.1836 | 589 | 43775.3594 | 589 | 44288.6797 | 589 | 44049.9609
[7 empty] | 590 | 43829.9531 | 590 | 44389.3633 | 590 | 44053.6367 | 590 | 43710.4805 | 590 | 44241 | 590 | 43992.1602
[7 empty] | 591 | 43847.1836 | 591 | 44389.4531 | 591 | 44072.6367 | 591 | 43658.2383 | 591 | 44268.7188 | 591 | 43956.7617
[7 empty] | 592 | 43874.7266 | 592 | 44424.8164 | 592 | 44057.7266 | 592 | 43684.8789 | 592 | 44265.9219 | 592 | 43918.9219
[7 empty] | 593 | 43872.7266 | 593 | 44438.2266 | 593 | 44007.0898 | 593 | 43750.6406 | 593 | 44294.8008 | 593 | 43931.8789
[7 empty] | 594 | 43855.6367 | 594 | 44420.3164 | 594 | 44023.9102 | 594 | 43762.2383 | 594 | 44385.5195 | 594 | 43923.6406
[7 empty] | 595 | 43832.6836 | 595 | 44389.3633 | 595 | 44014.6836 | 595 | 43766.1602 | 595 | 44389.1602 | 595 | 43921.9219
[7 empty] | 596 | 43820.5469 | 596 | 44436.6367 | 596 | 44016.0469 | 596 | 43803.2383 | 596 | 44361.1992 | 596 | 43933.2383
[7 empty] | 597 | 43814.5 | 597 | 44452.5898 | 597 | 44033.6367 | 597 | 43852.8398 | 597 | 44381.6016 | 597 | 43940.6797
[7 empty] | 598 | 43744.1367 | 598 | 44473.0469 | 598 | 44032.4531 | 598 | 43899.9609 | 598 | 44391.2383 | 598 | 43970.8398
[7 empty] | 599 | 43715.8633 | 599 | 44467.4102 | 599 | 44069.2734 | 599 | 43841.3594 | 599 | 44448.5195 | 599 | 43973.6016
[7 empty] | 600 | 43743.1836 | 600 | 44428.2734 | 600 | 44081.7266 | 600 | 43876.1602 | 600 | 44500.3984 | 600 | 43999
[7 empty] | 601 | 43801.5 | 601 | 44446.5898 | 601 | 44062.6836 | 601 | 43872.0781 | 601 | 44505.3594 | 601 | 43992.4414
[7 empty] | 602 | 43867.2734 | 602 | 44470.0469 | 602 | 44049.6836 | 602 | 43872.0391 | 602 | 44505.0391 | 602 | 43982.1992
[7 empty] | 603 | 43916 | 603 | 44427.6836 | 603 | 44048.6367 | 603 | 43857.6797 | 603 | 44442.9219 | 603 | 44004.8789
[7 empty] | 604 | 43944.6836 | 604 | 44383 | 604 | 43991.8164 | 604 | 43790.3203 | 604 | 44368.2383 | 604 | 43973.8398
[7 empty] | 605 | 43916.8164 | 605 | 44388.7734 | 605 | 43960 | 605 | 43740.8008 | 605 | 44373.7188 | 605 | 44016.4414
[7 empty] | 606 | 43898.3164 | 606 | 44385.1367 | 606 | 43954.3633 | 606 | 43728.6406 | 606 | 44407.4414 | 606 | 44025.9219
[7 empty] | 607 | 43926.9102 | 607 | 44387.5898 | 607 | 43949.2734 | 607 | 43671.2383 | 607 | 44408.6797 | 607 | 44019.7617
[7 empty] | 608 | 43823.8633 | 608 | 44416 | 608 | 43969.6836 | 608 | 43730.0781 | 608 | 44378.9609 | 608 | 43979.8398
[7 empty] | 609 | 43839.8164 | 609 | 44436.2734 | 609 | 44052.4102 | 609 | 43795.7188 | 609 | 44406.8008 | 609 | 43995.0391
[7 empty] | 610 | 43870.1367 | 610 | 44429 | 610 | 44025.9102 | 610 | 43818.4414 | 610 | 44382.9219 | 610 | 43985.4414
[7 empty] | 611 | 43837 | 611 | 44456.1367 | 611 | 44031.6367 | 611 | 43802.6797 | 611 | 44365.3203 | 611 | 43979.1992
[7 empty] | 612 | 43811.9531 | 612 | 44540.2266 | 612 | 43976.6836 | 612 | 43834.6016 | 612 | 44394.1211 | 612 | 43993.0781
[7 empty] | 613 | 43816.0469 | 613 | 44521.0898 | 613 | 43966.3164 | 613 | 43851.1211 | 613 | 44415.5195 | 613 | 44001.4414
[7 empty] | 614 | 43788.0469 | 614 | 44546.6367 | 614 | 43952.2266 | 614 | 43868.1602 | 614 | 44413.0781 | 614 | 44021.3984
[7 empty] | 615 | 43788.7266 | 615 | 44587.7734 | 615 | 43968.4102 | 615 | 43854.6797 | 615 | 44420.2812 | 615 | 44022.0391
[7 empty] | 616 | 43776.5 | 616 | 44573.2266 | 616 | 43949.1367 | 616 | 43848.1992 | 616 | 44413.2383 | 616 | 43987.2812
[7 empty] | 617 | 43762 | 617 | 44569.0898 | 617 | 43970.2266 | 617 | 43845.5586 | 617 | 44423.1992 | 617 | 43990.2383
[7 empty] | 618 | 43763.5898 | 618 | 44594.2734 | 618 | 43977.6836 | 618 | 43843.4805 | 618 | 44457.1211 | 618 | 43979.6406
[7 empty] | 619 | 43798.7266 | 619 | 44589.4531 | 619 | 44005.1367 | 619 | 43818.8008 | 619 | 44500.2812 | 619 | 43938.5586
[7 empty] | 620 | 43813.7266 | 620 | 44622.7734 | 620 | 44006.4531 | 620 | 43767.3984 | 620 | 44479.8008 | 620 | 43916.1602
[7 empty] | 621 | 43833.9102 | 621 | 44642.1836 | 621 | 43994.9102 | 621 | 43715.2383 | 621 | 44473.4805 | 621 | 43922.0781
[7 empty] | 622 | 43842.6367 | 622 | 44570.1367 | 622 | 44028 | 622 | 43750.9219 | 622 | 44496.4805 | 622 | 43911.9219
[7 empty] | 623 | 43883.2266 | 623 | 44511.1836 | 623 | 44015.6367 | 623 | 43799.9609 | 623 | 44487.2812 | 623 | 43915.4414
[7 empty] | 624 | 43855.2266 | 624 | 44473.2734 | 624 | 44003.6367 | 624 | 43839.2383 | 624 | 44441.9609 | 624 | 43964.2383
[7 empty] | 625 | 43882.3633 | 625 | 44450.7734 | 625 | 44010.0469 | 625 | 43878.3594 | 625 | 44442.3594 | 625 | 43968.3594
[7 empty] | 626 | 43870.9102 | 626 | 44456.2734 | 626 | 44004.7266 | 626 | 43898.4805 | 626 | 44423.0781 | 626 | 43981.0781
[7 empty] | 627 | 43868.5898 | 627 | 44455.2266 | 627 | 44010.4531 | 627 | 43896.6016 | 627 | 44391.8789 | 627 | 43936.3984
[7 empty] | 628 | 43897.3633 | 628 | 44503.8633 | 628 | 44028.0469 | 628 | 43903.6016 | 628 | 44445.5195 | 628 | 43901.0391
[7 empty] | 629 | 43930.1836 | 629 | 44525.0898 | 629 | 44044.0469 | 629 | 43889.1211 | 629 | 44465.4414 | 629 | 43956.4805
[7 empty] | 630 | 43934.8164 | 630 | 44564.6367 | 630 | 44036.4102 | 630 | 43784.4414 | 630 | 44492.9219 | 630 | 43954.3594
[7 empty] | 631 | 43926.1836 | 631 | 44593.8164 | 631 | 44020.6836 | 631 | 43850.3984 | 631 | 44528.5586 | 631 | 43917.5586
[7 empty] | 632 | 43910.4531 | 632 | 44615.8633 | 632 | 43992.5469 | 632 | 43831 | 632 | 44548.8789 | 632 | 43948.7617
[7 empty] | 633 | 43964.3633 | 633 | 44614.5898 | 633 | 43997.2266 | 633 | 43812.8008 | 633 | 44553.0781 | 633 | 43953.7188
[7 empty] | 634 | 43951.0469 | 634 | 44608.6367 | 634 | 43939.5469 | 634 | 43792.7188 | 634 | 44562.8008 | 634 | 43921.9219
[7 empty] | 635 | 43947 | 635 | 44631.5 | 635 | 43958.5898 | 635 | 43757.8008 | 635 | 44550.3984 | 635 | 43947.9609
[7 empty] | 636 | 43917.7266 | 636 | 44673.8633 | 636 | 43979.6367 | 636 | 43748.6406 | 636 | 44555.9609 | 636 | 43940.2383
[7 empty] | 637 | 43977.9102 | 637 | 44646.7266 | 637 | 43994.8164 | 637 | 43773.5195 | 637 | 44565.5195 | 637 | 43971.2812
[7 empty] | 638 | 44010.7266 | 638 | 44635.5 | 638 | 44032.9531 | 638 | 43773.1602 | 638 | 44596.0781 | 638 | 43969.3594
[7 empty] | 639 | 44029.0469 | 639 | 44613.5469 | 639 | 44059.3164 | 639 | 43767.7617 | 639 | 44616.5586 | 639 | 43958.5195
[7 empty] | 640 | 44042.4102 | 640 | 44644.7734 | 640 | 44081.0898 | 640 | 43784.6797 | 640 | 44633.1992 | 640 | 43950.1602
[7 empty] | 641 | 44019.5 | 641 | 44696.5469 | 641 | 44094.4102 | 641 | 43804.8789 | 641 | 44666.6797 | 641 | 43941.7617
[7 empty] | 642 | 44014.3164 | 642 | 44731.4531 | 642 | 44042.1836 | 642 | 43816.0391 | 642 | 44669.5586 | 642 | 43931.2812
[7 empty] | 643 | 44034.5898 | 643 | 44730.7266 | 643 | 44035.6836 | 643 | 43839.6016 | 643 | 44627.7617 | 643 | 43955.6016
[7 empty] | 644 | 44033.7734 | 644 | 44734.3164 | 644 | 44028.7734 | 644 | 43845.6406 | 644 | 44554.8789 | 644 | 43978.3594
[7 empty] | 645 | 44021.1836 | 645 | 44793.2266 | 645 | 44023.5469 | 645 | 43869.5586 | 645 | 44530.0781 | 645 | 43982.3984
[7 empty] | 646 | 44019.1836 | 646 | 44837.9531 | 646 | 44040.7734 | 646 | 43861.8789 | 646 | 44527.2383 | 646 | 43941.4805
[7 empty] | 647 | 43998 | 647 | 44841.1836 | 647 | 44078 | 647 | 43854.7188 | 647 | 44537.9219 | 647 | 43915.8789
[7 empty] | 648 | 44017.5469 | 648 | 44822.4102 | 648 | 44121.2734 | 648 | 43838.6797 | 648 | 44539.8398 | 648 | 43939.7188
[7 empty] | 649 | 44039.7266 | 649 | 44836.9531 | 649 | 44127 | 649 | 43878.8008 | 649 | 44558.7617 | 649 | 43887.1602
[7 empty] | 650 | 44078.0898 | 650 | 44845.9531 | 650 | 44118.8164 | 650 | 43880.8008 | 650 | 44572.8008 | 650 | 43881.1602
[7 empty] | 651 | 44104 | 651 | 44811.0469 | 651 | 44116 | 651 | 43895.9219 | 651 | 44604.5586 | 651 | 43914.0391
[7 empty] | 652 | 44142.2734 | 652 | 44805.5 | 652 | 44135.0469 | 652 | 43949.8789 | 652 | 44619.0781 | 652 | 43948.7188
[7 empty] | 653 | 44125.6836 | 653 | 44818.0469 | 653 | 44147.7266 | 653 | 43943.0391 | 653 | 44612.8008 | 653 | 43977.3594
[7 empty] | 654 | 44142.0898 | 654 | 44852.5898 | 654 | 44139.6836 | 654 | 43908.9219 | 654 | 44628.5195 | 654 | 43987.9609
[13 empty] | 655 | 43962.9609 | 655 | 44637.6016 | 655 | 44015.2383
[13 empty] | 656 | 43955.5586 | 656 | 44644.6406 | 656 | 43997.8789
[13 empty] | 657 | 43924.3594 | 657 | 44681.2812 | 657 | 43998.4805
[13 empty] | 658 | 43916.0391 | 658 | 44654.5586 | 658 | 44006.1992
[13 empty] | 659 | 43939.2383 | 659 | 44604.9219 | 659 | 43986.6797
[13 empty] | 660 | 44003.8789 | 660 | 44617 | 660 | 43998.6406
[13 empty] | 661 | 44026.7617 | 661 | 44657.3594 | 661 | 43990.4805
[13 empty] | 662 | 44049.8008 | 662 | 44663.8008 | 662 | 43997.1602
[13 empty] | 663 | 44029.1211 | 663 | 44714.6016 | 663 | 44030.0781
[13 empty] | 664 | 43989.5586 | 664 | 44735.1992 | 664 | 44042.1992
[13 empty] | 665 | 44012.4805 | 665 | 44755.8008 | 665 | 44045.6406
[13 empty] | 666 | 44025.2383 | 666 | 44797.7617 | 666 | 44042.8398
[13 empty] | 667 | 43999.5586 | 667 | 44813.1992 | 667 | 44049.7617
[13 empty] | 668 | 44005.1211 | 668 | 44840.0391 | 668 | 44076.8789
[13 empty] | 669 | 44013.2383 | 669 | 44840.5586 | 669 | 44088.7617
[13 empty] | 670 | 44026.7188 | 670 | 44863.4414 | 670 | 44087.5195
[13 empty] | 671 | 44039.7188 | 671 | 44885.6406 | 671 | 44089.1992
[13 empty] | 672 | 44068.6016 | 672 | 44870.4414 | 672 | 44058.8789
[13 empty] | 673 | 44085 | 673 | 44839.6797 | 673 | 44056.6406
[13 empty] | 674 | 44123.4805 | 674 | 44857.1992 | 674 | 43991.6797
[13 empty] | 675 | 44082.8008 | 675 | 44867.6016 | 675 | 43947.5195
[13 empty] | 676 | 44137.2383 | 676 | 44856.9219 | 676 | 43926.3203
[13 empty] | 677 | 44096 | 677 | 44878.7617 | 677 | 43928.7188
[13 empty] | 678 | 44089.3203 | 678 | 44891.2383 | 678 | 43946.5586
[13 empty] | 679 | 44025.1992 | 679 | 44897.0391 | 679 | 43904.5195
[13 empty] | 680 | 44028.5195 | 680 | 44916.7188 | 680 | 43931.6797
[13 empty] | 681 | 43991.7617 | 681 | 44860.9609 | 681 | 43950.5195
[13 empty] | 682 | 44024.3203 | 682 | 44859.9609 | 682 | 43975.6797
[13 empty] | 683 | 44066.9219 | 683 | 44915.8398 | 683 | 44025.3594
[13 empty] | 684 | 44066.5195 | 684 | 44902.8398 | 684 | 44039.1602
[13 empty] | 685 | 44045.6797 | 685 | 44913.0391 | 685 | 44061.1992
[13 empty] | 686 | 44071.5586 | 686 | 44921.7188 | 686 | 43997.5586
[13 empty] | 687 | 44071.6406 | 687 | 44957.5195 | 687 | 43987.4414
[13 empty] | 688 | 44077.1602 | 688 | 44959.5195 | 688 | 44013.0391
[13 empty] | 689 | 44055.9609 | 689 | 45006.9219 | 689 | 44002.3984

## unif profiles 0Gy
|  | dose=c+b/(X-a) |  | this equation uses the rational function to fit the dose data | [5 empty] | pixels to cm conversion | 1px = |  | 1/ | 1 inch = 2.54 centimetres
a |  | -0.040473 | [11 empty] | dpi | ppcm | cm/px
b |  | 2.7151 | [11 empty] | 72 | 28.3465 | 0.035278
c |  | -3.8345
centre across | dist cm | normalized film response | Gy | near to scan array, across | normalized film response | Gy | opposite to scan array, across |  |  |  | centre, down |  |  |  |  | left side, down |  |  | right side, down |  |  | centre, down repeated |  | left side, down, repeated |  | right side , down, repeated
0 | 0 | 0.66709 | 0.0026514 | 0 | 0.65984 | 0.042376 | 0 | 0.69445 | -0.14021 |  | 0 | 0 | 0.67062 | -0.016384 |  | 0 | 0.66018 | 0.040489 | 0 | 0.69872 | -0.16153 | 0 | 44048.9609 | 0 | 43328.2383 | 0 | 45893.4805
1 | 0.035278 | 0.66621 | 0.0074279 | 1 | 0.65973 | 0.042975 | 1 | 0.69273 | -0.13154 |  | 1 | 0.035278 | 0.67071 | -0.016898 |  | 1 | 0.66023 | 0.040217 | 1 | 0.69851 | -0.16047 | 1 | 44011.4805 | 1 | 43321.6797 | 1 | 45858.3594
2 | 0.070556 | 0.66709 | 0.0026621 | 2 | 0.66083 | 0.036937 | 2 | 0.69236 | -0.12965 |  | 2 | 0.070556 | 0.67061 | -0.016339 |  | 2 | 0.65959 | 0.043771 | 2 | 0.69838 | -0.15985 | 2 | 44005.9219 | 2 | 43315.1992 | 2 | 45851.1992
3 | 0.10583 | 0.66643 | 0.0062561 | 3 | 0.66117 | 0.035041 | 3 | 0.69147 | -0.12517 |  | 3 | 0.10583 | 0.67146 | -0.020911 |  | 3 | 0.66003 | 0.041325 | 3 | 0.69796 | -0.15775 | 3 | 44022.8789 | 3 | 43334.6406 | 3 | 45866.8008
4 | 0.14111 | 0.66564 | 0.01055 | 4 | 0.66009 | 0.040984 | 4 | 0.6902 | -0.11871 |  | 4 | 0.14111 | 0.67071 | -0.016894 |  | 4 | 0.66042 | 0.039197 | 4 | 0.69785 | -0.15723 | 4 | 44082 | 4 | 43351.2383 | 4 | 45872.4805
5 | 0.17639 | 0.66703 | 0.0029967 | 5 | 0.66133 | 0.034174 | 5 | 0.69108 | -0.1232 |  | 5 | 0.17639 | 0.6702 | -0.014148 |  | 5 | 0.66059 | 0.038258 | 5 | 0.6981 | -0.15844 | 5 | 44090.2812 | 5 | 43385.5586 | 5 | 45888.0781
6 | 0.21167 | 0.66681 | 0.0041593 | 6 | 0.66186 | 0.031206 | 6 | 0.69082 | -0.12188 |  | 6 | 0.21167 | 0.66949 | -0.010322 |  | 6 | 0.66111 | 0.035371 | 6 | 0.69758 | -0.15587 | 6 | 44102.8789 | 6 | 43363.6406 | 6 | 45904.4414
7 | 0.24694 | 0.66537 | 0.012011 | 7 | 0.66 | 0.041473 | 7 | 0.6891 | -0.11311 |  | 7 | 0.24694 | 0.67086 | -0.017653 |  | 7 | 0.66116 | 0.03508 | 7 | 0.69716 | -0.15377 | 7 | 44142.6797 | 7 | 43402.5586 | 7 | 45936
8 | 0.28222 | 0.66508 | 0.013573 | 8 | 0.65969 | 0.043192 | 8 | 0.68851 | -0.11007 |  | 8 | 0.28222 | 0.66865 | -0.0057778 |  | 8 | 0.66155 | 0.032958 | 8 | 0.69684 | -0.15218 | 8 | 44069.8789 | 8 | 43373.7617 | 8 | 45897.6016
9 | 0.3175 | 0.66593 | 0.008947 | 9 | 0.66036 | 0.039502 | 9 | 0.68775 | -0.10619 |  | 9 | 0.3175 | 0.66969 | -0.011379 |  | 9 | 0.66164 | 0.032431 | 9 | 0.69703 | -0.15309 | 9 | 44086.2812 | 9 | 43380.2383 | 9 | 45879.8008
10 | 0.35278 | 0.66629 | 0.0070206 | 10 | 0.66129 | 0.034353 | 10 | 0.68843 | -0.10969 |  | 10 | 0.35278 | 0.67015 | -0.013879 |  | 10 | 0.66074 | 0.037392 | 10 | 0.69723 | -0.15413 | 10 | 44048.8789 | 10 | 43423.6406 | 10 | 45870.8789
11 | 0.38806 | 0.66553 | 0.011157 | 11 | 0.66141 | 0.033702 | 11 | 0.68888 | -0.11197 |  | 11 | 0.38806 | 0.67154 | -0.021334 |  | 11 | 0.66135 | 0.03404 | 11 | 0.69713 | -0.15362 | 11 | 44017.4805 | 11 | 43472.8008 | 11 | 45822.2812
12 | 0.42333 | 0.66621 | 0.0074279 | 12 | 0.66109 | 0.035491 | 12 | 0.688 | -0.10747 |  | 12 | 0.42333 | 0.67059 | -0.016235 |  | 12 | 0.66093 | 0.036377 | 12 | 0.6964 | -0.14998 | 12 | 43993.3203 | 12 | 43458.4414 | 12 | 45802.6016
13 | 0.45861 | 0.66532 | 0.012304 | 13 | 0.66038 | 0.039389 | 13 | 0.68676 | -0.10113 |  | 13 | 0.45861 | 0.66999 | -0.01301 |  | 13 | 0.6613 | 0.034307 | 13 | 0.69661 | -0.15101 | 13 | 43971.6406 | 13 | 43490.0781 | 13 | 45805.3984
14 | 0.49389 | 0.66589 | 0.0091818 | 14 | 0.66098 | 0.036091 | 14 | 0.68778 | -0.10635 |  | 14 | 0.49389 | 0.67084 | -0.017568 |  | 14 | 0.66042 | 0.039158 | 14 | 0.6961 | -0.14848 | 14 | 43997.3984 | 14 | 43432.8398 | 14 | 45772.3203
15 | 0.52917 | 0.66538 | 0.011978 | 15 | 0.66078 | 0.037197 | 15 | 0.68689 | -0.10178 |  | 15 | 0.52917 | 0.6714 | -0.020584 |  | 15 | 0.66059 | 0.038216 | 15 | 0.69585 | -0.14724 | 15 | 43922.7617 | 15 | 43373.1992 | 15 | 45745.6016
16 | 0.56444 | 0.66424 | 0.01817 | 16 | 0.66047 | 0.038886 | 16 | 0.68638 | -0.099167 |  | 16 | 0.56444 | 0.67138 | -0.020483 |  | 16 | 0.65993 | 0.041886 | 16 | 0.69577 | -0.1468 | 16 | 43847.6406 | 16 | 43344.8789 | 16 | 45720.5586
17 | 0.59972 | 0.66338 | 0.022902 | 17 | 0.6613 | 0.034338 | 17 | 0.68612 | -0.097865 |  | 17 | 0.59972 | 0.67148 | -0.021 |  | 17 | 0.65994 | 0.04184 | 17 | 0.69519 | -0.14391 | 17 | 43876.7617 | 17 | 43293.7188 | 17 | 45724
18 | 0.635 | 0.66311 | 0.02435 | 18 | 0.66057 | 0.038318 | 18 | 0.68579 | -0.096161 |  | 18 | 0.635 | 0.67084 | -0.017597 |  | 18 | 0.65993 | 0.041901 | 18 | 0.696 | -0.14796 | 18 | 43926.0781 | 18 | 43244.3984 | 18 | 45723.2383
19 | 0.67028 | 0.66314 | 0.024208 | 19 | 0.65974 | 0.042942 | 19 | 0.68547 | -0.094511 |  | 19 | 0.67028 | 0.6702 | -0.014122 |  | 19 | 0.66006 | 0.041153 | 19 | 0.69662 | -0.15105 | 19 | 43908.7617 | 19 | 43223.1602 | 19 | 45690.4414
20 | 0.70556 | 0.66453 | 0.01659 | 20 | 0.66042 | 0.03918 | 20 | 0.6862 | -0.098243 |  | 20 | 0.70556 | 0.66973 | -0.011589 |  | 20 | 0.66028 | 0.03996 | 20 | 0.6975 | -0.15546 | 20 | 43878.5195 | 20 | 43161.3203 | 20 | 45709.6406
21 | 0.74083 | 0.66385 | 0.02029 | 21 | 0.66014 | 0.040706 | 21 | 0.68572 | -0.095765 |  | 21 | 0.74083 | 0.66934 | -0.0095301 |  | 21 | 0.66076 | 0.037277 | 21 | 0.69756 | -0.15574 | 21 | 43864.1602 | 21 | 43146.7188 | 21 | 45670.1992
22 | 0.77611 | 0.66369 | 0.021176 | 22 | 0.66023 | 0.040229 | 22 | 0.68502 | -0.092196 |  | 22 | 0.77611 | 0.66864 | -0.0057028 |  | 22 | 0.66065 | 0.037917 | 22 | 0.69777 | -0.15681 | 22 | 43880.4805 | 22 | 43148.3203 | 22 | 45640.0781
23 | 0.81139 | 0.66393 | 0.019898 | 23 | 0.66017 | 0.040544 | 23 | 0.68462 | -0.090117 |  | 23 | 0.81139 | 0.66775 | -0.00090292 |  | 23 | 0.6605 | 0.038721 | 23 | 0.6975 | -0.15546 | 23 | 43916.3984 | 23 | 43132 | 23 | 45648.3203
24 | 0.84667 | 0.66377 | 0.020733 | 24 | 0.66018 | 0.040489 | 24 | 0.68437 | -0.088795 |  | 24 | 0.84667 | 0.66765 | -0.00039242 |  | 24 | 0.6605 | 0.038729 | 24 | 0.69642 | -0.15009 | 24 | 43894.8789 | 24 | 43148.8789 | 24 | 45625.0781
25 | 0.88194 | 0.66331 | 0.023269 | 25 | 0.6598 | 0.042615 | 25 | 0.68371 | -0.085401 |  | 25 | 0.88194 | 0.66797 | -0.0021002 |  | 25 | 0.66019 | 0.040474 | 25 | 0.69541 | -0.14503 | 25 | 43917.5195 | 25 | 43186.2383 | 25 | 45627.5586
26 | 0.91722 | 0.66246 | 0.027956 | 26 | 0.65925 | 0.045648 | 26 | 0.68342 | -0.083913 |  | 26 | 0.91722 | 0.66784 | -0.0014175 |  | 26 | 0.66053 | 0.038591 | 26 | 0.69501 | -0.14303 | 26 | 43913.1602 | 26 | 43223.9219 | 26 | 45618.1602
27 | 0.9525 | 0.66273 | 0.026454 | 27 | 0.65949 | 0.04432 | 27 | 0.68355 | -0.084566 |  | 27 | 0.9525 | 0.66843 | -0.0045715 |  | 27 | 0.66012 | 0.040861 | 27 | 0.69439 | -0.1399 | 27 | 43860.8008 | 27 | 43237.8008 | 27 | 45565.7188
28 | 0.98778 | 0.66344 | 0.022546 | 28 | 0.66039 | 0.039323 | 28 | 0.68388 | -0.08631 |  | 28 | 0.98778 | 0.66871 | -0.0061298 |  | 28 | 0.66032 | 0.039741 | 28 | 0.69402 | -0.13803 | 28 | 43813.0391 | 28 | 43274.7188 | 28 | 45550.1992
29 | 1.0231 | 0.66282 | 0.025955 | 29 | 0.66048 | 0.038864 | 29 | 0.68331 | -0.083311 |  | 29 | 1.0231 | 0.66856 | -0.0052721 |  | 29 | 0.66047 | 0.038871 | 29 | 0.69416 | -0.13875 | 29 | 43888.9219 | 29 | 43270.8789 | 29 | 45558.3594
30 | 1.0583 | 0.66311 | 0.024354 | 30 | 0.66066 | 0.037827 | 30 | 0.68313 | -0.08241 |  | 30 | 1.0583 | 0.66921 | -0.0088272 |  | 30 | 0.66109 | 0.035455 | 30 | 0.69383 | -0.13706 | 30 | 43784.0781 | 30 | 43226.2383 | 30 | 45585.3594
31 | 1.0936 | 0.6633 | 0.023346 | 31 | 0.66056 | 0.038395 | 31 | 0.68285 | -0.080934 |  | 31 | 1.0936 | 0.66988 | -0.01244 |  | 31 | 0.66103 | 0.035807 | 31 | 0.69373 | -0.1366 | 31 | 43861.6406 | 31 | 43175.7188 | 31 | 45545.5586
32 | 1.1289 | 0.66245 | 0.027974 | 32 | 0.66007 | 0.04112 | 32 | 0.68142 | -0.073514 |  | 32 | 1.1289 | 0.66937 | -0.0096461 |  | 32 | 0.66167 | 0.032267 | 32 | 0.69379 | -0.13688 | 32 | 43937.1211 | 32 | 43225.6797 | 32 | 45538.6797
33 | 1.1642 | 0.66318 | 0.024008 | 33 | 0.66111 | 0.035367 | 33 | 0.68202 | -0.076611 |  | 33 | 1.1642 | 0.66958 | -0.010789 |  | 33 | 0.66213 | 0.029721 | 33 | 0.69357 | -0.13579 | 33 | 43989.1211 | 33 | 43197 | 33 | 45543.5586
34 | 1.1994 | 0.66215 | 0.029648 | 34 | 0.65972 | 0.043026 | 34 | 0.68053 | -0.068848 |  | 34 | 1.1994 | 0.66924 | -0.0089394 |  | 34 | 0.66136 | 0.03399 | 34 | 0.69383 | -0.1371 | 34 | 43923.1992 | 34 | 43270 | 34 | 45612.1211
35 | 1.2347 | 0.66245 | 0.027963 | 35 | 0.66014 | 0.040728 | 35 | 0.6813 | -0.072863 |  | 35 | 1.2347 | 0.66935 | -0.0095452 |  | 35 | 0.66089 | 0.036596 | 35 | 0.69397 | -0.13779 | 35 | 43868.3203 | 35 | 43254.8789 | 35 | 45668.8008
36 | 1.27 | 0.66273 | 0.026429 | 36 | 0.65985 | 0.042343 | 36 | 0.68144 | -0.073614 |  | 36 | 1.27 | 0.66907 | -0.0080382 |  | 36 | 0.66006 | 0.04116 | 36 | 0.69351 | -0.13545 | 36 | 43887 | 36 | 43239.1602 | 36 | 45661.1992
37 | 1.3053 | 0.66308 | 0.024524 | 37 | 0.661 | 0.035945 | 37 | 0.68214 | -0.077257 |  | 37 | 1.3053 | 0.66913 | -0.0083748 |  | 37 | 0.66041 | 0.039235 | 37 | 0.69323 | -0.13407 | 37 | 43934.0781 | 37 | 43193.8008 | 37 | 45677.9609
38 | 1.3406 | 0.66302 | 0.024863 | 38 | 0.66035 | 0.039546 | 38 | 0.6811 | -0.071843 |  | 38 | 1.3406 | 0.6691 | -0.0082253 |  | 38 | 0.65946 | 0.044502 | 38 | 0.69262 | -0.13098 | 38 | 43922.5195 | 38 | 43168.2383 | 38 | 45642.8008
39 | 1.3758 | 0.664 | 0.019494 | 39 | 0.66125 | 0.034609 | 39 | 0.68147 | -0.073766 |  | 39 | 1.3758 | 0.66887 | -0.0069534 |  | 39 | 0.65869 | 0.04875 | 39 | 0.69244 | -0.13009 | 39 | 43951.4414 | 39 | 43207.2812 | 39 | 45585.7188
40 | 1.4111 | 0.66329 | 0.023386 | 40 | 0.66042 | 0.039165 | 40 | 0.68024 | -0.067365 |  | 40 | 1.4111 | 0.66917 | -0.0086142 |  | 40 | 0.65971 | 0.043122 | 40 | 0.69288 | -0.13228 | 40 | 43911.4414 | 40 | 43178.7188 | 40 | 45541.5586
41 | 1.4464 | 0.66384 | 0.020377 | 41 | 0.66031 | 0.039756 | 41 | 0.68086 | -0.070604 |  | 41 | 1.4464 | 0.66839 | -0.0043953 |  | 41 | 0.66035 | 0.039588 | 41 | 0.69273 | -0.13153 | 41 | 43843.2383 | 41 | 43211.3594 | 41 | 45518.6797
42 | 1.4817 | 0.6642 | 0.018373 | 42 | 0.661 | 0.035945 | 42 | 0.68095 | -0.071044 |  | 42 | 1.4817 | 0.66803 | -0.002423 |  | 42 | 0.66079 | 0.037116 | 42 | 0.69321 | -0.13396 | 42 | 43825.9219 | 42 | 43235.8008 | 42 | 45474.6016
43 | 1.5169 | 0.66415 | 0.018646 | 43 | 0.66149 | 0.033256 | 43 | 0.68084 | -0.0705 |  | 43 | 1.5169 | 0.66886 | -0.0069344 |  | 43 | 0.6605 | 0.038729 | 43 | 0.69304 | -0.13311 | 43 | 43826.6016 | 43 | 43225.0391 | 43 | 45463.7188
44 | 1.5522 | 0.66381 | 0.020522 | 44 | 0.6612 | 0.034865 | 44 | 0.68063 | -0.069378 |  | 44 | 1.5522 | 0.66932 | -0.0093917 |  | 44 | 0.66013 | 0.040781 | 44 | 0.69272 | -0.13146 | 44 | 43776.8008 | 44 | 43227.8789 | 44 | 45478.6797
45 | 1.5875 | 0.66324 | 0.02364 | 45 | 0.66069 | 0.037677 | 45 | 0.67988 | -0.065453 |  | 45 | 1.5875 | 0.66942 | -0.0099486 |  | 45 | 0.6596 | 0.043702 | 45 | 0.69247 | -0.13019 | 45 | 43723.6797 | 45 | 43232.4414 | 45 | 45423.1992
46 | 1.6228 | 0.66309 | 0.024452 | 46 | 0.66064 | 0.037981 | 46 | 0.67962 | -0.064113 |  | 46 | 1.6228 | 0.66947 | -0.010225 |  | 46 | 0.66025 | 0.040113 | 46 | 0.69218 | -0.12875 | 46 | 43735.5195 | 46 | 43224.5195 | 46 | 45377.3594
47 | 1.6581 | 0.66333 | 0.023142 | 47 | 0.66084 | 0.036878 | 47 | 0.67996 | -0.065914 |  | 47 | 1.6581 | 0.66884 | -0.0068186 |  | 47 | 0.6598 | 0.042623 | 47 | 0.69193 | -0.12748 | 47 | 43735.9609 | 47 | 43227.9219 | 47 | 45406.2812
48 | 1.6933 | 0.66291 | 0.025456 | 48 | 0.6603 | 0.039862 | 48 | 0.67996 | -0.065904 |  | 48 | 1.6933 | 0.66849 | -0.0049048 |  | 48 | 0.65975 | 0.042892 | 48 | 0.69205 | -0.12809 | 48 | 43712.4414 | 48 | 43218.6016 | 48 | 45411
49 | 1.7286 | 0.66296 | 0.025194 | 49 | 0.66055 | 0.038458 | 49 | 0.67976 | -0.064821 |  | 49 | 1.7286 | 0.66875 | -0.0063393 |  | 49 | 0.6601 | 0.04093 | 49 | 0.69255 | -0.13062 | 49 | 43751 | 49 | 43206.7188 | 49 | 45405
50 | 1.7639 | 0.66269 | 0.026676 | 50 | 0.66075 | 0.037376 | 50 | 0.67912 | -0.061488 |  | 50 | 1.7639 | 0.66933 | -0.0094479 |  | 50 | 0.66041 | 0.039208 | 50 | 0.69311 | -0.13346 | 50 | 43771.9219 | 50 | 43192.5195 | 50 | 45428.6016
51 | 1.7992 | 0.66283 | 0.025908 | 51 | 0.66116 | 0.03511 | 51 | 0.67926 | -0.062219 |  | 51 | 1.7992 | 0.66913 | -0.0083899 |  | 51 | 0.66027 | 0.03999 | 51 | 0.69293 | -0.13252 | 51 | 43795.7617 | 51 | 43228.5195 | 51 | 45412.0781
52 | 1.8344 | 0.6621 | 0.029929 | 52 | 0.66125 | 0.034616 | 52 | 0.67944 | -0.063171 |  | 52 | 1.8344 | 0.66901 | -0.0077277 |  | 52 | 0.66046 | 0.038974 | 52 | 0.69277 | -0.13172 | 52 | 43811.8398 | 52 | 43245.5586 | 52 | 45394.6016
53 | 1.8697 | 0.66207 | 0.030075 | 53 | 0.66168 | 0.032207 | 53 | 0.67943 | -0.063108 |  | 53 | 1.8697 | 0.66899 | -0.0076081 |  | 53 | 0.66054 | 0.038507 | 53 | 0.69216 | -0.12865 | 53 | 43885.0391 | 53 | 43283.9609 | 53 | 45369.9609
54 | 1.905 | 0.66203 | 0.03029 | 54 | 0.66214 | 0.029688 | 54 | 0.67889 | -0.060274 |  | 54 | 1.905 | 0.66903 | -0.0078251 |  | 54 | 0.66039 | 0.039316 | 54 | 0.69214 | -0.12856 | 54 | 43846.9609 | 54 | 43300.7617 | 54 | 45356.7188
55 | 1.9403 | 0.66191 | 0.030969 | 55 | 0.66164 | 0.03243 | 55 | 0.67895 | -0.060594 |  | 55 | 1.9403 | 0.66956 | -0.010703 |  | 55 | 0.661 | 0.035972 | 55 | 0.69277 | -0.13175 | 55 | 43838.6406 | 55 | 43260.1211 | 55 | 45354.3594
56 | 1.9756 | 0.66186 | 0.031221 | 56 | 0.66098 | 0.036073 | 56 | 0.67907 | -0.06121 |  | 56 | 1.9756 | 0.67015 | -0.013879 |  | 56 | 0.6614 | 0.033749 | 56 | 0.69191 | -0.12737 | 56 | 43810.4414 | 56 | 43210.9219 | 56 | 45366.3594
57 | 2.0108 | 0.66193 | 0.030874 | 57 | 0.66078 | 0.037164 | 57 | 0.67885 | -0.060082 |  | 57 | 2.0108 | 0.66875 | -0.0063319 |  | 57 | 0.66161 | 0.03263 | 57 | 0.69154 | -0.12552 | 57 | 43814.7188 | 57 | 43166.0781 | 57 | 45380.7188
58 | 2.0461 | 0.66266 | 0.026859 | 58 | 0.66197 | 0.030611 | 58 | 0.6793 | -0.062417 |  | 58 | 2.0461 | 0.6692 | -0.0087562 |  | 58 | 0.66232 | 0.028721 | 58 | 0.69221 | -0.12888 | 58 | 43844.8008 | 58 | 43183.1992 | 58 | 45347.2812
59 | 2.0814 | 0.66274 | 0.02637 | 59 | 0.66249 | 0.02777 | 59 | 0.68008 | -0.066536 |  | 59 | 2.0814 | 0.66934 | -0.009515 |  | 59 | 0.66183 | 0.031392 | 59 | 0.69229 | -0.12932 | 59 | 43838.0391 | 59 | 43129.2383 | 59 | 45331.4414
60 | 2.1167 | 0.66071 | 0.037596 | 60 | 0.66194 | 0.030768 | 60 | 0.67888 | -0.060211 |  | 60 | 2.1167 | 0.66828 | -0.0037843 |  | 60 | 0.6612 | 0.034858 | 60 | 0.69206 | -0.12813 | 60 | 43823.6406 | 60 | 43133.3203 | 60 | 45319.6016
61 | 2.1519 | 0.66273 | 0.026429 | 61 | 0.66163 | 0.032511 | 61 | 0.67952 | -0.063588 |  | 61 | 2.1519 | 0.66829 | -0.0038368 |  | 61 | 0.6597 | 0.043176 | 61 | 0.69195 | -0.12757 | 61 | 43797.1211 | 61 | 43195.9219 | 61 | 45316.8789
62 | 2.1872 | 0.66224 | 0.029152 | 62 | 0.66109 | 0.035484 | 62 | 0.67885 | -0.060054 |  | 62 | 2.1872 | 0.66876 | -0.0063505 |  | 62 | 0.65992 | 0.041912 | 62 | 0.69259 | -0.13081 | 62 | 43830.7617 | 62 | 43223.2812 | 62 | 45333.8398
63 | 2.2225 | 0.6622 | 0.029363 | 63 | 0.66266 | 0.026826 | 63 | 0.67843 | -0.057895 |  | 63 | 2.2225 | 0.66837 | -0.0042716 |  | 63 | 0.65789 | 0.053212 | 63 | 0.69263 | -0.13101 | 63 | 43791.1992 | 63 | 43226 | 63 | 45321.5195
64 | 2.2578 | 0.66242 | 0.028152 | 64 | 0.66217 | 0.029516 | 64 | 0.67821 | -0.056706 |  | 64 | 2.2578 | 0.66818 | -0.003252 |  | 64 | 0.65944 | 0.044625 | 64 | 0.69191 | -0.12739 | 64 | 43741.1992 | 64 | 43185.8008 | 64 | 45327.6406
65 | 2.2931 | 0.6624 | 0.028236 | 65 | 0.66263 | 0.027001 | 65 | 0.67913 | -0.061526 |  | 65 | 2.2931 | 0.66871 | -0.0061298 |  | 65 | 0.65885 | 0.04786 | 65 | 0.69152 | -0.12539 | 65 | 43773.8398 | 65 | 43158.8398 | 65 | 45344.5586
66 | 2.3283 | 0.66263 | 0.027012 | 66 | 0.66271 | 0.026585 | 66 | 0.67896 | -0.06066 |  | 66 | 2.3283 | 0.66952 | -0.010472 |  | 66 | 0.6588 | 0.048134 | 66 | 0.6904 | -0.11973 | 66 | 43775.6016 | 66 | 43147.6797 | 66 | 45339.3594
67 | 2.3636 | 0.66267 | 0.026793 | 67 | 0.66251 | 0.027678 | 67 | 0.67834 | -0.05741 |  | 67 | 2.3636 | 0.66888 | -0.0070357 |  | 67 | 0.65901 | 0.046967 | 67 | 0.68937 | -0.11447 | 67 | 43794.6016 | 67 | 43187.4414 | 67 | 45346.6016
68 | 2.3989 | 0.66283 | 0.02589 | 68 | 0.66235 | 0.028561 | 68 | 0.67772 | -0.05412 |  | 68 | 2.3989 | 0.66941 | -0.009885 |  | 68 | 0.65887 | 0.047741 | 68 | 0.6889 | -0.11206 | 68 | 43792.1992 | 68 | 43165.4805 | 68 | 45317.8789
69 | 2.4342 | 0.66287 | 0.025704 | 69 | 0.66248 | 0.027824 | 69 | 0.67777 | -0.054424 |  | 69 | 2.4342 | 0.66964 | -0.011099 |  | 69 | 0.66002 | 0.041387 | 69 | 0.68947 | -0.11501 | 69 | 43754.4805 | 69 | 43158.0781 | 69 | 45316.4805
70 | 2.4694 | 0.66239 | 0.028342 | 70 | 0.66272 | 0.02652 | 70 | 0.67688 | -0.04969 |  | 70 | 2.4694 | 0.66966 | -0.011215 |  | 70 | 0.66168 | 0.032206 | 70 | 0.68999 | -0.11765 | 70 | 43745.0391 | 70 | 43199.3594 | 70 | 45332.1602
71 | 2.5047 | 0.66211 | 0.029834 | 71 | 0.66292 | 0.025416 | 71 | 0.67661 | -0.048261 |  | 71 | 2.5047 | 0.66953 | -0.010509 |  | 71 | 0.66293 | 0.025358 | 71 | 0.69041 | -0.11978 | 71 | 43788.1211 | 71 | 43239.5586 | 71 | 45308.0781
72 | 2.54 | 0.66219 | 0.029396 | 72 | 0.66322 | 0.023757 | 72 | 0.6769 | -0.04982 |  | 72 | 2.54 | 0.66975 | -0.011712 |  | 72 | 0.66322 | 0.023771 | 72 | 0.69109 | -0.12325 | 72 | 43841.3594 | 72 | 43235.1211 | 72 | 45301.1992
73 | 2.5753 | 0.66236 | 0.028502 | 73 | 0.66305 | 0.024677 | 73 | 0.67655 | -0.047977 |  | 73 | 2.5753 | 0.67068 | -0.016711 |  | 73 | 0.66369 | 0.021206 | 73 | 0.69141 | -0.12487 | 73 | 43838.6797 | 73 | 43258.4805 | 73 | 45303.0781
74 | 2.6106 | 0.66218 | 0.029447 | 74 | 0.66326 | 0.023531 | 74 | 0.67657 | -0.048062 |  | 74 | 2.6106 | 0.67141 | -0.020624 |  | 74 | 0.66404 | 0.019263 | 74 | 0.69136 | -0.12458 | 74 | 43824.0781 | 74 | 43269.3203 | 74 | 45309.3203
75 | 2.6458 | 0.66226 | 0.029053 | 75 | 0.66423 | 0.018239 | 75 | 0.67586 | -0.044313 |  | 75 | 2.6458 | 0.67211 | -0.024363 |  | 75 | 0.66393 | 0.01987 | 75 | 0.69096 | -0.12258 | 75 | 43803.9609 | 75 | 43241.2383 | 75 | 45320.3594
76 | 2.6811 | 0.66265 | 0.026869 | 76 | 0.66439 | 0.017355 | 76 | 0.67559 | -0.042915 |  | 76 | 2.6811 | 0.67117 | -0.01935 |  | 76 | 0.66343 | 0.022585 | 76 | 0.69099 | -0.12271 | 76 | 43802.0391 | 76 | 43243.0781 | 76 | 45322.1602
77 | 2.7164 | 0.66244 | 0.028054 | 77 | 0.66486 | 0.014767 | 77 | 0.67686 | -0.049596 |  | 77 | 2.7164 | 0.67145 | -0.020818 |  | 77 | 0.66256 | 0.027372 | 77 | 0.69107 | -0.12312 | 77 | 43819.9609 | 77 | 43260.6016 | 77 | 45328.0781
78 | 2.7517 | 0.66261 | 0.027085 | 78 | 0.6641 | 0.018969 | 78 | 0.6757 | -0.043463 |  | 78 | 2.7517 | 0.67268 | -0.027402 |  | 78 | 0.66226 | 0.029011 | 78 | 0.6912 | -0.12381 | 78 | 43871.6406 | 78 | 43293.6406 | 78 | 45328.0391
79 | 2.7869 | 0.66232 | 0.028681 | 79 | 0.66352 | 0.022139 | 79 | 0.67471 | -0.038234 |  | 79 | 2.7869 | 0.67113 | -0.019149 |  | 79 | 0.66184 | 0.031358 | 79 | 0.69112 | -0.12339 | 79 | 43818.5195 | 79 | 43302.0781 | 79 | 45304.4805
80 | 2.8222 | 0.66089 | 0.036567 | 80 | 0.66397 | 0.019669 | 80 | 0.67452 | -0.037216 |  | 80 | 2.8222 | 0.67137 | -0.020431 |  | 80 | 0.66179 | 0.03161 | 80 | 0.69181 | -0.12686 | 80 | 43807.0781 | 80 | 43287.7188 | 80 | 45261.3984
81 | 2.8575 | 0.66163 | 0.032489 | 81 | 0.66417 | 0.01858 | 81 | 0.67533 | -0.041516 |  | 81 | 2.8575 | 0.66913 | -0.0083526 |  | 81 | 0.6617 | 0.03211 | 81 | 0.69104 | -0.12297 | 81 | 43822.8398 | 81 | 43268 | 81 | 45240.8398
82 | 2.8928 | 0.66321 | 0.023804 | 82 | 0.66549 | 0.011342 | 82 | 0.67687 | -0.049645 |  | 82 | 2.8928 | 0.67038 | -0.01511 |  | 82 | 0.66179 | 0.031621 | 82 | 0.69116 | -0.12359 | 82 | 43783.8008 | 82 | 43215.9609 | 82 | 45149.5586
83 | 2.9281 | 0.66298 | 0.025059 | 83 | 0.66576 | 0.0098822 | 83 | 0.67727 | -0.051779 |  | 83 | 2.9281 | 0.66994 | -0.012716 |  | 83 | 0.66165 | 0.032404 | 83 | 0.69087 | -0.12213 | 83 | 43801.3984 | 83 | 43191.4414 | 83 | 45158.2383
84 | 2.9633 | 0.66138 | 0.033889 | 84 | 0.66373 | 0.020987 | 84 | 0.67519 | -0.040764 |  | 84 | 2.9633 | 0.66979 | -0.011943 |  | 84 | 0.66163 | 0.032477 | 84 | 0.69127 | -0.12412 | 84 | 43792.8398 | 84 | 43132.3203 | 84 | 45145.6406
85 | 2.9986 | 0.66161 | 0.032635 | 85 | 0.66348 | 0.022357 | 85 | 0.6741 | -0.034984 |  | 85 | 2.9986 | 0.66935 | -0.0095452 |  | 85 | 0.66107 | 0.035566 | 85 | 0.69051 | -0.12028 | 85 | 43749.2383 | 85 | 43163.3984 | 85 | 45177.3594
86 | 3.0339 | 0.66131 | 0.034269 | 86 | 0.66378 | 0.020718 | 86 | 0.67436 | -0.036356 |  | 86 | 3.0339 | 0.66998 | -0.012977 |  | 86 | 0.66125 | 0.034583 | 86 | 0.69034 | -0.11943 | 86 | 43750.8398 | 86 | 43163.8789 | 86 | 45204.3594
87 | 3.0692 | 0.66172 | 0.032028 | 87 | 0.6645 | 0.016753 | 87 | 0.67441 | -0.03662 |  | 87 | 3.0692 | 0.67006 | -0.013406 |  | 87 | 0.66097 | 0.036128 | 87 | 0.68933 | -0.11429 | 87 | 43804.3594 | 87 | 43121.6406 | 87 | 45236.5195
88 | 3.1044 | 0.66157 | 0.032825 | 88 | 0.66452 | 0.01664 | 88 | 0.67444 | -0.036807 |  | 88 | 3.1044 | 0.66946 | -0.010177 |  | 88 | 0.66076 | 0.0373 | 88 | 0.68962 | -0.11575 | 88 | 43854.8398 | 88 | 43150.6797 | 88 | 45265.5195
89 | 3.1397 | 0.66094 | 0.036303 | 89 | 0.66373 | 0.020954 | 89 | 0.67298 | -0.028999 |  | 89 | 3.1397 | 0.66902 | -0.0077653 |  | 89 | 0.66022 | 0.040274 | 89 | 0.68955 | -0.1154 | 89 | 43801.0781 | 89 | 43165.8398 | 89 | 45270.5195
90 | 3.175 | 0.6616 | 0.032682 | 90 | 0.66457 | 0.01635 | 90 | 0.6737 | -0.032863 |  | 90 | 3.175 | 0.66783 | -0.0013572 |  | 90 | 0.65969 | 0.043188 | 90 | 0.68985 | -0.11695 | 90 | 43819.8789 | 90 | 43230.6406 | 90 | 45250.7617
91 | 3.2103 | 0.66148 | 0.033304 | 91 | 0.66447 | 0.016923 | 91 | 0.67363 | -0.032478 |  | 91 | 3.2103 | 0.66703 | 0.0030019 |  | 91 | 0.65995 | 0.041763 | 91 | 0.68889 | -0.11203 | 91 | 43829.3984 | 91 | 43306.2812 | 91 | 45252.0781
92 | 3.2456 | 0.66133 | 0.034174 | 92 | 0.66373 | 0.020947 | 92 | 0.67238 | -0.025793 |  | 92 | 3.2456 | 0.66727 | 0.0016816 |  | 92 | 0.66002 | 0.041383 | 92 | 0.68875 | -0.11132 | 92 | 43835.9609 | 92 | 43347.2383 | 92 | 45220.5586
93 | 3.2808 | 0.66119 | 0.034902 | 93 | 0.66339 | 0.022807 | 93 | 0.67202 | -0.023914 |  | 93 | 3.2808 | 0.66718 | 0.0021666 |  | 93 | 0.66051 | 0.038691 | 93 | 0.68871 | -0.11113 | 93 | 43813.1992 | 93 | 43389.4414 | 93 | 45211.9219
94 | 3.3161 | 0.66149 | 0.033267 | 94 | 0.66326 | 0.023564 | 94 | 0.67207 | -0.024159 |  | 94 | 3.3161 | 0.66703 | 0.0030054 |  | 94 | 0.66049 | 0.038783 | 94 | 0.68841 | -0.10957 | 94 | 43829.7188 | 94 | 43417.5195 | 94 | 45238.1602
95 | 3.3514 | 0.66218 | 0.029462 | 95 | 0.66354 | 0.022015 | 95 | 0.67224 | -0.025081 |  | 95 | 3.3514 | 0.66656 | 0.0055166 |  | 95 | 0.66056 | 0.0384 | 95 | 0.68786 | -0.10676 | 95 | 43901.7617 | 95 | 43432.8789 | 95 | 45278.8008
96 | 3.3867 | 0.66173 | 0.031944 | 96 | 0.66396 | 0.019687 | 96 | 0.67233 | -0.025563 |  | 96 | 3.3867 | 0.66657 | 0.0054865 |  | 96 | 0.66073 | 0.037438 | 96 | 0.68742 | -0.10453 | 96 | 43982.1602 | 96 | 43422.5195 | 96 | 45290.3984
97 | 3.4219 | 0.66087 | 0.036714 | 97 | 0.66377 | 0.020743 | 97 | 0.6715 | -0.021126 |  | 97 | 3.4219 | 0.66683 | 0.0040594 |  | 97 | 0.66068 | 0.037717 | 97 | 0.68726 | -0.10371 | 97 | 44009.4805 | 97 | 43383.7188 | 97 | 45263.7617
98 | 3.4572 | 0.6607 | 0.037611 | 98 | 0.66332 | 0.023204 | 98 | 0.67182 | -0.022835 |  | 98 | 3.4572 | 0.66744 | 0.00077591 |  | 98 | 0.66055 | 0.038434 | 98 | 0.68726 | -0.10368 | 98 | 43972.0781 | 98 | 43334.2812 | 98 | 45224.7188
99 | 3.4925 | 0.66078 | 0.037208 | 99 | 0.66282 | 0.025944 | 99 | 0.6719 | -0.023247 |  | 99 | 3.4925 | 0.66818 | -0.0032594 |  | 99 | 0.66063 | 0.038024 | 99 | 0.6873 | -0.10389 | 99 | 43949.9219 | 99 | 43302.4805 | 99 | 45269.8008
100 | 3.5278 | 0.66073 | 0.037486 | 100 | 0.663 | 0.024947 | 100 | 0.67188 | -0.023147 |  | 100 | 3.5278 | 0.66745 | 0.00068194 |  | 100 | 0.65995 | 0.041771 | 100 | 0.68759 | -0.10536 | 100 | 43980.3203 | 100 | 43335.8008 | 100 | 45220.1602
101 | 3.5631 | 0.66058 | 0.038271 | 101 | 0.66304 | 0.024761 | 101 | 0.6718 | -0.022732 |  | 101 | 3.5631 | 0.66761 | -0.00017457 |  | 101 | 0.66035 | 0.039584 | 101 | 0.68784 | -0.10667 | 101 | 43919.1211 | 101 | 43348.0781 | 101 | 45159.4414
102 | 3.5983 | 0.6598 | 0.042607 | 102 | 0.66278 | 0.026181 | 102 | 0.67141 | -0.020636 |  | 102 | 3.5983 | 0.66921 | -0.0087973 |  | 102 | 0.66065 | 0.037882 | 102 | 0.68745 | -0.10469 | 102 | 43855.5586 | 102 | 43308.5586 | 102 | 45155.6406
103 | 3.6336 | 0.65999 | 0.041554 | 103 | 0.66223 | 0.029173 | 103 | 0.67116 | -0.019271 |  | 103 | 3.6336 | 0.66929 | -0.0092536 |  | 103 | 0.66029 | 0.039895 | 103 | 0.68717 | -0.10323 | 103 | 43861.5586 | 103 | 43314.1992 | 103 | 45129.6016
104 | 3.6689 | 0.66034 | 0.039594 | 104 | 0.66283 | 0.0259 | 104 | 0.67171 | -0.022217 |  | 104 | 3.6689 | 0.66887 | -0.0069755 |  | 104 | 0.66081 | 0.037013 | 104 | 0.68784 | -0.10666 | 104 | 43862.8008 | 104 | 43315.9219 | 104 | 45141.1211
105 | 3.7042 | 0.65987 | 0.042226 | 105 | 0.66243 | 0.028083 | 105 | 0.67078 | -0.017228 |  | 105 | 3.7042 | 0.66896 | -0.0074361 |  | 105 | 0.66063 | 0.038016 | 105 | 0.68803 | -0.10763 | 105 | 43841.9609 | 105 | 43285.6406 | 105 | 45139.1992
106 | 3.7394 | 0.66001 | 0.041443 | 106 | 0.66263 | 0.027022 | 106 | 0.67039 | -0.01514 |  | 106 | 3.7394 | 0.66898 | -0.0075708 |  | 106 | 0.66123 | 0.034709 | 106 | 0.68821 | -0.10855 | 106 | 43854.2812 | 106 | 43257.6797 | 106 | 45153.7617
107 | 3.7747 | 0.66006 | 0.041175 | 107 | 0.66309 | 0.024452 | 107 | 0.67203 | -0.023949 |  | 107 | 3.7747 | 0.66873 | -0.0062157 |  | 107 | 0.661 | 0.035968 | 107 | 0.68797 | -0.10733 | 107 | 43791.7188 | 107 | 43269.7188 | 107 | 45113.7617
108 | 3.81 | 0.66077 | 0.037252 | 108 | 0.66343 | 0.022619 | 108 | 0.67206 | -0.02413 |  | 108 | 3.81 | 0.66793 | -0.0018751 |  | 108 | 0.66068 | 0.037752 | 108 | 0.68767 | -0.1058 | 108 | 43859.7617 | 108 | 43242.7617 | 108 | 45114.8008
109 | 3.8453 | 0.65932 | 0.045287 | 109 | 0.66187 | 0.03117 | 109 | 0.67105 | -0.018691 |  | 109 | 3.8453 | 0.66701 | 0.0030846 |  | 109 | 0.66057 | 0.038342 | 109 | 0.68791 | -0.10703 | 109 | 43849.5195 | 109 | 43248.2812 | 109 | 45090.8789
110 | 3.8806 | 0.66015 | 0.040643 | 110 | 0.66276 | 0.026298 | 110 | 0.67122 | -0.019608 |  | 110 | 3.8806 | 0.66758 | -1.6724e-05 |  | 110 | 0.65989 | 0.042112 | 110 | 0.68858 | -0.11047 | 110 | 43813.7188 | 110 | 43206.2383 | 110 | 45037.0781
111 | 3.9158 | 0.66017 | 0.040581 | 111 | 0.66302 | 0.024841 | 111 | 0.67139 | -0.020515 |  | 111 | 3.9158 | 0.66752 | 0.00034743 |  | 111 | 0.66053 | 0.038564 | 111 | 0.68849 | -0.10997 | 111 | 43787.4414 | 111 | 43211.3203 | 111 | 45023.0781
112 | 3.9511 | 0.66075 | 0.037333 | 112 | 0.66404 | 0.019273 | 112 | 0.6721 | -0.024304 |  | 112 | 3.9511 | 0.66643 | 0.0062587 |  | 112 | 0.66081 | 0.037013 | 112 | 0.68799 | -0.10742 | 112 | 43716.1992 | 112 | 43177.4414 | 112 | 45020.8398
113 | 3.9864 | 0.66083 | 0.036915 | 113 | 0.66403 | 0.019331 | 113 | 0.67218 | -0.024761 |  | 113 | 3.9864 | 0.66717 | 0.0022232 |  | 113 | 0.66025 | 0.040113 | 113 | 0.68689 | -0.1018 | 113 | 43676.7188 | 113 | 43195.6016 | 113 | 44998.0781
114 | 4.0217 | 0.66034 | 0.039627 | 114 | 0.66363 | 0.021539 | 114 | 0.67121 | -0.019548 |  | 114 | 4.0217 | 0.66714 | 0.0023773 |  | 114 | 0.65971 | 0.043118 | 114 | 0.68673 | -0.10098 | 114 | 43639.0391 | 114 | 43207.8789 | 114 | 44992
115 | 4.0569 | 0.66021 | 0.040324 | 115 | 0.66332 | 0.023222 | 115 | 0.67131 | -0.020099 |  | 115 | 4.0569 | 0.6668 | 0.0042325 |  | 115 | 0.65904 | 0.046844 | 115 | 0.68726 | -0.10369 | 115 | 43631.7188 | 115 | 43204.3203 | 115 | 45000.7617
116 | 4.0922 | 0.66096 | 0.036172 | 116 | 0.6641 | 0.018928 | 116 | 0.67196 | -0.023573 |  | 116 | 4.0922 | 0.66677 | 0.0043755 |  | 116 | 0.65866 | 0.048904 | 116 | 0.68689 | -0.10182 | 116 | 43663.6797 | 116 | 43207.7188 | 116 | 45036.4414
117 | 4.1275 | 0.66159 | 0.032741 | 117 | 0.66381 | 0.020547 | 117 | 0.67193 | -0.023392 |  | 117 | 4.1275 | 0.66641 | 0.0063416 |  | 117 | 0.65888 | 0.047687 | 117 | 0.68706 | -0.10266 | 117 | 43693.6797 | 117 | 43214.9219 | 117 | 45025.6406
118 | 4.1628 | 0.66198 | 0.030589 | 118 | 0.66354 | 0.022001 | 118 | 0.67188 | -0.023122 |  | 118 | 4.1628 | 0.66691 | 0.0036075 |  | 118 | 0.65926 | 0.045621 | 118 | 0.68675 | -0.10105 | 118 | 43725 | 118 | 43214.3203 | 118 | 44988.5586
119 | 4.1981 | 0.66213 | 0.029772 | 119 | 0.66372 | 0.021002 | 119 | 0.67182 | -0.022803 |  | 119 | 4.1981 | 0.66842 | -0.0045303 |  | 119 | 0.65825 | 0.051221 | 119 | 0.68613 | -0.097883 | 119 | 43694.1602 | 119 | 43245.6797 | 119 | 45022.9219
120 | 4.2333 | 0.66149 | 0.033256 | 120 | 0.66391 | 0.019963 | 120 | 0.67124 | -0.01974 |  | 120 | 4.2333 | 0.66909 | -0.0081356 |  | 120 | 0.6582 | 0.051491 | 120 | 0.68565 | -0.095427 | 120 | 43651.7617 | 120 | 43236.1211 | 120 | 45013.0391
121 | 4.2686 | 0.66171 | 0.032036 | 121 | 0.66372 | 0.02102 | 121 | 0.67117 | -0.019321 |  | 121 | 4.2686 | 0.6692 | -0.0087601 |  | 121 | 0.65945 | 0.044555 | 121 | 0.68532 | -0.093737 | 121 | 43682.6797 | 121 | 43201.8008 | 121 | 45031.9609
122 | 4.3039 | 0.6608 | 0.037094 | 122 | 0.66319 | 0.02391 | 122 | 0.67049 | -0.015678 |  | 122 | 4.3039 | 0.67006 | -0.013365 |  | 122 | 0.66062 | 0.038055 | 122 | 0.68563 | -0.095317 | 122 | 43799.8789 | 122 | 43229.4805 | 122 | 45069.4414
123 | 4.3392 | 0.66093 | 0.036336 | 123 | 0.66388 | 0.020155 | 123 | 0.67187 | -0.023087 |  | 123 | 4.3392 | 0.66855 | -0.0052457 |  | 123 | 0.6615 | 0.033241 | 123 | 0.68572 | -0.095802 | 123 | 43762.1211 | 123 | 43236.3203 | 123 | 45036.1992
124 | 4.3744 | 0.66225 | 0.029108 | 124 | 0.66484 | 0.014894 | 124 | 0.67213 | -0.024457 |  | 124 | 4.3744 | 0.66957 | -0.010756 |  | 124 | 0.66182 | 0.031468 | 124 | 0.68612 | -0.097855 | 124 | 43759 | 124 | 43216.8398 | 124 | 45053.4805
125 | 4.4097 | 0.66179 | 0.031626 | 125 | 0.66521 | 0.012893 | 125 | 0.6719 | -0.023243 |  | 125 | 4.4097 | 0.66859 | -0.0054556 |  | 125 | 0.66156 | 0.032897 | 125 | 0.68588 | -0.096613 | 125 | 43800.6797 | 125 | 43236 | 125 | 45079.8398
126 | 4.445 | 0.6608 | 0.037076 | 126 | 0.66507 | 0.013671 | 126 | 0.67113 | -0.019146 |  | 126 | 4.445 | 0.66878 | -0.0064815 |  | 126 | 0.66148 | 0.03331 | 126 | 0.68588 | -0.096623 | 126 | 43812.6406 | 126 | 43246.9219 | 126 | 45080.5586
127 | 4.4803 | 0.66158 | 0.032759 | 127 | 0.66488 | 0.01467 | 127 | 0.67231 | -0.025432 |  | 127 | 4.4803 | 0.66981 | -0.012059 |  | 127 | 0.66152 | 0.033084 | 127 | 0.68642 | -0.099395 | 127 | 43772.6016 | 127 | 43260.5586 | 127 | 45015.2383
128 | 4.5156 | 0.66224 | 0.029166 | 128 | 0.66555 | 0.01102 | 128 | 0.67262 | -0.027105 |  | 128 | 4.5156 | 0.67 | -0.013037 |  | 128 | 0.66193 | 0.030865 | 128 | 0.68661 | -0.10035 | 128 | 43817.0391 | 128 | 43251.9219 | 128 | 44955.1211
129 | 4.5508 | 0.6617 | 0.03209 | 129 | 0.66418 | 0.018529 | 129 | 0.67217 | -0.024698 |  | 129 | 4.5508 | 0.66953 | -0.010546 |  | 129 | 0.66188 | 0.031102 | 129 | 0.6862 | -0.098229 | 129 | 43804.0391 | 129 | 43228.3594 | 129 | 44982.5586
130 | 4.5861 | 0.66316 | 0.024066 | 130 | 0.66594 | 0.0089038 | 130 | 0.67406 | -0.034779 |  | 130 | 4.5861 | 0.67042 | -0.015337 |  | 130 | 0.66211 | 0.029851 | 130 | 0.68586 | -0.096502 | 130 | 43765 | 130 | 43247.2812 | 130 | 45003.4414
131 | 4.6214 | 0.66189 | 0.031068 | 131 | 0.66573 | 0.010059 | 131 | 0.67236 | -0.02573 |  | 131 | 4.6214 | 0.66905 | -0.0079186 |  | 131 | 0.66302 | 0.024867 | 131 | 0.68597 | -0.097055 | 131 | 43726.3594 | 131 | 43226.2812 | 131 | 44969.5586
132 | 4.6567 | 0.6623 | 0.028823 | 132 | 0.66565 | 0.010493 | 132 | 0.67276 | -0.027866 |  | 132 | 4.6567 | 0.66807 | -0.002667 |  | 132 | 0.66227 | 0.02895 | 132 | 0.68631 | -0.098817 | 132 | 43711.6016 | 132 | 43243.1602 | 132 | 44978.1211
133 | 4.6919 | 0.66236 | 0.028459 | 133 | 0.66594 | 0.0089291 | 133 | 0.67213 | -0.024506 |  | 133 | 4.6919 | 0.668 | -0.0022805 |  | 133 | 0.66216 | 0.029602 | 133 | 0.68624 | -0.098464 | 133 | 43665.2812 | 133 | 43253.0781 | 133 | 44983.6016
134 | 4.7272 | 0.66299 | 0.025001 | 134 | 0.6659 | 0.0091383 | 134 | 0.67231 | -0.02546 |  | 134 | 4.7272 | 0.66767 | -0.00050893 |  | 134 | 0.66159 | 0.032714 | 134 | 0.68601 | -0.097284 | 134 | 43642.6406 | 134 | 43206.0781 | 134 | 44939.2812
135 | 4.7625 | 0.66345 | 0.022528 | 135 | 0.66572 | 0.01011 | 135 | 0.67275 | -0.027781 |  | 135 | 4.7625 | 0.66728 | 0.001648 |  | 135 | 0.66072 | 0.037518 | 135 | 0.68592 | -0.096838 | 135 | 43652.6016 | 135 | 43218.8789 | 135 | 44902.8789
136 | 4.7978 | 0.66258 | 0.027281 | 136 | 0.66594 | 0.0089181 | 136 | 0.67235 | -0.025648 |  | 136 | 4.7978 | 0.66849 | -0.0049125 |  | 136 | 0.66035 | 0.039538 | 136 | 0.68556 | -0.094945 | 136 | 43649.2812 | 136 | 43161.7188 | 136 | 44884.9609
137 | 4.8331 | 0.66305 | 0.024692 | 137 | 0.6658 | 0.0096801 | 137 | 0.67262 | -0.027087 |  | 137 | 4.8331 | 0.66799 | -0.0022318 |  | 137 | 0.66077 | 0.037254 | 137 | 0.68526 | -0.093426 | 137 | 43641.8398 | 137 | 43142.4805 | 137 | 44902.6016
138 | 4.8683 | 0.66239 | 0.028309 | 138 | 0.66604 | 0.0083876 | 138 | 0.67202 | -0.023875 |  | 138 | 4.8683 | 0.66718 | 0.0021779 |  | 138 | 0.66088 | 0.036661 | 138 | 0.6854 | -0.094123 | 138 | 43622.6406 | 138 | 43095.8789 | 138 | 44918.8398
139 | 4.9036 | 0.66212 | 0.029827 | 139 | 0.66692 | 0.0036014 | 139 | 0.6721 | -0.024343 |  | 139 | 4.9036 | 0.66659 | 0.0053657 |  | 139 | 0.66019 | 0.040439 | 139 | 0.68506 | -0.092407 | 139 | 43663.0781 | 139 | 43108.0391 | 139 | 44913.4414
140 | 4.9389 | 0.66213 | 0.029739 | 140 | 0.66682 | 0.0041487 | 140 | 0.67201 | -0.023836 |  | 140 | 4.9389 | 0.66674 | 0.0045337 |  | 140 | 0.66055 | 0.038461 | 140 | 0.68568 | -0.095595 | 140 | 43693.3594 | 140 | 43079.8008 | 140 | 44906.8398
141 | 4.9742 | 0.66121 | 0.034796 | 141 | 0.66534 | 0.012177 | 141 | 0.67119 | -0.019441 |  | 141 | 4.9742 | 0.66706 | 0.002795 |  | 141 | 0.66036 | 0.039492 | 141 | 0.68619 | -0.098215 | 141 | 43666.1602 | 141 | 43063.6406 | 141 | 44916.5586
142 | 5.0094 | 0.66205 | 0.03017 | 142 | 0.66683 | 0.0040623 | 142 | 0.67235 | -0.025652 |  | 142 | 5.0094 | 0.66656 | 0.0055467 |  | 142 | 0.66068 | 0.03776 | 142 | 0.68603 | -0.097355 | 142 | 43751.9219 | 142 | 43144.6016 | 142 | 44914.1992
143 | 5.0447 | 0.66265 | 0.02688 | 143 | 0.66735 | 0.0012378 | 143 | 0.67248 | -0.026357 |  | 143 | 5.0447 | 0.66707 | 0.0027458 |  | 143 | 0.66102 | 0.035883 | 143 | 0.68603 | -0.097394 | 143 | 43767.6406 | 143 | 43213.7188 | 143 | 44930.3984
144 | 5.08 | 0.66253 | 0.027522 | 144 | 0.66719 | 0.0021224 | 144 | 0.67186 | -0.023037 |  | 144 | 5.08 | 0.6674 | 0.00096742 |  | 144 | 0.6605 | 0.038729 | 144 | 0.68536 | -0.093941 | 144 | 43731.7188 | 144 | 43259.1211 | 144 | 44935.5195
145 | 5.1153 | 0.66225 | 0.029075 | 145 | 0.66642 | 0.006296 | 145 | 0.6707 | -0.016811 |  | 145 | 5.1153 | 0.66813 | -0.0029819 |  | 145 | 0.65968 | 0.043268 | 145 | 0.68473 | -0.090703 | 145 | 43864.6016 | 145 | 43264 | 145 | 44945.4805
146 | 5.1506 | 0.66248 | 0.02785 | 146 | 0.66706 | 0.0028418 | 146 | 0.67062 | -0.016402 |  | 146 | 5.1506 | 0.66708 | 0.0027271 |  | 146 | 0.65975 | 0.042857 | 146 | 0.68452 | -0.089589 | 146 | 43820.3203 | 146 | 43303.3203 | 146 | 44950.6406
147 | 5.1858 | 0.6614 | 0.033782 | 147 | 0.66605 | 0.0083228 | 147 | 0.67014 | -0.01381 |  | 147 | 5.1858 | 0.66603 | 0.0084343 |  | 147 | 0.65922 | 0.045798 | 147 | 0.68389 | -0.086362 | 147 | 43865.6016 | 147 | 43307.1992 | 147 | 44919.2812
148 | 5.2211 | 0.6622 | 0.029389 | 148 | 0.66696 | 0.0033495 | 148 | 0.67075 | -0.017082 |  | 148 | 5.2211 | 0.6664 | 0.0064171 |  | 148 | 0.65957 | 0.043864 | 148 | 0.68344 | -0.084006 | 148 | 43793.9609 | 148 | 43286.7617 | 148 | 44920.1992
149 | 5.2564 | 0.66103 | 0.035824 | 149 | 0.66578 | 0.0097884 | 149 | 0.66972 | -0.011576 |  | 149 | 5.2564 | 0.66673 | 0.00459 |  | 149 | 0.66082 | 0.036952 | 149 | 0.68347 | -0.084164 | 149 | 43825.9219 | 149 | 43296.3594 | 149 | 44922.3203
150 | 5.2917 | 0.66192 | 0.030925 | 150 | 0.66625 | 0.0072082 | 150 | 0.67039 | -0.01514 |  | 150 | 5.2917 | 0.6671 | 0.0026256 |  | 150 | 0.66053 | 0.038576 | 150 | 0.68348 | -0.084222 | 150 | 43827.0391 | 150 | 43357.2383 | 150 | 44907.9219
151 | 5.3269 | 0.6624 | 0.028265 | 151 | 0.6661 | 0.0080521 | 151 | 0.67094 | -0.018129 |  | 151 | 5.3269 | 0.66708 | 0.002697 |  | 151 | 0.65983 | 0.042446 | 151 | 0.68365 | -0.085077 | 151 | 43806.5586 | 151 | 43362.5195 | 151 | 44895.0781
152 | 5.3622 | 0.66224 | 0.02913 | 152 | 0.66621 | 0.0074531 | 152 | 0.67119 | -0.019438 |  | 152 | 5.3622 | 0.66573 | 0.01005 |  | 152 | 0.65921 | 0.045855 | 152 | 0.68322 | -0.082842 | 152 | 43748.9609 | 152 | 43330.3203 | 152 | 44842.1602
153 | 5.3975 | 0.6618 | 0.031557 | 153 | 0.66605 | 0.0083335 | 153 | 0.66986 | -0.012329 |  | 153 | 5.3975 | 0.66665 | 0.0050194 |  | 153 | 0.6591 | 0.046459 | 153 | 0.68323 | -0.082917 | 153 | 43787.3203 | 153 | 43314.2383 | 153 | 44858.5195
154 | 5.4328 | 0.66066 | 0.037868 | 154 | 0.66679 | 0.0042674 | 154 | 0.67033 | -0.014841 |  | 154 | 5.4328 | 0.66699 | 0.0032124 |  | 154 | 0.65919 | 0.04597 | 154 | 0.68405 | -0.087187 | 154 | 43774.4805 | 154 | 43290.6016 | 154 | 44859.9219
155 | 5.4681 | 0.66211 | 0.029881 | 155 | 0.66671 | 0.0047103 | 155 | 0.67048 | -0.015632 |  | 155 | 5.4681 | 0.66782 | -0.0012784 |  | 155 | 0.65903 | 0.046886 | 155 | 0.68359 | -0.084779 | 155 | 43690.0781 | 155 | 43283.2812 | 155 | 44859.5586
156 | 5.5033 | 0.66247 | 0.027861 | 156 | 0.66669 | 0.0048184 | 156 | 0.67065 | -0.016533 |  | 156 | 5.5033 | 0.66743 | 0.00083242 |  | 156 | 0.65943 | 0.044644 | 156 | 0.68322 | -0.08287 | 156 | 43677.4414 | 156 | 43251.7188 | 156 | 44903.0391
157 | 5.5386 | 0.66196 | 0.03068 | 157 | 0.66593 | 0.0089833 | 157 | 0.67058 | -0.016163 |  | 157 | 5.5386 | 0.66616 | 0.0076839 |  | 157 | 0.65926 | 0.045617 | 157 | 0.68384 | -0.086093 | 157 | 43695.1211 | 157 | 43210.1992 | 157 | 44884.6016
158 | 5.5739 | 0.66152 | 0.033128 | 158 | 0.66663 | 0.0051498 | 158 | 0.67047 | -0.0156 |  | 158 | 5.5739 | 0.66609 | 0.0080874 |  | 158 | 0.65913 | 0.046332 | 158 | 0.68447 | -0.089335 | 158 | 43709.9219 | 158 | 43226.9219 | 158 | 44874.1992
159 | 5.6092 | 0.66105 | 0.035721 | 159 | 0.66693 | 0.0035474 | 159 | 0.67079 | -0.017278 |  | 159 | 5.6092 | 0.66622 | 0.007401 |  | 159 | 0.65823 | 0.051337 | 159 | 0.68375 | -0.085598 | 159 | 43692.9609 | 159 | 43250.3594 | 159 | 44854.3203
160 | 5.6444 | 0.662 | 0.030469 | 160 | 0.66722 | 0.001975 | 160 | 0.67106 | -0.018726 |  | 160 | 5.6444 | 0.66665 | 0.0050194 |  | 160 | 0.65871 | 0.048631 | 160 | 0.6835 | -0.084333 | 160 | 43695.9219 | 160 | 43208 | 160 | 44832.9609
161 | 5.6797 | 0.66201 | 0.030392 | 161 | 0.66703 | 0.0030074 | 161 | 0.67094 | -0.018118 |  | 161 | 5.6797 | 0.66677 | 0.0043907 |  | 161 | 0.65861 | 0.049185 | 161 | 0.68385 | -0.086143 | 161 | 43667.6406 | 161 | 43217.0391 | 161 | 44804.4414
162 | 5.715 | 0.66158 | 0.032763 | 162 | 0.66664 | 0.0051103 | 162 | 0.67051 | -0.0158 |  | 162 | 5.715 | 0.6676 | -9.1935e-05 |  | 162 | 0.65905 | 0.046759 | 162 | 0.68324 | -0.082978 | 162 | 43641.1602 | 162 | 43243.5586 | 162 | 44787.9609
163 | 5.7503 | 0.66182 | 0.031425 | 163 | 0.66731 | 0.0014393 | 163 | 0.67112 | -0.019079 |  | 163 | 5.7503 | 0.66764 | -0.0003098 |  | 163 | 0.65942 | 0.044709 | 163 | 0.68277 | -0.080543 | 163 | 43635.9609 | 163 | 43278.3203 | 163 | 44757.8789
164 | 5.7856 | 0.66217 | 0.029546 | 164 | 0.66789 | -0.00169 | 164 | 0.67079 | -0.017324 |  | 164 | 5.7856 | 0.66658 | 0.0054149 |  | 164 | 0.65983 | 0.042431 | 164 | 0.68323 | -0.082942 | 164 | 43633.3984 | 164 | 43261.4414 | 164 | 44754.7617
165 | 5.8208 | 0.66258 | 0.027252 | 165 | 0.66826 | -0.0036921 | 165 | 0.67136 | -0.020362 |  | 165 | 5.8208 | 0.66765 | -0.00036241 |  | 165 | 0.65971 | 0.043111 | 165 | 0.68283 | -0.080841 | 165 | 43661.8398 | 165 | 43210.3984 | 165 | 44760.1602
166 | 5.8561 | 0.66285 | 0.025809 | 166 | 0.6676 | -8.1283e-05 | 166 | 0.67172 | -0.022281 |  | 166 | 5.8561 | 0.66865 | -0.0058077 |  | 166 | 0.65965 | 0.043418 | 166 | 0.68305 | -0.081975 | 166 | 43737.9609 | 166 | 43150.7188 | 166 | 44749.0391
167 | 5.8914 | 0.66258 | 0.027259 | 167 | 0.66631 | 0.0069052 | 167 | 0.6709 | -0.017872 |  | 167 | 5.8914 | 0.66778 | -0.0010681 |  | 167 | 0.65974 | 0.042938 | 167 | 0.68326 | -0.083072 | 167 | 43804.0781 | 167 | 43118.3203 | 167 | 44739.7188
168 | 5.9267 | 0.66212 | 0.029786 | 168 | 0.66671 | 0.0046997 | 168 | 0.67024 | -0.014327 |  | 168 | 5.9267 | 0.66687 | 0.0038333 |  | 168 | 0.65992 | 0.041951 | 168 | 0.68331 | -0.083341 | 168 | 43720 | 168 | 43116.2812 | 168 | 44739.1602
169 | 5.9619 | 0.66227 | 0.028995 | 169 | 0.66774 | -0.0008713 | 169 | 0.66992 | -0.012615 |  | 169 | 5.9619 | 0.66574 | 0.0099896 |  | 169 | 0.66057 | 0.03835 | 169 | 0.68338 | -0.083683 | 169 | 43622.8008 | 169 | 43172.3203 | 169 | 44760.8398
170 | 5.9972 | 0.66199 | 0.030495 | 170 | 0.6683 | -0.0039073 | 170 | 0.66923 | -0.0089384 |  | 170 | 5.9972 | 0.66527 | 0.012559 |  | 170 | 0.66087 | 0.036714 | 170 | 0.68292 | -0.081295 | 170 | 43626.1992 | 170 | 43245.4805 | 170 | 44766.5586
171 | 6.0325 | 0.66154 | 0.032971 | 171 | 0.6684 | -0.0044594 | 171 | 0.66882 | -0.0067128 |  | 171 | 6.0325 | 0.66539 | 0.011909 |  | 171 | 0.66077 | 0.037239 | 171 | 0.68218 | -0.077455 | 171 | 43672.8789 | 171 | 43226.9609 | 171 | 44711.9609
172 | 6.0678 | 0.66101 | 0.035905 | 172 | 0.66606 | 0.0082505 | 172 | 0.66864 | -0.0057279 |  | 172 | 6.0678 | 0.66478 | 0.015211 |  | 172 | 0.66042 | 0.039178 | 172 | 0.68206 | -0.076856 | 172 | 43653.5195 | 172 | 43187.3984 | 172 | 44776.7188
173 | 6.1031 | 0.6621 | 0.029885 | 173 | 0.6668 | 0.0042169 | 173 | 0.66944 | -0.01005 |  | 173 | 6.1031 | 0.66355 | 0.021943 |  | 173 | 0.65933 | 0.045228 | 173 | 0.6819 | -0.076005 | 173 | 43666.0781 | 173 | 43195.9219 | 173 | 44771.1211
174 | 6.1383 | 0.66174 | 0.031911 | 174 | 0.66654 | 0.0056506 | 174 | 0.66903 | -0.0078508 |  | 174 | 6.1383 | 0.66377 | 0.020735 |  | 174 | 0.65854 | 0.049594 | 174 | 0.6818 | -0.075481 | 174 | 43616.1211 | 174 | 43158.7188 | 174 | 44739.5586
175 | 6.1736 | 0.6619 | 0.03102 | 175 | 0.66702 | 0.0030326 | 175 | 0.66906 | -0.0079868 |  | 175 | 6.1736 | 0.66588 | 0.0092194 |  | 175 | 0.65824 | 0.051275 | 175 | 0.68146 | -0.073708 | 175 | 43645.0781 | 175 | 43158.1992 | 175 | 44733.7188
176 | 6.2089 | 0.66302 | 0.024885 | 176 | 0.66728 | 0.0016513 | 176 | 0.66981 | -0.012033 |  | 176 | 6.2089 | 0.66585 | 0.009393 |  | 176 | 0.65809 | 0.052112 | 176 | 0.6811 | -0.071832 | 176 | 43638.9219 | 176 | 43125.7188 | 176 | 44782.0781
177 | 6.2442 | 0.66107 | 0.035579 | 177 | 0.66628 | 0.0070313 | 177 | 0.66839 | -0.004377 |  | 177 | 6.2442 | 0.66654 | 0.0056221 |  | 177 | 0.6584 | 0.050373 | 177 | 0.68128 | -0.072772 | 177 | 43703.3984 | 177 | 43150.5586 | 177 | 44749.5195
178 | 6.2794 | 0.66122 | 0.034748 | 178 | 0.66632 | 0.0068547 | 178 | 0.6685 | -0.0049576 |  | 178 | 6.2794 | 0.66572 | 0.010095 |  | 178 | 0.65947 | 0.044409 | 178 | 0.68147 | -0.073766 | 178 | 43664.7188 | 178 | 43080.9609 | 178 | 44731.8008
179 | 6.3147 | 0.65896 | 0.047278 | 179 | 0.6672 | 0.0020361 | 179 | 0.66783 | -0.0013455 |  | 179 | 6.3147 | 0.66542 | 0.011742 |  | 179 | 0.66059 | 0.038254 | 179 | 0.68191 | -0.076073 | 179 | 43617.8789 | 179 | 43142.4414 | 179 | 44737.7188
180 | 6.35 | 0.66152 | 0.03311 | 180 | 0.66775 | -0.00093616 | 180 | 0.66907 | -0.0080549 |  | 180 | 6.35 | 0.6654 | 0.011848 |  | 180 | 0.66116 | 0.035095 | 180 | 0.68209 | -0.076983 | 180 | 43624.3203 | 180 | 43097.2383 | 180 | 44711.2812
181 | 6.3853 | 0.66175 | 0.031838 | 181 | 0.6679 | -0.0017045 | 181 | 0.66928 | -0.0091636 |  | 181 | 6.3853 | 0.66472 | 0.015555 |  | 181 | 0.66112 | 0.035336 | 181 | 0.68248 | -0.07903 | 181 | 43605.6016 | 181 | 43104.0781 | 181 | 44729.0781
182 | 6.4206 | 0.66144 | 0.033552 | 182 | 0.6672 | 0.0020542 | 182 | 0.66862 | -0.0055953 |  | 182 | 6.4206 | 0.66416 | 0.018595 |  | 182 | 0.65969 | 0.043226 | 182 | 0.68352 | -0.084445 | 182 | 43642.5195 | 182 | 43111.3594 | 182 | 44740.5195
183 | 6.4558 | 0.66108 | 0.035531 | 183 | 0.66694 | 0.00345 | 183 | 0.66848 | -0.0048751 |  | 183 | 6.4558 | 0.66453 | 0.016567 |  | 183 | 0.65844 | 0.050145 | 183 | 0.68309 | -0.082194 | 183 | 43669.5586 | 183 | 43138.8789 | 183 | 44716.6797
184 | 6.4911 | 0.6619 | 0.031006 | 184 | 0.66798 | -0.0021496 | 184 | 0.6694 | -0.0098035 |  | 184 | 6.4911 | 0.66489 | 0.014602 |  | 184 | 0.6592 | 0.045955 | 184 | 0.6826 | -0.079671 | 184 | 43676.3984 | 184 | 43136.5195 | 184 | 44690.3984
185 | 6.5264 | 0.66133 | 0.034163 | 185 | 0.66745 | 0.00071649 | 185 | 0.66906 | -0.0079833 |  | 185 | 6.5264 | 0.6649 | 0.014579 |  | 185 | 0.6583 | 0.050939 | 185 | 0.68182 | -0.075615 | 185 | 43660.4414 | 185 | 43150.4805 | 185 | 44694.1211
186 | 6.5617 | 0.66103 | 0.035791 | 186 | 0.66753 | 0.00029254 | 186 | 0.66863 | -0.0056992 |  | 186 | 6.5617 | 0.66499 | 0.014072 |  | 186 | 0.65855 | 0.049517 | 186 | 0.68095 | -0.071083 | 186 | 43660.2812 | 186 | 43154.4414 | 186 | 44667.8398
187 | 6.5969 | 0.6609 | 0.036512 | 187 | 0.66752 | 0.00031062 | 187 | 0.66849 | -0.0049324 |  | 187 | 6.5969 | 0.66524 | 0.012699 |  | 187 | 0.65901 | 0.046967 | 187 | 0.68106 | -0.071655 | 187 | 43685.2383 | 187 | 43155.6406 | 187 | 44657.3984
188 | 6.6322 | 0.66164 | 0.032452 | 188 | 0.66697 | 0.0033314 | 188 | 0.66909 | -0.0081549 |  | 188 | 6.6322 | 0.66574 | 0.01002 |  | 188 | 0.6591 | 0.046471 | 188 | 0.6803 | -0.067682 | 188 | 43705.3984 | 188 | 43181.2812 | 188 | 44655.2812
189 | 6.6675 | 0.66217 | 0.029505 | 189 | 0.66681 | 0.0041593 | 189 | 0.66981 | -0.012047 |  | 189 | 6.6675 | 0.6656 | 0.01073 |  | 189 | 0.65993 | 0.041897 | 189 | 0.67962 | -0.064098 | 189 | 43680.9609 | 189 | 43195.0391 | 189 | 44634.7617
190 | 6.7028 | 0.66185 | 0.031309 | 190 | 0.6666 | 0.0053372 | 190 | 0.66936 | -0.0096178 |  | 190 | 6.7028 | 0.66512 | 0.013349 |  | 190 | 0.66004 | 0.041287 | 190 | 0.67943 | -0.063095 | 190 | 43612.8789 | 190 | 43229.1211 | 190 | 44591.8398
191 | 6.7381 | 0.66234 | 0.028568 | 191 | 0.66764 | -0.00029657 | 191 | 0.66938 | -0.0097071 |  | 191 | 6.7381 | 0.66481 | 0.015037 |  | 191 | 0.65989 | 0.042101 | 191 | 0.67928 | -0.062306 | 191 | 43547.6797 | 191 | 43258.7617 | 191 | 44599.4805
192 | 6.7733 | 0.6623 | 0.028787 | 192 | 0.66763 | -0.00027882 | 192 | 0.66926 | -0.0090884 |  | 192 | 6.7733 | 0.66517 | 0.013107 |  | 192 | 0.65947 | 0.044429 | 192 | 0.67968 | -0.064414 | 192 | 43532.3984 | 192 | 43213.2383 | 192 | 44628.4805
193 | 6.8086 | 0.66245 | 0.027966 | 193 | 0.66724 | 0.0018635 | 193 | 0.66932 | -0.0093853 |  | 193 | 6.8086 | 0.66476 | 0.015355 |  | 193 | 0.65941 | 0.044778 | 193 | 0.68038 | -0.068088 | 193 | 43548.5195 | 193 | 43180.4414 | 193 | 44567.6797
194 | 6.8439 | 0.66291 | 0.025438 | 194 | 0.66805 | -0.0025226 | 194 | 0.66912 | -0.0083339 |  | 194 | 6.8439 | 0.66469 | 0.015699 |  | 194 | 0.65947 | 0.044421 | 194 | 0.68062 | -0.069317 | 194 | 43498.1602 | 194 | 43133.0781 | 194 | 44605.8008
195 | 6.8792 | 0.66257 | 0.027318 | 195 | 0.66788 | -0.0016255 | 195 | 0.66968 | -0.011344 |  | 195 | 6.8792 | 0.66508 | 0.013568 |  | 195 | 0.65967 | 0.043306 | 195 | 0.681 | -0.071347 | 195 | 43481.1211 | 195 | 43055.6016 | 195 | 44633.8008
196 | 6.9144 | 0.6626 | 0.027143 | 196 | 0.66762 | -0.00022847 | 196 | 0.66879 | -0.0065622 |  | 196 | 6.9144 | 0.66588 | 0.0092155 |  | 196 | 0.66029 | 0.039902 | 196 | 0.68103 | -0.071463 | 196 | 43445.3203 | 196 | 43051 | 196 | 44680.6016
197 | 6.9497 | 0.66349 | 0.022281 | 197 | 0.66775 | -0.00092164 | 197 | 0.66879 | -0.0065336 |  | 197 | 6.9497 | 0.66591 | 0.0090949 |  | 197 | 0.66094 | 0.036278 | 197 | 0.68059 | -0.069208 | 197 | 43547.1992 | 197 | 43085.6016 | 197 | 44751.3594
198 | 6.985 | 0.66294 | 0.025274 | 198 | 0.66877 | -0.0064335 | 198 | 0.66883 | -0.0067449 |  | 198 | 6.985 | 0.66515 | 0.013201 |  | 198 | 0.66207 | 0.03006 | 198 | 0.68047 | -0.068578 | 198 | 43534.2812 | 198 | 43127.2383 | 198 | 44721.4805
199 | 7.0203 | 0.66286 | 0.025711 | 199 | 0.66849 | -0.004936 | 199 | 0.66941 | -0.0098609 |  | 199 | 7.0203 | 0.66463 | 0.016029 |  | 199 | 0.66196 | 0.030675 | 199 | 0.68123 | -0.072534 | 199 | 43596.3984 | 199 | 43167.1602 | 199 | 44706.2383
200 | 7.0556 | 0.66247 | 0.027897 | 200 | 0.66731 | 0.00145 | 200 | 0.66923 | -0.0088883 |  | 200 | 7.0556 | 0.6648 | 0.015135 |  | 200 | 0.66133 | 0.034136 | 200 | 0.68163 | -0.074618 | 200 | 43539.8789 | 200 | 43235.0781 | 200 | 44647.8008
201 | 7.0908 | 0.66267 | 0.026778 | 201 | 0.66828 | -0.0038033 | 201 | 0.66898 | -0.0075396 |  | 201 | 7.0908 | 0.66374 | 0.020906 |  | 201 | 0.66108 | 0.035551 | 201 | 0.68134 | -0.073087 | 201 | 43534.4805 | 201 | 43205.7617 | 201 | 44605.3594
202 | 7.1261 | 0.66261 | 0.02711 | 202 | 0.66813 | -0.002989 | 202 | 0.66869 | -0.0059966 |  | 202 | 7.1261 | 0.6631 | 0.024395 |  | 202 | 0.66068 | 0.03776 | 202 | 0.68148 | -0.073809 | 202 | 43527.3594 | 202 | 43216.4414 | 202 | 44588.3203
203 | 7.1614 | 0.66195 | 0.030754 | 203 | 0.66701 | 0.003083 | 203 | 0.66716 | 0.0022665 |  | 203 | 7.1614 | 0.66389 | 0.020098 |  | 203 | 0.66016 | 0.040612 | 203 | 0.68131 | -0.072931 | 203 | 43483.9609 | 203 | 43116.8789 | 203 | 44551.7188
204 | 7.1967 | 0.66357 | 0.021863 | 204 | 0.66873 | -0.0062115 | 204 | 0.66821 | -0.0034195 |  | 204 | 7.1967 | 0.66487 | 0.014726 |  | 204 | 0.65967 | 0.043349 | 204 | 0.68087 | -0.070642 | 204 | 43460.3984 | 204 | 43050.3984 | 204 | 44489.7617
205 | 7.2319 | 0.66359 | 0.02171 | 205 | 0.66845 | -0.0046922 | 205 | 0.66869 | -0.0059786 |  | 205 | 7.2319 | 0.66537 | 0.011992 |  | 205 | 0.65974 | 0.042961 | 205 | 0.6811 | -0.071846 | 205 | 43490.3203 | 205 | 43104.2383 | 205 | 44467
206 | 7.2672 | 0.66358 | 0.021768 | 206 | 0.66839 | -0.0043985 | 206 | 0.66892 | -0.0072389 |  | 206 | 7.2672 | 0.66566 | 0.010424 |  | 206 | 0.65969 | 0.043222 | 206 | 0.68113 | -0.071984 | 206 | 43513.8008 | 206 | 43096.6016 | 206 | 44461.8398
207 | 7.3025 | 0.66321 | 0.02383 | 207 | 0.66786 | -0.0015068 | 207 | 0.66875 | -0.0063296 |  | 207 | 7.3025 | 0.66579 | 0.0097365 |  | 207 | 0.65969 | 0.043207 | 207 | 0.68039 | -0.06815 | 207 | 43524.5195 | 207 | 43101.1602 | 207 | 44501.5586
208 | 7.3378 | 0.66447 | 0.016941 | 208 | 0.66816 | -0.0031614 | 208 | 0.66994 | -0.012725 |  | 208 | 7.3378 | 0.66631 | 0.0069033 |  | 208 | 0.65974 | 0.042934 | 208 | 0.68049 | -0.068657 | 208 | 43539 | 208 | 43090.7188 | 208 | 44546.4414
209 | 7.3731 | 0.66393 | 0.019876 | 209 | 0.66856 | -0.0052837 | 209 | 0.66979 | -0.011904 |  | 209 | 7.3731 | 0.66669 | 0.0048123 |  | 209 | 0.65867 | 0.048892 | 209 | 0.68077 | -0.070127 | 209 | 43542.8398 | 209 | 43150.3203 | 209 | 44568.0391
210 | 7.4083 | 0.66331 | 0.023262 | 210 | 0.66829 | -0.0038249 | 210 | 0.66874 | -0.0062723 |  | 210 | 7.4083 | 0.66594 | 0.00891 |  | 210 | 0.6584 | 0.050349 | 210 | 0.67987 | -0.065408 | 210 | 43548.8789 | 210 | 43166.2812 | 210 | 44594
211 | 7.4436 | 0.66318 | 0.023979 | 211 | 0.66796 | -0.0020309 | 211 | 0.66903 | -0.0078257 |  | 211 | 7.4436 | 0.66561 | 0.010681 |  | 211 | 0.65878 | 0.048257 | 211 | 0.67967 | -0.064377 | 211 | 43540.8398 | 211 | 43182 | 211 | 44603.7617
212 | 7.4789 | 0.66365 | 0.021379 | 212 | 0.66888 | -0.0070386 | 212 | 0.66982 | -0.012065 |  | 212 | 7.4789 | 0.66505 | 0.013731 |  | 212 | 0.6597 | 0.043153 | 212 | 0.67989 | -0.06551 | 212 | 43535.8789 | 212 | 43173.3203 | 212 | 44576.9609
213 | 7.5142 | 0.66397 | 0.019651 | 213 | 0.669 | -0.0076541 | 213 | 0.66984 | -0.012204 |  | 213 | 7.5142 | 0.66518 | 0.01305 |  | 213 | 0.65943 | 0.044648 | 213 | 0.68101 | -0.071372 | 213 | 43528.7617 | 213 | 43183.8008 | 213 | 44543
214 | 7.5494 | 0.66347 | 0.022375 | 214 | 0.66862 | -0.005631 | 214 | 0.66935 | -0.0095783 |  | 214 | 7.5494 | 0.665 | 0.014049 |  | 214 | 0.65954 | 0.04404 | 214 | 0.6814 | -0.073401 | 214 | 43544.1992 | 214 | 43164.7617 | 214 | 44589.4805
215 | 7.5847 | 0.66408 | 0.019041 | 215 | 0.66936 | -0.0095928 | 215 | 0.6701 | -0.013589 |  | 215 | 7.5847 | 0.66435 | 0.017583 |  | 215 | 0.66016 | 0.040589 | 215 | 0.68132 | -0.072974 | 215 | 43515.2812 | 215 | 43140.5586 | 215 | 44599.8398
216 | 7.62 | 0.66373 | 0.020951 | 216 | 0.66989 | -0.012476 | 216 | 0.67051 | -0.015817 |  | 216 | 7.62 | 0.6646 | 0.016192 |  | 216 | 0.66034 | 0.039595 | 216 | 0.68153 | -0.074091 | 216 | 43507.9609 | 216 | 43145.3984 | 216 | 44584.8789
217 | 7.6553 | 0.66393 | 0.019876 | 217 | 0.66987 | -0.012336 | 217 | 0.67088 | -0.017783 |  | 217 | 7.6553 | 0.66482 | 0.015003 |  | 217 | 0.66022 | 0.040274 | 217 | 0.68096 | -0.071137 | 217 | 43564.6016 | 217 | 43175.5195 | 217 | 44592.2812
218 | 7.6906 | 0.66425 | 0.018112 | 218 | 0.66934 | -0.0094784 | 218 | 0.67027 | -0.014502 |  | 218 | 7.6906 | 0.66396 | 0.019684 |  | 218 | 0.65983 | 0.042442 | 218 | 0.68032 | -0.067791 | 218 | 43577.6797 | 218 | 43237.9609 | 218 | 44591.3984
219 | 7.7258 | 0.66398 | 0.0196 | 219 | 0.66888 | -0.0070421 | 219 | 0.66994 | -0.012722 |  | 219 | 7.7258 | 0.66362 | 0.021593 |  | 219 | 0.65967 | 0.043303 | 219 | 0.6804 | -0.068186 | 219 | 43590.0391 | 219 | 43289.0391 | 219 | 44589.7617
220 | 7.7611 | 0.66461 | 0.016158 | 220 | 0.66911 | -0.0082623 | 220 | 0.67049 | -0.015693 |  | 220 | 7.7611 | 0.6633 | 0.023327 |  | 220 | 0.65914 | 0.046275 | 220 | 0.67993 | -0.065728 | 220 | 43538.0391 | 220 | 43262.5586 | 220 | 44588.7617
221 | 7.7964 | 0.66466 | 0.015869 | 221 | 0.66902 | -0.007783 | 221 | 0.67086 | -0.017669 |  | 221 | 7.7964 | 0.66347 | 0.022399 |  | 221 | 0.65921 | 0.045897 | 221 | 0.67927 | -0.062303 | 221 | 43497.0781 | 221 | 43219.4805 | 221 | 44569.6016
222 | 7.8317 | 0.66389 | 0.020086 | 222 | 0.66871 | -0.0061218 | 222 | 0.6693 | -0.0092815 |  | 222 | 7.8317 | 0.6627 | 0.026626 |  | 222 | 0.65911 | 0.046405 | 222 | 0.67865 | -0.059006 | 222 | 43467.3594 | 222 | 43226.9219 | 222 | 44544
223 | 7.8669 | 0.66454 | 0.016542 | 223 | 0.6695 | -0.010347 | 223 | 0.66992 | -0.012618 |  | 223 | 7.8669 | 0.66267 | 0.026801 |  | 223 | 0.65887 | 0.047753 | 223 | 0.67832 | -0.057291 | 223 | 43428.6406 | 223 | 43200.8398 | 223 | 44561.7188
224 | 7.9022 | 0.66506 | 0.013703 | 224 | 0.67001 | -0.013097 | 224 | 0.67019 | -0.014085 |  | 224 | 7.9022 | 0.66277 | 0.026207 |  | 224 | 0.65946 | 0.044463 | 224 | 0.6784 | -0.057721 | 224 | 43434.8398 | 224 | 43174.6406 | 224 | 44527.7188
225 | 7.9375 | 0.66428 | 0.017949 | 225 | 0.66836 | -0.0042121 | 225 | 0.66923 | -0.0089098 |  | 225 | 7.9375 | 0.66284 | 0.02583 |  | 225 | 0.65856 | 0.049471 | 225 | 0.67886 | -0.06012 | 225 | 43466.3203 | 225 | 43152.8398 | 225 | 44478.5586
226 | 7.9728 | 0.66564 | 0.010554 | 226 | 0.66945 | -0.010075 | 226 | 0.66958 | -0.010794 |  | 226 | 7.9728 | 0.66364 | 0.021453 |  | 226 | 0.65811 | 0.051962 | 226 | 0.67863 | -0.058901 | 226 | 43510.1992 | 226 | 43130.8398 | 226 | 44481.5586
227 | 8.0081 | 0.66558 | 0.010869 | 227 | 0.66853 | -0.0051114 | 227 | 0.67 | -0.013065 |  | 227 | 8.0081 | 0.66312 | 0.024311 |  | 227 | 0.65884 | 0.047914 | 227 | 0.67837 | -0.057564 | 227 | 43559.5586 | 227 | 43165.3984 | 227 | 44512.9609
228 | 8.0433 | 0.66462 | 0.016115 | 228 | 0.66811 | -0.0028814 | 228 | 0.66864 | -0.0057459 |  | 228 | 8.0433 | 0.66285 | 0.025803 |  | 228 | 0.65945 | 0.044532 | 228 | 0.67872 | -0.059371 | 228 | 43593.1992 | 228 | 43153.2383 | 228 | 44573.9219
229 | 8.0786 | 0.66461 | 0.01614 | 229 | 0.6693 | -0.0092744 | 229 | 0.66878 | -0.0064834 |  | 229 | 8.0786 | 0.66287 | 0.025685 |  | 229 | 0.66031 | 0.03976 | 229 | 0.67946 | -0.063299 | 229 | 43608.6016 | 229 | 43152.8008 | 229 | 44588.1992
230 | 8.1139 | 0.66385 | 0.02029 | 230 | 0.66872 | -0.0061613 | 230 | 0.66912 | -0.0083053 |  | 230 | 8.1139 | 0.66232 | 0.028702 |  | 230 | 0.66004 | 0.04126 | 230 | 0.67949 | -0.063447 | 230 | 43617.5195 | 230 | 43099.9609 | 230 | 44604.3594
231 | 8.1492 | 0.6645 | 0.016782 | 231 | 0.66948 | -0.010258 | 231 | 0.66947 | -0.010179 |  | 231 | 8.1492 | 0.66225 | 0.02911 |  | 231 | 0.6605 | 0.038725 | 231 | 0.67886 | -0.06012 | 231 | 43627.6797 | 231 | 43120.8789 | 231 | 44602.3984
232 | 8.1844 | 0.66461 | 0.016129 | 232 | 0.67006 | -0.0134 | 232 | 0.67011 | -0.013646 |  | 232 | 8.1844 | 0.66264 | 0.026957 |  | 232 | 0.66052 | 0.03861 | 232 | 0.67891 | -0.06039 | 232 | 43610.7188 | 232 | 43130.9609 | 232 | 44568.7188
233 | 8.2197 | 0.66402 | 0.019375 | 233 | 0.66972 | -0.011562 | 233 | 0.66992 | -0.012629 |  | 233 | 8.2197 | 0.66329 | 0.023376 |  | 233 | 0.66085 | 0.036821 | 233 | 0.67795 | -0.055365 | 233 | 43578.0391 | 233 | 43189.1211 | 233 | 44538.3203
234 | 8.255 | 0.66464 | 0.015999 | 234 | 0.67034 | -0.014908 | 234 | 0.6702 | -0.014113 |  | 234 | 8.255 | 0.66398 | 0.019623 |  | 234 | 0.65983 | 0.042431 | 234 | 0.67798 | -0.055515 | 234 | 43530.1211 | 234 | 43176.0781 | 234 | 44535.8398
235 | 8.2903 | 0.66358 | 0.021786 | 235 | 0.66928 | -0.0091993 | 235 | 0.66925 | -0.0090457 |  | 235 | 8.2903 | 0.66442 | 0.017177 |  | 235 | 0.65981 | 0.042531 | 235 | 0.67817 | -0.056503 | 235 | 43539.8008 | 235 | 43196.8008 | 235 | 44503.2812
236 | 8.3256 | 0.66353 | 0.022055 | 236 | 0.66972 | -0.011533 | 236 | 0.66893 | -0.0073177 |  | 236 | 8.3256 | 0.66408 | 0.019039 |  | 236 | 0.65912 | 0.046375 | 236 | 0.67819 | -0.056591 | 236 | 43542.7188 | 236 | 43232.2383 | 236 | 44470.9609
237 | 8.3608 | 0.66405 | 0.019211 | 237 | 0.6704 | -0.015186 | 237 | 0.66912 | -0.0083089 |  | 237 | 8.3608 | 0.66416 | 0.018603 |  | 237 | 0.65909 | 0.046517 | 237 | 0.67796 | -0.05538 | 237 | 43538.3594 | 237 | 43232.0781 | 237 | 44448.5195
238 | 8.3961 | 0.6645 | 0.016771 | 238 | 0.67051 | -0.0158 | 238 | 0.66939 | -0.0097501 |  | 238 | 8.3961 | 0.66381 | 0.020538 |  | 238 | 0.65897 | 0.047214 | 238 | 0.67786 | -0.054894 | 238 | 43542.8008 | 238 | 43208.1211 | 238 | 44438.5195
239 | 8.4314 | 0.66369 | 0.021183 | 239 | 0.67078 | -0.017263 | 239 | 0.6693 | -0.0092815 |  | 239 | 8.4314 | 0.66378 | 0.020709 |  | 239 | 0.65868 | 0.048842 | 239 | 0.67742 | -0.052564 | 239 | 43510.3203 | 239 | 43187.9609 | 239 | 44426.9219
240 | 8.4667 | 0.66374 | 0.020929 | 240 | 0.66973 | -0.011604 | 240 | 0.66907 | -0.0080334 |  | 240 | 8.4667 | 0.66389 | 0.020086 |  | 240 | 0.65816 | 0.051719 | 240 | 0.67715 | -0.05111 | 240 | 43483.3203 | 240 | 43186.1602 | 240 | 44439.3203
241 | 8.5019 | 0.66433 | 0.01767 | 241 | 0.66996 | -0.012822 | 241 | 0.66909 | -0.0081337 |  | 241 | 8.5019 | 0.66355 | 0.021962 |  | 241 | 0.65785 | 0.05344 | 241 | 0.67716 | -0.051176 | 241 | 43440.7617 | 241 | 43167.4805 | 241 | 44448.5586
242 | 8.5372 | 0.66383 | 0.020424 | 242 | 0.66932 | -0.0094174 | 242 | 0.66921 | -0.0087919 |  | 242 | 8.5372 | 0.66349 | 0.022255 |  | 242 | 0.65784 | 0.053475 | 242 | 0.6779 | -0.055062 | 242 | 43442.3594 | 242 | 43152.3594 | 242 | 44429.8789
243 | 8.5725 | 0.66441 | 0.017242 | 243 | 0.66917 | -0.0086023 | 243 | 0.66983 | -0.012165 |  | 243 | 8.5725 | 0.6631 | 0.024433 |  | 243 | 0.65857 | 0.049401 | 243 | 0.67848 | -0.058122 | 243 | 43431.3984 | 243 | 43144.4414 | 243 | 44450.2383
244 | 8.6078 | 0.66419 | 0.018439 | 244 | 0.67044 | -0.0154 | 244 | 0.67042 | -0.015329 |  | 244 | 8.6078 | 0.66268 | 0.026713 |  | 244 | 0.65942 | 0.044725 | 244 | 0.67858 | -0.058668 | 244 | 43380.3984 | 244 | 43131.6016 | 244 | 44482.6797
245 | 8.6431 | 0.66465 | 0.015959 | 245 | 0.66938 | -0.0097286 | 245 | 0.67052 | -0.015878 |  | 245 | 8.6431 | 0.66334 | 0.023087 |  | 245 | 0.65955 | 0.043983 | 245 | 0.67897 | -0.060717 | 245 | 43386.8789 | 245 | 43130.2812 | 245 | 44447
246 | 8.6783 | 0.66442 | 0.01718 | 246 | 0.66852 | -0.0051008 | 246 | 0.67035 | -0.014926 |  | 246 | 8.6783 | 0.66327 | 0.023482 |  | 246 | 0.65971 | 0.043088 | 246 | 0.67889 | -0.060299 | 246 | 43390.8398 | 246 | 43085.4805 | 246 | 44438.9609
247 | 8.7136 | 0.66468 | 0.015753 | 247 | 0.66817 | -0.0032007 | 247 | 0.67003 | -0.013204 |  | 247 | 8.7136 | 0.66356 | 0.021913 |  | 247 | 0.65909 | 0.046521 | 247 | 0.67618 | -0.046027 | 247 | 43406.5195 | 247 | 43045.6797 | 247 | 44454.9609
248 | 8.7489 | 0.66449 | 0.016814 | 248 | 0.66987 | -0.012354 | 248 | 0.67041 | -0.015243 |  | 248 | 8.7489 | 0.66288 | 0.025639 |  | 248 | 0.65955 | 0.043964 | 248 | 0.67632 | -0.046753 | 248 | 43429.9219 | 248 | 43077.6406 | 248 | 44395.3984
249 | 8.7842 | 0.66487 | 0.014728 | 249 | 0.66946 | -0.010136 | 249 | 0.67024 | -0.014356 |  | 249 | 8.7842 | 0.66257 | 0.027315 |  | 249 | 0.65987 | 0.042231 | 249 | 0.67756 | -0.053273 | 249 | 43434.9609 | 249 | 43144.1602 | 249 | 44396.7188
250 | 8.8194 | 0.66396 | 0.019698 | 250 | 0.66921 | -0.0087989 | 250 | 0.66983 | -0.012122 |  | 250 | 8.8194 | 0.66301 | 0.024924 |  | 250 | 0.6602 | 0.040382 | 250 | 0.6778 | -0.054551 | 250 | 43433.9609 | 250 | 43196.0781 | 250 | 44395.4805
251 | 8.8547 | 0.66486 | 0.014764 | 251 | 0.66904 | -0.0079116 | 251 | 0.67052 | -0.015864 |  | 251 | 8.8547 | 0.66372 | 0.02102 |  | 251 | 0.65994 | 0.04184 | 251 | 0.6774 | -0.052433 | 251 | 43403.3594 | 251 | 43191.4805 | 251 | 44400.6016
252 | 8.89 | 0.66372 | 0.021002 | 252 | 0.66827 | -0.0037424 | 252 | 0.67079 | -0.017285 |  | 252 | 8.89 | 0.66364 | 0.021446 |  | 252 | 0.66071 | 0.037576 | 252 | 0.67773 | -0.054168 | 252 | 43363.8398 | 252 | 43180.7617 | 252 | 44397.2383
253 | 8.9253 | 0.66387 | 0.020188 | 253 | 0.66837 | -0.0042836 | 253 | 0.6708 | -0.017331 |  | 253 | 8.9253 | 0.66403 | 0.019312 |  | 253 | 0.66015 | 0.040646 | 253 | 0.67744 | -0.052678 | 253 | 43375.5586 | 253 | 43199.1602 | 253 | 44412.8789
254 | 8.9606 | 0.66413 | 0.018791 | 254 | 0.66935 | -0.0095391 | 254 | 0.67111 | -0.019014 |  | 254 | 8.9606 | 0.66426 | 0.018087 |  | 254 | 0.65949 | 0.04434 | 254 | 0.67816 | -0.056463 | 254 | 43395.2812 | 254 | 43244.3984 | 254 | 44377.4805
255 | 8.9958 | 0.66515 | 0.013204 | 255 | 0.66962 | -0.011036 | 255 | 0.6715 | -0.021133 |  | 255 | 8.9958 | 0.66447 | 0.016934 |  | 255 | 0.66055 | 0.038457 | 255 | 0.6781 | -0.056146 | 255 | 43417.2812 | 255 | 43209.4414 | 255 | 44350.3984
256 | 9.0311 | 0.66425 | 0.018145 | 256 | 0.66966 | -0.011204 | 256 | 0.67076 | -0.017132 |  | 256 | 9.0311 | 0.66408 | 0.019073 |  | 256 | 0.6611 | 0.035424 | 256 | 0.67824 | -0.056886 | 256 | 43438.5195 | 256 | 43168.1992 | 256 | 44367.6406
257 | 9.0664 | 0.66455 | 0.016484 | 257 | 0.66893 | -0.0073068 | 257 | 0.67076 | -0.017167 |  | 257 | 9.0664 | 0.66337 | 0.022958 |  | 257 | 0.66159 | 0.032729 | 257 | 0.67811 | -0.056178 | 257 | 43468.2812 | 257 | 43168.6016 | 257 | 44392.7617
258 | 9.1017 | 0.66505 | 0.013776 | 258 | 0.66973 | -0.011597 | 258 | 0.6713 | -0.020024 |  | 258 | 9.1017 | 0.66332 | 0.023216 |  | 258 | 0.66198 | 0.030591 | 258 | 0.67681 | -0.049354 | 258 | 43449.5586 | 258 | 43175.1602 | 258 | 44424.8008
259 | 9.1369 | 0.66469 | 0.015709 | 259 | 0.66943 | -0.0099859 | 259 | 0.67073 | -0.017 |  | 259 | 9.1369 | 0.66348 | 0.02235 |  | 259 | 0.66155 | 0.032951 | 259 | 0.67593 | -0.044685 | 259 | 43472.1602 | 259 | 43144.2383 | 259 | 44438.5195
260 | 9.1722 | 0.66513 | 0.01332 | 260 | 0.66942 | -0.0099322 | 260 | 0.67037 | -0.015065 |  | 260 | 9.1722 | 0.66369 | 0.021195 |  | 260 | 0.66184 | 0.031369 | 260 | 0.67497 | -0.039589 | 260 | 43471.5586 | 260 | 43131 | 260 | 44445.7617
261 | 9.2075 | 0.66556 | 0.010966 | 261 | 0.66987 | -0.012358 | 261 | 0.67066 | -0.016619 |  | 261 | 9.2075 | 0.66412 | 0.018826 |  | 261 | 0.66103 | 0.035822 | 261 | 0.67488 | -0.03914 | 261 | 43475.6797 | 261 | 43112.6797 | 261 | 44424.9609
262 | 9.2428 | 0.6656 | 0.010771 | 262 | 0.66988 | -0.0124 | 262 | 0.67072 | -0.016918 |  | 262 | 9.2428 | 0.66379 | 0.020614 |  | 262 | 0.66049 | 0.03881 | 262 | 0.67568 | -0.043341 | 262 | 43447.8789 | 262 | 43081.0781 | 262 | 44310.5586
263 | 9.2781 | 0.66515 | 0.013211 | 263 | 0.66926 | -0.0090884 | 263 | 0.67042 | -0.015315 |  | 263 | 9.2781 | 0.66433 | 0.017708 |  | 263 | 0.66085 | 0.036791 | 263 | 0.6768 | -0.049306 | 263 | 43423.8008 | 263 | 43064.2812 | 263 | 44290.7188
264 | 9.3133 | 0.6638 | 0.020602 | 264 | 0.66881 | -0.0066661 | 264 | 0.66955 | -0.010618 |  | 264 | 9.3133 | 0.66457 | 0.016397 |  | 264 | 0.66085 | 0.036802 | 264 | 0.67767 | -0.05389 | 264 | 43437.4414 | 264 | 43089.7617 | 264 | 44373.7188
265 | 9.3486 | 0.66406 | 0.019168 | 265 | 0.66915 | -0.008495 | 265 | 0.66959 | -0.010865 |  | 265 | 9.3486 | 0.6637 | 0.021115 |  | 265 | 0.66095 | 0.036224 | 265 | 0.67791 | -0.055143 | 265 | 43426.1992 | 265 | 43106.8008 | 265 | 44378.7188
266 | 9.3839 | 0.66427 | 0.018033 | 266 | 0.66962 | -0.010994 | 266 | 0.66975 | -0.011722 |  | 266 | 9.3839 | 0.66391 | 0.019964 |  | 266 | 0.66139 | 0.033841 | 266 | 0.67745 | -0.052704 | 266 | 43392.1602 | 266 | 43119.3594 | 266 | 44371.8008
267 | 9.4192 | 0.66406 | 0.019153 | 267 | 0.67014 | -0.013807 | 267 | 0.66987 | -0.012383 |  | 267 | 9.4192 | 0.66359 | 0.02173 |  | 267 | 0.66142 | 0.033662 | 267 | 0.6772 | -0.05142 | 267 | 43412.3594 | 267 | 43134.7188 | 267 | 44394.8008
268 | 9.4544 | 0.66333 | 0.023146 | 268 | 0.66941 | -0.009868 | 268 | 0.66891 | -0.0071637 |  | 268 | 9.4544 | 0.66344 | 0.022528 |  | 268 | 0.66125 | 0.034613 | 268 | 0.67773 | -0.054193 | 268 | 43388 | 268 | 43122.5195 | 268 | 44399.0781
269 | 9.4897 | 0.66347 | 0.022397 | 269 | 0.66902 | -0.007783 | 269 | 0.66955 | -0.010643 |  | 269 | 9.4897 | 0.66372 | 0.021028 |  | 269 | 0.66091 | 0.036465 | 269 | 0.67779 | -0.054482 | 269 | 43403.5195 | 269 | 43167.8398 | 269 | 44411.9609
270 | 9.525 | 0.66457 | 0.016365 | 270 | 0.67067 | -0.016683 | 270 | 0.67009 | -0.013564 |  | 270 | 9.525 | 0.66408 | 0.019065 |  | 270 | 0.66086 | 0.036718 | 270 | 0.67709 | -0.050839 | 270 | 43312.8008 | 270 | 43211.0781 | 270 | 44385.7617
271 | 9.5603 | 0.66181 | 0.031491 | 271 | 0.67002 | -0.013147 | 271 | 0.66895 | -0.0074177 |  | 271 | 9.5603 | 0.66447 | 0.016946 |  | 271 | 0.6614 | 0.033746 | 271 | 0.67728 | -0.051797 | 271 | 43411.0781 | 271 | 43200.3203 | 271 | 44384.4414
272 | 9.5956 | 0.66446 | 0.016992 | 272 | 0.67005 | -0.013318 | 272 | 0.66999 | -0.01299 |  | 272 | 9.5956 | 0.66408 | 0.019046 |  | 272 | 0.66179 | 0.031644 | 272 | 0.67761 | -0.053551 | 272 | 43413.7188 | 272 | 43212.0391 | 272 | 44400.2383
273 | 9.6308 | 0.66419 | 0.01845 | 273 | 0.66989 | -0.012443 | 273 | 0.6697 | -0.011451 |  | 273 | 9.6308 | 0.66346 | 0.022471 |  | 273 | 0.66254 | 0.027475 | 273 | 0.67763 | -0.053679 | 273 | 43420.1992 | 273 | 43220.5195 | 273 | 44334.3594
274 | 9.6661 | 0.66439 | 0.017347 | 274 | 0.67008 | -0.013493 | 274 | 0.66996 | -0.012839 |  | 274 | 9.6661 | 0.66385 | 0.020306 |  | 274 | 0.66188 | 0.031129 | 274 | 0.67736 | -0.052232 | 274 | 43432.4414 | 274 | 43143.6797 | 274 | 44267.7188
275 | 9.7014 | 0.66456 | 0.016426 | 275 | 0.66982 | -0.012115 | 275 | 0.66993 | -0.012683 |  | 275 | 9.7014 | 0.66374 | 0.020902 |  | 275 | 0.66185 | 0.031278 | 275 | 0.67742 | -0.052532 | 275 | 43438.9219 | 275 | 43162.6016 | 275 | 44226.9609
276 | 9.7367 | 0.66468 | 0.015771 | 276 | 0.6702 | -0.014127 | 276 | 0.67101 | -0.018498 |  | 276 | 9.7367 | 0.6636 | 0.021654 |  | 276 | 0.66176 | 0.031789 | 276 | 0.67719 | -0.051358 | 276 | 43516.2383 | 276 | 43237.6797 | 276 | 44205.9219
277 | 9.7719 | 0.66404 | 0.019269 | 277 | 0.6691 | -0.0082337 | 277 | 0.67039 | -0.015169 |  | 277 | 9.7719 | 0.66432 | 0.017715 |  | 277 | 0.66247 | 0.027852 | 277 | 0.67694 | -0.05002 | 277 | 43499.3594 | 277 | 43266.3203 | 277 | 44233.9609
278 | 9.8072 | 0.66508 | 0.013613 | 278 | 0.67006 | -0.013403 | 278 | 0.6709 | -0.017879 |  | 278 | 9.8072 | 0.66525 | 0.012672 |  | 278 | 0.66209 | 0.029973 | 278 | 0.67565 | -0.043191 | 278 | 43485 | 278 | 43299.2812 | 278 | 44304.0781
279 | 9.8425 | 0.66466 | 0.015887 | 279 | 0.66972 | -0.011526 | 279 | 0.67091 | -0.017947 |  | 279 | 9.8425 | 0.66539 | 0.011878 |  | 279 | 0.66248 | 0.027803 | 279 | 0.67616 | -0.045888 | 279 | 43476 | 279 | 43348.9609 | 279 | 44346.7188
280 | 9.8778 | 0.66466 | 0.015905 | 280 | 0.66984 | -0.012172 | 280 | 0.67134 | -0.020274 |  | 280 | 9.8778 | 0.66602 | 0.0084645 |  | 280 | 0.66332 | 0.023205 | 280 | 0.67612 | -0.04569 | 280 | 43458.4414 | 280 | 43326.2383 | 280 | 44357.5195
281 | 9.9131 | 0.66544 | 0.01162 | 281 | 0.67048 | -0.015628 | 281 | 0.67215 | -0.024609 |  | 281 | 9.9131 | 0.66637 | 0.0065755 |  | 281 | 0.66301 | 0.024939 | 281 | 0.67572 | -0.043577 | 281 | 43445.8008 | 281 | 43300.1602 | 281 | 44341.7188
282 | 9.9483 | 0.66502 | 0.013939 | 282 | 0.67069 | -0.016779 | 282 | 0.67183 | -0.022881 |  | 282 | 9.9483 | 0.66619 | 0.0075481 |  | 282 | 0.66384 | 0.020344 | 282 | 0.67599 | -0.045033 | 282 | 43442.6406 | 282 | 43254.0391 | 282 | 44326.7617
283 | 9.9836 | 0.66562 | 0.010641 | 283 | 0.67055 | -0.016013 | 283 | 0.67216 | -0.024655 |  | 283 | 9.9836 | 0.66588 | 0.0092269 |  | 283 | 0.66443 | 0.017162 | 283 | 0.67662 | -0.048361 | 283 | 43478 | 283 | 43246.8789 | 283 | 44335.9219
284 | 10.0189 | 0.66627 | 0.0070889 | 284 | 0.67091 | -0.017922 | 284 | 0.6722 | -0.024857 |  | 284 | 10.0189 | 0.66541 | 0.011803 |  | 284 | 0.66341 | 0.022741 | 284 | 0.677 | -0.050327 | 284 | 43482.0391 | 284 | 43250.1992 | 284 | 44339.3984
285 | 10.0542 | 0.66694 | 0.0034539 | 285 | 0.67124 | -0.019715 | 285 | 0.67197 | -0.023644 |  | 285 | 10.0542 | 0.66532 | 0.012275 |  | 285 | 0.6626 | 0.027186 | 285 | 0.67797 | -0.055438 | 285 | 43514.0781 | 285 | 43228.3203 | 285 | 44315.4414
286 | 10.0894 | 0.6668 | 0.0042457 | 286 | 0.67107 | -0.018805 | 286 | 0.67169 | -0.022103 |  | 286 | 10.0894 | 0.66553 | 0.011145 |  | 286 | 0.66225 | 0.029099 | 286 | 0.67842 | -0.057822 | 286 | 43492.1992 | 286 | 43252.1211 | 286 | 44344.8398
287 | 10.1247 | 0.66665 | 0.00502 | 287 | 0.67141 | -0.02064 | 287 | 0.67138 | -0.020476 |  | 287 | 10.1247 | 0.66582 | 0.0095704 |  | 287 | 0.66166 | 0.032343 | 287 | 0.67887 | -0.060175 | 287 | 43482.3984 | 287 | 43275.8398 | 287 | 44395.8789
288 | 10.16 | 0.6668 | 0.0042205 | 288 | 0.67203 | -0.023946 | 288 | 0.6718 | -0.022704 |  | 288 | 10.16 | 0.66617 | 0.0076385 |  | 288 | 0.66178 | 0.031678 | 288 | 0.67902 | -0.060979 | 288 | 43490.5195 | 288 | 43228.7617 | 288 | 44393
289 | 10.1953 | 0.66723 | 0.0019104 | 289 | 0.67233 | -0.02557 | 289 | 0.67178 | -0.022608 |  | 289 | 10.1953 | 0.66613 | 0.00788 |  | 289 | 0.66273 | 0.026462 | 289 | 0.67922 | -0.062026 | 289 | 43451.2812 | 289 | 43228.1602 | 289 | 44365.7188
290 | 10.2306 | 0.66718 | 0.0021944 | 290 | 0.67238 | -0.025793 | 290 | 0.67104 | -0.018623 |  | 290 | 10.2306 | 0.66624 | 0.0072691 |  | 290 | 0.6637 | 0.021111 | 290 | 0.67954 | -0.063716 | 290 | 43438.9219 | 290 | 43223 | 290 | 44354.3594
291 | 10.2658 | 0.66718 | 0.0021799 | 291 | 0.67185 | -0.023002 | 291 | 0.66968 | -0.011312 |  | 291 | 10.2658 | 0.6659 | 0.0090984 |  | 291 | 0.66444 | 0.017109 | 291 | 0.67946 | -0.063295 | 291 | 43491.3203 | 291 | 43210.4805 | 291 | 44353.2383
292 | 10.3011 | 0.66711 | 0.002547 | 292 | 0.67187 | -0.023087 | 292 | 0.67029 | -0.014591 |  | 292 | 10.3011 | 0.66459 | 0.01629 |  | 292 | 0.6642 | 0.018386 | 292 | 0.67882 | -0.059924 | 292 | 43482.8008 | 292 | 43210.3594 | 292 | 44324.6406
293 | 10.3364 | 0.66721 | 0.0019967 | 293 | 0.67205 | -0.024049 | 293 | 0.6712 | -0.019495 |  | 293 | 10.3364 | 0.66328 | 0.023406 |  | 293 | 0.66324 | 0.023665 | 293 | 0.67762 | -0.053599 | 293 | 43564.6797 | 293 | 43255.5195 | 293 | 44227.3203
294 | 10.3717 | 0.66686 | 0.0039326 | 294 | 0.67232 | -0.02551 | 294 | 0.67069 | -0.016761 |  | 294 | 10.3717 | 0.66093 | 0.036335 |  | 294 | 0.66303 | 0.024798 | 294 | 0.67825 | -0.056919 | 294 | 43470.1992 | 294 | 43262.4414 | 294 | 44281.4805
295 | 10.4069 | 0.66658 | 0.0054308 | 295 | 0.67133 | -0.020177 | 295 | 0.67134 | -0.02027 |  | 295 | 10.4069 | 0.66347 | 0.022403 |  | 295 | 0.66284 | 0.025845 | 295 | 0.67849 | -0.05818 | 295 | 43430.2383 | 295 | 43299.3203 | 295 | 44265.6016
296 | 10.4422 | 0.66709 | 0.0026586 | 296 | 0.67185 | -0.022966 | 296 | 0.67138 | -0.020441 |  | 296 | 10.4422 | 0.66441 | 0.01723 |  | 296 | 0.66224 | 0.029126 | 296 | 0.67872 | -0.059374 | 296 | 43460.0781 | 296 | 43302.5586 | 296 | 44262.0781
297 | 10.4775 | 0.66621 | 0.0074531 | 297 | 0.67057 | -0.016138 | 297 | 0.67017 | -0.013967 |  | 297 | 10.4775 | 0.66367 | 0.021282 |  | 297 | 0.66163 | 0.032511 | 297 | 0.67846 | -0.058045 | 297 | 43457.9609 | 297 | 43366 | 297 | 44260.0391
298 | 10.5128 | 0.66672 | 0.004678 | 298 | 0.67086 | -0.017662 | 298 | 0.6707 | -0.016829 |  | 298 | 10.5128 | 0.66423 | 0.018231 |  | 298 | 0.66213 | 0.029763 | 298 | 0.67832 | -0.057283 | 298 | 43490.1992 | 298 | 43355.2812 | 298 | 44288.2812
299 | 10.5481 | 0.66617 | 0.0076337 | 299 | 0.6716 | -0.021652 | 299 | 0.67208 | -0.024212 |  | 299 | 10.5481 | 0.66347 | 0.022376 |  | 299 | 0.66232 | 0.028714 | 299 | 0.67741 | -0.052499 | 299 | 43521.2812 | 299 | 43333.9219 | 299 | 44352.8398
300 | 10.5833 | 0.66588 | 0.0092577 | 300 | 0.672 | -0.023782 | 300 | 0.67199 | -0.023719 |  | 300 | 10.5833 | 0.66361 | 0.021609 |  | 300 | 0.66255 | 0.027414 | 300 | 0.67863 | -0.058904 | 300 | 43577.7188 | 300 | 43353.8398 | 300 | 44378.9609
301 | 10.6186 | 0.66598 | 0.0086872 | 301 | 0.67156 | -0.021442 | 301 | 0.67168 | -0.022057 |  | 301 | 10.6186 | 0.66264 | 0.02693 |  | 301 | 0.66212 | 0.02982 | 301 | 0.67884 | -0.060019 | 301 | 43553.0391 | 301 | 43368.6406 | 301 | 44409.7617
302 | 10.6539 | 0.66616 | 0.0077059 | 302 | 0.6727 | -0.027526 | 302 | 0.6729 | -0.028616 |  | 302 | 10.6539 | 0.6633 | 0.023319 |  | 302 | 0.66347 | 0.02241 | 302 | 0.67872 | -0.059414 | 302 | 43582.3594 | 302 | 43379.8008 | 302 | 44450.2812
303 | 10.6892 | 0.66569 | 0.010287 | 303 | 0.67204 | -0.024006 | 303 | 0.67191 | -0.023279 |  | 303 | 10.6892 | 0.66413 | 0.018762 |  | 303 | 0.66326 | 0.023547 | 303 | 0.67793 | -0.055263 | 303 | 43633.1211 | 303 | 43474.3203 | 303 | 44437.9219
304 | 10.7244 | 0.66694 | 0.00345 | 304 | 0.67322 | -0.030325 | 304 | 0.67195 | -0.023545 |  | 304 | 10.7244 | 0.66402 | 0.019399 |  | 304 | 0.66303 | 0.024806 | 304 | 0.67755 | -0.053255 | 304 | 43636.2383 | 304 | 43479.4414 | 304 | 44439.6797
305 | 10.7597 | 0.66624 | 0.0072513 | 305 | 0.67238 | -0.025811 | 305 | 0.67205 | -0.024031 |  | 305 | 10.7597 | 0.66427 | 0.018037 |  | 305 | 0.66319 | 0.02395 | 305 | 0.67636 | -0.046977 | 305 | 43604.6797 | 305 | 43418.1211 | 305 | 44457.8008
306 | 10.795 | 0.66551 | 0.011262 | 306 | 0.67146 | -0.020881 | 306 | 0.67145 | -0.020863 |  | 306 | 10.795 | 0.66486 | 0.014802 |  | 306 | 0.66267 | 0.026778 | 306 | 0.67651 | -0.047772 | 306 | 43573.8789 | 306 | 43390.8398 | 306 | 44466.1602
307 | 10.8303 | 0.66541 | 0.011783 | 307 | 0.67142 | -0.020671 | 307 | 0.67092 | -0.018025 |  | 307 | 10.8303 | 0.66501 | 0.013977 |  | 307 | 0.66288 | 0.025639 | 307 | 0.67564 | -0.043165 | 307 | 43561.2812 | 307 | 43347.3203 | 307 | 44433.1992
308 | 10.8656 | 0.66543 | 0.011663 | 308 | 0.67161 | -0.021716 | 308 | 0.67088 | -0.017805 |  | 308 | 10.8656 | 0.66539 | 0.011882 |  | 308 | 0.66324 | 0.023642 | 308 | 0.67597 | -0.044927 | 308 | 43557.0781 | 308 | 43310.0781 | 308 | 44371.3984
309 | 10.9008 | 0.66536 | 0.012083 | 309 | 0.6716 | -0.021642 | 309 | 0.67099 | -0.01836 |  | 309 | 10.9008 | 0.66532 | 0.012286 |  | 309 | 0.66385 | 0.020295 | 309 | 0.67592 | -0.044663 | 309 | 43578.6406 | 309 | 43340.2812 | 309 | 44388.4414
310 | 10.9361 | 0.66598 | 0.0086691 | 310 | 0.67254 | -0.026658 | 310 | 0.6718 | -0.022693 |  | 310 | 10.9361 | 0.6655 | 0.011285 |  | 310 | 0.66395 | 0.019775 | 310 | 0.67632 | -0.046753 | 310 | 43607.3984 | 310 | 43416.9609 | 310 | 44409.4805
311 | 10.9714 | 0.66646 | 0.0060686 | 311 | 0.67334 | -0.030923 | 311 | 0.67231 | -0.025442 |  | 311 | 10.9714 | 0.66572 | 0.010114 |  | 311 | 0.66404 | 0.019263 | 311 | 0.67651 | -0.047783 | 311 | 43638.3594 | 311 | 43489.5586 | 311 | 44434.0781
312 | 11.0067 | 0.66632 | 0.0068437 | 312 | 0.67315 | -0.029929 | 312 | 0.67231 | -0.025428 |  | 312 | 11.0067 | 0.6655 | 0.011311 |  | 312 | 0.66407 | 0.019118 | 312 | 0.67652 | -0.04783 | 312 | 43658.4805 | 312 | 43523.2812 | 312 | 44428.2383
313 | 11.0419 | 0.66612 | 0.0079185 | 313 | 0.67259 | -0.026917 | 313 | 0.67188 | -0.023162 |  | 313 | 11.0419 | 0.66516 | 0.013168 |  | 313 | 0.66443 | 0.017113 | 313 | 0.67628 | -0.04653 | 313 | 43595.5195 | 313 | 43469.6406 | 313 | 44416.4805
314 | 11.0772 | 0.66603 | 0.0084022 | 314 | 0.67208 | -0.024223 | 314 | 0.67218 | -0.024761 |  | 314 | 11.0772 | 0.66566 | 0.010454 |  | 314 | 0.66462 | 0.016074 | 314 | 0.67703 | -0.050506 | 314 | 43540.0781 | 314 | 43439.3594 | 314 | 44407.3594
315 | 11.1125 | 0.66661 | 0.0052832 | 315 | 0.6729 | -0.028588 | 315 | 0.67234 | -0.025613 |  | 315 | 11.1125 | 0.6659 | 0.00914 |  | 315 | 0.66484 | 0.014916 | 315 | 0.67747 | -0.052798 | 315 | 43451.3984 | 315 | 43387.4805 | 315 | 44405.8789
316 | 11.1478 | 0.6668 | 0.0042169 | 316 | 0.67264 | -0.027225 | 316 | 0.67228 | -0.025294 |  | 316 | 11.1478 | 0.66551 | 0.011232 |  | 316 | 0.66511 | 0.013406 | 316 | 0.67755 | -0.053259 | 316 | 43289.1211 | 316 | 43369.3984 | 316 | 44433.7188
317 | 11.1831 | 0.66635 | 0.00666 | 317 | 0.67197 | -0.023616 | 317 | 0.67174 | -0.022398 |  | 317 | 11.1831 | 0.66623 | 0.0073482 |  | 317 | 0.66571 | 0.010182 | 317 | 0.67749 | -0.052912 | 317 | 43455.6016 | 317 | 43310.6406 | 317 | 44457.7617
318 | 11.2183 | 0.66691 | 0.0036337 | 318 | 0.67288 | -0.028493 | 318 | 0.67171 | -0.022259 |  | 318 | 11.2183 | 0.66649 | 0.0059158 |  | 318 | 0.66551 | 0.01124 | 318 | 0.67728 | -0.051834 | 318 | 43472.0391 | 318 | 43280.2383 | 318 | 44366.2383
319 | 11.2536 | 0.66682 | 0.0040982 | 319 | 0.67213 | -0.024478 | 319 | 0.67194 | -0.023463 |  | 319 | 11.2536 | 0.66528 | 0.012506 |  | 319 | 0.66506 | 0.013705 | 319 | 0.67752 | -0.053058 | 319 | 43467.7188 | 319 | 43301.3984 | 319 | 44364.2812
320 | 11.2889 | 0.6668 | 0.0042457 | 320 | 0.67116 | -0.01931 | 320 | 0.67073 | -0.016986 |  | 320 | 11.2889 | 0.66613 | 0.0078686 |  | 320 | 0.66472 | 0.015529 | 320 | 0.67776 | -0.054365 | 320 | 43436.7617 | 320 | 43346.7617 | 320 | 44290
321 | 11.3242 | 0.66793 | -0.0018768 | 321 | 0.67213 | -0.024485 | 321 | 0.67261 | -0.027055 |  | 321 | 11.3242 | 0.66479 | 0.015173 |  | 321 | 0.66425 | 0.018117 | 321 | 0.67779 | -0.054518 | 321 | 43450.7188 | 321 | 43345.2812 | 321 | 44256.9219
322 | 11.3594 | 0.66752 | 0.00034646 | 322 | 0.67197 | -0.02364 | 322 | 0.67254 | -0.026697 |  | 322 | 11.3594 | 0.66487 | 0.014734 |  | 322 | 0.66397 | 0.019669 | 322 | 0.67796 | -0.05542 | 322 | 43457.6016 | 322 | 43388.7617 | 322 | 44242.5586
323 | 11.3947 | 0.66614 | 0.0078213 | 323 | 0.67115 | -0.019221 | 323 | 0.67088 | -0.017773 |  | 323 | 11.3947 | 0.66462 | 0.016082 |  | 323 | 0.66399 | 0.019521 | 323 | 0.67794 | -0.055292 | 323 | 43422.3594 | 323 | 43407.5586 | 323 | 44258.6406
324 | 11.43 | 0.66606 | 0.0082686 | 324 | 0.67182 | -0.022835 | 324 | 0.67121 | -0.019545 |  | 324 | 11.43 | 0.66366 | 0.021358 |  | 324 | 0.66414 | 0.018728 | 324 | 0.67795 | -0.055329 | 324 | 43407.1602 | 324 | 43396.8008 | 324 | 44249.5586
325 | 11.4653 | 0.66733 | 0.0013492 | 325 | 0.67282 | -0.028156 | 325 | 0.6726 | -0.026998 |  | 325 | 11.4653 | 0.66296 | 0.02516 |  | 325 | 0.66484 | 0.014916 | 325 | 0.67742 | -0.052535 | 325 | 43459.2812 | 325 | 43380.8008 | 325 | 44278.6406
326 | 11.5006 | 0.6666 | 0.0053049 | 326 | 0.67198 | -0.023694 | 326 | 0.67149 | -0.02108 |  | 326 | 11.5006 | 0.66302 | 0.024844 |  | 326 | 0.6638 | 0.020591 | 326 | 0.67775 | -0.05427 | 326 | 43481.8398 | 326 | 43393.5586 | 326 | 44294.7188
327 | 11.5358 | 0.66819 | -0.0033013 | 327 | 0.67238 | -0.025793 | 327 | 0.67241 | -0.025953 |  | 327 | 11.5358 | 0.66338 | 0.022878 |  | 327 | 0.66381 | 0.020515 | 327 | 0.67801 | -0.055664 | 327 | 43471.1992 | 327 | 43373.8398 | 327 | 44315.8008
328 | 11.5711 | 0.66735 | 0.0012594 | 328 | 0.67164 | -0.021862 | 328 | 0.67226 | -0.025194 |  | 328 | 11.5711 | 0.66392 | 0.019942 |  | 328 | 0.66381 | 0.02053 | 328 | 0.67736 | -0.052217 | 328 | 43496.0781 | 328 | 43368 | 328 | 44295.0391
329 | 11.6064 | 0.66744 | 0.00075976 | 329 | 0.67167 | -0.022047 | 329 | 0.67189 | -0.0232 |  | 329 | 11.6064 | 0.66388 | 0.020124 |  | 329 | 0.66327 | 0.023494 | 329 | 0.67766 | -0.053796 | 329 | 43525.0781 | 329 | 43392 | 329 | 44295.5195
330 | 11.6417 | 0.66728 | 0.0016296 | 330 | 0.67148 | -0.020998 | 330 | 0.67167 | -0.022018 |  | 330 | 11.6417 | 0.66411 | 0.018906 |  | 330 | 0.66264 | 0.026946 | 330 | 0.67752 | -0.053083 | 330 | 43554.8789 | 330 | 43413.2383 | 330 | 44320.8008
331 | 11.6769 | 0.66719 | 0.0021405 | 331 | 0.67229 | -0.025318 | 331 | 0.67186 | -0.02303 |  | 331 | 11.6769 | 0.66408 | 0.01905 |  | 331 | 0.66204 | 0.030247 | 331 | 0.67704 | -0.050539 | 331 | 43572.2812 | 331 | 43427.4805 | 331 | 44347.6797
332 | 11.7122 | 0.66804 | -0.0024617 | 332 | 0.67231 | -0.025432 | 332 | 0.67166 | -0.021965 |  | 332 | 11.7122 | 0.66341 | 0.022737 |  | 332 | 0.66177 | 0.031717 | 332 | 0.67584 | -0.044219 | 332 | 43568.6016 | 332 | 43448 | 332 | 44324.3594
333 | 11.7475 | 0.66844 | -0.0046707 | 333 | 0.67272 | -0.027639 | 333 | 0.67176 | -0.022487 |  | 333 | 11.7475 | 0.66371 | 0.021062 |  | 333 | 0.66247 | 0.027879 | 333 | 0.6753 | -0.041339 | 333 | 43582.8789 | 333 | 43459.8789 | 333 | 44309.0391
334 | 11.7828 | 0.66766 | -0.00042244 | 334 | 0.6727 | -0.027529 | 334 | 0.67217 | -0.024698 |  | 334 | 11.7828 | 0.66428 | 0.017947 |  | 334 | 0.66368 | 0.021237 | 334 | 0.67482 | -0.038805 | 334 | 43590.1211 | 334 | 43492.3984 | 334 | 44340.9219
335 | 11.8181 | 0.6674 | 0.00095741 | 335 | 0.67283 | -0.02822 | 335 | 0.67204 | -0.024013 |  | 335 | 11.8181 | 0.66445 | 0.017037 |  | 335 | 0.66465 | 0.015961 | 335 | 0.67479 | -0.038665 | 335 | 43559.9219 | 335 | 43526.6406 | 335 | 44345.8008
336 | 11.8533 | 0.66739 | 0.0010546 | 336 | 0.67306 | -0.029445 | 336 | 0.67158 | -0.021563 |  | 336 | 11.8533 | 0.66501 | 0.013996 |  | 336 | 0.66491 | 0.014499 | 336 | 0.67491 | -0.039273 | 336 | 43618.0391 | 336 | 43550.9609 | 336 | 44343.6016
337 | 11.8886 | 0.6673 | 0.0015362 | 337 | 0.67243 | -0.026105 | 337 | 0.67104 | -0.018652 |  | 337 | 11.8886 | 0.66517 | 0.013077 |  | 337 | 0.66524 | 0.012736 | 337 | 0.67484 | -0.038923 | 337 | 43632.7617 | 337 | 43560.9609 | 337 | 44354.0781
338 | 11.9239 | 0.66721 | 0.0019857 | 338 | 0.67257 | -0.026814 | 338 | 0.67102 | -0.018556 |  | 338 | 11.9239 | 0.66514 | 0.013285 |  | 338 | 0.66486 | 0.014817 | 338 | 0.67462 | -0.03773 | 338 | 43560.0391 | 338 | 43556.2812 | 338 | 44353.8398
339 | 11.9592 | 0.66699 | 0.0032269 | 339 | 0.672 | -0.023765 | 339 | 0.67154 | -0.021304 |  | 339 | 11.9592 | 0.66525 | 0.012653 |  | 339 | 0.6652 | 0.01291 | 339 | 0.67446 | -0.036916 | 339 | 43591.8008 | 339 | 43537.8398 | 339 | 44355.5195
340 | 11.9944 | 0.66722 | 0.0019282 | 340 | 0.67273 | -0.027689 | 340 | 0.6718 | -0.022689 |  | 340 | 11.9944 | 0.66591 | 0.0090647 |  | 340 | 0.66535 | 0.012116 | 340 | 0.67448 | -0.037005 | 340 | 43599.4805 | 340 | 43520.1211 | 340 | 44367
341 | 12.0297 | 0.66629 | 0.0069989 | 341 | 0.67244 | -0.026134 | 341 | 0.67137 | -0.020391 |  | 341 | 12.0297 | 0.6651 | 0.01347 |  | 341 | 0.66546 | 0.011496 | 341 | 0.67469 | -0.038098 | 341 | 43612.3984 | 341 | 43470.4414 | 341 | 44363.2812
342 | 12.065 | 0.66603 | 0.0084094 | 342 | 0.67184 | -0.022927 | 342 | 0.67088 | -0.017769 |  | 342 | 12.065 | 0.66546 | 0.0115 |  | 342 | 0.6649 | 0.014598 | 342 | 0.67438 | -0.036499 | 342 | 43597.4805 | 342 | 43482.4414 | 342 | 44381.4414
343 | 12.1003 | 0.66726 | 0.0017197 | 343 | 0.67198 | -0.023687 | 343 | 0.67123 | -0.019651 |  | 343 | 12.1003 | 0.66596 | 0.0088079 |  | 343 | 0.66509 | 0.013557 | 343 | 0.67456 | -0.037406 | 343 | 43544.1211 | 343 | 43509.5586 | 343 | 44335.0781
344 | 12.1356 | 0.66767 | -0.00046181 | 344 | 0.67299 | -0.029087 | 344 | 0.67137 | -0.020409 |  | 344 | 12.1356 | 0.66526 | 0.012634 |  | 344 | 0.66534 | 0.012162 | 344 | 0.67495 | -0.039493 | 344 | 43545.0781 | 344 | 43482.9219 | 344 | 44365.3984
345 | 12.1708 | 0.66704 | 0.0029534 | 345 | 0.67279 | -0.02799 | 345 | 0.67129 | -0.02 |  | 345 | 12.1708 | 0.66549 | 0.01138 |  | 345 | 0.66548 | 0.01141 | 345 | 0.67615 | -0.045844 | 345 | 43520.8398 | 345 | 43497.7617 | 345 | 44340.2812
346 | 12.2061 | 0.6666 | 0.0053262 | 346 | 0.67245 | -0.026169 | 346 | 0.67077 | -0.017213 |  | 346 | 12.2061 | 0.66579 | 0.0096989 |  | 346 | 0.66581 | 0.0095967 | 346 | 0.67759 | -0.053449 | 346 | 43486.9609 | 346 | 43487.1602 | 346 | 44300.3984
347 | 12.2414 | 0.66744 | 0.00078495 | 347 | 0.67381 | -0.033456 | 347 | 0.67145 | -0.020821 |  | 347 | 12.2414 | 0.66662 | 0.0051887 |  | 347 | 0.6656 | 0.010741 | 347 | 0.67829 | -0.05713 | 347 | 43419.6016 | 347 | 43487.6406 | 347 | 44231.5586
348 | 12.2767 | 0.66739 | 0.001033 | 348 | 0.67323 | -0.030364 | 348 | 0.6709 | -0.017911 |  | 348 | 12.2767 | 0.66701 | 0.0031073 |  | 348 | 0.66534 | 0.012196 | 348 | 0.67804 | -0.055807 | 348 | 43431.6016 | 348 | 43462.8398 | 348 | 44196.6016
349 | 12.3119 | 0.66744 | 0.00077785 | 349 | 0.67303 | -0.029296 | 349 | 0.67132 | -0.020145 |  | 349 | 12.3119 | 0.66762 | -0.00020846 |  | 349 | 0.66675 | 0.004511 | 349 | 0.67932 | -0.062525 | 349 | 43436.5195 | 349 | 43443.0781 | 349 | 44186.5195
350 | 12.3472 | 0.66712 | 0.0024714 | 350 | 0.67372 | -0.03299 | 350 | 0.67153 | -0.021279 |  | 350 | 12.3472 | 0.66765 | -0.00035854 |  | 350 | 0.66648 | 0.0059611 | 350 | 0.67892 | -0.060423 | 350 | 43462.4805 | 350 | 43421.5586 | 350 | 44179.3984
351 | 12.3825 | 0.66705 | 0.0028635 | 351 | 0.67324 | -0.030386 | 351 | 0.67194 | -0.023488 |  | 351 | 12.3825 | 0.66745 | 0.00071584 |  | 351 | 0.6668 | 0.0042172 | 351 | 0.67889 | -0.06027 | 351 | 43454.5195 | 351 | 43408.4805 | 351 | 44144.8789
352 | 12.4178 | 0.66644 | 0.0061768 | 352 | 0.67256 | -0.026804 | 352 | 0.67065 | -0.016537 |  | 352 | 12.4178 | 0.66682 | 0.0041231 |  | 352 | 0.66631 | 0.0068881 | 352 | 0.67816 | -0.056438 | 352 | 43475.6797 | 352 | 43384 | 352 | 44134.3594
353 | 12.4531 | 0.66606 | 0.0082505 | 353 | 0.67305 | -0.029388 | 353 | 0.67071 | -0.016861 |  | 353 | 12.4531 | 0.66643 | 0.0062137 |  | 353 | 0.66693 | 0.0035322 | 353 | 0.67766 | -0.053847 | 353 | 43473.3984 | 353 | 43378.6406 | 353 | 44142.5195
354 | 12.4883 | 0.66591 | 0.0090589 | 354 | 0.67323 | -0.030353 | 354 | 0.67118 | -0.019409 |  | 354 | 12.4883 | 0.66555 | 0.011006 |  | 354 | 0.66673 | 0.0045977 | 354 | 0.67784 | -0.054756 | 354 | 43445.2383 | 354 | 43409.4805 | 354 | 44141.7617
355 | 12.5236 | 0.66602 | 0.0084489 | 355 | 0.67255 | -0.026701 | 355 | 0.67153 | -0.021297 |  | 355 | 12.5236 | 0.6649 | 0.014556 |  | 355 | 0.66622 | 0.0073822 | 355 | 0.67737 | -0.052309 | 355 | 43440.5195 | 355 | 43452.6406 | 355 | 44127.8008
356 | 12.5589 | 0.6667 | 0.0047715 | 356 | 0.67301 | -0.0292 | 356 | 0.67281 | -0.028124 |  | 356 | 12.5589 | 0.6655 | 0.011304 |  | 356 | 0.6661 | 0.0080194 | 356 | 0.67708 | -0.050773 | 356 | 43448.8398 | 356 | 43505.7617 | 356 | 44131.4805
357 | 12.5942 | 0.66666 | 0.0049699 | 357 | 0.67313 | -0.029827 | 357 | 0.67278 | -0.027979 |  | 357 | 12.5942 | 0.66548 | 0.011413 |  | 357 | 0.66568 | 0.010329 | 357 | 0.67773 | -0.054182 | 357 | 43524 | 357 | 43494.5195 | 357 | 44151.8789
358 | 12.6294 | 0.66657 | 0.0054994 | 358 | 0.67268 | -0.027438 | 358 | 0.67229 | -0.025343 |  | 358 | 12.6294 | 0.66495 | 0.014272 |  | 358 | 0.66546 | 0.011504 | 358 | 0.67777 | -0.054394 | 358 | 43547.4414 | 358 | 43514.5195 | 358 | 44172.1602
359 | 12.6647 | 0.6666 | 0.0052974 | 359 | 0.67266 | -0.027321 | 359 | 0.67211 | -0.02435 |  | 359 | 12.6647 | 0.66509 | 0.013561 |  | 359 | 0.66485 | 0.014855 | 359 | 0.67731 | -0.051983 | 359 | 43551.9219 | 359 | 43498.0781 | 359 | 44198.5586
360 | 12.7 | 0.66633 | 0.0067789 | 360 | 0.67275 | -0.027806 | 360 | 0.67201 | -0.023853 |  | 360 | 12.7 | 0.66469 | 0.015699 |  | 360 | 0.66486 | 0.014802 | 360 | 0.67733 | -0.052057 | 360 | 43586.8789 | 360 | 43518.5195 | 360 | 44282.2383
361 | 12.7353 | 0.66617 | 0.0076301 | 361 | 0.67283 | -0.028195 | 361 | 0.6718 | -0.022721 |  | 361 | 12.7353 | 0.66471 | 0.015635 |  | 361 | 0.66521 | 0.012884 | 361 | 0.6773 | -0.05194 | 361 | 43567.4805 | 361 | 43561.3984 | 361 | 44361.3203
362 | 12.7706 | 0.66685 | 0.0039468 | 362 | 0.67315 | -0.029933 | 362 | 0.67215 | -0.024581 |  | 362 | 12.7706 | 0.66548 | 0.011394 |  | 362 | 0.66548 | 0.011413 | 362 | 0.67675 | -0.049013 | 362 | 43628.8398 | 362 | 43589.6797 | 362 | 44387.1602
363 | 12.8058 | 0.66598 | 0.0086837 | 363 | 0.67259 | -0.026952 | 363 | 0.67122 | -0.019623 |  | 363 | 12.8058 | 0.66599 | 0.0086081 |  | 363 | 0.66518 | 0.013066 | 363 | 0.67683 | -0.049427 | 363 | 43586.2812 | 363 | 43580.0391 | 363 | 44413.4805
364 | 12.8411 | 0.66571 | 0.010178 | 364 | 0.67331 | -0.030781 | 364 | 0.67146 | -0.02087 |  | 364 | 12.8411 | 0.66633 | 0.0067902 |  | 364 | 0.66464 | 0.015995 | 364 | 0.67722 | -0.051483 | 364 | 43609.8398 | 364 | 43585.4805 | 364 | 44471.3594
365 | 12.8764 | 0.66553 | 0.011154 | 365 | 0.67337 | -0.031075 | 365 | 0.67124 | -0.019737 |  | 365 | 12.8764 | 0.66652 | 0.0057464 |  | 365 | 0.66455 | 0.016461 | 365 | 0.67661 | -0.048281 | 365 | 43599.3594 | 365 | 43622.6016 | 365 | 44448.9219
366 | 12.9117 | 0.66564 | 0.01055 | 366 | 0.67294 | -0.028832 | 366 | 0.67077 | -0.01721 |  | 366 | 12.9117 | 0.66712 | 0.0024902 |  | 366 | 0.6656 | 0.01073 | 366 | 0.67562 | -0.043029 | 366 | 43621.1211 | 366 | 43577.6797 | 366 | 44456.6016
367 | 12.9469 | 0.66647 | 0.0060362 | 367 | 0.67388 | -0.033798 | 367 | 0.67133 | -0.020213 |  | 367 | 12.9469 | 0.66672 | 0.0046767 |  | 367 | 0.6658 | 0.0096723 | 367 | 0.67445 | -0.036839 | 367 | 43581.9219 | 367 | 43634.6406 | 367 | 44403.3984
368 | 12.9822 | 0.66474 | 0.015456 | 368 | 0.67223 | -0.025017 | 368 | 0.6702 | -0.01412 |  | 368 | 12.9822 | 0.66611 | 0.0080044 |  | 368 | 0.66684 | 0.0040141 | 368 | 0.67373 | -0.03304 | 368 | 43625 | 368 | 43588.1211 | 368 | 44379.5586
369 | 13.0175 | 0.66482 | 0.015028 | 369 | 0.67305 | -0.029374 | 369 | 0.6705 | -0.015743 |  | 369 | 13.0175 | 0.66741 | 0.00092995 |  | 369 | 0.6669 | 0.003713 | 369 | 0.67395 | -0.034169 | 369 | 43690.6406 | 369 | 43580.7617 | 369 | 44363.9219
370 | 13.0528 | 0.66493 | 0.014398 | 370 | 0.67365 | -0.032612 | 370 | 0.67067 | -0.016679 |  | 370 | 13.0528 | 0.66674 | 0.0045337 |  | 370 | 0.66687 | 0.0038559 | 370 | 0.67408 | -0.034867 | 370 | 43700.9219 | 370 | 43657.3984 | 370 | 44385.3203
371 | 13.0881 | 0.66535 | 0.012112 | 371 | 0.67507 | -0.040138 | 371 | 0.67066 | -0.01663 |  | 371 | 13.0881 | 0.66669 | 0.0048158 |  | 371 | 0.66673 | 0.0045977 | 371 | 0.67487 | -0.039085 | 371 | 43707.9609 | 371 | 43640.5586 | 371 | 44336.3203
372 | 13.1233 | 0.66617 | 0.0076479 | 372 | 0.67438 | -0.036479 | 372 | 0.6704 | -0.015208 |  | 372 | 13.1233 | 0.66595 | 0.0088721 |  | 372 | 0.66629 | 0.0070164 | 372 | 0.67634 | -0.046838 | 372 | 43701.1602 | 372 | 43679.6016 | 372 | 44342.3203
373 | 13.1586 | 0.66622 | 0.0073955 | 373 | 0.67374 | -0.033075 | 373 | 0.67021 | -0.014199 |  | 373 | 13.1586 | 0.66677 | 0.0044169 |  | 373 | 0.66613 | 0.0078498 | 373 | 0.67689 | -0.049775 | 373 | 43696.2383 | 373 | 43623.9609 | 373 | 44350.7188
374 | 13.1939 | 0.66588 | 0.0092249 | 374 | 0.67366 | -0.032668 | 374 | 0.6699 | -0.012522 |  | 374 | 13.1939 | 0.66678 | 0.0043528 |  | 374 | 0.66707 | 0.0027458 | 374 | 0.67673 | -0.048929 | 374 | 43653.0781 | 374 | 43659.3203 | 374 | 44345.1602
375 | 13.2292 | 0.666 | 0.0085572 | 375 | 0.67368 | -0.032746 | 375 | 0.66935 | -0.0095783 |  | 375 | 13.2292 | 0.66666 | 0.0049666 |  | 375 | 0.66674 | 0.0045638 | 375 | 0.6766 | -0.048255 | 375 | 43620.7188 | 375 | 43612.1602 | 375 | 44335.8789
376 | 13.2644 | 0.66637 | 0.0065483 | 376 | 0.6738 | -0.03341 | 376 | 0.66897 | -0.0075361 |  | 376 | 13.2644 | 0.66685 | 0.0039652 |  | 376 | 0.66728 | 0.0016441 | 376 | 0.67651 | -0.047743 | 376 | 43586.1992 | 376 | 43609.5195 | 376 | 44335.4805
377 | 13.2997 | 0.66634 | 0.0067248 | 377 | 0.67377 | -0.033219 | 377 | 0.6693 | -0.0092995 |  | 377 | 13.2997 | 0.66688 | 0.003811 |  | 377 | 0.6675 | 0.0004191 | 377 | 0.67651 | -0.047746 | 377 | 43567.2383 | 377 | 43605.8398 | 377 | 44281.6016
378 | 13.335 | 0.66608 | 0.0081315 | 378 | 0.67338 | -0.031138 | 378 | 0.66997 | -0.012875 |  | 378 | 13.335 | 0.66633 | 0.0067828 |  | 378 | 0.66687 | 0.0038446 | 378 | 0.67713 | -0.051026 | 378 | 43563.7617 | 378 | 43598.0391 | 378 | 44278.5195
379 | 13.3703 | 0.66672 | 0.0046922 | 379 | 0.67405 | -0.034716 | 379 | 0.66975 | -0.011711 |  | 379 | 13.3703 | 0.66653 | 0.0056972 |  | 379 | 0.66768 | -0.00054249 | 379 | 0.67773 | -0.054208 | 379 | 43567.9219 | 379 | 43525.1602 | 379 | 44303.9609
380 | 13.4056 | 0.66614 | 0.0078177 | 380 | 0.67427 | -0.035912 | 380 | 0.66924 | -0.0089921 |  | 380 | 13.4056 | 0.66707 | 0.0027423 |  | 380 | 0.66808 | -0.0026931 | 380 | 0.67735 | -0.052188 | 380 | 43557.3203 | 380 | 43502.4805 | 380 | 44285.3594
381 | 13.4408 | 0.6661 | 0.0080557 | 381 | 0.67444 | -0.036775 | 381 | 0.66916 | -0.0085342 |  | 381 | 13.4408 | 0.66746 | 0.00065933 |  | 381 | 0.6679 | -0.0017139 | 381 | 0.6779 | -0.055081 | 381 | 43519.8789 | 381 | 43522.6406 | 381 | 44240.5586
382 | 13.4761 | 0.66648 | 0.0059822 | 382 | 0.67459 | -0.037578 | 382 | 0.66919 | -0.0086772 |  | 382 | 13.4761 | 0.66802 | -0.002363 |  | 382 | 0.66717 | 0.0022119 | 382 | 0.67739 | -0.052393 | 382 | 43540.3203 | 382 | 43558.1992 | 382 | 44157.4805
383 | 13.5114 | 0.66638 | 0.006512 | 383 | 0.67441 | -0.036648 | 383 | 0.66975 | -0.011711 |  | 383 | 13.5114 | 0.66683 | 0.0040555 |  | 383 | 0.66743 | 0.00082112 | 383 | 0.67715 | -0.05115 | 383 | 43564.7188 | 383 | 43567.1992 | 383 | 44086
384 | 13.5467 | 0.66683 | 0.0040623 | 384 | 0.675 | -0.039772 | 384 | 0.66998 | -0.012932 |  | 384 | 13.5467 | 0.66617 | 0.0076346 |  | 384 | 0.66773 | -0.00081677 | 384 | 0.67682 | -0.049409 | 384 | 43561.2383 | 384 | 43547.0391 | 384 | 44143.4805
385 | 13.5819 | 0.66573 | 0.010041 | 385 | 0.67462 | -0.037748 | 385 | 0.66979 | -0.011926 |  | 385 | 13.5819 | 0.66606 | 0.0082268 |  | 385 | 0.66685 | 0.0039614 | 385 | 0.67598 | -0.044964 | 385 | 43571.0391 | 385 | 43523.2812 | 385 | 44152.6797
386 | 13.6172 | 0.66466 | 0.015887 | 386 | 0.67278 | -0.027951 | 386 | 0.66991 | -0.012554 |  | 386 | 13.6172 | 0.66601 | 0.0085251 |  | 386 | 0.66594 | 0.0088948 | 386 | 0.67591 | -0.044586 | 386 | 43627.5586 | 386 | 43526.5195 | 386 | 44215.1992
387 | 13.6525 | 0.66602 | 0.0084742 | 387 | 0.67356 | -0.032125 | 387 | 0.66937 | -0.0096891 |  | 387 | 13.6525 | 0.66564 | 0.010545 |  | 387 | 0.66542 | 0.011712 | 387 | 0.67555 | -0.042691 | 387 | 43644.0391 | 387 | 43569.4805 | 387 | 44275.1992
388 | 13.6878 | 0.66632 | 0.0068149 | 388 | 0.67428 | -0.035929 | 388 | 0.66938 | -0.0097463 |  | 388 | 13.6878 | 0.66593 | 0.0089778 |  | 388 | 0.66605 | 0.0083325 | 388 | 0.676 | -0.045077 | 388 | 43645.9219 | 388 | 43581.1211 | 388 | 44323.3594
389 | 13.7231 | 0.66634 | 0.0067394 | 389 | 0.67486 | -0.039019 | 389 | 0.66941 | -0.0099001 |  | 389 | 13.7231 | 0.66528 | 0.012513 |  | 389 | 0.66736 | 0.0011932 | 389 | 0.67652 | -0.04779 | 389 | 43654.6406 | 389 | 43648.5195 | 389 | 44328.9219
390 | 13.7583 | 0.66516 | 0.01315 | 390 | 0.67431 | -0.036084 | 390 | 0.67036 | -0.01498 |  | 390 | 13.7583 | 0.6658 | 0.0096762 |  | 390 | 0.66842 | -0.0045264 | 390 | 0.67653 | -0.047856 | 390 | 43650.1992 | 390 | 43661.5586 | 390 | 44290.8008
391 | 13.7936 | 0.66565 | 0.010504 | 391 | 0.6745 | -0.037131 | 391 | 0.67169 | -0.022103 |  | 391 | 13.7936 | 0.66607 | 0.0081967 |  | 391 | 0.66886 | -0.0069196 | 391 | 0.67637 | -0.046995 | 391 | 43656.8008 | 391 | 43666.1602 | 391 | 44276.5586
392 | 13.8289 | 0.6658 | 0.009669 | 392 | 0.67511 | -0.040339 | 392 | 0.67108 | -0.018844 |  | 392 | 13.8289 | 0.66547 | 0.011459 |  | 392 | 0.66863 | -0.0056616 | 392 | 0.67654 | -0.047915 | 392 | 43685.6406 | 392 | 43659.9219 | 392 | 44275.6797
393 | 13.8642 | 0.66597 | 0.0087595 | 393 | 0.67594 | -0.044762 | 393 | 0.66992 | -0.012626 |  | 393 | 13.8642 | 0.66627 | 0.0071181 |  | 393 | 0.66884 | -0.0067961 | 393 | 0.67606 | -0.045375 | 393 | 43649.8008 | 393 | 43651.6016 | 393 | 44307
394 | 13.8994 | 0.66606 | 0.0082327 | 394 | 0.67497 | -0.039593 | 394 | 0.67059 | -0.016252 |  | 394 | 13.8994 | 0.66687 | 0.0038672 |  | 394 | 0.66838 | -0.0043277 | 394 | 0.67602 | -0.045173 | 394 | 43598.1211 | 394 | 43647.2383 | 394 | 44351.8008
395 | 13.9347 | 0.66663 | 0.0051751 | 395 | 0.67622 | -0.046204 | 395 | 0.66934 | -0.0095282 |  | 395 | 13.9347 | 0.66589 | 0.0091666 |  | 395 | 0.66743 | 0.00080207 | 395 | 0.67603 | -0.045243 | 395 | 43621.9609 | 395 | 43685.6406 | 395 | 44329
396 | 13.97 | 0.66622 | 0.0074029 | 396 | 0.6747 | -0.038195 | 396 | 0.66881 | -0.0066269 |  | 396 | 13.97 | 0.66627 | 0.0071369 |  | 396 | 0.66678 | 0.004334 | 396 | 0.67594 | -0.044744 | 396 | 43645.5586 | 396 | 43692.4414 | 396 | 44385.8789
397 | 14.0053 | 0.66662 | 0.0052291 | 397 | 0.67619 | -0.046043 | 397 | 0.67178 | -0.0226 |  | 397 | 14.0053 | 0.6651 | 0.013497 |  | 397 | 0.66651 | 0.0058329 | 397 | 0.67602 | -0.045173 | 397 | 43688.3984 | 397 | 43672.5586 | 397 | 44349.7188
398 | 14.0406 | 0.6673 | 0.0015146 | 398 | 0.6765 | -0.047722 | 398 | 0.67279 | -0.028032 |  | 398 | 14.0406 | 0.66469 | 0.015722 |  | 398 | 0.66704 | 0.0029417 | 398 | 0.67592 | -0.044641 | 398 | 43660.2383 | 398 | 43712.0391 | 398 | 44322.0391
399 | 14.0758 | 0.66672 | 0.0046492 | 399 | 0.67621 | -0.046183 | 399 | 0.67031 | -0.014737 |  | 399 | 14.0758 | 0.66356 | 0.021905 |  | 399 | 0.6674 | 0.00095644 | 399 | 0.67505 | -0.040038 | 399 | 43642.3984 | 399 | 43694.9219 | 399 | 44303.2383
400 | 14.1111 | 0.66651 | 0.0057986 | 400 | 0.67641 | -0.047228 | 400 | 0.67081 | -0.017434 |  | 400 | 14.1111 | 0.66306 | 0.02462 |  | 400 | 0.66772 | -0.00074546 | 400 | 0.67573 | -0.043635 | 400 | 43666.7188 | 400 | 43743.5195 | 400 | 44254.5195
401 | 14.1464 | 0.66662 | 0.0051822 | 401 | 0.67635 | -0.04693 | 401 | 0.67102 | -0.018556 |  | 401 | 14.1464 | 0.66331 | 0.023288 |  | 401 | 0.66731 | 0.0014674 | 401 | 0.67554 | -0.042647 | 401 | 43688 | 401 | 43787.6406 | 401 | 44242.7617
402 | 14.1817 | 0.6666 | 0.0053444 | 402 | 0.67625 | -0.0464 | 402 | 0.67221 | -0.024889 |  | 402 | 14.1817 | 0.66305 | 0.024677 |  | 402 | 0.66717 | 0.0022158 | 402 | 0.67558 | -0.042846 | 402 | 43680.1211 | 402 | 43786.3203 | 402 | 44236.6797
403 | 14.2169 | 0.66669 | 0.004833 | 403 | 0.6772 | -0.051373 | 403 | 0.67244 | -0.026155 |  | 403 | 14.2169 | 0.66373 | 0.020982 |  | 403 | 0.66682 | 0.0041008 | 403 | 0.67509 | -0.040255 | 403 | 43632.9219 | 403 | 43733.6406 | 403 | 44230.3203
404 | 14.2522 | 0.66704 | 0.0029498 | 404 | 0.67718 | -0.051307 | 404 | 0.67254 | -0.026686 |  | 404 | 14.2522 | 0.66394 | 0.019828 |  | 404 | 0.66641 | 0.006353 | 404 | 0.67567 | -0.043323 | 404 | 43631.4805 | 404 | 43719.7617 | 404 | 44276.2383
405 | 14.2875 | 0.66687 | 0.0038245 | 405 | 0.6769 | -0.049823 | 405 | 0.67293 | -0.028754 |  | 405 | 14.2875 | 0.66391 | 0.019968 |  | 405 | 0.66624 | 0.0072953 | 405 | 0.67533 | -0.041494 | 405 | 43647.5195 | 405 | 43697.6406 | 405 | 44302
406 | 14.3228 | 0.66704 | 0.0029463 | 406 | 0.67744 | -0.052656 | 406 | 0.67303 | -0.029289 |  | 406 | 14.3228 | 0.66505 | 0.013735 |  | 406 | 0.66673 | 0.0046165 | 406 | 0.67534 | -0.041552 | 406 | 43678.1992 | 406 | 43665.3203 | 406 | 44284.4414
407 | 14.3581 | 0.66784 | -0.0013884 | 407 | 0.67784 | -0.054763 | 407 | 0.67311 | -0.029692 |  | 407 | 14.3581 | 0.66501 | 0.013947 |  | 407 | 0.66719 | 0.0020991 | 407 | 0.67541 | -0.041953 | 407 | 43732.3203 | 407 | 43627.5586 | 407 | 44261.6016
408 | 14.3933 | 0.66798 | -0.0021853 | 408 | 0.67759 | -0.053432 | 408 | 0.67245 | -0.026208 |  | 408 | 14.3933 | 0.66596 | 0.0087777 |  | 408 | 0.66757 | 5.0742e-05 | 408 | 0.67509 | -0.040233 | 408 | 43656.9609 | 408 | 43618.3984 | 408 | 44246.3203
409 | 14.4286 | 0.66856 | -0.0052801 | 409 | 0.67886 | -0.060145 | 409 | 0.67328 | -0.030594 |  | 409 | 14.4286 | 0.66689 | 0.0037356 |  | 409 | 0.66796 | -0.0020441 | 409 | 0.6748 | -0.038687 | 409 | 43620.6406 | 409 | 43648.8398 | 409 | 44233.8398
410 | 14.4639 | 0.66872 | -0.0061613 | 410 | 0.67862 | -0.058877 | 410 | 0.6722 | -0.024871 |  | 410 | 14.4639 | 0.66701 | 0.0031147 |  | 410 | 0.6682 | -0.0033719 | 410 | 0.67481 | -0.038779 | 410 | 43669 | 410 | 43723.2383 | 410 | 44248.2812
411 | 14.4992 | 0.66931 | -0.0093316 | 411 | 0.67836 | -0.057491 | 411 | 0.67222 | -0.024967 |  | 411 | 14.4992 | 0.66854 | -0.0052158 |  | 411 | 0.66836 | -0.0042153 | 411 | 0.67508 | -0.040166 | 411 | 43647.2383 | 411 | 43775 | 411 | 44234.5195
412 | 14.5344 | 0.66944 | -0.010054 | 412 | 0.67868 | -0.059184 | 412 | 0.67209 | -0.024268 |  | 412 | 14.5344 | 0.66803 | -0.0024568 |  | 412 | 0.6684 | -0.0044291 | 412 | 0.67524 | -0.041056 | 412 | 43657.0391 | 412 | 43780.0781 | 412 | 44242.8789
413 | 14.5697 | 0.66901 | -0.0077078 | 413 | 0.6787 | -0.059306 | 413 | 0.67196 | -0.02357 |  | 413 | 14.5697 | 0.66804 | -0.0024794 |  | 413 | 0.66841 | -0.0044739 | 413 | 0.67562 | -0.043051 | 413 | 43642.7617 | 413 | 43803 | 413 | 44250.9219
414 | 14.605 | 0.66993 | -0.012679 | 414 | 0.67907 | -0.061227 | 414 | 0.67178 | -0.022583 |  | 414 | 14.605 | 0.66716 | 0.0022607 |  | 414 | 0.66859 | -0.0054404 | 414 | 0.67563 | -0.043095 | 414 | 43676.7617 | 414 | 43807.9219 | 414 | 44229.8008
415 | 14.6403 | 0.67026 | -0.014455 | 415 | 0.6783 | -0.057205 | 415 | 0.67188 | -0.023137 |  | 415 | 14.6403 | 0.66764 | -0.0003098 |  | 415 | 0.66913 | -0.0083674 | 415 | 0.67564 | -0.04318 | 415 | 43657.4414 | 415 | 43780.5195 | 415 | 44255.7617
416 | 14.6756 | 0.67023 | -0.014295 | 416 | 0.67826 | -0.056978 | 416 | 0.67141 | -0.020608 |  | 416 | 14.6756 | 0.66719 | 0.0021405 |  | 416 | 0.66944 | -0.010031 | 416 | 0.67565 | -0.043224 | 416 | 43710.3984 | 416 | 43748.1602 | 416 | 44249.1602
417 | 14.7108 | 0.66972 | -0.011565 | 417 | 0.67751 | -0.053044 | 417 | 0.67102 | -0.018563 |  | 417 | 14.7108 | 0.66678 | 0.0043567 |  | 417 | 0.66944 | -0.010042 | 417 | 0.67555 | -0.042684 | 417 | 43688.7617 | 417 | 43717.5586 | 417 | 44238.1992
418 | 14.7461 | 0.66997 | -0.012918 | 418 | 0.67875 | -0.059525 | 418 | 0.6712 | -0.019491 |  | 418 | 14.7461 | 0.6676 | -0.00011066 |  | 418 | 0.66915 | -0.0084982 | 418 | 0.67617 | -0.045947 | 418 | 43602.8008 | 418 | 43685.1992 | 418 | 44193.5586
419 | 14.7814 | 0.66952 | -0.010501 | 419 | 0.67853 | -0.0584 | 419 | 0.67106 | -0.018744 |  | 419 | 14.7814 | 0.66762 | -0.00020458 |  | 419 | 0.66849 | -0.0049237 | 419 | 0.67632 | -0.04676 | 419 | 43627.2383 | 419 | 43700.8008 | 419 | 44233.6016
420 | 14.8167 | 0.66873 | -0.0062151 | 420 | 0.67784 | -0.054769 | 420 | 0.67093 | -0.018068 |  | 420 | 14.8167 | 0.66754 | 0.0002425 |  | 420 | 0.66904 | -0.0078849 | 420 | 0.67719 | -0.051355 | 420 | 43552.0391 | 420 | 43736.1992 | 420 | 44225.3594
421 | 14.8519 | 0.66908 | -0.0081086 | 421 | 0.67889 | -0.060274 | 421 | 0.67136 | -0.020337 |  | 421 | 14.8519 | 0.66763 | -0.00024202 |  | 421 | 0.66883 | -0.0067285 | 421 | 0.67834 | -0.057403 | 421 | 43533.7617 | 421 | 43756.0391 | 421 | 44213.4805
422 | 14.8872 | 0.66788 | -0.0016468 | 422 | 0.6781 | -0.056113 | 422 | 0.67024 | -0.01437 |  | 422 | 14.8872 | 0.66729 | 0.0015989 |  | 422 | 0.66855 | -0.0052608 | 422 | 0.67806 | -0.055934 | 422 | 43511.3984 | 422 | 43709.7188 | 422 | 44210.5195
423 | 14.9225 | 0.66882 | -0.0066947 | 423 | 0.67907 | -0.061248 | 423 | 0.6713 | -0.020024 |  | 423 | 14.9225 | 0.66762 | -0.0002233 |  | 423 | 0.66924 | -0.0089731 | 423 | 0.67764 | -0.053737 | 423 | 43455.8789 | 423 | 43706.1602 | 423 | 44208.6016
424 | 14.9578 | 0.66886 | -0.0069167 | 424 | 0.67891 | -0.060392 | 424 | 0.67086 | -0.017659 |  | 424 | 14.9578 | 0.66708 | 0.0027045 |  | 424 | 0.67026 | -0.014465 | 424 | 0.67768 | -0.053912 | 424 | 43473.6016 | 424 | 43693.4805 | 424 | 44189.6797
425 | 14.9931 | 0.66944 | -0.010039 | 425 | 0.68006 | -0.066439 | 425 | 0.67128 | -0.01995 |  | 425 | 14.9931 | 0.66685 | 0.003984 |  | 425 | 0.67064 | -0.016522 | 425 | 0.67741 | -0.052503 | 425 | 43491.7617 | 425 | 43656.3203 | 425 | 44150.8008
426 | 15.0283 | 0.66813 | -0.0029497 | 426 | 0.67955 | -0.063758 | 426 | 0.67033 | -0.014819 |  | 426 | 15.0283 | 0.66615 | 0.0077516 |  | 426 | 0.67113 | -0.019142 | 426 | 0.67768 | -0.053923 | 426 | 43523.3594 | 426 | 43635.1211 | 426 | 44162.5195
427 | 15.0636 | 0.66831 | -0.0039325 | 427 | 0.68033 | -0.067847 | 427 | 0.67093 | -0.018061 |  | 427 | 15.0636 | 0.66574 | 0.0099896 |  | 427 | 0.67086 | -0.017687 | 427 | 0.67713 | -0.051033 | 427 | 43539.0781 | 427 | 43662.0781 | 427 | 44196.0781
428 | 15.0989 | 0.66762 | -0.00019264 | 428 | 0.67964 | -0.064224 | 428 | 0.67029 | -0.014644 |  | 428 | 15.0989 | 0.66524 | 0.012695 |  | 428 | 0.67036 | -0.015013 | 428 | 0.67626 | -0.046438 | 428 | 43590.7188 | 428 | 43707.3594 | 428 | 44234.0391
429 | 15.1342 | 0.66838 | -0.0043122 | 429 | 0.68057 | -0.069091 | 429 | 0.67062 | -0.016377 |  | 429 | 15.1342 | 0.66581 | 0.0095853 |  | 429 | 0.67021 | -0.014185 | 429 | 0.67597 | -0.044912 | 429 | 43599.8398 | 429 | 43731.2383 | 429 | 44239.6797
430 | 15.1694 | 0.66888 | -0.0070026 | 430 | 0.68116 | -0.072175 | 430 | 0.67087 | -0.017733 |  | 430 | 15.1694 | 0.66583 | 0.009525 |  | 430 | 0.66936 | -0.0096236 | 430 | 0.67563 | -0.043092 | 430 | 43629.3594 | 430 | 43751.1211 | 430 | 44210.1602
431 | 15.2047 | 0.66901 | -0.0077184 | 431 | 0.68148 | -0.073811 | 431 | 0.67122 | -0.019608 |  | 431 | 15.2047 | 0.66615 | 0.0077895 |  | 431 | 0.66881 | -0.0066462 | 431 | 0.67513 | -0.04045 | 431 | 43701.8398 | 431 | 43766.9609 | 431 | 44242.8398
432 | 15.24 | 0.66855 | -0.0052335 | 432 | 0.68044 | -0.068377 | 432 | 0.67026 | -0.014473 |  | 432 | 15.24 | 0.66658 | 0.0054111 |  | 432 | 0.66898 | -0.0075894 | 432 | 0.67528 | -0.041226 | 432 | 43728.7188 | 432 | 43767.8398 | 432 | 44241.2812
433 | 15.2753 | 0.66754 | 0.00020989 | 433 | 0.67979 | -0.065002 | 433 | 0.66913 | -0.0083519 |  | 433 | 15.2753 | 0.66655 | 0.0056144 |  | 433 | 0.66885 | -0.0068746 | 433 | 0.67598 | -0.044971 | 433 | 43736.1992 | 433 | 43775.7188 | 433 | 44276.9219
434 | 15.3106 | 0.66718 | 0.0021873 | 434 | 0.67935 | -0.062674 | 434 | 0.66928 | -0.0091957 |  | 434 | 15.3106 | 0.66606 | 0.008246 |  | 434 | 0.66889 | -0.0070691 | 434 | 0.67653 | -0.047864 | 434 | 43779.8789 | 434 | 43813.4805 | 434 | 44292.0781
435 | 15.3458 | 0.66859 | -0.005452 | 435 | 0.68113 | -0.071985 | 435 | 0.67104 | -0.01863 |  | 435 | 15.3458 | 0.6662 | 0.0074687 |  | 435 | 0.66943 | -0.0099672 | 435 | 0.67682 | -0.049372 | 435 | 43766.2383 | 435 | 43835.0781 | 435 | 44354.0781
436 | 15.3811 | 0.66849 | -0.0048967 | 436 | 0.68101 | -0.071373 | 436 | 0.67081 | -0.017399 |  | 436 | 15.3811 | 0.66623 | 0.0073255 |  | 436 | 0.66994 | -0.012749 | 436 | 0.67657 | -0.048054 | 436 | 43799.7617 | 436 | 43872.6797 | 436 | 44413.3984
437 | 15.4164 | 0.66901 | -0.0077293 | 437 | 0.68167 | -0.074809 | 437 | 0.67122 | -0.019619 |  | 437 | 15.4164 | 0.66479 | 0.015166 |  | 437 | 0.67101 | -0.018502 | 437 | 0.67583 | -0.044138 | 437 | 43738.4805 | 437 | 43857.3203 | 437 | 44414
438 | 15.4517 | 0.66883 | -0.0067449 | 438 | 0.68221 | -0.077622 | 438 | 0.67127 | -0.019882 |  | 438 | 15.4517 | 0.6655 | 0.011319 |  | 438 | 0.6719 | -0.023276 | 438 | 0.67565 | -0.043195 | 438 | 43705.6016 | 438 | 43842.4805 | 438 | 44362.8398
439 | 15.4869 | 0.66839 | -0.0043554 | 439 | 0.68227 | -0.077953 | 439 | 0.67023 | -0.014281 |  | 439 | 15.4869 | 0.66531 | 0.012339 |  | 439 | 0.67268 | -0.027405 | 439 | 0.67532 | -0.041449 | 439 | 43703.6797 | 439 | 43861 | 439 | 44384.4805
440 | 15.5222 | 0.66915 | -0.0084914 | 440 | 0.68224 | -0.07777 | 440 | 0.67002 | -0.013154 |  | 440 | 15.5222 | 0.6666 | 0.0053169 |  | 440 | 0.6737 | -0.032859 | 440 | 0.67491 | -0.039313 | 440 | 43732.9609 | 440 | 43821.4805 | 440 | 44375.4805
441 | 15.5575 | 0.67019 | -0.014074 | 441 | 0.68353 | -0.084449 | 441 | 0.67054 | -0.015974 |  | 441 | 15.5575 | 0.66659 | 0.0053657 |  | 441 | 0.6735 | -0.031773 | 441 | 0.675 | -0.039747 | 441 | 43701.4414 | 441 | 43799.3203 | 441 | 44367.6406
442 | 15.5928 | 0.66907 | -0.0080476 | 442 | 0.68181 | -0.075538 | 442 | 0.66902 | -0.0077579 |  | 442 | 15.5928 | 0.66691 | 0.0036602 |  | 442 | 0.67259 | -0.026931 | 442 | 0.67523 | -0.040964 | 442 | 43704.3594 | 442 | 43846.0781 | 442 | 44346.6797
443 | 15.6281 | 0.669 | -0.0076971 | 443 | 0.68195 | -0.076252 | 443 | 0.66894 | -0.0073499 |  | 443 | 15.6281 | 0.66723 | 0.0018997 |  | 443 | 0.67239 | -0.025883 | 443 | 0.67567 | -0.043305 | 443 | 43685.6797 | 443 | 43795.7617 | 443 | 44310.8398
444 | 15.6633 | 0.6693 | -0.0092638 | 444 | 0.68286 | -0.080996 | 444 | 0.66911 | -0.0082658 |  | 444 | 15.6633 | 0.66718 | 0.0021666 |  | 444 | 0.67261 | -0.027068 | 444 | 0.67519 | -0.040777 | 444 | 43744.1211 | 444 | 43896.8789 | 444 | 44260.4414
445 | 15.6986 | 0.66941 | -0.0098895 | 445 | 0.6829 | -0.081192 | 445 | 0.66953 | -0.010547 |  | 445 | 15.6986 | 0.66805 | -0.0025242 |  | 445 | 0.67196 | -0.023565 | 445 | 0.67505 | -0.040012 | 445 | 43723.2812 | 445 | 43905.3984 | 445 | 44231.6797
446 | 15.7339 | 0.66973 | -0.011608 | 446 | 0.68351 | -0.084363 | 446 | 0.66957 | -0.010761 |  | 446 | 15.7339 | 0.66808 | -0.0026818 |  | 446 | 0.67217 | -0.024686 | 446 | 0.67294 | -0.028793 | 446 | 43706.2383 | 446 | 43935.7617 | 446 | 44208.5586
447 | 15.7692 | 0.66991 | -0.012565 | 447 | 0.68424 | -0.08816 | 447 | 0.66919 | -0.0086701 |  | 447 | 15.7692 | 0.6673 | 0.0015427 |  | 447 | 0.67133 | -0.020205 | 447 | 0.67384 | -0.033608 | 447 | 43678.6797 | 447 | 43990.8398 | 447 | 44229.3594
448 | 15.8044 | 0.6699 | -0.012543 | 448 | 0.68364 | -0.085033 | 448 | 0.66936 | -0.0095963 |  | 448 | 15.8044 | 0.66632 | 0.0068544 |  | 448 | 0.67153 | -0.021249 | 448 | 0.67363 | -0.032508 | 448 | 43665.1992 | 448 | 43985.8789 | 448 | 44270.2812
449 | 15.8397 | 0.67058 | -0.016184 | 449 | 0.68312 | -0.082362 | 449 | 0.66902 | -0.0077794 |  | 449 | 15.8397 | 0.66637 | 0.0065528 |  | 449 | 0.67124 | -0.019703 | 449 | 0.67302 | -0.029256 | 449 | 43616 | 449 | 43952.2383 | 449 | 44305.9219
450 | 15.875 | 0.67029 | -0.014591 | 450 | 0.68268 | -0.080076 | 450 | 0.66827 | -0.003735 |  | 450 | 15.875 | 0.666 | 0.0085666 |  | 450 | 0.67059 | -0.016216 | 450 | 0.67358 | -0.032228 | 450 | 43560.9609 | 450 | 43950.4805 | 450 | 44301.3984
451 | 15.9103 | 0.67104 | -0.01863 | 451 | 0.6835 | -0.084322 | 451 | 0.66964 | -0.011136 |  | 451 | 15.9103 | 0.66579 | 0.009744 |  | 451 | 0.66984 | -0.012182 | 451 | 0.67355 | -0.032061 | 451 | 43604.2812 | 451 | 43885.6406 | 451 | 44303
452 | 15.9456 | 0.67061 | -0.016327 | 452 | 0.68367 | -0.085174 | 452 | 0.66914 | -0.0084091 |  | 452 | 15.9456 | 0.66572 | 0.010106 |  | 452 | 0.66945 | -0.010094 | 452 | 0.67358 | -0.032194 | 452 | 43612.2812 | 452 | 43853.5195 | 452 | 44280.5586
453 | 15.9808 | 0.6703 | -0.014659 | 453 | 0.68334 | -0.083469 | 453 | 0.66877 | -0.0064297 |  | 453 | 15.9808 | 0.66594 | 0.0088909 |  | 453 | 0.66967 | -0.011297 | 453 | 0.67409 | -0.034948 | 453 | 43630.3594 | 453 | 43843.5195 | 453 | 44246
454 | 16.0161 | 0.66997 | -0.012907 | 454 | 0.68369 | -0.085325 | 454 | 0.66887 | -0.0069919 |  | 454 | 16.0161 | 0.66588 | 0.009257 |  | 454 | 0.67018 | -0.014014 | 454 | 0.67422 | -0.035607 | 454 | 43669.7188 | 454 | 43827.6797 | 454 | 44231.5195
455 | 16.0514 | 0.6695 | -0.010343 | 455 | 0.68377 | -0.085693 | 455 | 0.66849 | -0.004918 |  | 455 | 16.0514 | 0.66558 | 0.010851 |  | 455 | 0.67021 | -0.014182 | 455 | 0.67435 | -0.036304 | 455 | 43673.2812 | 455 | 43833.1211 | 455 | 44200.2812
456 | 16.0867 | 0.66947 | -0.010197 | 456 | 0.68347 | -0.084188 | 456 | 0.66863 | -0.0056957 |  | 456 | 16.0867 | 0.66558 | 0.010862 |  | 456 | 0.6695 | -0.010389 | 456 | 0.67416 | -0.035283 | 456 | 43651.3203 | 456 | 43861.1602 | 456 | 44213.8789
457 | 16.1219 | 0.6698 | -0.011983 | 457 | 0.68387 | -0.086211 | 457 | 0.6689 | -0.0071135 |  | 457 | 16.1219 | 0.6656 | 0.01076 |  | 457 | 0.66999 | -0.013007 | 457 | 0.67451 | -0.037188 | 457 | 43659 | 457 | 43889 | 457 | 44212.2383
458 | 16.1572 | 0.67072 | -0.016918 | 458 | 0.68458 | -0.089915 | 458 | 0.66936 | -0.009639 |  | 458 | 16.1572 | 0.6655 | 0.011274 |  | 458 | 0.67046 | -0.015505 | 458 | 0.67503 | -0.039935 | 458 | 43678.9219 | 458 | 43956.1992 | 458 | 44230.1602
459 | 16.1925 | 0.67007 | -0.013425 | 459 | 0.68389 | -0.086359 | 459 | 0.66959 | -0.010876 |  | 459 | 16.1925 | 0.66543 | 0.011682 |  | 459 | 0.67017 | -0.013969 | 459 | 0.67546 | -0.042185 | 459 | 43572.0781 | 459 | 43977.4805 | 459 | 44185.6406
460 | 16.2278 | 0.67013 | -0.013771 | 460 | 0.68347 | -0.08414 | 460 | 0.66925 | -0.0090242 |  | 460 | 16.2278 | 0.6656 | 0.01076 |  | 460 | 0.66974 | -0.011641 | 460 | 0.67494 | -0.039431 | 460 | 43592.3594 | 460 | 44012.9609 | 460 | 44197.9219
461 | 16.2631 | 0.67003 | -0.013221 | 461 | 0.68292 | -0.081333 | 461 | 0.66863 | -0.0056957 |  | 461 | 16.2631 | 0.66614 | 0.007797 |  | 461 | 0.67025 | -0.014413 | 461 | 0.67398 | -0.034369 | 461 | 43601.2812 | 461 | 44072.9609 | 461 | 44048.0391
462 | 16.2983 | 0.66975 | -0.011733 | 462 | 0.68368 | -0.08527 | 462 | 0.66905 | -0.007926 |  | 462 | 16.2983 | 0.66616 | 0.0077328 |  | 462 | 0.67011 | -0.013674 | 462 | 0.67342 | -0.031359 | 462 | 43604.8398 | 462 | 44060.4805 | 462 | 44106.4805
463 | 16.3336 | 0.67113 | -0.01915 | 463 | 0.68535 | -0.093875 | 463 | 0.67065 | -0.016576 |  | 463 | 16.3336 | 0.66603 | 0.0083893 |  | 463 | 0.66993 | -0.012667 | 463 | 0.67242 | -0.026024 | 463 | 43643.0781 | 463 | 44046.0781 | 463 | 44107.6797
464 | 16.3689 | 0.67102 | -0.018552 | 464 | 0.68478 | -0.09095 | 464 | 0.67008 | -0.013482 |  | 464 | 16.3689 | 0.66552 | 0.011213 |  | 464 | 0.66964 | -0.011096 | 464 | 0.6722 | -0.024867 | 464 | 43707.9609 | 464 | 44021.3594 | 464 | 44055.6406
465 | 16.4042 | 0.67082 | -0.017445 | 465 | 0.68394 | -0.086578 | 465 | 0.67002 | -0.013139 |  | 465 | 16.4042 | 0.66632 | 0.0068618 |  | 465 | 0.66947 | -0.010203 | 465 | 0.67232 | -0.025497 | 465 | 43676.8398 | 465 | 44030.8789 | 465 | 44068.0781
466 | 16.4394 | 0.67032 | -0.014794 | 466 | 0.68375 | -0.085627 | 466 | 0.66895 | -0.0073965 |  | 466 | 16.4394 | 0.66634 | 0.0067073 |  | 466 | 0.66942 | -0.0099261 | 466 | 0.67282 | -0.028146 | 466 | 43665.1211 | 466 | 44037.6016 | 466 | 44091.4805
467 | 16.4747 | 0.6708 | -0.01737 | 467 | 0.68478 | -0.090957 | 467 | 0.67038 | -0.015083 |  | 467 | 16.4747 | 0.66617 | 0.0076725 |  | 467 | 0.66893 | -0.0072676 | 467 | 0.67325 | -0.030469 | 467 | 43689.3203 | 467 | 44039.9219 | 467 | 44097.4414
468 | 16.51 | 0.67067 | -0.01664 | 468 | 0.68477 | -0.090898 | 468 | 0.66946 | -0.010157 |  | 468 | 16.51 | 0.66661 | 0.0052379 |  | 468 | 0.66838 | -0.0043167 | 468 | 0.67349 | -0.031714 | 468 | 43715.0391 | 468 | 44022.1211 | 468 | 44104.2812
469 | 16.5453 | 0.67142 | -0.020675 | 469 | 0.68568 | -0.095553 | 469 | 0.67049 | -0.015718 |  | 469 | 16.5453 | 0.66691 | 0.0036263 |  | 469 | 0.66889 | -0.0070769 | 469 | 0.67328 | -0.030635 | 469 | 43694.0391 | 469 | 43973.6797 | 469 | 44102.7188
470 | 16.5806 | 0.67132 | -0.02017 | 470 | 0.68636 | -0.099052 | 470 | 0.67039 | -0.015129 |  | 470 | 16.5806 | 0.6672 | 0.0020729 |  | 470 | 0.66864 | -0.0057253 | 470 | 0.67355 | -0.032058 | 470 | 43675.0781 | 470 | 43925.9219 | 470 | 44133.4805
471 | 16.6158 | 0.67154 | -0.021325 | 471 | 0.68572 | -0.095772 | 471 | 0.66971 | -0.011504 |  | 471 | 16.6158 | 0.66764 | -0.00029851 |  | 471 | 0.66987 | -0.012361 | 471 | 0.67413 | -0.035132 | 471 | 43676.8789 | 471 | 43901.2812 | 471 | 44126.9219
472 | 16.6511 | 0.67174 | -0.022402 | 472 | 0.68644 | -0.099508 | 472 | 0.66991 | -0.012583 |  | 472 | 16.6511 | 0.66738 | 0.0010692 |  | 472 | 0.67061 | -0.01635 | 472 | 0.67399 | -0.034405 | 472 | 43640.5195 | 472 | 43891.1992 | 472 | 44144.7617
473 | 16.6864 | 0.67275 | -0.027788 | 473 | 0.6872 | -0.10337 | 473 | 0.67066 | -0.016598 |  | 473 | 16.6864 | 0.66724 | 0.0018435 |  | 473 | 0.6713 | -0.020045 | 473 | 0.67414 | -0.035176 | 473 | 43627 | 473 | 43874.9219 | 473 | 44182.9219
474 | 16.7217 | 0.67238 | -0.025829 | 474 | 0.68685 | -0.10161 | 474 | 0.66986 | -0.012311 |  | 474 | 16.7217 | 0.6679 | -0.0017065 |  | 474 | 0.67071 | -0.016849 | 474 | 0.67513 | -0.04045 | 474 | 43603.1992 | 474 | 43884.8789 | 474 | 44192.3203
475 | 16.7569 | 0.67187 | -0.023101 | 475 | 0.68602 | -0.097336 | 475 | 0.66919 | -0.0087058 |  | 475 | 16.7569 | 0.66771 | -0.00068157 |  | 475 | 0.67082 | -0.01746 | 475 | 0.6751 | -0.040321 | 475 | 43624.3203 | 475 | 43894.3203 | 475 | 44175.0391
476 | 16.7922 | 0.67231 | -0.02545 | 476 | 0.68616 | -0.098039 | 476 | 0.66965 | -0.011179 |  | 476 | 16.7922 | 0.66729 | 0.0015876 |  | 476 | 0.67036 | -0.014983 | 476 | 0.67476 | -0.038514 | 476 | 43596.3984 | 476 | 43911.6016 | 476 | 44144.1992
477 | 16.8275 | 0.67297 | -0.028978 | 477 | 0.68697 | -0.1022 | 477 | 0.6701 | -0.013596 |  | 477 | 16.8275 | 0.6669 | 0.0037016 |  | 477 | 0.67015 | -0.013887 | 477 | 0.6743 | -0.036039 | 477 | 43547.9219 | 477 | 43894.5586 | 477 | 44080.7617
478 | 16.8628 | 0.673 | -0.029108 | 478 | 0.68679 | -0.1013 | 478 | 0.66947 | -0.010229 |  | 478 | 16.8628 | 0.66621 | 0.0074425 |  | 478 | 0.67034 | -0.014871 | 478 | 0.6742 | -0.035505 | 478 | 43547.1211 | 478 | 43923.2812 | 478 | 44017.8789
479 | 16.8981 | 0.67296 | -0.028893 | 479 | 0.68702 | -0.10248 | 479 | 0.66947 | -0.010232 |  | 479 | 16.8981 | 0.66622 | 0.0073669 |  | 479 | 0.67084 | -0.017583 | 479 | 0.67384 | -0.033597 | 479 | 43576.4805 | 479 | 43933.9219 | 479 | 44005.7617
480 | 16.9333 | 0.67281 | -0.028135 | 480 | 0.68745 | -0.10469 | 480 | 0.66966 | -0.011251 |  | 480 | 16.9333 | 0.66638 | 0.0065 |  | 480 | 0.67078 | -0.017248 | 480 | 0.67371 | -0.032896 | 480 | 43584.3203 | 480 | 43923.1992 | 480 | 44025.4414
481 | 16.9686 | 0.67283 | -0.028217 | 481 | 0.68729 | -0.10384 | 481 | 0.66991 | -0.012593 |  | 481 | 16.9686 | 0.66668 | 0.0049065 |  | 481 | 0.67106 | -0.018766 | 481 | 0.67366 | -0.032667 | 481 | 43630.0781 | 481 | 43920.6797 | 481 | 44029.5195
482 | 17.0039 | 0.67398 | -0.034345 | 482 | 0.68837 | -0.10937 | 482 | 0.66999 | -0.012982 |  | 482 | 17.0039 | 0.66688 | 0.0037958 |  | 482 | 0.67178 | -0.022626 | 482 | 0.67358 | -0.032213 | 482 | 43610.7188 | 482 | 43933 | 482 | 44068.6797
483 | 17.0392 | 0.67362 | -0.032453 | 483 | 0.68834 | -0.10921 | 483 | 0.67007 | -0.013457 |  | 483 | 17.0392 | 0.66694 | 0.0034458 |  | 483 | 0.67277 | -0.027913 | 483 | 0.67393 | -0.034106 | 483 | 43635.0781 | 483 | 43891.4414 | 483 | 44080.1602
484 | 17.0744 | 0.67308 | -0.029561 | 484 | 0.68757 | -0.10527 | 484 | 0.66933 | -0.0094318 |  | 484 | 17.0744 | 0.66632 | 0.0068165 |  | 484 | 0.67297 | -0.028941 | 484 | 0.67409 | -0.034929 | 484 | 43596.7617 | 484 | 43899.1602 | 484 | 44067.3203
485 | 17.1097 | 0.67375 | -0.033117 | 485 | 0.68935 | -0.11438 | 485 | 0.67062 | -0.016369 |  | 485 | 17.1097 | 0.66546 | 0.011538 |  | 485 | 0.67292 | -0.028723 | 485 | 0.67459 | -0.037601 | 485 | 43559.6016 | 485 | 43842.2812 | 485 | 44104.8008
486 | 17.145 | 0.67351 | -0.031831 | 486 | 0.68885 | -0.11185 | 486 | 0.66529 | 0.012441 |  | 486 | 17.145 | 0.66657 | 0.0054564 |  | 486 | 0.67333 | -0.030908 | 486 | 0.67482 | -0.038798 | 486 | 43578.3984 | 486 | 43857.0781 | 486 | 44123.5586
487 | 17.1803 | 0.67373 | -0.032993 | 487 | 0.68806 | -0.1078 | 487 | 0.66857 | -0.0053551 |  | 487 | 17.1803 | 0.66711 | 0.0025729 |  | 487 | 0.67296 | -0.02889 | 487 | 0.67492 | -0.039361 | 487 | 43654.9609 | 487 | 43859.1211 | 487 | 44130
488 | 17.2156 | 0.67393 | -0.034059 | 488 | 0.68854 | -0.11024 | 488 | 0.66984 | -0.012179 |  | 488 | 17.2156 | 0.66822 | -0.0034843 |  | 488 | 0.6726 | -0.026998 | 488 | 0.67464 | -0.037829 | 488 | 43631.6797 | 488 | 43808.1211 | 488 | 44132.3203
489 | 17.2508 | 0.67402 | -0.034536 | 489 | 0.68889 | -0.11203 | 489 | 0.67041 | -0.015236 |  | 489 | 17.2508 | 0.66806 | -0.0025954 |  | 489 | 0.67242 | -0.026053 | 489 | 0.67414 | -0.035198 | 489 | 43666.1211 | 489 | 43785.7617 | 489 | 44188.3594
490 | 17.2861 | 0.67418 | -0.03539 | 490 | 0.68899 | -0.11256 | 490 | 0.67037 | -0.015037 |  | 490 | 17.2861 | 0.66797 | -0.0020893 |  | 490 | 0.67314 | -0.02987 | 490 | 0.67443 | -0.036724 | 490 | 43673.0391 | 490 | 43818.6406 | 490 | 44219.6016
491 | 17.3214 | 0.67375 | -0.033149 | 491 | 0.68839 | -0.10947 | 491 | 0.66949 | -0.010297 |  | 491 | 17.3214 | 0.6679 | -0.0017365 |  | 491 | 0.67347 | -0.031659 | 491 | 0.67441 | -0.036647 | 491 | 43698.8398 | 491 | 43784.4805 | 491 | 44196.2383
492 | 17.3567 | 0.67434 | -0.036285 | 492 | 0.6885 | -0.11002 | 492 | 0.66925 | -0.0090422 |  | 492 | 17.3567 | 0.66775 | -0.00094422 |  | 492 | 0.67363 | -0.032471 | 492 | 0.67481 | -0.038772 | 492 | 43722 | 492 | 43866.6406 | 492 | 44156.3203
493 | 17.3919 | 0.67397 | -0.034278 | 493 | 0.68912 | -0.11319 | 493 | 0.66965 | -0.011186 |  | 493 | 17.3919 | 0.66723 | 0.001911 |  | 493 | 0.67381 | -0.033442 | 493 | 0.67468 | -0.038043 | 493 | 43719.1602 | 493 | 43949.1211 | 493 | 44148.9219
494 | 17.4272 | 0.67483 | -0.03886 | 494 | 0.69058 | -0.12064 | 494 | 0.67036 | -0.014994 |  | 494 | 17.4272 | 0.66709 | 0.002667 |  | 494 | 0.673 | -0.029115 | 494 | 0.67402 | -0.034557 | 494 | 43712.6406 | 494 | 43960.3984 | 494 | 44140.1211
495 | 17.4625 | 0.67397 | -0.034275 | 495 | 0.68986 | -0.11696 | 495 | 0.66959 | -0.010876 |  | 495 | 17.4625 | 0.66731 | 0.00149 |  | 495 | 0.67227 | -0.025223 | 495 | 0.67402 | -0.034579 | 495 | 43705.1602 | 495 | 43944.5195 | 495 | 44116.8789
496 | 17.4978 | 0.67429 | -0.036003 | 496 | 0.68998 | -0.11757 | 496 | 0.67025 | -0.014416 |  | 496 | 17.4978 | 0.66815 | -0.0030718 |  | 496 | 0.6725 | -0.026453 | 496 | 0.67432 | -0.036168 | 496 | 43720.4414 | 496 | 43946.9609 | 496 | 44107.3594
497 | 17.5331 | 0.67531 | -0.041432 | 497 | 0.69118 | -0.12369 | 497 | 0.67095 | -0.018136 |  | 497 | 17.5331 | 0.66818 | -0.0032256 |  | 497 | 0.67233 | -0.025564 | 497 | 0.67422 | -0.035622 | 497 | 43724.8398 | 497 | 43901.7617 | 497 | 44107.2812
498 | 17.5683 | 0.67545 | -0.042125 | 498 | 0.69172 | -0.1264 | 498 | 0.67142 | -0.020707 |  | 498 | 17.5683 | 0.66871 | -0.0061221 |  | 498 | 0.67191 | -0.023328 | 498 | 0.67455 | -0.037402 | 498 | 43703.1211 | 498 | 43892.9609 | 498 | 44118.6797
499 | 17.6036 | 0.67594 | -0.044759 | 499 | 0.69236 | -0.12964 | 499 | 0.67131 | -0.020067 |  | 499 | 17.6036 | 0.66915 | -0.0084795 |  | 499 | 0.67181 | -0.022753 | 499 | 0.67479 | -0.038654 | 499 | 43674.3594 | 499 | 43898.8398 | 499 | 44127.7188
500 | 17.6389 | 0.67625 | -0.046376 | 500 | 0.69307 | -0.13324 | 500 | 0.67115 | -0.019253 |  | 500 | 17.6389 | 0.66972 | -0.011562 |  | 500 | 0.672 | -0.023788 | 500 | 0.67515 | -0.040534 | 500 | 43653.4414 | 500 | 43908.8789 | 500 | 44147
501 | 17.6742 | 0.67649 | -0.047645 | 501 | 0.69308 | -0.13332 | 501 | 0.67148 | -0.020984 |  | 501 | 17.6742 | 0.66995 | -0.012813 |  | 501 | 0.67224 | -0.025064 | 501 | 0.67571 | -0.043532 | 501 | 43634.8008 | 501 | 43938.1211 | 501 | 44159.6016
502 | 17.7094 | 0.67681 | -0.04934 | 502 | 0.6932 | -0.1339 | 502 | 0.67164 | -0.021872 |  | 502 | 17.7094 | 0.67008 | -0.013507 |  | 502 | 0.67193 | -0.023406 | 502 | 0.67522 | -0.040942 | 502 | 43648.3984 | 502 | 43971.2383 | 502 | 44154.1992
503 | 17.7447 | 0.67673 | -0.04892 | 503 | 0.69353 | -0.13559 | 503 | 0.67134 | -0.020263 |  | 503 | 17.7447 | 0.66984 | -0.012201 |  | 503 | 0.67189 | -0.023198 | 503 | 0.67533 | -0.041538 | 503 | 43662.9609 | 503 | 44019.8008 | 503 | 44171.1211
504 | 17.78 | 0.67672 | -0.048871 | 504 | 0.69375 | -0.13667 | 504 | 0.67167 | -0.021993 |  | 504 | 17.78 | 0.66977 | -0.011838 |  | 504 | 0.67181 | -0.022782 | 504 | 0.67543 | -0.042063 | 504 | 43687.7188 | 504 | 44064.5586 | 504 | 44138.1602
505 | 17.8153 | 0.67697 | -0.050187 | 505 | 0.69426 | -0.13922 | 505 | 0.67196 | -0.02358 |  | 505 | 17.8153 | 0.66979 | -0.011928 |  | 505 | 0.67167 | -0.022044 | 505 | 0.67558 | -0.042864 | 505 | 43678.7188 | 505 | 44082.6797 | 505 | 44141.9219
506 | 17.8506 | 0.67714 | -0.051097 | 506 | 0.69477 | -0.14181 | 506 | 0.67224 | -0.025063 |  | 506 | 17.8506 | 0.66981 | -0.01204 |  | 506 | 0.67148 | -0.021019 | 506 | 0.67558 | -0.042838 | 506 | 43673.9609 | 506 | 44088.0391 | 506 | 44175.0781
507 | 17.8858 | 0.67666 | -0.048563 | 507 | 0.69393 | -0.13759 | 507 | 0.67177 | -0.022562 |  | 507 | 17.8858 | 0.66904 | -0.0078662 |  | 507 | 0.67178 | -0.022589 | 507 | 0.67552 | -0.042537 | 507 | 43622.8398 | 507 | 44101.4805 | 507 | 44151.9609
508 | 17.9211 | 0.67685 | -0.049543 | 508 | 0.69424 | -0.13914 | 508 | 0.67121 | -0.019545 |  | 508 | 17.9211 | 0.66906 | -0.0080196 |  | 508 | 0.67258 | -0.026887 | 508 | 0.67542 | -0.042001 | 508 | 43677.6406 | 508 | 44087.7617 | 508 | 44143.3984
509 | 17.9564 | 0.67737 | -0.052296 | 509 | 0.69547 | -0.14531 | 509 | 0.67179 | -0.022668 |  | 509 | 17.9564 | 0.66821 | -0.0034057 |  | 509 | 0.67342 | -0.031344 | 509 | 0.67498 | -0.03964 | 509 | 43726.4414 | 509 | 44088.2812 | 509 | 44129.0391
510 | 17.9917 | 0.67631 | -0.046674 | 510 | 0.69423 | -0.13907 | 510 | 0.67072 | -0.016929 |  | 510 | 17.9917 | 0.66786 | -0.0015336 |  | 510 | 0.67361 | -0.03239 | 510 | 0.6754 | -0.041887 | 510 | 43757.3984 | 510 | 44075.6016 | 510 | 44123.7617
511 | 18.0269 | 0.67735 | -0.052212 | 511 | 0.69598 | -0.14788 | 511 | 0.67189 | -0.023218 |  | 511 | 18.0269 | 0.6673 | 0.0015201 |  | 511 | 0.67328 | -0.030609 | 511 | 0.67563 | -0.043099 | 511 | 43749.6406 | 511 | 44067.6797 | 511 | 44180.9219
512 | 18.0622 | 0.67764 | -0.053694 | 512 | 0.69607 | -0.14833 | 512 | 0.67226 | -0.025173 |  | 512 | 18.0622 | 0.66783 | -0.0013649 |  | 512 | 0.67259 | -0.026931 | 512 | 0.67517 | -0.040656 | 512 | 43741.7188 | 512 | 44113.8008 | 512 | 44172.8789
513 | 18.0975 | 0.67789 | -0.055042 | 513 | 0.69653 | -0.1506 | 513 | 0.67258 | -0.02691 |  | 513 | 18.0975 | 0.66732 | 0.0014112 |  | 513 | 0.67243 | -0.026098 | 513 | 0.67509 | -0.040218 | 513 | 43737.9219 | 513 | 44136.8398 | 513 | 44184.1992
514 | 18.1328 | 0.67781 | -0.054602 | 514 | 0.69648 | -0.15036 | 514 | 0.67265 | -0.02725 |  | 514 | 18.1328 | 0.66754 | 0.0002425 |  | 514 | 0.67204 | -0.024022 | 514 | 0.67448 | -0.036979 | 514 | 43714.8008 | 514 | 44136.3203 | 514 | 44201.0391
515 | 18.1681 | 0.67849 | -0.058184 | 515 | 0.6964 | -0.14997 | 515 | 0.67281 | -0.028124 |  | 515 | 18.1681 | 0.66747 | 0.00058797 |  | 515 | 0.67173 | -0.022348 | 515 | 0.67405 | -0.034737 | 515 | 43689.6797 | 515 | 44087.9219 | 515 | 44194.5586
516 | 18.2033 | 0.67867 | -0.059121 | 516 | 0.69639 | -0.1499 | 516 | 0.6723 | -0.025368 |  | 516 | 18.2033 | 0.66727 | 0.0016593 |  | 516 | 0.67145 | -0.020833 | 516 | 0.67384 | -0.033608 | 516 | 43684.2812 | 516 | 44070.3984 | 516 | 44241.7188
517 | 18.2386 | 0.67797 | -0.055436 | 517 | 0.69656 | -0.15076 | 517 | 0.67174 | -0.022402 |  | 517 | 18.2386 | 0.6671 | 0.0025994 |  | 517 | 0.67156 | -0.021442 | 517 | 0.67393 | -0.034088 | 517 | 43705.4414 | 517 | 44078.6406 | 517 | 44246.3203
518 | 18.2739 | 0.67797 | -0.055478 | 518 | 0.69687 | -0.1523 | 518 | 0.67129 | -0.019982 |  | 518 | 18.2739 | 0.6671 | 0.0026256 |  | 518 | 0.67121 | -0.019536 | 518 | 0.67323 | -0.03035 | 518 | 43747.4414 | 518 | 44052.6016 | 518 | 44254.8789
519 | 18.3092 | 0.6775 | -0.052964 | 519 | 0.69721 | -0.154 | 519 | 0.67089 | -0.017865 |  | 519 | 18.3092 | 0.66726 | 0.0017269 |  | 519 | 0.67126 | -0.019826 | 519 | 0.67272 | -0.027624 | 519 | 43781.6016 | 519 | 44051.8398 | 519 | 44255
520 | 18.3444 | 0.67787 | -0.054912 | 520 | 0.69874 | -0.16161 | 520 | 0.67151 | -0.021144 |  | 520 | 18.3444 | 0.66728 | 0.001648 |  | 520 | 0.67141 | -0.020624 | 520 | 0.67257 | -0.026824 | 520 | 43810.9219 | 520 | 44034.6406 | 520 | 44235.6406
521 | 18.3797 | 0.67867 | -0.059111 | 521 | 0.69846 | -0.16023 | 521 | 0.67162 | -0.021773 |  | 521 | 18.3797 | 0.66646 | 0.0061042 |  | 521 | 0.67142 | -0.020684 | 521 | 0.67267 | -0.027376 | 521 | 43853.9219 | 521 | 44051.3203 | 521 | 44237.9219
522 | 18.415 | 0.67854 | -0.058442 | 522 | 0.69811 | -0.15852 | 522 | 0.67111 | -0.019022 |  | 522 | 18.415 | 0.66559 | 0.010805 |  | 522 | 0.67182 | -0.022812 | 522 | 0.67126 | -0.019837 | 522 | 43891.3203 | 522 | 44046.0391 | 522 | 44241.6016
523 | 18.4503 | 0.67808 | -0.056022 | 523 | 0.69903 | -0.16309 | 523 | 0.67107 | -0.018826 |  | 523 | 18.4503 | 0.66551 | 0.01124 |  | 523 | 0.67225 | -0.025127 | 523 | 0.67192 | -0.023372 | 523 | 43907.0391 | 523 | 44049.5195 | 523 | 44224.7617
524 | 18.4856 | 0.67803 | -0.05575 | 524 | 0.69866 | -0.16125 | 524 | 0.67083 | -0.017538 |  | 524 | 18.4856 | 0.66587 | 0.0092684 |  | 524 | 0.67214 | -0.024541 | 524 | 0.67279 | -0.028005 | 524 | 43895.8008 | 524 | 44049.8789 | 524 | 44195.1992
525 | 18.5208 | 0.67827 | -0.057034 | 525 | 0.69889 | -0.16239 | 525 | 0.6709 | -0.017872 |  | 525 | 18.5208 | 0.6663 | 0.0069522 |  | 525 | 0.6729 | -0.028579 | 525 | 0.6736 | -0.032327 | 525 | 43892.7188 | 525 | 44060 | 525 | 44186.3594
526 | 18.5561 | 0.67813 | -0.056294 | 526 | 0.69901 | -0.16297 | 526 | 0.67075 | -0.017082 |  | 526 | 18.5561 | 0.66675 | 0.0045223 |  | 526 | 0.67286 | -0.028375 | 526 | 0.67364 | -0.032523 | 526 | 43867.6016 | 526 | 44048.8398 | 526 | 44205.3594
527 | 18.5914 | 0.67808 | -0.056026 | 527 | 0.69923 | -0.16409 | 527 | 0.67046 | -0.015521 |  | 527 | 18.5914 | 0.6675 | 0.00042653 |  | 527 | 0.67292 | -0.028686 | 527 | 0.67363 | -0.032464 | 527 | 43879.1602 | 527 | 44039.8008 | 527 | 44200.5195
528 | 18.6267 | 0.67847 | -0.058059 | 528 | 0.69962 | -0.16599 | 528 | 0.67087 | -0.017741 |  | 528 | 18.6267 | 0.66796 | -0.0020328 |  | 528 | 0.67339 | -0.031204 | 528 | 0.67372 | -0.032944 | 528 | 43868.5195 | 528 | 44055.7188 | 528 | 44206.6797
529 | 18.6619 | 0.67894 | -0.060552 | 529 | 0.69992 | -0.16749 | 529 | 0.67121 | -0.019573 |  | 529 | 18.6619 | 0.6684 | -0.0044066 |  | 529 | 0.67346 | -0.031588 | 529 | 0.67375 | -0.033125 | 529 | 43814.2383 | 529 | 44081.8008 | 529 | 44180.9219
530 | 18.6972 | 0.67931 | -0.062486 | 530 | 0.70002 | -0.168 | 530 | 0.67108 | -0.018876 |  | 530 | 18.6972 | 0.66834 | -0.0041215 |  | 530 | 0.67399 | -0.034398 | 530 | 0.67348 | -0.031707 | 530 | 43817.5586 | 530 | 44122.5195 | 530 | 44144.8789
531 | 18.7325 | 0.67981 | -0.065106 | 531 | 0.70047 | -0.17023 | 531 | 0.67132 | -0.020156 |  | 531 | 18.7325 | 0.66757 | 5.8166e-05 |  | 531 | 0.67401 | -0.034498 | 531 | 0.67323 | -0.030361 | 531 | 43761.6016 | 531 | 44149.3984 | 531 | 44148.1211
532 | 18.7678 | 0.68054 | -0.068921 | 532 | 0.701 | -0.17283 | 532 | 0.67187 | -0.02309 |  | 532 | 18.7678 | 0.66784 | -0.0013875 |  | 532 | 0.67367 | -0.032671 | 532 | 0.67315 | -0.029903 | 532 | 43732.6406 | 532 | 44147.5586 | 532 | 44130.8008
533 | 18.8031 | 0.6792 | -0.061891 | 533 | 0.70065 | -0.17112 | 533 | 0.67121 | -0.019534 |  | 533 | 18.8031 | 0.66724 | 0.0018396 |  | 533 | 0.67371 | -0.032896 | 533 | 0.67309 | -0.029626 | 533 | 43717.7617 | 533 | 44108.8789 | 533 | 44061.7617
534 | 18.8383 | 0.67923 | -0.062045 | 534 | 0.70052 | -0.17043 | 534 | 0.67114 | -0.019206 |  | 534 | 18.8383 | 0.66705 | 0.0028777 |  | 534 | 0.67354 | -0.031995 | 534 | 0.67294 | -0.028804 | 534 | 43717.1992 | 534 | 44098.2812 | 534 | 44054
535 | 18.8736 | 0.67879 | -0.059758 | 535 | 0.70042 | -0.16996 | 535 | 0.67123 | -0.019669 |  | 535 | 18.8736 | 0.66726 | 0.0017117 |  | 535 | 0.67356 | -0.032113 | 535 | 0.67249 | -0.026424 | 535 | 43671.7617 | 535 | 44051 | 535 | 44049.7188
536 | 18.9089 | 0.67836 | -0.057476 | 536 | 0.70021 | -0.16894 | 536 | 0.67043 | -0.01539 |  | 536 | 18.9089 | 0.66694 | 0.0034798 |  | 536 | 0.67319 | -0.030147 | 536 | 0.67294 | -0.028823 | 536 | 43698.8008 | 536 | 44036.8789 | 536 | 44029.4414
537 | 18.9442 | 0.67905 | -0.061147 | 537 | 0.70144 | -0.175 | 537 | 0.67121 | -0.01953 |  | 537 | 18.9442 | 0.66657 | 0.0054826 |  | 537 | 0.67231 | -0.025449 | 537 | 0.67276 | -0.027842 | 537 | 43708.5586 | 537 | 44003.8398 | 537 | 43985.9609
538 | 18.9794 | 0.67973 | -0.064679 | 538 | 0.7025 | -0.18021 | 538 | 0.67194 | -0.023488 |  | 538 | 18.9794 | 0.66586 | 0.0093628 |  | 538 | 0.67206 | -0.024133 | 538 | 0.67234 | -0.025616 | 538 | 43672.7617 | 538 | 44003.7617 | 538 | 44008.1992
539 | 19.0147 | 0.67901 | -0.06089 | 539 | 0.70254 | -0.18043 | 539 | 0.67123 | -0.019676 |  | 539 | 19.0147 | 0.66605 | 0.0083098 |  | 539 | 0.67147 | -0.020941 | 539 | 0.67271 | -0.027601 | 539 | 43653.4414 | 539 | 43963.4414 | 539 | 44039.1602
540 | 19.05 | 0.6795 | -0.063505 | 540 | 0.70261 | -0.18079 | 540 | 0.67142 | -0.020686 |  | 540 | 19.05 | 0.66595 | 0.0088344 |  | 540 | 0.67166 | -0.021951 | 540 | 0.67235 | -0.025664 | 540 | 43674.1211 | 540 | 43937.9609 | 540 | 44097.3203
541 | 19.0853 | 0.6801 | -0.066598 | 541 | 0.70272 | -0.1813 | 541 | 0.67123 | -0.019658 |  | 541 | 19.0853 | 0.66609 | 0.0080988 |  | 541 | 0.67177 | -0.022575 | 541 | 0.67199 | -0.023755 | 541 | 43704.7188 | 541 | 43985.6406 | 541 | 44122.9609
542 | 19.1206 | 0.68017 | -0.067014 | 542 | 0.70299 | -0.18263 | 542 | 0.67059 | -0.016248 |  | 542 | 19.1206 | 0.66744 | 0.000742 |  | 542 | 0.67242 | -0.026005 | 542 | 0.67284 | -0.028257 | 542 | 43695 | 542 | 44000.2812 | 542 | 44114.4805
543 | 19.1558 | 0.68234 | -0.078302 | 543 | 0.7045 | -0.19002 | 543 | 0.67248 | -0.02635 |  | 543 | 19.1558 | 0.66779 | -0.0011433 |  | 543 | 0.67309 | -0.029592 | 543 | 0.67315 | -0.029947 | 543 | 43682.3984 | 543 | 44043.3984 | 543 | 44108.5195
544 | 19.1911 | 0.68127 | -0.072722 | 544 | 0.70402 | -0.18768 | 544 | 0.67175 | -0.022441 |  | 544 | 19.1911 | 0.66794 | -0.001939 |  | 544 | 0.67358 | -0.032217 | 544 | 0.67308 | -0.029555 | 544 | 43647.1602 | 544 | 44041.8398 | 544 | 44100.2812
545 | 19.2264 | 0.68241 | -0.078646 | 545 | 0.70486 | -0.19177 | 545 | 0.67206 | -0.024098 |  | 545 | 19.2264 | 0.66806 | -0.0025845 |  | 545 | 0.67427 | -0.035914 | 545 | 0.67328 | -0.030635 | 545 | 43627.3203 | 545 | 44033.2383 | 545 | 44098.8008
546 | 19.2617 | 0.68226 | -0.077881 | 546 | 0.70561 | -0.19548 | 546 | 0.67255 | -0.026733 |  | 546 | 19.2617 | 0.6684 | -0.0044591 |  | 546 | 0.67464 | -0.037859 | 546 | 0.67348 | -0.031662 | 546 | 43652.8789 | 546 | 44104.3203 | 546 | 44098.3594
547 | 19.2969 | 0.68277 | -0.080534 | 547 | 0.70644 | -0.19949 | 547 | 0.67308 | -0.029547 |  | 547 | 19.2969 | 0.6689 | -0.007118 |  | 547 | 0.67541 | -0.041935 | 547 | 0.67259 | -0.026935 | 547 | 43687.1211 | 547 | 44124 | 547 | 44091.2812
548 | 19.3322 | 0.68257 | -0.079518 | 548 | 0.70635 | -0.19905 | 548 | 0.67302 | -0.029243 |  | 548 | 19.3322 | 0.66931 | -0.0093506 |  | 548 | 0.67579 | -0.043925 | 548 | 0.67282 | -0.028186 | 548 | 43711.5195 | 548 | 44120.8008 | 548 | 44096.0781
549 | 19.3675 | 0.68214 | -0.077263 | 549 | 0.70596 | -0.19714 | 549 | 0.67214 | -0.02451 |  | 549 | 19.3675 | 0.66945 | -0.010094 |  | 549 | 0.67527 | -0.041196 | 549 | 0.67287 | -0.028431 | 549 | 43709.0781 | 549 | 44133.0391 | 549 | 44073.4805
550 | 19.4028 | 0.68274 | -0.080372 | 550 | 0.70672 | -0.20086 | 550 | 0.67275 | -0.027799 |  | 550 | 19.4028 | 0.66941 | -0.0098963 |  | 550 | 0.67553 | -0.042581 | 550 | 0.67269 | -0.027479 | 550 | 43752.8789 | 550 | 44122.7617 | 550 | 44044.3594
551 | 19.4381 | 0.6835 | -0.084298 | 551 | 0.7082 | -0.20805 | 551 | 0.6739 | -0.033911 |  | 551 | 19.4381 | 0.66976 | -0.011749 |  | 551 | 0.6762 | -0.046101 | 551 | 0.67256 | -0.026805 | 551 | 43776.8789 | 551 | 44146.9219 | 551 | 44056.3984
552 | 19.4733 | 0.68355 | -0.084566 | 552 | 0.70806 | -0.20738 | 552 | 0.67381 | -0.033449 |  | 552 | 19.4733 | 0.66989 | -0.012443 |  | 552 | 0.67745 | -0.0527 | 552 | 0.67265 | -0.027283 | 552 | 43792.6797 | 552 | 44172.0781 | 552 | 44039.9609
[11 empty] | 553 | 19.5086 | 0.66998 | -0.012965 |  | 553 | 0.67769 | -0.053971 | 553 | 0.67265 | -0.027257 | 553 | 43760.5195 | 553 | 44176.2812 | 553 | 44005.8008
[11 empty] | 554 | 19.5439 | 0.67058 | -0.016157 |  | 554 | 0.67766 | -0.053843 | 554 | 0.67279 | -0.027983 | 554 | 43728.0781 | 554 | 44175 | 554 | 44038.4805
[11 empty] | 555 | 19.5792 | 0.67097 | -0.018267 |  | 555 | 0.67798 | -0.055482 | 555 | 0.67268 | -0.027394 | 555 | 43740.2383 | 555 | 44168.3203 | 555 | 44051.2383
[11 empty] | 556 | 19.6144 | 0.67082 | -0.017456 |  | 556 | 0.67775 | -0.054285 | 556 | 0.67314 | -0.029884 | 556 | 43711.0781 | 556 | 44157.4805 | 556 | 44029.5586
[11 empty] | 557 | 19.6497 | 0.6705 | -0.015721 |  | 557 | 0.67752 | -0.053061 | 557 | 0.67344 | -0.031463 | 557 | 43699.9609 | 557 | 44137.4414 | 557 | 44039.2812
[11 empty] | 558 | 19.685 | 0.67031 | -0.014741 |  | 558 | 0.67761 | -0.053551 | 558 | 0.67346 | -0.031563 | 558 | 43673.9609 | 558 | 44106.7188 | 558 | 44052.7188
[11 empty] | 559 | 19.7203 | 0.67074 | -0.017028 |  | 559 | 0.67739 | -0.0524 | 559 | 0.67419 | -0.035486 | 559 | 43672.8789 | 559 | 44133.5586 | 559 | 44065.6016
[11 empty] | 560 | 19.7556 | 0.67089 | -0.017839 |  | 560 | 0.67708 | -0.05074 | 560 | 0.67488 | -0.03914 | 560 | 43642.0781 | 560 | 44097.2812 | 560 | 44098.0391
[11 empty] | 561 | 19.7908 | 0.6705 | -0.015743 |  | 561 | 0.6766 | -0.048233 | 561 | 0.67502 | -0.039887 | 561 | 43615.5195 | 561 | 44068.8008 | 561 | 44076.2383
[11 empty] | 562 | 19.8261 | 0.67025 | -0.014394 |  | 562 | 0.67655 | -0.047988 | 562 | 0.67474 | -0.038374 | 562 | 43641.3984 | 562 | 44049.2812 | 562 | 44058.6016
[11 empty] | 563 | 19.8614 | 0.66987 | -0.012335 |  | 563 | 0.67615 | -0.045859 | 563 | 0.67408 | -0.034881 | 563 | 43644.6797 | 563 | 44109.9219 | 563 | 44049.9609
[11 empty] | 564 | 19.8967 | 0.66884 | -0.0067848 |  | 564 | 0.67627 | -0.046471 | 564 | 0.67422 | -0.035648 | 564 | 43716.1992 | 564 | 44136.8789 | 564 | 44035.1602
[11 empty] | 565 | 19.9319 | 0.66891 | -0.0071926 |  | 565 | 0.67621 | -0.046189 | 565 | 0.67396 | -0.034243 | 565 | 43739.1211 | 565 | 44164.8008 | 565 | 44032.3594
[11 empty] | 566 | 19.9672 | 0.66888 | -0.0070466 |  | 566 | 0.67655 | -0.047955 | 566 | 0.67368 | -0.032734 | 566 | 43752 | 566 | 44193.0781 | 566 | 44043.3984
[11 empty] | 567 | 20.0025 | 0.66849 | -0.0049199 |  | 567 | 0.67579 | -0.043943 | 567 | 0.67349 | -0.031762 | 567 | 43756.4805 | 567 | 44228.1992 | 567 | 44047.0781
[11 empty] | 568 | 20.0378 | 0.66739 | 0.0010465 |  | 568 | 0.67612 | -0.045694 | 568 | 0.67423 | -0.035674 | 568 | 43789.0781 | 568 | 44264.1211 | 568 | 44053
[11 empty] | 569 | 20.0731 | 0.66649 | 0.0059009 |  | 569 | 0.67602 | -0.04518 | 569 | 0.67427 | -0.035902 | 569 | 43848.7617 | 569 | 44318.6016 | 569 | 44058.8789
[11 empty] | 570 | 20.1083 | 0.66713 | 0.0024452 |  | 570 | 0.67574 | -0.043705 | 570 | 0.67422 | -0.035648 | 570 | 43867.6797 | 570 | 44325.8789 | 570 | 44064.8398
[11 empty] | 571 | 20.1436 | 0.66787 | -0.00156 |  | 571 | 0.67565 | -0.043191 | 571 | 0.67464 | -0.037874 | 571 | 43864.5586 | 571 | 44329.9219 | 571 | 44070.3203
[11 empty] | 572 | 20.1789 | 0.66805 | -0.0025468 |  | 572 | 0.67588 | -0.044435 | 572 | 0.67488 | -0.039122 | 572 | 43874.4805 | 572 | 44350.6016 | 572 | 44096.1211
[11 empty] | 573 | 20.2142 | 0.66817 | -0.0032033 |  | 573 | 0.67743 | -0.052601 | 573 | 0.67424 | -0.035737 | 573 | 43872.7188 | 573 | 44388.4414 | 573 | 44115.8789
[11 empty] | 574 | 20.2494 | 0.66868 | -0.0059573 |  | 574 | 0.67704 | -0.050554 | 574 | 0.67263 | -0.027135 | 574 | 43909.8789 | 574 | 44426.0391 | 574 | 44140.5586
[11 empty] | 575 | 20.2847 | 0.66921 | -0.0087787 |  | 575 | 0.67663 | -0.048402 | 575 | 0.67182 | -0.022808 | 575 | 43908.1992 | 575 | 44434.3203 | 575 | 44182.1211
[11 empty] | 576 | 20.32 | 0.66984 | -0.012178 |  | 576 | 0.67749 | -0.052912 | 576 | 0.67158 | -0.021557 | 576 | 43928.6016 | 576 | 44441.1992 | 576 | 44197.7188
[11 empty] | 577 | 20.3553 | 0.66998 | -0.012965 |  | 577 | 0.67785 | -0.054821 | 577 | 0.671 | -0.018427 | 577 | 43939.4414 | 577 | 44437.6016 | 577 | 44192.9609
[11 empty] | 578 | 20.3906 | 0.66986 | -0.012324 |  | 578 | 0.67834 | -0.057382 | 578 | 0.67121 | -0.019558 | 578 | 43947.7617 | 578 | 44440.5195 | 578 | 44155.8008
[11 empty] | 579 | 20.4258 | 0.66959 | -0.010838 |  | 579 | 0.67861 | -0.058795 | 579 | 0.67107 | -0.018814 | 579 | 43915.8789 | 579 | 44435.7188 | 579 | 44144.9609
[11 empty] | 580 | 20.4611 | 0.66974 | -0.011678 |  | 580 | 0.67851 | -0.058303 | 580 | 0.67118 | -0.019376 | 580 | 43881.3203 | 580 | 44462.5586 | 580 | 44104.4414
[11 empty] | 581 | 20.4964 | 0.66956 | -0.010666 |  | 581 | 0.67814 | -0.056372 | 581 | 0.67108 | -0.018851 | 581 | 43893.2383 | 581 | 44421.7617 | 581 | 44069.8398
[11 empty] | 582 | 20.5317 | 0.66845 | -0.0047025 |  | 582 | 0.67736 | -0.052261 | 582 | 0.67104 | -0.018654 | 582 | 43913.5195 | 582 | 44399.9219 | 582 | 44077.9609
[11 empty] | 583 | 20.5669 | 0.66815 | -0.0030796 |  | 583 | 0.67746 | -0.052762 | 583 | 0.6714 | -0.020591 | 583 | 43921.7617 | 583 | 44370.2812 | 583 | 44099
[11 empty] | 584 | 20.6022 | 0.66805 | -0.0025168 |  | 584 | 0.67725 | -0.051669 | 584 | 0.67179 | -0.022649 | 584 | 43868.6016 | 584 | 44342.1992 | 584 | 44091.2812
[11 empty] | 585 | 20.6375 | 0.66707 | 0.0027837 |  | 585 | 0.67743 | -0.05259 | 585 | 0.67182 | -0.022805 | 585 | 43882.1211 | 585 | 44327.4805 | 585 | 44131.6797
[11 empty] | 586 | 20.6728 | 0.6679 | -0.0017026 |  | 586 | 0.67758 | -0.053387 | 586 | 0.67183 | -0.022897 | 586 | 43792.0781 | 586 | 44325.0781 | 586 | 44169.3203
[11 empty] | 587 | 20.7081 | 0.66823 | -0.003522 |  | 587 | 0.6773 | -0.051911 | 587 | 0.67179 | -0.022656 | 587 | 43779.6797 | 587 | 44309 | 587 | 44208.9219
[11 empty] | 588 | 20.7433 | 0.66891 | -0.0071701 |  | 588 | 0.67758 | -0.05339 | 588 | 0.67187 | -0.023098 | 588 | 43813.4414 | 588 | 44295.7617 | 588 | 44124.4414
[11 empty] | 589 | 20.7786 | 0.66867 | -0.0059087 |  | 589 | 0.67687 | -0.049661 | 589 | 0.6716 | -0.021635 | 589 | 43775.3594 | 589 | 44288.6797 | 589 | 44049.9609
[11 empty] | 590 | 20.8139 | 0.6688 | -0.006605 |  | 590 | 0.67734 | -0.052126 | 590 | 0.67222 | -0.024938 | 590 | 43710.4805 | 590 | 44241 | 590 | 43992.1602
[11 empty] | 591 | 20.8492 | 0.66907 | -0.0080234 |  | 591 | 0.67734 | -0.052133 | 591 | 0.67251 | -0.026487 | 591 | 43658.2383 | 591 | 44268.7188 | 591 | 43956.7617
[11 empty] | 592 | 20.8844 | 0.66949 | -0.010289 |  | 592 | 0.67788 | -0.054975 | 592 | 0.67228 | -0.025271 | 592 | 43684.8789 | 592 | 44265.9219 | 592 | 43918.9219
[11 empty] | 593 | 20.9197 | 0.66945 | -0.010124 |  | 593 | 0.67808 | -0.056051 | 593 | 0.67151 | -0.021137 | 593 | 43750.6406 | 593 | 44294.8008 | 593 | 43931.8789
[11 empty] | 594 | 20.955 | 0.66919 | -0.0087189 |  | 594 | 0.67781 | -0.054613 | 594 | 0.67176 | -0.022511 | 594 | 43762.2383 | 594 | 44385.5195 | 594 | 43923.6406
[11 empty] | 595 | 20.9903 | 0.66884 | -0.0068298 |  | 595 | 0.67734 | -0.052126 | 595 | 0.67162 | -0.021758 | 595 | 43766.1602 | 595 | 44389.1602 | 595 | 43921.9219
[11 empty] | 596 | 21.0256 | 0.66866 | -0.0058302 |  | 596 | 0.67806 | -0.055923 | 596 | 0.67164 | -0.021869 | 596 | 43803.2383 | 596 | 44361.1992 | 596 | 43933.2383
[11 empty] | 597 | 21.0608 | 0.66857 | -0.005332 |  | 597 | 0.6783 | -0.057203 | 597 | 0.67191 | -0.023306 | 597 | 43852.8398 | 597 | 44381.6016 | 597 | 43940.6797
[11 empty] | 598 | 21.0961 | 0.66749 | 0.00047528 |  | 598 | 0.67862 | -0.058843 | 598 | 0.67189 | -0.023209 | 598 | 43899.9609 | 598 | 44391.2383 | 598 | 43970.8398
[11 empty] | 599 | 21.1314 | 0.66706 | 0.0028137 |  | 599 | 0.67853 | -0.058391 | 599 | 0.67245 | -0.026213 | 599 | 43841.3594 | 599 | 44448.5195 | 599 | 43973.6016
[11 empty] | 600 | 21.1667 | 0.66748 | 0.00055407 |  | 600 | 0.67793 | -0.055252 | 600 | 0.67264 | -0.027227 | 600 | 43876.1602 | 600 | 44500.3984 | 600 | 43999
[11 empty] | 601 | 21.2019 | 0.66837 | -0.0042604 |  | 601 | 0.67821 | -0.056722 | 601 | 0.67235 | -0.025675 | 601 | 43872.0781 | 601 | 44505.3594 | 601 | 43992.4414
[11 empty] | 602 | 21.2372 | 0.66937 | -0.009676 |  | 602 | 0.67857 | -0.058602 | 602 | 0.67216 | -0.024615 | 602 | 43872.0391 | 602 | 44505.0391 | 602 | 43982.1992
[11 empty] | 603 | 21.2725 | 0.67012 | -0.013678 |  | 603 | 0.67792 | -0.055205 | 603 | 0.67214 | -0.02453 | 603 | 43857.6797 | 603 | 44442.9219 | 603 | 44004.8789
[11 empty] | 604 | 21.3078 | 0.67055 | -0.01603 |  | 604 | 0.67724 | -0.051614 | 604 | 0.67127 | -0.019889 | 604 | 43790.3203 | 604 | 44368.2383 | 604 | 43973.8398
[11 empty] | 605 | 21.3431 | 0.67013 | -0.013745 |  | 605 | 0.67733 | -0.052079 | 605 | 0.67079 | -0.017285 | 605 | 43740.8008 | 605 | 44373.7188 | 605 | 44016.4414
[11 empty] | 606 | 21.3783 | 0.66985 | -0.012227 |  | 606 | 0.67727 | -0.051786 | 606 | 0.6707 | -0.016823 | 606 | 43728.6406 | 606 | 44407.4414 | 606 | 44025.9219
[11 empty] | 607 | 21.4136 | 0.67028 | -0.014573 |  | 607 | 0.67731 | -0.051983 | 607 | 0.67062 | -0.016406 | 607 | 43671.2383 | 607 | 44408.6797 | 607 | 44019.7617
[11 empty] | 608 | 21.4489 | 0.66871 | -0.0061034 |  | 608 | 0.67774 | -0.054267 | 608 | 0.67093 | -0.018078 | 608 | 43730.0781 | 608 | 44378.9609 | 608 | 43979.8398
[11 empty] | 609 | 21.4842 | 0.66895 | -0.0074171 |  | 609 | 0.67805 | -0.055894 | 609 | 0.6722 | -0.024838 | 609 | 43795.7188 | 609 | 44406.8008 | 609 | 43995.0391
[11 empty] | 610 | 21.5194 | 0.66942 | -0.0099114 |  | 610 | 0.67794 | -0.05531 | 610 | 0.67179 | -0.022675 | 610 | 43818.4414 | 610 | 44382.9219 | 610 | 43985.4414
[11 empty] | 611 | 21.5547 | 0.66891 | -0.0071852 |  | 611 | 0.67836 | -0.057487 | 611 | 0.67188 | -0.023142 | 611 | 43802.6797 | 611 | 44365.3203 | 611 | 43979.1992
[11 empty] | 612 | 21.59 | 0.66853 | -0.0051221 |  | 612 | 0.67964 | -0.064218 | 612 | 0.67104 | -0.018651 | 612 | 43834.6016 | 612 | 44394.1211 | 612 | 43993.0781
[11 empty] | 613 | 21.6253 | 0.66859 | -0.0054594 |  | 613 | 0.67935 | -0.062688 | 613 | 0.67088 | -0.017802 | 613 | 43851.1211 | 613 | 44415.5195 | 613 | 44001.4414
[11 empty] | 614 | 21.6606 | 0.66816 | -0.0031508 |  | 614 | 0.67974 | -0.06473 | 614 | 0.67067 | -0.016648 | 614 | 43868.1602 | 614 | 44413.0781 | 614 | 44021.3984
[11 empty] | 615 | 21.6958 | 0.66817 | -0.0032069 |  | 615 | 0.68037 | -0.068012 | 615 | 0.67091 | -0.017973 | 615 | 43854.6797 | 615 | 44420.2812 | 615 | 44022.0391
[11 empty] | 616 | 21.7311 | 0.66799 | -0.0021979 |  | 616 | 0.68014 | -0.066852 | 616 | 0.67062 | -0.016395 | 616 | 43848.1992 | 616 | 44413.2383 | 616 | 43987.2812
[11 empty] | 617 | 21.7664 | 0.66777 | -0.0010007 |  | 617 | 0.68008 | -0.066522 | 617 | 0.67094 | -0.018122 | 617 | 43845.5586 | 617 | 44423.1992 | 617 | 43990.2383
[11 empty] | 618 | 21.8017 | 0.66779 | -0.001132 |  | 618 | 0.68046 | -0.068531 | 618 | 0.67106 | -0.018732 | 618 | 43843.4805 | 618 | 44457.1211 | 618 | 43979.6406
[11 empty] | 619 | 21.8369 | 0.66833 | -0.0040317 |  | 619 | 0.68039 | -0.068146 | 619 | 0.67148 | -0.020978 | 619 | 43818.8008 | 619 | 44500.2812 | 619 | 43938.5586
[11 empty] | 620 | 21.8722 | 0.66855 | -0.0052682 |  | 620 | 0.6809 | -0.070801 | 620 | 0.6715 | -0.021085 | 620 | 43767.3984 | 620 | 44479.8008 | 620 | 43916.1602
[11 empty] | 621 | 21.9075 | 0.66886 | -0.0069308 |  | 621 | 0.6812 | -0.072346 | 621 | 0.67132 | -0.020142 | 621 | 43715.2383 | 621 | 44473.4805 | 621 | 43922.0781
[11 empty] | 622 | 21.9428 | 0.669 | -0.0076492 |  | 622 | 0.6801 | -0.066606 | 622 | 0.67182 | -0.022845 | 622 | 43750.9219 | 622 | 44496.4805 | 622 | 43911.9219
[11 empty] | 623 | 21.9781 | 0.66962 | -0.010987 |  | 623 | 0.6792 | -0.061896 | 623 | 0.67164 | -0.021836 | 623 | 43799.9609 | 623 | 44487.2812 | 623 | 43915.4414
[11 empty] | 624 | 22.0133 | 0.66919 | -0.0086852 |  | 624 | 0.67862 | -0.058861 | 624 | 0.67145 | -0.020855 | 624 | 43839.2383 | 624 | 44441.9609 | 624 | 43964.2383
[11 empty] | 625 | 22.0486 | 0.6696 | -0.010916 |  | 625 | 0.67828 | -0.057057 | 625 | 0.67155 | -0.021379 | 625 | 43878.3594 | 625 | 44442.3594 | 625 | 43968.3594
[11 empty] | 626 | 22.0839 | 0.66943 | -0.0099749 |  | 626 | 0.67836 | -0.057498 | 626 | 0.67147 | -0.020944 | 626 | 43898.4805 | 626 | 44423.0781 | 626 | 43981.0781
[11 empty] | 627 | 22.1192 | 0.66939 | -0.0097842 |  | 627 | 0.67834 | -0.057414 | 627 | 0.67156 | -0.021412 | 627 | 43896.6016 | 627 | 44391.8789 | 627 | 43936.3984
[11 empty] | 628 | 22.1544 | 0.66983 | -0.012148 |  | 628 | 0.67909 | -0.06131 | 628 | 0.67182 | -0.022849 | 628 | 43903.6016 | 628 | 44445.5195 | 628 | 43901.0391
[11 empty] | 629 | 22.1897 | 0.67033 | -0.014841 |  | 629 | 0.67941 | -0.063008 | 629 | 0.67207 | -0.024155 | 629 | 43889.1211 | 629 | 44465.4414 | 629 | 43956.4805
[11 empty] | 630 | 22.225 | 0.6704 | -0.015221 |  | 630 | 0.68001 | -0.066167 | 630 | 0.67195 | -0.023532 | 630 | 43784.4414 | 630 | 44492.9219 | 630 | 43954.3594
[11 empty] | 631 | 22.2603 | 0.67027 | -0.014513 |  | 631 | 0.68046 | -0.068494 | 631 | 0.67171 | -0.022248 | 631 | 43850.3984 | 631 | 44528.5586 | 631 | 43917.5586
[11 empty] | 632 | 22.2956 | 0.67003 | -0.013223 |  | 632 | 0.68079 | -0.070251 | 632 | 0.67128 | -0.019948 | 632 | 43831 | 632 | 44548.8789 | 632 | 43948.7617
[11 empty] | 633 | 22.3308 | 0.67085 | -0.017642 |  | 633 | 0.68078 | -0.070149 | 633 | 0.67135 | -0.020331 | 633 | 43812.8008 | 633 | 44553.0781 | 633 | 43953.7188
[11 empty] | 634 | 22.3661 | 0.67065 | -0.016551 |  | 634 | 0.68068 | -0.069675 | 634 | 0.67047 | -0.015609 | 634 | 43792.7188 | 634 | 44562.8008 | 634 | 43921.9219
[11 empty] | 635 | 22.4014 | 0.67059 | -0.01622 |  | 635 | 0.68103 | -0.071495 | 635 | 0.67077 | -0.017169 | 635 | 43757.8008 | 635 | 44550.3984 | 635 | 43947.9609
[11 empty] | 636 | 22.4367 | 0.67014 | -0.01382 |  | 636 | 0.68168 | -0.074864 | 636 | 0.67109 | -0.018892 | 636 | 43748.6406 | 636 | 44555.9609 | 636 | 43940.2383
[11 empty] | 637 | 22.4719 | 0.67106 | -0.018751 |  | 637 | 0.68127 | -0.072707 | 637 | 0.67132 | -0.020134 | 637 | 43773.5195 | 637 | 44565.5195 | 637 | 43971.2812
[11 empty] | 638 | 22.5072 | 0.67156 | -0.021434 |  | 638 | 0.68109 | -0.071814 | 638 | 0.6719 | -0.02325 | 638 | 43773.1602 | 638 | 44596.0781 | 638 | 43969.3594
[11 empty] | 639 | 22.5425 | 0.67184 | -0.022931 |  | 639 | 0.68076 | -0.070066 | 639 | 0.6723 | -0.025401 | 639 | 43767.7617 | 639 | 44616.5586 | 639 | 43958.5195
[11 empty] | 640 | 22.5778 | 0.67204 | -0.024022 |  | 640 | 0.68124 | -0.072552 | 640 | 0.67263 | -0.027176 | 640 | 43784.6797 | 640 | 44633.1992 | 640 | 43950.1602
[11 empty] | 641 | 22.6131 | 0.67169 | -0.022151 |  | 641 | 0.68203 | -0.076665 | 641 | 0.67284 | -0.028261 | 641 | 43804.8789 | 641 | 44666.6797 | 641 | 43941.7617
[11 empty] | 642 | 22.6483 | 0.67162 | -0.021728 |  | 642 | 0.68256 | -0.079433 | 642 | 0.67204 | -0.024003 | 642 | 43816.0391 | 642 | 44669.5586 | 642 | 43931.2812
[11 empty] | 643 | 22.6836 | 0.67192 | -0.023383 |  | 643 | 0.68255 | -0.079376 | 643 | 0.67194 | -0.023473 | 643 | 43839.6016 | 643 | 44627.7617 | 643 | 43955.6016
[11 empty] | 644 | 22.7189 | 0.67191 | -0.023317 |  | 644 | 0.6826 | -0.07966 | 644 | 0.67184 | -0.022909 | 644 | 43845.6406 | 644 | 44554.8789 | 644 | 43978.3594
[11 empty] | 645 | 22.7542 | 0.67172 | -0.022289 |  | 645 | 0.6835 | -0.084323 | 645 | 0.67176 | -0.022482 | 645 | 43869.5586 | 645 | 44530.0781 | 645 | 43982.3984
[11 empty] | 646 | 22.7894 | 0.67169 | -0.022125 |  | 646 | 0.68418 | -0.087854 | 646 | 0.67202 | -0.023888 | 646 | 43861.8789 | 646 | 44527.2383 | 646 | 43941.4805
[11 empty] | 647 | 22.8247 | 0.67137 | -0.020394 |  | 647 | 0.68423 | -0.088109 | 647 | 0.67259 | -0.026924 | 647 | 43854.7188 | 647 | 44537.9219 | 647 | 43915.8789
[11 empty] | 648 | 22.86 | 0.67166 | -0.021992 |  | 648 | 0.68395 | -0.086628 | 648 | 0.67325 | -0.030447 | 648 | 43838.6797 | 648 | 44539.8398 | 648 | 43939.7188
[11 empty] | 649 | 22.8953 | 0.672 | -0.023803 |  | 649 | 0.68417 | -0.087776 | 649 | 0.67333 | -0.030912 | 649 | 43878.8008 | 649 | 44558.7617 | 649 | 43887.1602
[11 empty] | 650 | 22.9306 | 0.67259 | -0.026931 |  | 650 | 0.68431 | -0.088486 | 650 | 0.67321 | -0.030247 | 650 | 43880.8008 | 650 | 44572.8008 | 650 | 43881.1602
[11 empty] | 651 | 22.9658 | 0.67298 | -0.029041 |  | 651 | 0.68377 | -0.085731 | 651 | 0.67317 | -0.030018 | 651 | 43895.9219 | 651 | 44604.5586 | 651 | 43914.0391
[11 empty] | 652 | 23.0011 | 0.67357 | -0.032154 |  | 652 | 0.68369 | -0.085292 | 652 | 0.67346 | -0.031566 | 652 | 43949.8789 | 652 | 44619.0781 | 652 | 43948.7188
[11 empty] | 653 | 23.0364 | 0.67331 | -0.030805 |  | 653 | 0.68388 | -0.086283 | 653 | 0.67365 | -0.032597 | 653 | 43943.0391 | 653 | 44612.8008 | 653 | 43977.3594
[11 empty] | 654 | 23.0717 | 0.67357 | -0.032139 |  | 654 | 0.68441 | -0.089009 | 654 | 0.67353 | -0.031943 | 654 | 43908.9219 | 654 | 44628.5195 | 654 | 43987.9609
[22 empty] | 655 | 43962.9609 | 655 | 44637.6016 | 655 | 44015.2383
[22 empty] | 656 | 43955.5586 | 656 | 44644.6406 | 656 | 43997.8789
[22 empty] | 657 | 43924.3594 | 657 | 44681.2812 | 657 | 43998.4805
[22 empty] | 658 | 43916.0391 | 658 | 44654.5586 | 658 | 44006.1992
[22 empty] | 659 | 43939.2383 | 659 | 44604.9219 | 659 | 43986.6797
[22 empty] | 660 | 44003.8789 | 660 | 44617 | 660 | 43998.6406
[22 empty] | 661 | 44026.7617 | 661 | 44657.3594 | 661 | 43990.4805
[22 empty] | 662 | 44049.8008 | 662 | 44663.8008 | 662 | 43997.1602
[22 empty] | 663 | 44029.1211 | 663 | 44714.6016 | 663 | 44030.0781
[22 empty] | 664 | 43989.5586 | 664 | 44735.1992 | 664 | 44042.1992
[22 empty] | 665 | 44012.4805 | 665 | 44755.8008 | 665 | 44045.6406
[22 empty] | 666 | 44025.2383 | 666 | 44797.7617 | 666 | 44042.8398
[22 empty] | 667 | 43999.5586 | 667 | 44813.1992 | 667 | 44049.7617
[22 empty] | 668 | 44005.1211 | 668 | 44840.0391 | 668 | 44076.8789
[22 empty] | 669 | 44013.2383 | 669 | 44840.5586 | 669 | 44088.7617
[22 empty] | 670 | 44026.7188 | 670 | 44863.4414 | 670 | 44087.5195
[22 empty] | 671 | 44039.7188 | 671 | 44885.6406 | 671 | 44089.1992
[22 empty] | 672 | 44068.6016 | 672 | 44870.4414 | 672 | 44058.8789
[22 empty] | 673 | 44085 | 673 | 44839.6797 | 673 | 44056.6406
[22 empty] | 674 | 44123.4805 | 674 | 44857.1992 | 674 | 43991.6797
[22 empty] | 675 | 44082.8008 | 675 | 44867.6016 | 675 | 43947.5195
[22 empty] | 676 | 44137.2383 | 676 | 44856.9219 | 676 | 43926.3203
[22 empty] | 677 | 44096 | 677 | 44878.7617 | 677 | 43928.7188
[22 empty] | 678 | 44089.3203 | 678 | 44891.2383 | 678 | 43946.5586
[22 empty] | 679 | 44025.1992 | 679 | 44897.0391 | 679 | 43904.5195
[22 empty] | 680 | 44028.5195 | 680 | 44916.7188 | 680 | 43931.6797
[22 empty] | 681 | 43991.7617 | 681 | 44860.9609 | 681 | 43950.5195
[22 empty] | 682 | 44024.3203 | 682 | 44859.9609 | 682 | 43975.6797
[22 empty] | 683 | 44066.9219 | 683 | 44915.8398 | 683 | 44025.3594
[22 empty] | 684 | 44066.5195 | 684 | 44902.8398 | 684 | 44039.1602
[22 empty] | 685 | 44045.6797 | 685 | 44913.0391 | 685 | 44061.1992
[22 empty] | 686 | 44071.5586 | 686 | 44921.7188 | 686 | 43997.5586
[22 empty] | 687 | 44071.6406 | 687 | 44957.5195 | 687 | 43987.4414
[22 empty] | 688 | 44077.1602 | 688 | 44959.5195 | 688 | 44013.0391
[22 empty] | 689 | 44055.9609 | 689 | 45006.9219 | 689 | 44002.3984

## unif profiles 0Gy flat
|  |  | dose=c+b/(X-a) |  | this equation uses the rational function to fit the dose data | [7 empty] | pixels to cm conversion | 1px = |  | 1/ | 1 inch = 2.54 centimetres
a |  |  | -0.040473 |  | using the file mephysto flattened using same film to flatten but film moved between flatteneing and scanning | [11 empty] | dpi | ppcm | cm/px
b |  |  | 2.7151 | [13 empty] | 72 | 28.3465 | 0.035278
c |  |  | -3.8345
centre across | dist cm |  | normalized film response | Gy | near to scan array, across |  | normalized film response | Gy | opposite to scan array, across |  |  |  |  | centre, down |  |  |  |  | left side, down |  |  | right side, down |  |  | centre, down repeated |  | left side, down, repeated |  | right side , down, repeated
0 | 0 | 43469.8789 | 0.66331 | 0.02328 | 0 | 43449.6094 | 0.663 | 0.024976 | 0 | 43461.6602 | 0.66318 | 0.023967 |  | 0 | 0 | 0.67062 | -0.016384 |  | 0 | 0.66018 | 0.040489 | 0 | 0.69872 | -0.16153 | 0 | 44048.9609 | 0 | 43328.2383 | 0 | 45893.4805
1 | 0.035278 | 43479.4102 | 0.66345 | 0.022482 | 1 | 43433.9922 | 0.66276 | 0.026284 | 1 | 43462.4688 | 0.66319 | 0.023899 |  | 1 | 0.035278 | 0.67071 | -0.016898 |  | 1 | 0.66023 | 0.040217 | 1 | 0.69851 | -0.16047 | 1 | 44011.4805 | 1 | 43321.6797 | 1 | 45858.3594
2 | 0.070556 | 43493.1914 | 0.66366 | 0.021331 | 2 | 43423.7734 | 0.6626 | 0.02714 | 2 | 43447.9219 | 0.66297 | 0.025117 |  | 2 | 0.070556 | 0.67061 | -0.016339 |  | 2 | 0.65959 | 0.043771 | 2 | 0.69838 | -0.15985 | 2 | 44005.9219 | 2 | 43315.1992 | 2 | 45851.1992
3 | 0.10583 | 43432.4688 | 0.66274 | 0.026411 | 3 | 43441.2891 | 0.66287 | 0.025672 | 3 | 43397.0508 | 0.6622 | 0.029381 |  | 3 | 0.10583 | 0.67146 | -0.020911 |  | 3 | 0.66003 | 0.041325 | 3 | 0.69796 | -0.15775 | 3 | 44022.8789 | 3 | 43334.6406 | 3 | 45866.8008
4 | 0.14111 | 43444.9336 | 0.66293 | 0.025367 | 4 | 43444.9141 | 0.66293 | 0.025369 | 4 | 43437.0664 | 0.66281 | 0.026026 |  | 4 | 0.14111 | 0.67071 | -0.016894 |  | 4 | 0.66042 | 0.039197 | 4 | 0.69785 | -0.15723 | 4 | 44082 | 4 | 43351.2383 | 4 | 45872.4805
5 | 0.17639 | 43466.2344 | 0.66325 | 0.023584 | 5 | 43491.3203 | 0.66364 | 0.021487 | 5 | 43473.5195 | 0.66336 | 0.022975 |  | 5 | 0.17639 | 0.6702 | -0.014148 |  | 5 | 0.66059 | 0.038258 | 5 | 0.6981 | -0.15844 | 5 | 44090.2812 | 5 | 43385.5586 | 5 | 45888.0781
6 | 0.21167 | 43401.2812 | 0.66226 | 0.029026 | 6 | 43454.1953 | 0.66307 | 0.024592 | 6 | 43449.043 | 0.66299 | 0.025023 |  | 6 | 0.21167 | 0.66949 | -0.010322 |  | 6 | 0.66111 | 0.035371 | 6 | 0.69758 | -0.15587 | 6 | 44102.8789 | 6 | 43363.6406 | 6 | 45904.4414
7 | 0.24694 | 43397.625 | 0.66221 | 0.029333 | 7 | 43416.5938 | 0.66249 | 0.027742 | 7 | 43403.7734 | 0.6623 | 0.028817 |  | 7 | 0.24694 | 0.67086 | -0.017653 |  | 7 | 0.66116 | 0.03508 | 7 | 0.69716 | -0.15377 | 7 | 44142.6797 | 7 | 43402.5586 | 7 | 45936
8 | 0.28222 | 43409.7344 | 0.66239 | 0.028317 | 8 | 43400.8438 | 0.66225 | 0.029063 | 8 | 43415.4883 | 0.66248 | 0.027834 |  | 8 | 0.28222 | 0.66865 | -0.0057778 |  | 8 | 0.66155 | 0.032958 | 8 | 0.69684 | -0.15218 | 8 | 44069.8789 | 8 | 43373.7617 | 8 | 45897.6016
9 | 0.3175 | 43409.8789 | 0.66239 | 0.028305 | 9 | 43415.1602 | 0.66247 | 0.027862 | 9 | 43459.1133 | 0.66314 | 0.02418 |  | 9 | 0.3175 | 0.66969 | -0.011379 |  | 9 | 0.66164 | 0.032431 | 9 | 0.69703 | -0.15309 | 9 | 44086.2812 | 9 | 43380.2383 | 9 | 45879.8008
10 | 0.35278 | 43369.3828 | 0.66177 | 0.031704 | 10 | 43420.4492 | 0.66255 | 0.027418 | 10 | 43393.8594 | 0.66215 | 0.029649 |  | 10 | 0.35278 | 0.67015 | -0.013879 |  | 10 | 0.66074 | 0.037392 | 10 | 0.69723 | -0.15413 | 10 | 44048.8789 | 10 | 43423.6406 | 10 | 45870.8789
11 | 0.38806 | 43421.4023 | 0.66257 | 0.027339 | 11 | 43398.2266 | 0.66221 | 0.029282 | 11 | 43474.6172 | 0.66338 | 0.022883 |  | 11 | 0.38806 | 0.67154 | -0.021334 |  | 11 | 0.66135 | 0.03404 | 11 | 0.69713 | -0.15362 | 11 | 44017.4805 | 11 | 43472.8008 | 11 | 45822.2812
12 | 0.42333 | 43377.0508 | 0.66189 | 0.03106 | 12 | 43444.8828 | 0.66293 | 0.025371 | 12 | 43432.3984 | 0.66274 | 0.026417 |  | 12 | 0.42333 | 0.67059 | -0.016235 |  | 12 | 0.66093 | 0.036377 | 12 | 0.6964 | -0.14998 | 12 | 43993.3203 | 12 | 43458.4414 | 12 | 45802.6016
13 | 0.45861 | 43393.2383 | 0.66214 | 0.029701 | 13 | 43432.0312 | 0.66273 | 0.026448 | 13 | 43392.1484 | 0.66212 | 0.029792 |  | 13 | 0.45861 | 0.66999 | -0.01301 |  | 13 | 0.6613 | 0.034307 | 13 | 0.69661 | -0.15101 | 13 | 43971.6406 | 13 | 43490.0781 | 13 | 45805.3984
14 | 0.49389 | 43470.8359 | 0.66332 | 0.0232 | 14 | 43366.0156 | 0.66172 | 0.031987 | 14 | 43471.0664 | 0.66333 | 0.02318 |  | 14 | 0.49389 | 0.67084 | -0.017568 |  | 14 | 0.66042 | 0.039158 | 14 | 0.6961 | -0.14848 | 14 | 43997.3984 | 14 | 43432.8398 | 14 | 45772.3203
15 | 0.52917 | 43417.4219 | 0.66251 | 0.027672 | 15 | 43454.043 | 0.66307 | 0.024605 | 15 | 43405.0703 | 0.66232 | 0.028708 |  | 15 | 0.52917 | 0.6714 | -0.020584 |  | 15 | 0.66059 | 0.038216 | 15 | 0.69585 | -0.14724 | 15 | 43922.7617 | 15 | 43373.1992 | 15 | 45745.6016
16 | 0.56444 | 43427.8516 | 0.66267 | 0.026798 | 16 | 43409.4219 | 0.66239 | 0.028343 | 16 | 43409.75 | 0.66239 | 0.028316 |  | 16 | 0.56444 | 0.67138 | -0.020483 |  | 16 | 0.65993 | 0.041886 | 16 | 0.69577 | -0.1468 | 16 | 43847.6406 | 16 | 43344.8789 | 16 | 45720.5586
17 | 0.59972 | 43454.2344 | 0.66307 | 0.024589 | 17 | 43398.7734 | 0.66222 | 0.029236 | 17 | 43452.3125 | 0.66304 | 0.024749 |  | 17 | 0.59972 | 0.67148 | -0.021 |  | 17 | 0.65994 | 0.04184 | 17 | 0.69519 | -0.14391 | 17 | 43876.7617 | 17 | 43293.7188 | 17 | 45724
18 | 0.635 | 43409.5039 | 0.66239 | 0.028336 | 18 | 43508.8047 | 0.6639 | 0.020026 | 18 | 43386.3984 | 0.66203 | 0.030275 |  | 18 | 0.635 | 0.67084 | -0.017597 |  | 18 | 0.65993 | 0.041901 | 18 | 0.696 | -0.14796 | 18 | 43926.0781 | 18 | 43244.3984 | 18 | 45723.2383
19 | 0.67028 | 43416.4297 | 0.66249 | 0.027755 | 19 | 43408.7422 | 0.66237 | 0.0284 | 19 | 43458.6523 | 0.66314 | 0.024219 |  | 19 | 0.67028 | 0.6702 | -0.014122 |  | 19 | 0.66006 | 0.041153 | 19 | 0.69662 | -0.15105 | 19 | 43908.7617 | 19 | 43223.1602 | 19 | 45690.4414
20 | 0.70556 | 43473.625 | 0.66336 | 0.022966 | 20 | 43492.7812 | 0.66366 | 0.021365 | 20 | 43488.5703 | 0.66359 | 0.021717 |  | 20 | 0.70556 | 0.66973 | -0.011589 |  | 20 | 0.66028 | 0.03996 | 20 | 0.6975 | -0.15546 | 20 | 43878.5195 | 20 | 43161.3203 | 20 | 45709.6406
21 | 0.74083 | 43411.918 | 0.66242 | 0.028134 | 21 | 43469.0234 | 0.66329 | 0.023351 | 21 | 43460.5508 | 0.66317 | 0.02406 |  | 21 | 0.74083 | 0.66934 | -0.0095301 |  | 21 | 0.66076 | 0.037277 | 21 | 0.69756 | -0.15574 | 21 | 43864.1602 | 21 | 43146.7188 | 21 | 45670.1992
22 | 0.77611 | 43382.7891 | 0.66198 | 0.030578 | 22 | 43422.2422 | 0.66258 | 0.027268 | 22 | 43464.2227 | 0.66322 | 0.023753 |  | 22 | 0.77611 | 0.66864 | -0.0057028 |  | 22 | 0.66065 | 0.037917 | 22 | 0.69777 | -0.15681 | 22 | 43880.4805 | 22 | 43148.3203 | 22 | 45640.0781
23 | 0.81139 | 43404.9141 | 0.66232 | 0.028721 | 23 | 43424.8477 | 0.66262 | 0.02705 | 23 | 43475.5195 | 0.66339 | 0.022808 |  | 23 | 0.81139 | 0.66775 | -0.00090292 |  | 23 | 0.6605 | 0.038721 | 23 | 0.6975 | -0.15546 | 23 | 43916.3984 | 23 | 43132 | 23 | 45648.3203
24 | 0.84667 | 43453.7695 | 0.66306 | 0.024627 | 24 | 43468.957 | 0.66329 | 0.023357 | 24 | 43508.6953 | 0.6639 | 0.020036 |  | 24 | 0.84667 | 0.66765 | -0.00039242 |  | 24 | 0.6605 | 0.038729 | 24 | 0.69642 | -0.15009 | 24 | 43894.8789 | 24 | 43148.8789 | 24 | 45625.0781
25 | 0.88194 | 43504.6172 | 0.66384 | 0.020376 | 25 | 43474.2148 | 0.66337 | 0.022917 | 25 | 43500.8828 | 0.66378 | 0.020688 |  | 25 | 0.88194 | 0.66797 | -0.0021002 |  | 25 | 0.66019 | 0.040474 | 25 | 0.69541 | -0.14503 | 25 | 43917.5195 | 25 | 43186.2383 | 25 | 45627.5586
26 | 0.91722 | 43529.1406 | 0.66421 | 0.018329 | 26 | 43460.332 | 0.66316 | 0.024078 | 26 | 43472.3555 | 0.66335 | 0.023072 |  | 26 | 0.91722 | 0.66784 | -0.0014175 |  | 26 | 0.66053 | 0.038591 | 26 | 0.69501 | -0.14303 | 26 | 43913.1602 | 26 | 43223.9219 | 26 | 45618.1602
27 | 0.9525 | 43458.5625 | 0.66314 | 0.024226 | 27 | 43496.0742 | 0.66371 | 0.02109 | 27 | 43456.5195 | 0.6631 | 0.024397 |  | 27 | 0.9525 | 0.66843 | -0.0045715 |  | 27 | 0.66012 | 0.040861 | 27 | 0.69439 | -0.1399 | 27 | 43860.8008 | 27 | 43237.8008 | 27 | 45565.7188
28 | 0.98778 | 43399.3789 | 0.66223 | 0.029185 | 28 | 43477.2148 | 0.66342 | 0.022666 | 28 | 43437.707 | 0.66282 | 0.025972 |  | 28 | 0.98778 | 0.66871 | -0.0061298 |  | 28 | 0.66032 | 0.039741 | 28 | 0.69402 | -0.13803 | 28 | 43813.0391 | 28 | 43274.7188 | 28 | 45550.1992
29 | 1.0231 | 43426.8633 | 0.66265 | 0.026881 | 29 | 43451.1484 | 0.66302 | 0.024847 | 29 | 43446.2461 | 0.66295 | 0.025257 |  | 29 | 1.0231 | 0.66856 | -0.0052721 |  | 29 | 0.66047 | 0.038871 | 29 | 0.69416 | -0.13875 | 29 | 43888.9219 | 29 | 43270.8789 | 29 | 45558.3594
30 | 1.0583 | 43467.2734 | 0.66327 | 0.023498 | 30 | 43443.1953 | 0.6629 | 0.025513 | 30 | 43500.6836 | 0.66378 | 0.020705 |  | 30 | 1.0583 | 0.66921 | -0.0088272 |  | 30 | 0.66109 | 0.035455 | 30 | 0.69383 | -0.13706 | 30 | 43784.0781 | 30 | 43226.2383 | 30 | 45585.3594
31 | 1.0936 | 43409.8906 | 0.66239 | 0.028304 | 31 | 43451.7383 | 0.66303 | 0.024798 | 31 | 43442.3555 | 0.66289 | 0.025583 |  | 31 | 1.0936 | 0.66988 | -0.01244 |  | 31 | 0.66103 | 0.035807 | 31 | 0.69373 | -0.1366 | 31 | 43861.6406 | 31 | 43175.7188 | 31 | 45545.5586
32 | 1.1289 | 43416.8047 | 0.6625 | 0.027724 | 32 | 43415.2656 | 0.66247 | 0.027853 | 32 | 43427.7461 | 0.66266 | 0.026807 |  | 32 | 1.1289 | 0.66937 | -0.0096461 |  | 32 | 0.66167 | 0.032267 | 32 | 0.69379 | -0.13688 | 32 | 43937.1211 | 32 | 43225.6797 | 32 | 45538.6797
33 | 1.1642 | 43451.5625 | 0.66303 | 0.024812 | 33 | 43412.1211 | 0.66243 | 0.028117 | 33 | 43446.9766 | 0.66296 | 0.025196 |  | 33 | 1.1642 | 0.66958 | -0.010789 |  | 33 | 0.66213 | 0.029721 | 33 | 0.69357 | -0.13579 | 33 | 43989.1211 | 33 | 43197 | 33 | 45543.5586
34 | 1.1994 | 43489.2539 | 0.6636 | 0.02166 | 34 | 43448.3438 | 0.66298 | 0.025082 | 34 | 43478.1953 | 0.66343 | 0.022584 |  | 34 | 1.1994 | 0.66924 | -0.0089394 |  | 34 | 0.66136 | 0.03399 | 34 | 0.69383 | -0.1371 | 34 | 43923.1992 | 34 | 43270 | 34 | 45612.1211
35 | 1.2347 | 43480.793 | 0.66347 | 0.022367 | 35 | 43467.5039 | 0.66327 | 0.023478 | 35 | 43423.5742 | 0.6626 | 0.027157 |  | 35 | 1.2347 | 0.66935 | -0.0095452 |  | 35 | 0.66089 | 0.036596 | 35 | 0.69397 | -0.13779 | 35 | 43868.3203 | 35 | 43254.8789 | 35 | 45668.8008
36 | 1.27 | 43394.8203 | 0.66216 | 0.029568 | 36 | 43469.9297 | 0.66331 | 0.023275 | 36 | 43378.832 | 0.66192 | 0.03091 |  | 36 | 1.27 | 0.66907 | -0.0080382 |  | 36 | 0.66006 | 0.04116 | 36 | 0.69351 | -0.13545 | 36 | 43887 | 36 | 43239.1602 | 36 | 45661.1992
37 | 1.3053 | 43423.8164 | 0.6626 | 0.027136 | 37 | 43369.6875 | 0.66178 | 0.031678 | 37 | 43381.8047 | 0.66196 | 0.03066 |  | 37 | 1.3053 | 0.66913 | -0.0083748 |  | 37 | 0.66041 | 0.039235 | 37 | 0.69323 | -0.13407 | 37 | 43934.0781 | 37 | 43193.8008 | 37 | 45677.9609
38 | 1.3406 | 43470.6016 | 0.66332 | 0.023219 | 38 | 43390.0039 | 0.66209 | 0.029972 | 38 | 43456.25 | 0.6631 | 0.02442 |  | 38 | 1.3406 | 0.6691 | -0.0082253 |  | 38 | 0.65946 | 0.044502 | 38 | 0.69262 | -0.13098 | 38 | 43922.5195 | 38 | 43168.2383 | 38 | 45642.8008
39 | 1.3758 | 43287.0156 | 0.66052 | 0.038636 | 39 | 43478.2031 | 0.66343 | 0.022583 | 39 | 43362.707 | 0.66167 | 0.032265 |  | 39 | 1.3758 | 0.66887 | -0.0069534 |  | 39 | 0.65869 | 0.04875 | 39 | 0.69244 | -0.13009 | 39 | 43951.4414 | 39 | 43207.2812 | 39 | 45585.7188
40 | 1.4111 | 43383.4102 | 0.66199 | 0.030526 | 40 | 43363.7266 | 0.66169 | 0.032179 | 40 | 43454.457 | 0.66307 | 0.02457 |  | 40 | 1.4111 | 0.66917 | -0.0086142 |  | 40 | 0.65971 | 0.043122 | 40 | 0.69288 | -0.13228 | 40 | 43911.4414 | 40 | 43178.7188 | 40 | 45541.5586
41 | 1.4464 | 43424.4727 | 0.66261 | 0.027081 | 41 | 43433.1445 | 0.66275 | 0.026355 | 41 | 43494.6445 | 0.66369 | 0.021209 |  | 41 | 1.4464 | 0.66839 | -0.0043953 |  | 41 | 0.66035 | 0.039588 | 41 | 0.69273 | -0.13153 | 41 | 43843.2383 | 41 | 43211.3594 | 41 | 45518.6797
42 | 1.4817 | 43419.0352 | 0.66253 | 0.027537 | 42 | 43484.9102 | 0.66354 | 0.022023 | 42 | 43369.9766 | 0.66178 | 0.031654 |  | 42 | 1.4817 | 0.66803 | -0.002423 |  | 42 | 0.66079 | 0.037116 | 42 | 0.69321 | -0.13396 | 42 | 43825.9219 | 42 | 43235.8008 | 42 | 45474.6016
43 | 1.5169 | 43437.9141 | 0.66282 | 0.025955 | 43 | 43408.4609 | 0.66237 | 0.028424 | 43 | 43469.2148 | 0.6633 | 0.023335 |  | 43 | 1.5169 | 0.66886 | -0.0069344 |  | 43 | 0.6605 | 0.038729 | 43 | 0.69304 | -0.13311 | 43 | 43826.6016 | 43 | 43225.0391 | 43 | 45463.7188
44 | 1.5522 | 43425.0352 | 0.66262 | 0.027034 | 44 | 43448.1406 | 0.66298 | 0.025099 | 44 | 43445.9414 | 0.66294 | 0.025283 |  | 44 | 1.5522 | 0.66932 | -0.0093917 |  | 44 | 0.66013 | 0.040781 | 44 | 0.69272 | -0.13146 | 44 | 43776.8008 | 44 | 43227.8789 | 44 | 45478.6797
45 | 1.5875 | 43399.0078 | 0.66223 | 0.029217 | 45 | 43451.957 | 0.66303 | 0.024779 | 45 | 43378.875 | 0.66192 | 0.030906 |  | 45 | 1.5875 | 0.66942 | -0.0099486 |  | 45 | 0.6596 | 0.043702 | 45 | 0.69247 | -0.13019 | 45 | 43723.6797 | 45 | 43232.4414 | 45 | 45423.1992
46 | 1.6228 | 43434.2227 | 0.66276 | 0.026264 | 46 | 43435.3984 | 0.66278 | 0.026166 | 46 | 43376.4141 | 0.66188 | 0.031113 |  | 46 | 1.6228 | 0.66947 | -0.010225 |  | 46 | 0.66025 | 0.040113 | 46 | 0.69218 | -0.12875 | 46 | 43735.5195 | 46 | 43224.5195 | 46 | 45377.3594
47 | 1.6581 | 43386.4258 | 0.66203 | 0.030273 | 47 | 43431.5859 | 0.66272 | 0.026485 | 47 | 43364.3906 | 0.6617 | 0.032123 |  | 47 | 1.6581 | 0.66884 | -0.0068186 |  | 47 | 0.6598 | 0.042623 | 47 | 0.69193 | -0.12748 | 47 | 43735.9609 | 47 | 43227.9219 | 47 | 45406.2812
48 | 1.6933 | 43397.1094 | 0.6622 | 0.029376 | 48 | 43407.1758 | 0.66235 | 0.028531 | 48 | 43362.3867 | 0.66167 | 0.032292 |  | 48 | 1.6933 | 0.66849 | -0.0049048 |  | 48 | 0.65975 | 0.042892 | 48 | 0.69205 | -0.12809 | 48 | 43712.4414 | 48 | 43218.6016 | 48 | 45411
49 | 1.7286 | 43485.6836 | 0.66355 | 0.021958 | 49 | 43391.1094 | 0.66211 | 0.029879 | 49 | 43431.793 | 0.66273 | 0.026468 |  | 49 | 1.7286 | 0.66875 | -0.0063393 |  | 49 | 0.6601 | 0.04093 | 49 | 0.69255 | -0.13062 | 49 | 43751 | 49 | 43206.7188 | 49 | 45405
50 | 1.7639 | 43503.3672 | 0.66382 | 0.020481 | 50 | 43428.332 | 0.66267 | 0.026758 | 50 | 43470.1953 | 0.66331 | 0.023253 |  | 50 | 1.7639 | 0.66933 | -0.0094479 |  | 50 | 0.66041 | 0.039208 | 50 | 0.69311 | -0.13346 | 50 | 43771.9219 | 50 | 43192.5195 | 50 | 45428.6016
51 | 1.7992 | 43460.0859 | 0.66316 | 0.024099 | 51 | 43464.2031 | 0.66322 | 0.023754 | 51 | 43480.4062 | 0.66347 | 0.022399 |  | 51 | 1.7992 | 0.66913 | -0.0083899 |  | 51 | 0.66027 | 0.03999 | 51 | 0.69293 | -0.13252 | 51 | 43795.7617 | 51 | 43228.5195 | 51 | 45412.0781
52 | 1.8344 | 43444.6367 | 0.66292 | 0.025392 | 52 | 43474.582 | 0.66338 | 0.022886 | 52 | 43480.9648 | 0.66348 | 0.022353 |  | 52 | 1.8344 | 0.66901 | -0.0077277 |  | 52 | 0.66046 | 0.038974 | 52 | 0.69277 | -0.13172 | 52 | 43811.8398 | 52 | 43245.5586 | 52 | 45394.6016
53 | 1.8697 | 43434.0039 | 0.66276 | 0.026283 | 53 | 43430.6562 | 0.66271 | 0.026563 | 53 | 43457.4023 | 0.66312 | 0.024323 |  | 53 | 1.8697 | 0.66899 | -0.0076081 |  | 53 | 0.66054 | 0.038507 | 53 | 0.69216 | -0.12865 | 53 | 43885.0391 | 53 | 43283.9609 | 53 | 45369.9609
54 | 1.905 | 43428.8555 | 0.66268 | 0.026714 | 54 | 43434.375 | 0.66277 | 0.026252 | 54 | 43436.5117 | 0.6628 | 0.026073 |  | 54 | 1.905 | 0.66903 | -0.0078251 |  | 54 | 0.66039 | 0.039316 | 54 | 0.69214 | -0.12856 | 54 | 43846.9609 | 54 | 43300.7617 | 54 | 45356.7188
55 | 1.9403 | 43429.6641 | 0.66269 | 0.026646 | 55 | 43453.5156 | 0.66306 | 0.024649 | 55 | 43436.0938 | 0.66279 | 0.026108 |  | 55 | 1.9403 | 0.66956 | -0.010703 |  | 55 | 0.661 | 0.035972 | 55 | 0.69277 | -0.13175 | 55 | 43838.6406 | 55 | 43260.1211 | 55 | 45354.3594
56 | 1.9756 | 43368.5352 | 0.66176 | 0.031775 | 56 | 43439.9648 | 0.66285 | 0.025783 | 56 | 43416.4727 | 0.66249 | 0.027752 |  | 56 | 1.9756 | 0.67015 | -0.013879 |  | 56 | 0.6614 | 0.033749 | 56 | 0.69191 | -0.12737 | 56 | 43810.4414 | 56 | 43210.9219 | 56 | 45366.3594
57 | 2.0108 | 43401.8203 | 0.66227 | 0.028981 | 57 | 43374.4922 | 0.66185 | 0.031275 | 57 | 43458.7656 | 0.66314 | 0.024209 |  | 57 | 2.0108 | 0.66875 | -0.0063319 |  | 57 | 0.66161 | 0.03263 | 57 | 0.69154 | -0.12552 | 57 | 43814.7188 | 57 | 43166.0781 | 57 | 45380.7188
58 | 2.0461 | 43448 | 0.66297 | 0.02511 | 58 | 43445.0234 | 0.66293 | 0.02536 | 58 | 43461.3633 | 0.66318 | 0.023992 |  | 58 | 2.0461 | 0.6692 | -0.0087562 |  | 58 | 0.66232 | 0.028721 | 58 | 0.69221 | -0.12888 | 58 | 43844.8008 | 58 | 43183.1992 | 58 | 45347.2812
59 | 2.0814 | 43491.9414 | 0.66364 | 0.021435 | 59 | 43474.043 | 0.66337 | 0.022931 | 59 | 43480.2891 | 0.66347 | 0.022409 |  | 59 | 2.0814 | 0.66934 | -0.009515 |  | 59 | 0.66183 | 0.031392 | 59 | 0.69229 | -0.12932 | 59 | 43838.0391 | 59 | 43129.2383 | 59 | 45331.4414
60 | 2.1167 | 43486.0273 | 0.66355 | 0.021929 | 60 | 43469.8164 | 0.66331 | 0.023285 | 60 | 43453.2617 | 0.66305 | 0.02467 |  | 60 | 2.1167 | 0.66828 | -0.0037843 |  | 60 | 0.6612 | 0.034858 | 60 | 0.69206 | -0.12813 | 60 | 43823.6406 | 60 | 43133.3203 | 60 | 45319.6016
61 | 2.1519 | 43352.6445 | 0.66152 | 0.03311 | 61 | 43420.6211 | 0.66256 | 0.027404 | 61 | 43377.3633 | 0.6619 | 0.031033 |  | 61 | 2.1519 | 0.66829 | -0.0038368 |  | 61 | 0.6597 | 0.043176 | 61 | 0.69195 | -0.12757 | 61 | 43797.1211 | 61 | 43195.9219 | 61 | 45316.8789
62 | 2.1872 | 43465.1328 | 0.66324 | 0.023677 | 62 | 43351.2656 | 0.6615 | 0.033226 | 62 | 43414.1094 | 0.66246 | 0.02795 |  | 62 | 2.1872 | 0.66876 | -0.0063505 |  | 62 | 0.65992 | 0.041912 | 62 | 0.69259 | -0.13081 | 62 | 43830.7617 | 62 | 43223.2812 | 62 | 45333.8398
63 | 2.2225 | 43504.207 | 0.66383 | 0.02041 | 63 | 43452.5352 | 0.66304 | 0.024731 | 63 | 43491.8906 | 0.66364 | 0.021439 |  | 63 | 2.2225 | 0.66837 | -0.0042716 |  | 63 | 0.65789 | 0.053212 | 63 | 0.69263 | -0.13101 | 63 | 43791.1992 | 63 | 43226 | 63 | 45321.5195
64 | 2.2578 | 43469.6797 | 0.6633 | 0.023296 | 64 | 43474.5469 | 0.66338 | 0.022889 | 64 | 43449.8438 | 0.663 | 0.024956 |  | 64 | 2.2578 | 0.66818 | -0.003252 |  | 64 | 0.65944 | 0.044625 | 64 | 0.69191 | -0.12739 | 64 | 43741.1992 | 64 | 43185.8008 | 64 | 45327.6406
65 | 2.2931 | 43468.0938 | 0.66328 | 0.023429 | 65 | 43444.5 | 0.66292 | 0.025404 | 65 | 43433.8906 | 0.66276 | 0.026292 |  | 65 | 2.2931 | 0.66871 | -0.0061298 |  | 65 | 0.65885 | 0.04786 | 65 | 0.69152 | -0.12539 | 65 | 43773.8398 | 65 | 43158.8398 | 65 | 45344.5586
66 | 2.3283 | 43394.7695 | 0.66216 | 0.029572 | 66 | 43430.4492 | 0.66271 | 0.02658 | 66 | 43433.543 | 0.66275 | 0.026321 |  | 66 | 2.3283 | 0.66952 | -0.010472 |  | 66 | 0.6588 | 0.048134 | 66 | 0.6904 | -0.11973 | 66 | 43775.6016 | 66 | 43147.6797 | 66 | 45339.3594
67 | 2.3636 | 43431.082 | 0.66272 | 0.026527 | 67 | 43445.1641 | 0.66293 | 0.025348 | 67 | 43480.668 | 0.66347 | 0.022377 |  | 67 | 2.3636 | 0.66888 | -0.0070357 |  | 67 | 0.65901 | 0.046967 | 67 | 0.68937 | -0.11447 | 67 | 43794.6016 | 67 | 43187.4414 | 67 | 45346.6016
68 | 2.3989 | 43426.6289 | 0.66265 | 0.026901 | 68 | 43464.6406 | 0.66323 | 0.023718 | 68 | 43506.0391 | 0.66386 | 0.020257 |  | 68 | 2.3989 | 0.66941 | -0.009885 |  | 68 | 0.65887 | 0.047741 | 68 | 0.6889 | -0.11206 | 68 | 43792.1992 | 68 | 43165.4805 | 68 | 45317.8789
69 | 2.4342 | 43433.3555 | 0.66275 | 0.026337 | 69 | 43434.2852 | 0.66276 | 0.026259 | 69 | 43510.5859 | 0.66393 | 0.019878 |  | 69 | 2.4342 | 0.66964 | -0.011099 |  | 69 | 0.66002 | 0.041387 | 69 | 0.68947 | -0.11501 | 69 | 43754.4805 | 69 | 43158.0781 | 69 | 45316.4805
70 | 2.4694 | 43398.8047 | 0.66222 | 0.029234 | 70 | 43466.9336 | 0.66326 | 0.023526 | 70 | 43454.3672 | 0.66307 | 0.024577 |  | 70 | 2.4694 | 0.66966 | -0.011215 |  | 70 | 0.66168 | 0.032206 | 70 | 0.68999 | -0.11765 | 70 | 43745.0391 | 70 | 43199.3594 | 70 | 45332.1602
71 | 2.5047 | 43449.8398 | 0.663 | 0.024956 | 71 | 43442.8672 | 0.6629 | 0.02554 | 71 | 43456.7656 | 0.66311 | 0.024377 |  | 71 | 2.5047 | 0.66953 | -0.010509 |  | 71 | 0.66293 | 0.025358 | 71 | 0.69041 | -0.11978 | 71 | 43788.1211 | 71 | 43239.5586 | 71 | 45308.0781
72 | 2.54 | 43447.1094 | 0.66296 | 0.025185 | 72 | 43474.9102 | 0.66338 | 0.022859 | 72 | 43439.9922 | 0.66285 | 0.025781 |  | 72 | 2.54 | 0.66975 | -0.011712 |  | 72 | 0.66322 | 0.023771 | 72 | 0.69109 | -0.12325 | 72 | 43841.3594 | 72 | 43235.1211 | 72 | 45301.1992
73 | 2.5753 | 43452.2891 | 0.66304 | 0.024751 | 73 | 43419.8359 | 0.66254 | 0.02747 | 73 | 43439.7227 | 0.66285 | 0.025804 |  | 73 | 2.5753 | 0.67068 | -0.016711 |  | 73 | 0.66369 | 0.021206 | 73 | 0.69141 | -0.12487 | 73 | 43838.6797 | 73 | 43258.4805 | 73 | 45303.0781
74 | 2.6106 | 43411.5977 | 0.66242 | 0.028161 | 74 | 43406.5781 | 0.66234 | 0.028582 | 74 | 43464.7695 | 0.66323 | 0.023707 |  | 74 | 2.6106 | 0.67141 | -0.020624 |  | 74 | 0.66404 | 0.019263 | 74 | 0.69136 | -0.12458 | 74 | 43824.0781 | 74 | 43269.3203 | 74 | 45309.3203
75 | 2.6458 | 43429.3398 | 0.66269 | 0.026673 | 75 | 43424.6289 | 0.66262 | 0.027068 | 75 | 43498.2656 | 0.66374 | 0.020907 |  | 75 | 2.6458 | 0.67211 | -0.024363 |  | 75 | 0.66393 | 0.01987 | 75 | 0.69096 | -0.12258 | 75 | 43803.9609 | 75 | 43241.2383 | 75 | 45320.3594
76 | 2.6811 | 43389.4414 | 0.66208 | 0.030019 | 76 | 43448.3867 | 0.66298 | 0.025078 | 76 | 43426.6758 | 0.66265 | 0.026897 |  | 76 | 2.6811 | 0.67117 | -0.01935 |  | 76 | 0.66343 | 0.022585 | 76 | 0.69099 | -0.12271 | 76 | 43802.0391 | 76 | 43243.0781 | 76 | 45322.1602
77 | 2.7164 | 43371.75 | 0.66181 | 0.031505 | 77 | 43427.6562 | 0.66266 | 0.026814 | 77 | 43400.2617 | 0.66225 | 0.029111 |  | 77 | 2.7164 | 0.67145 | -0.020818 |  | 77 | 0.66256 | 0.027372 | 77 | 0.69107 | -0.12312 | 77 | 43819.9609 | 77 | 43260.6016 | 77 | 45328.0781
78 | 2.7517 | 43382.0156 | 0.66197 | 0.030643 | 78 | 43465.582 | 0.66324 | 0.023639 | 78 | 43421.5469 | 0.66257 | 0.027326 |  | 78 | 2.7517 | 0.67268 | -0.027402 |  | 78 | 0.66226 | 0.029011 | 78 | 0.6912 | -0.12381 | 78 | 43871.6406 | 78 | 43293.6406 | 78 | 45328.0391
79 | 2.7869 | 43418.5742 | 0.66252 | 0.027576 | 79 | 43441.9609 | 0.66288 | 0.025616 | 79 | 43419.9648 | 0.66255 | 0.027459 |  | 79 | 2.7869 | 0.67113 | -0.019149 |  | 79 | 0.66184 | 0.031358 | 79 | 0.69112 | -0.12339 | 79 | 43818.5195 | 79 | 43302.0781 | 79 | 45304.4805
80 | 2.8222 | 43397.5352 | 0.6622 | 0.02934 | 80 | 43431.043 | 0.66272 | 0.026531 | 80 | 43385.2656 | 0.66202 | 0.03037 |  | 80 | 2.8222 | 0.67137 | -0.020431 |  | 80 | 0.66179 | 0.03161 | 80 | 0.69181 | -0.12686 | 80 | 43807.0781 | 80 | 43287.7188 | 80 | 45261.3984
81 | 2.8575 | 43429.6328 | 0.66269 | 0.026649 | 81 | 43390.9688 | 0.6621 | 0.029891 | 81 | 43409.3945 | 0.66238 | 0.028345 |  | 81 | 2.8575 | 0.66913 | -0.0083526 |  | 81 | 0.6617 | 0.03211 | 81 | 0.69104 | -0.12297 | 81 | 43822.8398 | 81 | 43268 | 81 | 45240.8398
82 | 2.8928 | 43442.4336 | 0.66289 | 0.025577 | 82 | 43435.918 | 0.66279 | 0.026122 | 82 | 43421.0625 | 0.66256 | 0.027367 |  | 82 | 2.8928 | 0.67038 | -0.01511 |  | 82 | 0.66179 | 0.031621 | 82 | 0.69116 | -0.12359 | 82 | 43783.8008 | 82 | 43215.9609 | 82 | 45149.5586
83 | 2.9281 | 43488.8555 | 0.6636 | 0.021693 | 83 | 43395.9961 | 0.66218 | 0.029469 | 83 | 43449.4727 | 0.663 | 0.024987 |  | 83 | 2.9281 | 0.66994 | -0.012716 |  | 83 | 0.66165 | 0.032404 | 83 | 0.69087 | -0.12213 | 83 | 43801.3984 | 83 | 43191.4414 | 83 | 45158.2383
84 | 2.9633 | 43488.9766 | 0.6636 | 0.021683 | 84 | 43466.6641 | 0.66326 | 0.023548 | 84 | 43493.1055 | 0.66366 | 0.021338 |  | 84 | 2.9633 | 0.66979 | -0.011943 |  | 84 | 0.66163 | 0.032477 | 84 | 0.69127 | -0.12412 | 84 | 43792.8398 | 84 | 43132.3203 | 84 | 45145.6406
85 | 2.9986 | 43402.1602 | 0.66227 | 0.028952 | 85 | 43464.0234 | 0.66322 | 0.023769 | 85 | 43450.9648 | 0.66302 | 0.024862 |  | 85 | 2.9986 | 0.66935 | -0.0095452 |  | 85 | 0.66107 | 0.035566 | 85 | 0.69051 | -0.12028 | 85 | 43749.2383 | 85 | 43163.3984 | 85 | 45177.3594
86 | 3.0339 | 43431.4609 | 0.66272 | 0.026496 | 86 | 43445.5938 | 0.66294 | 0.025312 | 86 | 43475.7344 | 0.6634 | 0.02279 |  | 86 | 3.0339 | 0.66998 | -0.012977 |  | 86 | 0.66125 | 0.034583 | 86 | 0.69034 | -0.11943 | 86 | 43750.8398 | 86 | 43163.8789 | 86 | 45204.3594
87 | 3.0692 | 43461.1719 | 0.66317 | 0.024008 | 87 | 43459.9453 | 0.66316 | 0.024111 | 87 | 43487.6602 | 0.66358 | 0.021793 |  | 87 | 3.0692 | 0.67006 | -0.013406 |  | 87 | 0.66097 | 0.036128 | 87 | 0.68933 | -0.11429 | 87 | 43804.3594 | 87 | 43121.6406 | 87 | 45236.5195
88 | 3.1044 | 43366.4062 | 0.66173 | 0.031954 | 88 | 43477.7461 | 0.66343 | 0.022622 | 88 | 43432.6992 | 0.66274 | 0.026392 |  | 88 | 3.1044 | 0.66946 | -0.010177 |  | 88 | 0.66076 | 0.0373 | 88 | 0.68962 | -0.11575 | 88 | 43854.8398 | 88 | 43150.6797 | 88 | 45265.5195
89 | 3.1397 | 43368.4844 | 0.66176 | 0.031779 | 89 | 43483.75 | 0.66352 | 0.02212 | 89 | 43369.2812 | 0.66177 | 0.031712 |  | 89 | 3.1397 | 0.66902 | -0.0077653 |  | 89 | 0.66022 | 0.040274 | 89 | 0.68955 | -0.1154 | 89 | 43801.0781 | 89 | 43165.8398 | 89 | 45270.5195
90 | 3.175 | 43413.8008 | 0.66245 | 0.027976 | 90 | 43327.543 | 0.66114 | 0.035222 | 90 | 43461.2383 | 0.66318 | 0.024002 |  | 90 | 3.175 | 0.66783 | -0.0013572 |  | 90 | 0.65969 | 0.043188 | 90 | 0.68985 | -0.11695 | 90 | 43819.8789 | 90 | 43230.6406 | 90 | 45250.7617
91 | 3.2103 | 43434.8281 | 0.66277 | 0.026214 | 91 | 43474.4219 | 0.66338 | 0.0229 | 91 | 43477.7227 | 0.66343 | 0.022624 |  | 91 | 3.2103 | 0.66703 | 0.0030019 |  | 91 | 0.65995 | 0.041763 | 91 | 0.68889 | -0.11203 | 91 | 43829.3984 | 91 | 43306.2812 | 91 | 45252.0781
92 | 3.2456 | 43448.7539 | 0.66299 | 0.025047 | 92 | 43449.1289 | 0.66299 | 0.025016 | 92 | 43492.3516 | 0.66365 | 0.021401 |  | 92 | 3.2456 | 0.66727 | 0.0016816 |  | 92 | 0.66002 | 0.041383 | 92 | 0.68875 | -0.11132 | 92 | 43835.9609 | 92 | 43347.2383 | 92 | 45220.5586
93 | 3.2808 | 43466.293 | 0.66325 | 0.02358 | 93 | 43484.6523 | 0.66353 | 0.022044 | 93 | 43478.1953 | 0.66343 | 0.022584 |  | 93 | 3.2808 | 0.66718 | 0.0021666 |  | 93 | 0.66051 | 0.038691 | 93 | 0.68871 | -0.11113 | 93 | 43813.1992 | 93 | 43389.4414 | 93 | 45211.9219
94 | 3.3161 | 43554.0664 | 0.66459 | 0.016251 | 94 | 43424.5781 | 0.66262 | 0.027072 | 94 | 43498.0352 | 0.66374 | 0.020926 |  | 94 | 3.3161 | 0.66703 | 0.0030054 |  | 94 | 0.66049 | 0.038783 | 94 | 0.68841 | -0.10957 | 94 | 43829.7188 | 94 | 43417.5195 | 94 | 45238.1602
95 | 3.3514 | 43595.4102 | 0.66522 | 0.012808 | 95 | 43487.8633 | 0.66358 | 0.021776 | 95 | 43473.0117 | 0.66336 | 0.023018 |  | 95 | 3.3514 | 0.66656 | 0.0055166 |  | 95 | 0.66056 | 0.0384 | 95 | 0.68786 | -0.10676 | 95 | 43901.7617 | 95 | 43432.8789 | 95 | 45278.8008
96 | 3.3867 | 43495.6328 | 0.6637 | 0.021127 | 96 | 43500.3359 | 0.66377 | 0.020734 | 96 | 43431.8516 | 0.66273 | 0.026463 |  | 96 | 3.3867 | 0.66657 | 0.0054865 |  | 96 | 0.66073 | 0.037438 | 96 | 0.68742 | -0.10453 | 96 | 43982.1602 | 96 | 43422.5195 | 96 | 45290.3984
97 | 3.4219 | 43445.7578 | 0.66294 | 0.025298 | 97 | 43425.0039 | 0.66262 | 0.027037 | 97 | 43445.7812 | 0.66294 | 0.025296 |  | 97 | 3.4219 | 0.66683 | 0.0040594 |  | 97 | 0.66068 | 0.037717 | 97 | 0.68726 | -0.10371 | 97 | 44009.4805 | 97 | 43383.7188 | 97 | 45263.7617
98 | 3.4572 | 43457.2148 | 0.66311 | 0.024339 | 98 | 43452.4336 | 0.66304 | 0.024739 | 98 | 43504.4531 | 0.66384 | 0.02039 |  | 98 | 3.4572 | 0.66744 | 0.00077591 |  | 98 | 0.66055 | 0.038434 | 98 | 0.68726 | -0.10368 | 98 | 43972.0781 | 98 | 43334.2812 | 98 | 45224.7188
99 | 3.4925 | 43476.6445 | 0.66341 | 0.022714 | 99 | 43480.9492 | 0.66348 | 0.022354 | 99 | 43509.3008 | 0.66391 | 0.019985 |  | 99 | 3.4925 | 0.66818 | -0.0032594 |  | 99 | 0.66063 | 0.038024 | 99 | 0.6873 | -0.10389 | 99 | 43949.9219 | 99 | 43302.4805 | 99 | 45269.8008
100 | 3.5278 | 43479.2148 | 0.66345 | 0.022499 | 100 | 43437.3984 | 0.66281 | 0.025998 | 100 | 43479.5742 | 0.66346 | 0.022469 |  | 100 | 3.5278 | 0.66745 | 0.00068194 |  | 100 | 0.65995 | 0.041771 | 100 | 0.68759 | -0.10536 | 100 | 43980.3203 | 100 | 43335.8008 | 100 | 45220.1602
101 | 3.5631 | 43452.5664 | 0.66304 | 0.024728 | 101 | 43468.0039 | 0.66328 | 0.023436 | 101 | 43455.418 | 0.66309 | 0.02449 |  | 101 | 3.5631 | 0.66761 | -0.00017457 |  | 101 | 0.66035 | 0.039584 | 101 | 0.68784 | -0.10667 | 101 | 43919.1211 | 101 | 43348.0781 | 101 | 45159.4414
102 | 3.5983 | 43439.6875 | 0.66285 | 0.025807 | 102 | 43486.9531 | 0.66357 | 0.021852 | 102 | 43449.6641 | 0.663 | 0.024971 |  | 102 | 3.5983 | 0.66921 | -0.0087973 |  | 102 | 0.66065 | 0.037882 | 102 | 0.68745 | -0.10469 | 102 | 43855.5586 | 102 | 43308.5586 | 102 | 45155.6406
103 | 3.6336 | 43454.3828 | 0.66307 | 0.024576 | 103 | 43482.1523 | 0.6635 | 0.022253 | 103 | 43445.4883 | 0.66294 | 0.025321 |  | 103 | 3.6336 | 0.66929 | -0.0092536 |  | 103 | 0.66029 | 0.039895 | 103 | 0.68717 | -0.10323 | 103 | 43861.5586 | 103 | 43314.1992 | 103 | 45129.6016
104 | 3.6689 | 43450.2734 | 0.66301 | 0.02492 | 104 | 43475.3398 | 0.66339 | 0.022823 | 104 | 43428.4688 | 0.66268 | 0.026746 |  | 104 | 3.6689 | 0.66887 | -0.0069755 |  | 104 | 0.66081 | 0.037013 | 104 | 0.68784 | -0.10666 | 104 | 43862.8008 | 104 | 43315.9219 | 104 | 45141.1211
105 | 3.7042 | 43457.7656 | 0.66312 | 0.024293 | 105 | 43425.6953 | 0.66263 | 0.026979 | 105 | 43438.6523 | 0.66283 | 0.025893 |  | 105 | 3.7042 | 0.66896 | -0.0074361 |  | 105 | 0.66063 | 0.038016 | 105 | 0.68803 | -0.10763 | 105 | 43841.9609 | 105 | 43285.6406 | 105 | 45139.1992
106 | 3.7394 | 43387.0508 | 0.66204 | 0.03022 | 106 | 43443.4141 | 0.6629 | 0.025494 | 106 | 43379.6367 | 0.66193 | 0.030843 |  | 106 | 3.7394 | 0.66898 | -0.0075708 |  | 106 | 0.66123 | 0.034709 | 106 | 0.68821 | -0.10855 | 106 | 43854.2812 | 106 | 43257.6797 | 106 | 45153.7617
107 | 3.7747 | 43435.3516 | 0.66278 | 0.02617 | 107 | 43387.2383 | 0.66205 | 0.030204 | 107 | 43390.4492 | 0.6621 | 0.029935 |  | 107 | 3.7747 | 0.66873 | -0.0062157 |  | 107 | 0.661 | 0.035968 | 107 | 0.68797 | -0.10733 | 107 | 43791.7188 | 107 | 43269.7188 | 107 | 45113.7617
108 | 3.81 | 43461.0039 | 0.66317 | 0.024022 | 108 | 43394.7461 | 0.66216 | 0.029574 | 108 | 43424.3047 | 0.66261 | 0.027095 |  | 108 | 3.81 | 0.66793 | -0.0018751 |  | 108 | 0.66068 | 0.037752 | 108 | 0.68767 | -0.1058 | 108 | 43859.7617 | 108 | 43242.7617 | 108 | 45114.8008
109 | 3.8453 | 43442.168 | 0.66288 | 0.025599 | 109 | 43436.2422 | 0.66279 | 0.026095 | 109 | 43355.4062 | 0.66156 | 0.032878 |  | 109 | 3.8453 | 0.66701 | 0.0030846 |  | 109 | 0.66057 | 0.038342 | 109 | 0.68791 | -0.10703 | 109 | 43849.5195 | 109 | 43248.2812 | 109 | 45090.8789
110 | 3.8806 | 43456.4062 | 0.6631 | 0.024407 | 110 | 43408.6055 | 0.66237 | 0.028412 | 110 | 43498.2148 | 0.66374 | 0.020911 |  | 110 | 3.8806 | 0.66758 | -1.6724e-05 |  | 110 | 0.65989 | 0.042112 | 110 | 0.68858 | -0.11047 | 110 | 43813.7188 | 110 | 43206.2383 | 110 | 45037.0781
111 | 3.9158 | 43417.7695 | 0.66251 | 0.027643 | 111 | 43463.3164 | 0.66321 | 0.023829 | 111 | 43371.1484 | 0.6618 | 0.031555 |  | 111 | 3.9158 | 0.66752 | 0.00034743 |  | 111 | 0.66053 | 0.038564 | 111 | 0.68849 | -0.10997 | 111 | 43787.4414 | 111 | 43211.3203 | 111 | 45023.0781
112 | 3.9511 | 43440.9609 | 0.66287 | 0.0257 | 112 | 43406.1797 | 0.66234 | 0.028615 | 112 | 43449.4219 | 0.663 | 0.024991 |  | 112 | 3.9511 | 0.66643 | 0.0062587 |  | 112 | 0.66081 | 0.037013 | 112 | 0.68799 | -0.10742 | 112 | 43716.1992 | 112 | 43177.4414 | 112 | 45020.8398
113 | 3.9864 | 43400.0195 | 0.66224 | 0.029132 | 113 | 43421.7305 | 0.66257 | 0.027311 | 113 | 43440.6406 | 0.66286 | 0.025727 |  | 113 | 3.9864 | 0.66717 | 0.0022232 |  | 113 | 0.66025 | 0.040113 | 113 | 0.68689 | -0.1018 | 113 | 43676.7188 | 113 | 43195.6016 | 113 | 44998.0781
114 | 4.0217 | 43460.4844 | 0.66316 | 0.024066 | 114 | 43390.8789 | 0.6621 | 0.029899 | 114 | 43434.6445 | 0.66277 | 0.026229 |  | 114 | 4.0217 | 0.66714 | 0.0023773 |  | 114 | 0.65971 | 0.043118 | 114 | 0.68673 | -0.10098 | 114 | 43639.0391 | 114 | 43207.8789 | 114 | 44992
115 | 4.0569 | 43454.9297 | 0.66308 | 0.02453 | 115 | 43390.6367 | 0.6621 | 0.029919 | 115 | 43445.0391 | 0.66293 | 0.025358 |  | 115 | 4.0569 | 0.6668 | 0.0042325 |  | 115 | 0.65904 | 0.046844 | 115 | 0.68726 | -0.10369 | 115 | 43631.7188 | 115 | 43204.3203 | 115 | 45000.7617
116 | 4.0922 | 43484.6797 | 0.66353 | 0.022042 | 116 | 43437.3164 | 0.66281 | 0.026005 | 116 | 43464.75 | 0.66323 | 0.023709 |  | 116 | 4.0922 | 0.66677 | 0.0043755 |  | 116 | 0.65866 | 0.048904 | 116 | 0.68689 | -0.10182 | 116 | 43663.6797 | 116 | 43207.7188 | 116 | 45036.4414
117 | 4.1275 | 43431.5938 | 0.66272 | 0.026485 | 117 | 43445.0039 | 0.66293 | 0.025361 | 117 | 43444.5039 | 0.66292 | 0.025403 |  | 117 | 4.1275 | 0.66641 | 0.0063416 |  | 117 | 0.65888 | 0.047687 | 117 | 0.68706 | -0.10266 | 117 | 43693.6797 | 117 | 43214.9219 | 117 | 45025.6406
118 | 4.1628 | 43461.6602 | 0.66318 | 0.023967 | 118 | 43432.1289 | 0.66273 | 0.02644 | 118 | 43461.4414 | 0.66318 | 0.023985 |  | 118 | 4.1628 | 0.66691 | 0.0036075 |  | 118 | 0.65926 | 0.045621 | 118 | 0.68675 | -0.10105 | 118 | 43725 | 118 | 43214.3203 | 118 | 44988.5586
119 | 4.1981 | 43507.8516 | 0.66389 | 0.020106 | 119 | 43470.1602 | 0.66331 | 0.023256 | 119 | 43474.6602 | 0.66338 | 0.02288 |  | 119 | 4.1981 | 0.66842 | -0.0045303 |  | 119 | 0.65825 | 0.051221 | 119 | 0.68613 | -0.097883 | 119 | 43694.1602 | 119 | 43245.6797 | 119 | 45022.9219
120 | 4.2333 | 43481.2227 | 0.66348 | 0.022331 | 120 | 43488.3398 | 0.66359 | 0.021736 | 120 | 43485.5117 | 0.66355 | 0.021972 |  | 120 | 4.2333 | 0.66909 | -0.0081356 |  | 120 | 0.6582 | 0.051491 | 120 | 0.68565 | -0.095427 | 120 | 43651.7617 | 120 | 43236.1211 | 120 | 45013.0391
121 | 4.2686 | 43424.6328 | 0.66262 | 0.027068 | 121 | 43516.3008 | 0.66402 | 0.019401 | 121 | 43438.125 | 0.66282 | 0.025937 |  | 121 | 4.2686 | 0.6692 | -0.0087601 |  | 121 | 0.65945 | 0.044555 | 121 | 0.68532 | -0.093737 | 121 | 43682.6797 | 121 | 43201.8008 | 121 | 45031.9609
122 | 4.3039 | 43442.3633 | 0.66289 | 0.025582 | 122 | 43423.9258 | 0.66261 | 0.027127 | 122 | 43461.9844 | 0.66319 | 0.02394 |  | 122 | 4.3039 | 0.67006 | -0.013365 |  | 122 | 0.66062 | 0.038055 | 122 | 0.68563 | -0.095317 | 122 | 43799.8789 | 122 | 43229.4805 | 122 | 45069.4414
123 | 4.3392 | 43477.1484 | 0.66342 | 0.022672 | 123 | 43418.4688 | 0.66252 | 0.027584 | 123 | 43536.3398 | 0.66432 | 0.017729 |  | 123 | 4.3392 | 0.66855 | -0.0052457 |  | 123 | 0.6615 | 0.033241 | 123 | 0.68572 | -0.095802 | 123 | 43762.1211 | 123 | 43236.3203 | 123 | 45036.1992
124 | 4.3744 | 43444.2344 | 0.66292 | 0.025426 | 124 | 43463.7852 | 0.66321 | 0.023789 | 124 | 43486.8281 | 0.66357 | 0.021862 |  | 124 | 4.3744 | 0.66957 | -0.010756 |  | 124 | 0.66182 | 0.031468 | 124 | 0.68612 | -0.097855 | 124 | 43759 | 124 | 43216.8398 | 124 | 45053.4805
125 | 4.4097 | 43415.5352 | 0.66248 | 0.02783 | 125 | 43439.0703 | 0.66284 | 0.025858 | 125 | 43467.7773 | 0.66328 | 0.023455 |  | 125 | 4.4097 | 0.66859 | -0.0054556 |  | 125 | 0.66156 | 0.032897 | 125 | 0.68588 | -0.096613 | 125 | 43800.6797 | 125 | 43236 | 125 | 45079.8398
126 | 4.445 | 43484.6055 | 0.66353 | 0.022048 | 126 | 43430.3633 | 0.6627 | 0.026588 | 126 | 43511.9883 | 0.66395 | 0.019761 |  | 126 | 4.445 | 0.66878 | -0.0064815 |  | 126 | 0.66148 | 0.03331 | 126 | 0.68588 | -0.096623 | 126 | 43812.6406 | 126 | 43246.9219 | 126 | 45080.5586
127 | 4.4803 | 43522.1289 | 0.66411 | 0.018914 | 127 | 43500.7188 | 0.66378 | 0.020702 | 127 | 43499.7539 | 0.66376 | 0.020782 |  | 127 | 4.4803 | 0.66981 | -0.012059 |  | 127 | 0.66152 | 0.033084 | 127 | 0.68642 | -0.099395 | 127 | 43772.6016 | 127 | 43260.5586 | 127 | 45015.2383
128 | 4.5156 | 43524.2188 | 0.66414 | 0.01874 | 128 | 43469.4141 | 0.6633 | 0.023318 | 128 | 43468.8359 | 0.66329 | 0.023367 |  | 128 | 4.5156 | 0.67 | -0.013037 |  | 128 | 0.66193 | 0.030865 | 128 | 0.68661 | -0.10035 | 128 | 43817.0391 | 128 | 43251.9219 | 128 | 44955.1211
129 | 4.5508 | 43483.0195 | 0.66351 | 0.022181 | 129 | 43480.8906 | 0.66348 | 0.022359 | 129 | 43410.2305 | 0.6624 | 0.028275 |  | 129 | 4.5508 | 0.66953 | -0.010546 |  | 129 | 0.66188 | 0.031102 | 129 | 0.6862 | -0.098229 | 129 | 43804.0391 | 129 | 43228.3594 | 129 | 44982.5586
130 | 4.5861 | 43469.8242 | 0.66331 | 0.023284 | 130 | 43421.5039 | 0.66257 | 0.02733 | 130 | 43425.7734 | 0.66263 | 0.026972 |  | 130 | 4.5861 | 0.67042 | -0.015337 |  | 130 | 0.66211 | 0.029851 | 130 | 0.68586 | -0.096502 | 130 | 43765 | 130 | 43247.2812 | 130 | 45003.4414
131 | 4.6214 | 43454.5391 | 0.66307 | 0.024563 | 131 | 43440.7617 | 0.66286 | 0.025717 | 131 | 43425.5312 | 0.66263 | 0.026993 |  | 131 | 4.6214 | 0.66905 | -0.0079186 |  | 131 | 0.66302 | 0.024867 | 131 | 0.68597 | -0.097055 | 131 | 43726.3594 | 131 | 43226.2812 | 131 | 44969.5586
132 | 4.6567 | 43473.2656 | 0.66336 | 0.022996 | 132 | 43437.4844 | 0.66281 | 0.025991 | 132 | 43440.9688 | 0.66287 | 0.025699 |  | 132 | 4.6567 | 0.66807 | -0.002667 |  | 132 | 0.66227 | 0.02895 | 132 | 0.68631 | -0.098817 | 132 | 43711.6016 | 132 | 43243.1602 | 132 | 44978.1211
133 | 4.6919 | 43445.1602 | 0.66293 | 0.025348 | 133 | 43432.6992 | 0.66274 | 0.026392 | 133 | 43449.2773 | 0.66299 | 0.025004 |  | 133 | 4.6919 | 0.668 | -0.0022805 |  | 133 | 0.66216 | 0.029602 | 133 | 0.68624 | -0.098464 | 133 | 43665.2812 | 133 | 43253.0781 | 133 | 44983.6016
134 | 4.7272 | 43411.0586 | 0.66241 | 0.028206 | 134 | 43455.6758 | 0.66309 | 0.024468 | 134 | 43436.0898 | 0.66279 | 0.026108 |  | 134 | 4.7272 | 0.66767 | -0.00050893 |  | 134 | 0.66159 | 0.032714 | 134 | 0.68601 | -0.097284 | 134 | 43642.6406 | 134 | 43206.0781 | 134 | 44939.2812
135 | 4.7625 | 43431.0039 | 0.66271 | 0.026534 | 135 | 43418.7305 | 0.66253 | 0.027563 | 135 | 43435.2227 | 0.66278 | 0.026181 |  | 135 | 4.7625 | 0.66728 | 0.001648 |  | 135 | 0.66072 | 0.037518 | 135 | 0.68592 | -0.096838 | 135 | 43652.6016 | 135 | 43218.8789 | 135 | 44902.8789
136 | 4.7978 | 43499.3789 | 0.66376 | 0.020814 | 136 | 43441.5117 | 0.66287 | 0.025654 | 136 | 43466.0547 | 0.66325 | 0.023599 |  | 136 | 4.7978 | 0.66849 | -0.0049125 |  | 136 | 0.66035 | 0.039538 | 136 | 0.68556 | -0.094945 | 136 | 43649.2812 | 136 | 43161.7188 | 136 | 44884.9609
137 | 4.8331 | 43424.5352 | 0.66262 | 0.027076 | 137 | 43466.1094 | 0.66325 | 0.023595 | 137 | 43434.7266 | 0.66277 | 0.026222 |  | 137 | 4.8331 | 0.66799 | -0.0022318 |  | 137 | 0.66077 | 0.037254 | 137 | 0.68526 | -0.093426 | 137 | 43641.8398 | 137 | 43142.4805 | 137 | 44902.6016
138 | 4.8683 | 43358.1992 | 0.6616 | 0.032644 | 138 | 43476.2539 | 0.66341 | 0.022746 | 138 | 43391.9219 | 0.66212 | 0.029811 |  | 138 | 4.8683 | 0.66718 | 0.0021779 |  | 138 | 0.66088 | 0.036661 | 138 | 0.6854 | -0.094123 | 138 | 43622.6406 | 138 | 43095.8789 | 138 | 44918.8398
139 | 4.9036 | 43393.3594 | 0.66214 | 0.029691 | 139 | 43434.5664 | 0.66277 | 0.026235 | 139 | 43404.6172 | 0.66231 | 0.028746 |  | 139 | 4.9036 | 0.66659 | 0.0053657 |  | 139 | 0.66019 | 0.040439 | 139 | 0.68506 | -0.092407 | 139 | 43663.0781 | 139 | 43108.0391 | 139 | 44913.4414
140 | 4.9389 | 43347.6289 | 0.66144 | 0.033532 | 140 | 43354.5859 | 0.66155 | 0.032947 | 140 | 43392 | 0.66212 | 0.029805 |  | 140 | 4.9389 | 0.66674 | 0.0045337 |  | 140 | 0.66055 | 0.038461 | 140 | 0.68568 | -0.095595 | 140 | 43693.3594 | 140 | 43079.8008 | 140 | 44906.8398
141 | 4.9742 | 43418.7852 | 0.66253 | 0.027558 | 141 | 43358.707 | 0.66161 | 0.032601 | 141 | 43436.8594 | 0.6628 | 0.026043 |  | 141 | 4.9742 | 0.66706 | 0.002795 |  | 141 | 0.66036 | 0.039492 | 141 | 0.68619 | -0.098215 | 141 | 43666.1602 | 141 | 43063.6406 | 141 | 44916.5586
142 | 5.0094 | 43432.9453 | 0.66274 | 0.026371 | 142 | 43421.0117 | 0.66256 | 0.027371 | 142 | 43426.2773 | 0.66264 | 0.02693 |  | 142 | 5.0094 | 0.66656 | 0.0055467 |  | 142 | 0.66068 | 0.03776 | 142 | 0.68603 | -0.097355 | 142 | 43751.9219 | 142 | 43144.6016 | 142 | 44914.1992
143 | 5.0447 | 43432.2227 | 0.66273 | 0.026432 | 143 | 43447.0234 | 0.66296 | 0.025192 | 143 | 43431.6562 | 0.66272 | 0.026479 |  | 143 | 5.0447 | 0.66707 | 0.0027458 |  | 143 | 0.66102 | 0.035883 | 143 | 0.68603 | -0.097394 | 143 | 43767.6406 | 143 | 43213.7188 | 143 | 44930.3984
144 | 5.08 | 43431.9648 | 0.66273 | 0.026453 | 144 | 43424.3828 | 0.66261 | 0.027089 | 144 | 43409.3633 | 0.66238 | 0.028348 |  | 144 | 5.08 | 0.6674 | 0.00096742 |  | 144 | 0.6605 | 0.038729 | 144 | 0.68536 | -0.093941 | 144 | 43731.7188 | 144 | 43259.1211 | 144 | 44935.5195
145 | 5.1153 | 43406.3672 | 0.66234 | 0.028599 | 145 | 43428.8789 | 0.66268 | 0.026712 | 145 | 43404.4102 | 0.66231 | 0.028763 |  | 145 | 5.1153 | 0.66813 | -0.0029819 |  | 145 | 0.65968 | 0.043268 | 145 | 0.68473 | -0.090703 | 145 | 43864.6016 | 145 | 43264 | 145 | 44945.4805
146 | 5.1506 | 43404.6953 | 0.66231 | 0.028739 | 146 | 43427.6367 | 0.66266 | 0.026816 | 146 | 43390.2461 | 0.66209 | 0.029952 |  | 146 | 5.1506 | 0.66708 | 0.0027271 |  | 146 | 0.65975 | 0.042857 | 146 | 0.68452 | -0.089589 | 146 | 43820.3203 | 146 | 43303.3203 | 146 | 44950.6406
147 | 5.1858 | 43390.457 | 0.6621 | 0.029934 | 147 | 43397.1758 | 0.6622 | 0.02937 | 147 | 43367.9961 | 0.66175 | 0.03182 |  | 147 | 5.1858 | 0.66603 | 0.0084343 |  | 147 | 0.65922 | 0.045798 | 147 | 0.68389 | -0.086362 | 147 | 43865.6016 | 147 | 43307.1992 | 147 | 44919.2812
148 | 5.2211 | 43429.0234 | 0.66268 | 0.0267 | 148 | 43401.5547 | 0.66227 | 0.029003 | 148 | 43406.1367 | 0.66234 | 0.028619 |  | 148 | 5.2211 | 0.6664 | 0.0064171 |  | 148 | 0.65957 | 0.043864 | 148 | 0.68344 | -0.084006 | 148 | 43793.9609 | 148 | 43286.7617 | 148 | 44920.1992
149 | 5.2564 | 43437.1172 | 0.66281 | 0.026022 | 149 | 43431.3594 | 0.66272 | 0.026504 | 149 | 43425.3203 | 0.66263 | 0.02701 |  | 149 | 5.2564 | 0.66673 | 0.00459 |  | 149 | 0.66082 | 0.036952 | 149 | 0.68347 | -0.084164 | 149 | 43825.9219 | 149 | 43296.3594 | 149 | 44922.3203
150 | 5.2917 | 43399.7812 | 0.66224 | 0.029152 | 150 | 43433.4297 | 0.66275 | 0.026331 | 150 | 43436.9766 | 0.66281 | 0.026034 |  | 150 | 5.2917 | 0.6671 | 0.0026256 |  | 150 | 0.66053 | 0.038576 | 150 | 0.68348 | -0.084222 | 150 | 43827.0391 | 150 | 43357.2383 | 150 | 44907.9219
151 | 5.3269 | 43393.0195 | 0.66214 | 0.029719 | 151 | 43410.6758 | 0.6624 | 0.028238 | 151 | 43416.6445 | 0.6625 | 0.027737 |  | 151 | 5.3269 | 0.66708 | 0.002697 |  | 151 | 0.65983 | 0.042446 | 151 | 0.68365 | -0.085077 | 151 | 43806.5586 | 151 | 43362.5195 | 151 | 44895.0781
152 | 5.3622 | 43407.293 | 0.66235 | 0.028522 | 152 | 43407.0312 | 0.66235 | 0.028544 | 152 | 43418.2617 | 0.66252 | 0.027602 |  | 152 | 5.3622 | 0.66573 | 0.01005 |  | 152 | 0.65921 | 0.045855 | 152 | 0.68322 | -0.082842 | 152 | 43748.9609 | 152 | 43330.3203 | 152 | 44842.1602
153 | 5.3975 | 43432.7969 | 0.66274 | 0.026384 | 153 | 43401.8789 | 0.66227 | 0.028976 | 153 | 43413.7031 | 0.66245 | 0.027984 |  | 153 | 5.3975 | 0.66665 | 0.0050194 |  | 153 | 0.6591 | 0.046459 | 153 | 0.68323 | -0.082917 | 153 | 43787.3203 | 153 | 43314.2383 | 153 | 44858.5195
154 | 5.4328 | 43460.5625 | 0.66317 | 0.024059 | 154 | 43427.2734 | 0.66266 | 0.026847 | 154 | 43420.9648 | 0.66256 | 0.027375 |  | 154 | 5.4328 | 0.66699 | 0.0032124 |  | 154 | 0.65919 | 0.04597 | 154 | 0.68405 | -0.087187 | 154 | 43774.4805 | 154 | 43290.6016 | 154 | 44859.9219
155 | 5.4681 | 43476.4414 | 0.66341 | 0.022731 | 155 | 43454.1953 | 0.66307 | 0.024592 | 155 | 43465.2305 | 0.66324 | 0.023668 |  | 155 | 5.4681 | 0.66782 | -0.0012784 |  | 155 | 0.65903 | 0.046886 | 155 | 0.68359 | -0.084779 | 155 | 43690.0781 | 155 | 43283.2812 | 155 | 44859.5586
156 | 5.5033 | 43363.1562 | 0.66168 | 0.032227 | 156 | 43458.5742 | 0.66314 | 0.024225 | 156 | 43374.9922 | 0.66186 | 0.031233 |  | 156 | 5.5033 | 0.66743 | 0.00083242 |  | 156 | 0.65943 | 0.044644 | 156 | 0.68322 | -0.08287 | 156 | 43677.4414 | 156 | 43251.7188 | 156 | 44903.0391
157 | 5.5386 | 43426.8398 | 0.66265 | 0.026883 | 157 | 43355.8555 | 0.66157 | 0.032841 | 157 | 43449.4883 | 0.663 | 0.024986 |  | 157 | 5.5386 | 0.66616 | 0.0076839 |  | 157 | 0.65926 | 0.045617 | 157 | 0.68384 | -0.086093 | 157 | 43695.1211 | 157 | 43210.1992 | 157 | 44884.6016
158 | 5.5739 | 43517.5234 | 0.66403 | 0.019299 | 158 | 43470.5156 | 0.66332 | 0.023226 | 158 | 43477.6016 | 0.66343 | 0.022634 |  | 158 | 5.5739 | 0.66609 | 0.0080874 |  | 158 | 0.65913 | 0.046332 | 158 | 0.68447 | -0.089335 | 158 | 43709.9219 | 158 | 43226.9219 | 158 | 44874.1992
159 | 5.6092 | 43406.9375 | 0.66235 | 0.028551 | 159 | 43513.5977 | 0.66397 | 0.019626 | 159 | 43425.5586 | 0.66263 | 0.02699 |  | 159 | 5.6092 | 0.66622 | 0.007401 |  | 159 | 0.65823 | 0.051337 | 159 | 0.68375 | -0.085598 | 159 | 43692.9609 | 159 | 43250.3594 | 159 | 44854.3203
160 | 5.6444 | 43453.8281 | 0.66306 | 0.024623 | 160 | 43386.9727 | 0.66204 | 0.030227 | 160 | 43407.4336 | 0.66235 | 0.02851 |  | 160 | 5.6444 | 0.66665 | 0.0050194 |  | 160 | 0.65871 | 0.048631 | 160 | 0.6835 | -0.084333 | 160 | 43695.9219 | 160 | 43208 | 160 | 44832.9609
161 | 5.6797 | 43392.7109 | 0.66213 | 0.029745 | 161 | 43492.8242 | 0.66366 | 0.021361 | 161 | 43382.2461 | 0.66197 | 0.030623 |  | 161 | 5.6797 | 0.66677 | 0.0043907 |  | 161 | 0.65861 | 0.049185 | 161 | 0.68385 | -0.086143 | 161 | 43667.6406 | 161 | 43217.0391 | 161 | 44804.4414
162 | 5.715 | 43482.2695 | 0.6635 | 0.022243 | 162 | 43428.4023 | 0.66267 | 0.026752 | 162 | 43421.6328 | 0.66257 | 0.027319 |  | 162 | 5.715 | 0.6676 | -9.1935e-05 |  | 162 | 0.65905 | 0.046759 | 162 | 0.68324 | -0.082978 | 162 | 43641.1602 | 162 | 43243.5586 | 162 | 44787.9609
163 | 5.7503 | 43435.5039 | 0.66278 | 0.026157 | 163 | 43454.293 | 0.66307 | 0.024584 | 163 | 43454.3633 | 0.66307 | 0.024578 |  | 163 | 5.7503 | 0.66764 | -0.0003098 |  | 163 | 0.65942 | 0.044709 | 163 | 0.68277 | -0.080543 | 163 | 43635.9609 | 163 | 43278.3203 | 163 | 44757.8789
164 | 5.7856 | 43401.0898 | 0.66226 | 0.029042 | 164 | 43461.6914 | 0.66318 | 0.023965 | 164 | 43389.1172 | 0.66208 | 0.030047 |  | 164 | 5.7856 | 0.66658 | 0.0054149 |  | 164 | 0.65983 | 0.042431 | 164 | 0.68323 | -0.082942 | 164 | 43633.3984 | 164 | 43261.4414 | 164 | 44754.7617
165 | 5.8208 | 43488.0117 | 0.66358 | 0.021763 | 165 | 43418.5781 | 0.66253 | 0.027575 | 165 | 43452.9883 | 0.66305 | 0.024693 |  | 165 | 5.8208 | 0.66765 | -0.00036241 |  | 165 | 0.65971 | 0.043111 | 165 | 0.68283 | -0.080841 | 165 | 43661.8398 | 165 | 43210.3984 | 165 | 44760.1602
166 | 5.8561 | 43433.5742 | 0.66275 | 0.026319 | 166 | 43429.6914 | 0.66269 | 0.026644 | 166 | 43427.4648 | 0.66266 | 0.02683 |  | 166 | 5.8561 | 0.66865 | -0.0058077 |  | 166 | 0.65965 | 0.043418 | 166 | 0.68305 | -0.081975 | 166 | 43737.9609 | 166 | 43150.7188 | 166 | 44749.0391
167 | 5.8914 | 43469.7695 | 0.66331 | 0.023289 | 167 | 43453.418 | 0.66306 | 0.024657 | 167 | 43464.7305 | 0.66323 | 0.02371 |  | 167 | 5.8914 | 0.66778 | -0.0010681 |  | 167 | 0.65974 | 0.042938 | 167 | 0.68326 | -0.083072 | 167 | 43804.0781 | 167 | 43118.3203 | 167 | 44739.7188
168 | 5.9267 | 43442.5156 | 0.66289 | 0.02557 | 168 | 43437.2305 | 0.66281 | 0.026012 | 168 | 43438.0977 | 0.66282 | 0.02594 |  | 168 | 5.9267 | 0.66687 | 0.0038333 |  | 168 | 0.65992 | 0.041951 | 168 | 0.68331 | -0.083341 | 168 | 43720 | 168 | 43116.2812 | 168 | 44739.1602
169 | 5.9619 | 43470.4805 | 0.66332 | 0.023229 | 169 | 43426.4609 | 0.66265 | 0.026915 | 169 | 43458.5273 | 0.66313 | 0.024229 |  | 169 | 5.9619 | 0.66574 | 0.0099896 |  | 169 | 0.66057 | 0.03835 | 169 | 0.68338 | -0.083683 | 169 | 43622.8008 | 169 | 43172.3203 | 169 | 44760.8398
170 | 5.9972 | 43481.6836 | 0.66349 | 0.022292 | 170 | 43396.6836 | 0.66219 | 0.029412 | 170 | 43461.9688 | 0.66319 | 0.023941 |  | 170 | 5.9972 | 0.66527 | 0.012559 |  | 170 | 0.66087 | 0.036714 | 170 | 0.68292 | -0.081295 | 170 | 43626.1992 | 170 | 43245.4805 | 170 | 44766.5586
171 | 6.0325 | 43478.668 | 0.66344 | 0.022545 | 171 | 43415.2031 | 0.66247 | 0.027858 | 171 | 43490.0586 | 0.66362 | 0.021592 |  | 171 | 6.0325 | 0.66539 | 0.011909 |  | 171 | 0.66077 | 0.037239 | 171 | 0.68218 | -0.077455 | 171 | 43672.8789 | 171 | 43226.9609 | 171 | 44711.9609
172 | 6.0678 | 43524.9844 | 0.66415 | 0.018676 | 172 | 43495.9414 | 0.66371 | 0.021101 | 172 | 43479.3398 | 0.66345 | 0.022488 |  | 172 | 6.0678 | 0.66478 | 0.015211 |  | 172 | 0.66042 | 0.039178 | 172 | 0.68206 | -0.076856 | 172 | 43653.5195 | 172 | 43187.3984 | 172 | 44776.7188
173 | 6.1031 | 43545.7422 | 0.66447 | 0.016945 | 173 | 43489.1445 | 0.6636 | 0.021669 | 173 | 43476.4688 | 0.66341 | 0.022728 |  | 173 | 6.1031 | 0.66355 | 0.021943 |  | 173 | 0.65933 | 0.045228 | 173 | 0.6819 | -0.076005 | 173 | 43666.0781 | 173 | 43195.9219 | 173 | 44771.1211
174 | 6.1383 | 43438.3945 | 0.66283 | 0.025915 | 174 | 43483.4062 | 0.66351 | 0.022148 | 174 | 43435.2891 | 0.66278 | 0.026175 |  | 174 | 6.1383 | 0.66377 | 0.020735 |  | 174 | 0.65854 | 0.049594 | 174 | 0.6818 | -0.075481 | 174 | 43616.1211 | 174 | 43158.7188 | 174 | 44739.5586
175 | 6.1736 | 43427.125 | 0.66266 | 0.026859 | 175 | 43442.9922 | 0.6629 | 0.02553 | 175 | 43447.1953 | 0.66296 | 0.025178 |  | 175 | 6.1736 | 0.66588 | 0.0092194 |  | 175 | 0.65824 | 0.051275 | 175 | 0.68146 | -0.073708 | 175 | 43645.0781 | 175 | 43158.1992 | 175 | 44733.7188
176 | 6.2089 | 43394.0977 | 0.66215 | 0.029629 | 176 | 43442.6094 | 0.66289 | 0.025562 | 176 | 43397.1016 | 0.6622 | 0.029377 |  | 176 | 6.2089 | 0.66585 | 0.009393 |  | 176 | 0.65809 | 0.052112 | 176 | 0.6811 | -0.071832 | 176 | 43638.9219 | 176 | 43125.7188 | 176 | 44782.0781
177 | 6.2442 | 43436.4844 | 0.6628 | 0.026075 | 177 | 43404.9414 | 0.66232 | 0.028719 | 177 | 43415.7656 | 0.66248 | 0.027811 |  | 177 | 6.2442 | 0.66654 | 0.0056221 |  | 177 | 0.6584 | 0.050373 | 177 | 0.68128 | -0.072772 | 177 | 43703.3984 | 177 | 43150.5586 | 177 | 44749.5195
178 | 6.2794 | 43425.4648 | 0.66263 | 0.026998 | 178 | 43433.8086 | 0.66276 | 0.026299 | 178 | 43407.3164 | 0.66235 | 0.02852 |  | 178 | 6.2794 | 0.66572 | 0.010095 |  | 178 | 0.65947 | 0.044409 | 178 | 0.68147 | -0.073766 | 178 | 43664.7188 | 178 | 43080.9609 | 178 | 44731.8008
179 | 6.3147 | 43403.2852 | 0.66229 | 0.028858 | 179 | 43439.9648 | 0.66285 | 0.025783 | 179 | 43441.8203 | 0.66288 | 0.025628 |  | 179 | 6.3147 | 0.66542 | 0.011742 |  | 179 | 0.66059 | 0.038254 | 179 | 0.68191 | -0.076073 | 179 | 43617.8789 | 179 | 43142.4414 | 179 | 44737.7188
180 | 6.35 | 43448.25 | 0.66298 | 0.02509 | 180 | 43406.4297 | 0.66234 | 0.028594 | 180 | 43447.4453 | 0.66297 | 0.025157 |  | 180 | 6.35 | 0.6654 | 0.011848 |  | 180 | 0.66116 | 0.035095 | 180 | 0.68209 | -0.076983 | 180 | 43624.3203 | 180 | 43097.2383 | 180 | 44711.2812
181 | 6.3853 | 43519.6836 | 0.66407 | 0.019118 | 181 | 43497.4492 | 0.66373 | 0.020975 | 181 | 43488.0859 | 0.66359 | 0.021757 |  | 181 | 6.3853 | 0.66472 | 0.015555 |  | 181 | 0.66112 | 0.035336 | 181 | 0.68248 | -0.07903 | 181 | 43605.6016 | 181 | 43104.0781 | 181 | 44729.0781
182 | 6.4206 | 43464.6445 | 0.66323 | 0.023717 | 182 | 43541.4062 | 0.6644 | 0.017306 | 182 | 43439.9023 | 0.66285 | 0.025789 |  | 182 | 6.4206 | 0.66416 | 0.018595 |  | 182 | 0.65969 | 0.043226 | 182 | 0.68352 | -0.084445 | 182 | 43642.5195 | 182 | 43111.3594 | 182 | 44740.5195
183 | 6.4558 | 43462.9414 | 0.6632 | 0.02386 | 183 | 43455.5586 | 0.66309 | 0.024478 | 183 | 43403.3984 | 0.66229 | 0.028848 |  | 183 | 6.4558 | 0.66453 | 0.016567 |  | 183 | 0.65844 | 0.050145 | 183 | 0.68309 | -0.082194 | 183 | 43669.5586 | 183 | 43138.8789 | 183 | 44716.6797
184 | 6.4911 | 43462.4883 | 0.6632 | 0.023898 | 184 | 43378.0625 | 0.66191 | 0.030975 | 184 | 43482.332 | 0.6635 | 0.022238 |  | 184 | 6.4911 | 0.66489 | 0.014602 |  | 184 | 0.6592 | 0.045955 | 184 | 0.6826 | -0.079671 | 184 | 43676.3984 | 184 | 43136.5195 | 184 | 44690.3984
185 | 6.5264 | 43387.8359 | 0.66206 | 0.030154 | 185 | 43472.0625 | 0.66334 | 0.023097 | 185 | 43392.2188 | 0.66212 | 0.029786 |  | 185 | 6.5264 | 0.6649 | 0.014579 |  | 185 | 0.6583 | 0.050939 | 185 | 0.68182 | -0.075615 | 185 | 43660.4414 | 185 | 43150.4805 | 185 | 44694.1211
186 | 6.5617 | 43423.5508 | 0.6626 | 0.027159 | 186 | 43446.293 | 0.66295 | 0.025253 | 186 | 43375.4766 | 0.66187 | 0.031192 |  | 186 | 6.5617 | 0.66499 | 0.014072 |  | 186 | 0.65855 | 0.049517 | 186 | 0.68095 | -0.071083 | 186 | 43660.2812 | 186 | 43154.4414 | 186 | 44667.8398
187 | 6.5969 | 43495.3945 | 0.6637 | 0.021147 | 187 | 43446.0391 | 0.66294 | 0.025275 | 187 | 43430.0273 | 0.6627 | 0.026616 |  | 187 | 6.5969 | 0.66524 | 0.012699 |  | 187 | 0.65901 | 0.046967 | 187 | 0.68106 | -0.071655 | 187 | 43685.2383 | 187 | 43155.6406 | 187 | 44657.3984
188 | 6.6322 | 43481.2773 | 0.66348 | 0.022326 | 188 | 43459.6719 | 0.66315 | 0.024134 | 188 | 43427.7617 | 0.66267 | 0.026806 |  | 188 | 6.6322 | 0.66574 | 0.01002 |  | 188 | 0.6591 | 0.046471 | 188 | 0.6803 | -0.067682 | 188 | 43705.3984 | 188 | 43181.2812 | 188 | 44655.2812
189 | 6.6675 | 43431.3828 | 0.66272 | 0.026502 | 189 | 43444.0391 | 0.66291 | 0.025442 | 189 | 43435.0195 | 0.66278 | 0.026198 |  | 189 | 6.6675 | 0.6656 | 0.01073 |  | 189 | 0.65993 | 0.041897 | 189 | 0.67962 | -0.064098 | 189 | 43680.9609 | 189 | 43195.0391 | 189 | 44634.7617
190 | 6.7028 | 43422.0117 | 0.66258 | 0.027287 | 190 | 43412.9727 | 0.66244 | 0.028045 | 190 | 43416.0352 | 0.66249 | 0.027789 |  | 190 | 6.7028 | 0.66512 | 0.013349 |  | 190 | 0.66004 | 0.041287 | 190 | 0.67943 | -0.063095 | 190 | 43612.8789 | 190 | 43229.1211 | 190 | 44591.8398
191 | 6.7381 | 43454.7031 | 0.66308 | 0.024549 | 191 | 43406.7578 | 0.66234 | 0.028566 | 191 | 43406.4258 | 0.66234 | 0.028594 |  | 191 | 6.7381 | 0.66481 | 0.015037 |  | 191 | 0.65989 | 0.042101 | 191 | 0.67928 | -0.062306 | 191 | 43547.6797 | 191 | 43258.7617 | 191 | 44599.4805
192 | 6.7733 | 43455.793 | 0.66309 | 0.024458 | 192 | 43477.7852 | 0.66343 | 0.022618 | 192 | 43399.1055 | 0.66223 | 0.029208 |  | 192 | 6.7733 | 0.66517 | 0.013107 |  | 192 | 0.65947 | 0.044429 | 192 | 0.67968 | -0.064414 | 192 | 43532.3984 | 192 | 43213.2383 | 192 | 44628.4805
193 | 6.8086 | 43446.6406 | 0.66295 | 0.025224 | 193 | 43458.1289 | 0.66313 | 0.024263 | 193 | 43408.7227 | 0.66237 | 0.028402 |  | 193 | 6.8086 | 0.66476 | 0.015355 |  | 193 | 0.65941 | 0.044778 | 193 | 0.68038 | -0.068088 | 193 | 43548.5195 | 193 | 43180.4414 | 193 | 44567.6797
194 | 6.8439 | 43450.1289 | 0.66301 | 0.024932 | 194 | 43420.8633 | 0.66256 | 0.027384 | 194 | 43414.8789 | 0.66247 | 0.027885 |  | 194 | 6.8439 | 0.66469 | 0.015699 |  | 194 | 0.65947 | 0.044421 | 194 | 0.68062 | -0.069317 | 194 | 43498.1602 | 194 | 43133.0781 | 194 | 44605.8008
195 | 6.8792 | 43440.6367 | 0.66286 | 0.025727 | 195 | 43443.3477 | 0.6629 | 0.0255 | 195 | 43423.7383 | 0.6626 | 0.027143 |  | 195 | 6.8792 | 0.66508 | 0.013568 |  | 195 | 0.65967 | 0.043306 | 195 | 0.681 | -0.071347 | 195 | 43481.1211 | 195 | 43055.6016 | 195 | 44633.8008
196 | 6.9144 | 43429.207 | 0.66269 | 0.026684 | 196 | 43427.7109 | 0.66266 | 0.02681 | 196 | 43442.3906 | 0.66289 | 0.02558 |  | 196 | 6.9144 | 0.66588 | 0.0092155 |  | 196 | 0.66029 | 0.039902 | 196 | 0.68103 | -0.071463 | 196 | 43445.3203 | 196 | 43051 | 196 | 44680.6016
197 | 6.9497 | 43447.7383 | 0.66297 | 0.025132 | 197 | 43409.7617 | 0.66239 | 0.028315 | 197 | 43441.1055 | 0.66287 | 0.025688 |  | 197 | 6.9497 | 0.66591 | 0.0090949 |  | 197 | 0.66094 | 0.036278 | 197 | 0.68059 | -0.069208 | 197 | 43547.1992 | 197 | 43085.6016 | 197 | 44751.3594
198 | 6.985 | 43466.1016 | 0.66325 | 0.023596 | 198 | 43378.9219 | 0.66192 | 0.030903 | 198 | 43445.1133 | 0.66293 | 0.025352 |  | 198 | 6.985 | 0.66515 | 0.013201 |  | 198 | 0.66207 | 0.03006 | 198 | 0.68047 | -0.068578 | 198 | 43534.2812 | 198 | 43127.2383 | 198 | 44721.4805
199 | 7.0203 | 43436.8867 | 0.6628 | 0.026041 | 199 | 43429.6445 | 0.66269 | 0.026648 | 199 | 43442.5703 | 0.66289 | 0.025565 |  | 199 | 7.0203 | 0.66463 | 0.016029 |  | 199 | 0.66196 | 0.030675 | 199 | 0.68123 | -0.072534 | 199 | 43596.3984 | 199 | 43167.1602 | 199 | 44706.2383
200 | 7.0556 | 43418.1211 | 0.66252 | 0.027614 | 200 | 43440.4688 | 0.66286 | 0.025741 | 200 | 43437.75 | 0.66282 | 0.025969 |  | 200 | 7.0556 | 0.6648 | 0.015135 |  | 200 | 0.66133 | 0.034136 | 200 | 0.68163 | -0.074618 | 200 | 43539.8789 | 200 | 43235.0781 | 200 | 44647.8008
201 | 7.0908 | 43479.8984 | 0.66346 | 0.022442 | 201 | 43418.375 | 0.66252 | 0.027592 | 201 | 43443.2891 | 0.6629 | 0.025505 |  | 201 | 7.0908 | 0.66374 | 0.020906 |  | 201 | 0.66108 | 0.035551 | 201 | 0.68134 | -0.073087 | 201 | 43534.4805 | 201 | 43205.7617 | 201 | 44605.3594
202 | 7.1261 | 43436.707 | 0.6628 | 0.026056 | 202 | 43434.3086 | 0.66277 | 0.026257 | 202 | 43401.2891 | 0.66226 | 0.029025 |  | 202 | 7.1261 | 0.6631 | 0.024395 |  | 202 | 0.66068 | 0.03776 | 202 | 0.68148 | -0.073809 | 202 | 43527.3594 | 202 | 43216.4414 | 202 | 44588.3203
203 | 7.1614 | 43442.8516 | 0.6629 | 0.025542 | 203 | 43452.2422 | 0.66304 | 0.024755 | 203 | 43419.6523 | 0.66254 | 0.027485 |  | 203 | 7.1614 | 0.66389 | 0.020098 |  | 203 | 0.66016 | 0.040612 | 203 | 0.68131 | -0.072931 | 203 | 43483.9609 | 203 | 43116.8789 | 203 | 44551.7188
204 | 7.1967 | 43446.2031 | 0.66295 | 0.025261 | 204 | 43450.9023 | 0.66302 | 0.024867 | 204 | 43457.2891 | 0.66312 | 0.024333 |  | 204 | 7.1967 | 0.66487 | 0.014726 |  | 204 | 0.65967 | 0.043349 | 204 | 0.68087 | -0.070642 | 204 | 43460.3984 | 204 | 43050.3984 | 204 | 44489.7617
205 | 7.2319 | 43384.6758 | 0.66201 | 0.030419 | 205 | 43419.8477 | 0.66254 | 0.027469 | 205 | 43438.0469 | 0.66282 | 0.025944 |  | 205 | 7.2319 | 0.66537 | 0.011992 |  | 205 | 0.65974 | 0.042961 | 205 | 0.6811 | -0.071846 | 205 | 43490.3203 | 205 | 43104.2383 | 205 | 44467
206 | 7.2672 | 43423.5234 | 0.6626 | 0.027161 | 206 | 43390.2891 | 0.66209 | 0.029948 | 206 | 43409.293 | 0.66238 | 0.028354 |  | 206 | 7.2672 | 0.66566 | 0.010424 |  | 206 | 0.65969 | 0.043222 | 206 | 0.68113 | -0.071984 | 206 | 43513.8008 | 206 | 43096.6016 | 206 | 44461.8398
207 | 7.3025 | 43456.2617 | 0.6631 | 0.024419 | 207 | 43383.9492 | 0.662 | 0.03048 | 207 | 43480.2227 | 0.66347 | 0.022415 |  | 207 | 7.3025 | 0.66579 | 0.0097365 |  | 207 | 0.65969 | 0.043207 | 207 | 0.68039 | -0.06815 | 207 | 43524.5195 | 207 | 43101.1602 | 207 | 44501.5586
208 | 7.3378 | 43401.2656 | 0.66226 | 0.029027 | 208 | 43510.2578 | 0.66392 | 0.019905 | 208 | 43456.9453 | 0.66311 | 0.024362 |  | 208 | 7.3378 | 0.66631 | 0.0069033 |  | 208 | 0.65974 | 0.042934 | 208 | 0.68049 | -0.068657 | 208 | 43539 | 208 | 43090.7188 | 208 | 44546.4414
209 | 7.3731 | 43409.4258 | 0.66239 | 0.028343 | 209 | 43403.3086 | 0.66229 | 0.028856 | 209 | 43405.4102 | 0.66232 | 0.028679 |  | 209 | 7.3731 | 0.66669 | 0.0048123 |  | 209 | 0.65867 | 0.048892 | 209 | 0.68077 | -0.070127 | 209 | 43542.8398 | 209 | 43150.3203 | 209 | 44568.0391
210 | 7.4083 | 43399.8828 | 0.66224 | 0.029143 | 210 | 43434.9375 | 0.66277 | 0.026204 | 210 | 43434.2109 | 0.66276 | 0.026265 |  | 210 | 7.4083 | 0.66594 | 0.00891 |  | 210 | 0.6584 | 0.050349 | 210 | 0.67987 | -0.065408 | 210 | 43548.8789 | 210 | 43166.2812 | 210 | 44594
211 | 7.4436 | 43439.332 | 0.66284 | 0.025836 | 211 | 43417.5742 | 0.66251 | 0.027659 | 211 | 43482.5156 | 0.6635 | 0.022223 |  | 211 | 7.4436 | 0.66561 | 0.010681 |  | 211 | 0.65878 | 0.048257 | 211 | 0.67967 | -0.064377 | 211 | 43540.8398 | 211 | 43182 | 211 | 44603.7617
212 | 7.4789 | 43423.9961 | 0.66261 | 0.027121 | 212 | 43457.9883 | 0.66313 | 0.024274 | 212 | 43376.0547 | 0.66188 | 0.031143 |  | 212 | 7.4789 | 0.66505 | 0.013731 |  | 212 | 0.6597 | 0.043153 | 212 | 0.67989 | -0.06551 | 212 | 43535.8789 | 212 | 43173.3203 | 212 | 44576.9609
213 | 7.5142 | 43478.9062 | 0.66345 | 0.022525 | 213 | 43381.793 | 0.66196 | 0.030661 | 213 | 43420.6172 | 0.66256 | 0.027404 |  | 213 | 7.5142 | 0.66518 | 0.01305 |  | 213 | 0.65943 | 0.044648 | 213 | 0.68101 | -0.071372 | 213 | 43528.7617 | 213 | 43183.8008 | 213 | 44543
214 | 7.5494 | 43469.3594 | 0.6633 | 0.023323 | 214 | 43397.0586 | 0.6622 | 0.02938 | 214 | 43408.7266 | 0.66237 | 0.028401 |  | 214 | 7.5494 | 0.665 | 0.014049 |  | 214 | 0.65954 | 0.04404 | 214 | 0.6814 | -0.073401 | 214 | 43544.1992 | 214 | 43164.7617 | 214 | 44589.4805
215 | 7.5847 | 43469.2109 | 0.6633 | 0.023335 | 215 | 43436.3281 | 0.6628 | 0.026088 | 215 | 43383.0938 | 0.66198 | 0.030552 |  | 215 | 7.5847 | 0.66435 | 0.017583 |  | 215 | 0.66016 | 0.040589 | 215 | 0.68132 | -0.072974 | 215 | 43515.2812 | 215 | 43140.5586 | 215 | 44599.8398
216 | 7.62 | 43482.3711 | 0.6635 | 0.022235 | 216 | 43444.8438 | 0.66293 | 0.025375 | 216 | 43422.2695 | 0.66258 | 0.027266 |  | 216 | 7.62 | 0.6646 | 0.016192 |  | 216 | 0.66034 | 0.039595 | 216 | 0.68153 | -0.074091 | 216 | 43507.9609 | 216 | 43145.3984 | 216 | 44584.8789
217 | 7.6553 | 43475.8359 | 0.6634 | 0.022781 | 217 | 43468.5742 | 0.66329 | 0.023389 | 217 | 43457.2305 | 0.66311 | 0.024338 |  | 217 | 7.6553 | 0.66482 | 0.015003 |  | 217 | 0.66022 | 0.040274 | 217 | 0.68096 | -0.071137 | 217 | 43564.6016 | 217 | 43175.5195 | 217 | 44592.2812
218 | 7.6906 | 43472.0156 | 0.66334 | 0.023101 | 218 | 43458.5078 | 0.66313 | 0.024231 | 218 | 43455.1875 | 0.66308 | 0.024509 |  | 218 | 7.6906 | 0.66396 | 0.019684 |  | 218 | 0.65983 | 0.042442 | 218 | 0.68032 | -0.067791 | 218 | 43577.6797 | 218 | 43237.9609 | 218 | 44591.3984
219 | 7.7258 | 43474.8789 | 0.66338 | 0.022861 | 219 | 43436.8906 | 0.6628 | 0.026041 | 219 | 43464.9023 | 0.66323 | 0.023696 |  | 219 | 7.7258 | 0.66362 | 0.021593 |  | 219 | 0.65967 | 0.043303 | 219 | 0.6804 | -0.068186 | 219 | 43590.0391 | 219 | 43289.0391 | 219 | 44589.7617
220 | 7.7611 | 43460.9727 | 0.66317 | 0.024025 | 220 | 43451.9219 | 0.66303 | 0.024782 | 220 | 43459.0117 | 0.66314 | 0.024189 |  | 220 | 7.7611 | 0.6633 | 0.023327 |  | 220 | 0.65914 | 0.046275 | 220 | 0.67993 | -0.065728 | 220 | 43538.0391 | 220 | 43262.5586 | 220 | 44588.7617
221 | 7.7964 | 43504.0625 | 0.66383 | 0.020422 | 221 | 43432.3281 | 0.66273 | 0.026423 | 221 | 43449.6641 | 0.663 | 0.024971 |  | 221 | 7.7964 | 0.66347 | 0.022399 |  | 221 | 0.65921 | 0.045897 | 221 | 0.67927 | -0.062303 | 221 | 43497.0781 | 221 | 43219.4805 | 221 | 44569.6016
222 | 7.8317 | 43497.2109 | 0.66372 | 0.020995 | 222 | 43473.9727 | 0.66337 | 0.022937 | 222 | 43443.2734 | 0.6629 | 0.025506 |  | 222 | 7.8317 | 0.6627 | 0.026626 |  | 222 | 0.65911 | 0.046405 | 222 | 0.67865 | -0.059006 | 222 | 43467.3594 | 222 | 43226.9219 | 222 | 44544
223 | 7.8669 | 43426.3242 | 0.66264 | 0.026926 | 223 | 43478.4727 | 0.66344 | 0.022561 | 223 | 43413.9414 | 0.66245 | 0.027964 |  | 223 | 7.8669 | 0.66267 | 0.026801 |  | 223 | 0.65887 | 0.047753 | 223 | 0.67832 | -0.057291 | 223 | 43428.6406 | 223 | 43200.8398 | 223 | 44561.7188
224 | 7.9022 | 43406.3828 | 0.66234 | 0.028598 | 224 | 43425.1875 | 0.66263 | 0.027021 | 224 | 43446.4062 | 0.66295 | 0.025244 |  | 224 | 7.9022 | 0.66277 | 0.026207 |  | 224 | 0.65946 | 0.044463 | 224 | 0.6784 | -0.057721 | 224 | 43434.8398 | 224 | 43174.6406 | 224 | 44527.7188
225 | 7.9375 | 43437.9258 | 0.66282 | 0.025954 | 225 | 43440.5391 | 0.66286 | 0.025735 | 225 | 43417.5312 | 0.66251 | 0.027663 |  | 225 | 7.9375 | 0.66284 | 0.02583 |  | 225 | 0.65856 | 0.049471 | 225 | 0.67886 | -0.06012 | 225 | 43466.3203 | 225 | 43152.8398 | 225 | 44478.5586
226 | 7.9728 | 43459.0156 | 0.66314 | 0.024188 | 226 | 43433.7539 | 0.66276 | 0.026304 | 226 | 43382.0586 | 0.66197 | 0.030639 |  | 226 | 7.9728 | 0.66364 | 0.021453 |  | 226 | 0.65811 | 0.051962 | 226 | 0.67863 | -0.058901 | 226 | 43510.1992 | 226 | 43130.8398 | 226 | 44481.5586
227 | 8.0081 | 43433.1289 | 0.66275 | 0.026356 | 227 | 43397.293 | 0.6622 | 0.02936 | 227 | 43432.6953 | 0.66274 | 0.026392 |  | 227 | 8.0081 | 0.66312 | 0.024311 |  | 227 | 0.65884 | 0.047914 | 227 | 0.67837 | -0.057564 | 227 | 43559.5586 | 227 | 43165.3984 | 227 | 44512.9609
228 | 8.0433 | 43416.8398 | 0.6625 | 0.027721 | 228 | 43414.6484 | 0.66247 | 0.027905 | 228 | 43391.4258 | 0.66211 | 0.029853 |  | 228 | 8.0433 | 0.66285 | 0.025803 |  | 228 | 0.65945 | 0.044532 | 228 | 0.67872 | -0.059371 | 228 | 43593.1992 | 228 | 43153.2383 | 228 | 44573.9219
229 | 8.0786 | 43441.7773 | 0.66288 | 0.025632 | 229 | 43377.8438 | 0.6619 | 0.030993 | 229 | 43422.668 | 0.66259 | 0.027232 |  | 229 | 8.0786 | 0.66287 | 0.025685 |  | 229 | 0.66031 | 0.03976 | 229 | 0.67946 | -0.063299 | 229 | 43608.6016 | 229 | 43152.8008 | 229 | 44588.1992
230 | 8.1139 | 43429.8047 | 0.6627 | 0.026634 | 230 | 43456.1211 | 0.6631 | 0.024431 | 230 | 43410.3398 | 0.6624 | 0.028266 |  | 230 | 8.1139 | 0.66232 | 0.028702 |  | 230 | 0.66004 | 0.04126 | 230 | 0.67949 | -0.063447 | 230 | 43617.5195 | 230 | 43099.9609 | 230 | 44604.3594
231 | 8.1492 | 43446.8867 | 0.66296 | 0.025204 | 231 | 43461.3477 | 0.66318 | 0.023993 | 231 | 43421.1602 | 0.66256 | 0.027359 |  | 231 | 8.1492 | 0.66225 | 0.02911 |  | 231 | 0.6605 | 0.038725 | 231 | 0.67886 | -0.06012 | 231 | 43627.6797 | 231 | 43120.8789 | 231 | 44602.3984
232 | 8.1844 | 43433.0625 | 0.66275 | 0.026361 | 232 | 43447.3555 | 0.66296 | 0.025164 | 232 | 43468.8828 | 0.66329 | 0.023363 |  | 232 | 8.1844 | 0.66264 | 0.026957 |  | 232 | 0.66052 | 0.03861 | 232 | 0.67891 | -0.06039 | 232 | 43610.7188 | 232 | 43130.9609 | 232 | 44568.7188
233 | 8.2197 | 43426.6016 | 0.66265 | 0.026903 | 233 | 43435.3164 | 0.66278 | 0.026173 | 233 | 43446.4531 | 0.66295 | 0.02524 |  | 233 | 8.2197 | 0.66329 | 0.023376 |  | 233 | 0.66085 | 0.036821 | 233 | 0.67795 | -0.055365 | 233 | 43578.0391 | 233 | 43189.1211 | 233 | 44538.3203
234 | 8.255 | 43438.8633 | 0.66283 | 0.025876 | 234 | 43476.5273 | 0.66341 | 0.022724 | 234 | 43473.9023 | 0.66337 | 0.022943 |  | 234 | 8.255 | 0.66398 | 0.019623 |  | 234 | 0.65983 | 0.042431 | 234 | 0.67798 | -0.055515 | 234 | 43530.1211 | 234 | 43176.0781 | 234 | 44535.8398
235 | 8.2903 | 43433.9531 | 0.66276 | 0.026287 | 235 | 43499.1797 | 0.66375 | 0.02083 | 235 | 43388.9297 | 0.66207 | 0.030062 |  | 235 | 8.2903 | 0.66442 | 0.017177 |  | 235 | 0.65981 | 0.042531 | 235 | 0.67817 | -0.056503 | 235 | 43539.8008 | 235 | 43196.8008 | 235 | 44503.2812
236 | 8.3256 | 43468.6562 | 0.66329 | 0.023382 | 236 | 43479.5352 | 0.66346 | 0.022472 | 236 | 43399.0508 | 0.66223 | 0.029213 |  | 236 | 8.3256 | 0.66408 | 0.019039 |  | 236 | 0.65912 | 0.046375 | 236 | 0.67819 | -0.056591 | 236 | 43542.7188 | 236 | 43232.2383 | 236 | 44470.9609
237 | 8.3608 | 43498.0586 | 0.66374 | 0.020924 | 237 | 43431.8828 | 0.66273 | 0.02646 | 237 | 43424.4258 | 0.66261 | 0.027085 |  | 237 | 8.3608 | 0.66416 | 0.018603 |  | 237 | 0.65909 | 0.046517 | 237 | 0.67796 | -0.05538 | 237 | 43538.3594 | 237 | 43232.0781 | 237 | 44448.5195
238 | 8.3961 | 43439.9688 | 0.66285 | 0.025783 | 238 | 43445.0039 | 0.66293 | 0.025361 | 238 | 43386.668 | 0.66204 | 0.030252 |  | 238 | 8.3961 | 0.66381 | 0.020538 |  | 238 | 0.65897 | 0.047214 | 238 | 0.67786 | -0.054894 | 238 | 43542.8008 | 238 | 43208.1211 | 238 | 44438.5195
239 | 8.4314 | 43474.1992 | 0.66337 | 0.022918 | 239 | 43436.9023 | 0.6628 | 0.02604 | 239 | 43416.8008 | 0.6625 | 0.027724 |  | 239 | 8.4314 | 0.66378 | 0.020709 |  | 239 | 0.65868 | 0.048842 | 239 | 0.67742 | -0.052564 | 239 | 43510.3203 | 239 | 43187.9609 | 239 | 44426.9219
240 | 8.4667 | 43421.4766 | 0.66257 | 0.027332 | 240 | 43465.9844 | 0.66325 | 0.023605 | 240 | 43419.3633 | 0.66254 | 0.027509 |  | 240 | 8.4667 | 0.66389 | 0.020086 |  | 240 | 0.65816 | 0.051719 | 240 | 0.67715 | -0.05111 | 240 | 43483.3203 | 240 | 43186.1602 | 240 | 44439.3203
241 | 8.5019 | 43474.1602 | 0.66337 | 0.022921 | 241 | 43445.9062 | 0.66294 | 0.025286 | 241 | 43439.3984 | 0.66284 | 0.025831 |  | 241 | 8.5019 | 0.66355 | 0.021962 |  | 241 | 0.65785 | 0.05344 | 241 | 0.67716 | -0.051176 | 241 | 43440.7617 | 241 | 43167.4805 | 241 | 44448.5586
242 | 8.5372 | 43469.3281 | 0.6633 | 0.023326 | 242 | 43428.7852 | 0.66268 | 0.02672 | 242 | 43409.8789 | 0.66239 | 0.028305 |  | 242 | 8.5372 | 0.66349 | 0.022255 |  | 242 | 0.65784 | 0.053475 | 242 | 0.6779 | -0.055062 | 242 | 43442.3594 | 242 | 43152.3594 | 242 | 44429.8789
243 | 8.5725 | 43473.4727 | 0.66336 | 0.022979 | 243 | 43362.7383 | 0.66167 | 0.032262 | 243 | 43407.9141 | 0.66236 | 0.028469 |  | 243 | 8.5725 | 0.6631 | 0.024433 |  | 243 | 0.65857 | 0.049401 | 243 | 0.67848 | -0.058122 | 243 | 43431.3984 | 243 | 43144.4414 | 243 | 44450.2383
244 | 8.6078 | 43498.5938 | 0.66375 | 0.020879 | 244 | 43408.3711 | 0.66237 | 0.028431 | 244 | 43448.7539 | 0.66299 | 0.025047 |  | 244 | 8.6078 | 0.66268 | 0.026713 |  | 244 | 0.65942 | 0.044725 | 244 | 0.67858 | -0.058668 | 244 | 43380.3984 | 244 | 43131.6016 | 244 | 44482.6797
245 | 8.6431 | 43511.9688 | 0.66395 | 0.019762 | 245 | 43455.7344 | 0.66309 | 0.024463 | 245 | 43441.4375 | 0.66287 | 0.02566 |  | 245 | 8.6431 | 0.66334 | 0.023087 |  | 245 | 0.65955 | 0.043983 | 245 | 0.67897 | -0.060717 | 245 | 43386.8789 | 245 | 43130.2812 | 245 | 44447
246 | 8.6783 | 43461.2539 | 0.66318 | 0.024001 | 246 | 43457.5977 | 0.66312 | 0.024307 | 246 | 43380.6914 | 0.66195 | 0.030754 |  | 246 | 8.6783 | 0.66327 | 0.023482 |  | 246 | 0.65971 | 0.043088 | 246 | 0.67889 | -0.060299 | 246 | 43390.8398 | 246 | 43085.4805 | 246 | 44438.9609
247 | 8.7136 | 43452.5234 | 0.66304 | 0.024732 | 247 | 43427.9531 | 0.66267 | 0.02679 | 247 | 43401.4531 | 0.66226 | 0.029011 |  | 247 | 8.7136 | 0.66356 | 0.021913 |  | 247 | 0.65909 | 0.046521 | 247 | 0.67618 | -0.046027 | 247 | 43406.5195 | 247 | 43045.6797 | 247 | 44454.9609
248 | 8.7489 | 43409.4805 | 0.66239 | 0.028338 | 248 | 43429.9688 | 0.6627 | 0.026621 | 248 | 43387.6602 | 0.66205 | 0.030169 |  | 248 | 8.7489 | 0.66288 | 0.025639 |  | 248 | 0.65955 | 0.043964 | 248 | 0.67632 | -0.046753 | 248 | 43429.9219 | 248 | 43077.6406 | 248 | 44395.3984
249 | 8.7842 | 43432.7539 | 0.66274 | 0.026387 | 249 | 43439.8242 | 0.66285 | 0.025795 | 249 | 43413.2031 | 0.66244 | 0.028026 |  | 249 | 8.7842 | 0.66257 | 0.027315 |  | 249 | 0.65987 | 0.042231 | 249 | 0.67756 | -0.053273 | 249 | 43434.9609 | 249 | 43144.1602 | 249 | 44396.7188
250 | 8.8194 | 43455.1562 | 0.66308 | 0.024511 | 250 | 43447.6523 | 0.66297 | 0.02514 | 250 | 43469.9961 | 0.66331 | 0.02327 |  | 250 | 8.8194 | 0.66301 | 0.024924 |  | 250 | 0.6602 | 0.040382 | 250 | 0.6778 | -0.054551 | 250 | 43433.9609 | 250 | 43196.0781 | 250 | 44395.4805
251 | 8.8547 | 43411.0977 | 0.66241 | 0.028203 | 251 | 43435.2969 | 0.66278 | 0.026174 | 251 | 43365.9961 | 0.66172 | 0.031988 |  | 251 | 8.8547 | 0.66372 | 0.02102 |  | 251 | 0.65994 | 0.04184 | 251 | 0.6774 | -0.052433 | 251 | 43403.3594 | 251 | 43191.4805 | 251 | 44400.6016
252 | 8.89 | 43455.6484 | 0.66309 | 0.02447 | 252 | 43396.3906 | 0.66219 | 0.029436 | 252 | 43372.6953 | 0.66182 | 0.031425 |  | 252 | 8.89 | 0.66364 | 0.021446 |  | 252 | 0.66071 | 0.037576 | 252 | 0.67773 | -0.054168 | 252 | 43363.8398 | 252 | 43180.7617 | 252 | 44397.2383
253 | 8.9253 | 43413.4258 | 0.66245 | 0.028007 | 253 | 43449.8242 | 0.663 | 0.024958 | 253 | 43397.8984 | 0.66221 | 0.02931 |  | 253 | 8.9253 | 0.66403 | 0.019312 |  | 253 | 0.66015 | 0.040646 | 253 | 0.67744 | -0.052678 | 253 | 43375.5586 | 253 | 43199.1602 | 253 | 44412.8789
254 | 8.9606 | 43436.3633 | 0.6628 | 0.026085 | 254 | 43442.5547 | 0.66289 | 0.025566 | 254 | 43430.293 | 0.6627 | 0.026594 |  | 254 | 8.9606 | 0.66426 | 0.018087 |  | 254 | 0.65949 | 0.04434 | 254 | 0.67816 | -0.056463 | 254 | 43395.2812 | 254 | 43244.3984 | 254 | 44377.4805
255 | 8.9958 | 43362.8164 | 0.66167 | 0.032255 | 255 | 43443.6328 | 0.66291 | 0.025476 | 255 | 43333.9531 | 0.66123 | 0.034682 |  | 255 | 8.9958 | 0.66447 | 0.016934 |  | 255 | 0.66055 | 0.038457 | 255 | 0.6781 | -0.056146 | 255 | 43417.2812 | 255 | 43209.4414 | 255 | 44350.3984
256 | 9.0311 | 43442.8242 | 0.6629 | 0.025544 | 256 | 43393.0469 | 0.66214 | 0.029717 | 256 | 43393.3438 | 0.66214 | 0.029692 |  | 256 | 9.0311 | 0.66408 | 0.019073 |  | 256 | 0.6611 | 0.035424 | 256 | 0.67824 | -0.056886 | 256 | 43438.5195 | 256 | 43168.1992 | 256 | 44367.6406
257 | 9.0664 | 43443.7734 | 0.66291 | 0.025464 | 257 | 43445.1055 | 0.66293 | 0.025353 | 257 | 43482.2031 | 0.6635 | 0.022249 |  | 257 | 9.0664 | 0.66337 | 0.022958 |  | 257 | 0.66159 | 0.032729 | 257 | 0.67811 | -0.056178 | 257 | 43468.2812 | 257 | 43168.6016 | 257 | 44392.7617
258 | 9.1017 | 43446.5625 | 0.66295 | 0.025231 | 258 | 43447.2305 | 0.66296 | 0.025175 | 258 | 43428.2617 | 0.66267 | 0.026764 |  | 258 | 9.1017 | 0.66332 | 0.023216 |  | 258 | 0.66198 | 0.030591 | 258 | 0.67681 | -0.049354 | 258 | 43449.5586 | 258 | 43175.1602 | 258 | 44424.8008
259 | 9.1369 | 43481.8242 | 0.66349 | 0.022281 | 259 | 43406.2461 | 0.66234 | 0.028609 | 259 | 43462.2656 | 0.66319 | 0.023916 |  | 259 | 9.1369 | 0.66348 | 0.02235 |  | 259 | 0.66155 | 0.032951 | 259 | 0.67593 | -0.044685 | 259 | 43472.1602 | 259 | 43144.2383 | 259 | 44438.5195
260 | 9.1722 | 43429.4336 | 0.66269 | 0.026666 | 260 | 43457.3008 | 0.66312 | 0.024332 | 260 | 43473.125 | 0.66336 | 0.023008 |  | 260 | 9.1722 | 0.66369 | 0.021195 |  | 260 | 0.66184 | 0.031369 | 260 | 0.67497 | -0.039589 | 260 | 43471.5586 | 260 | 43131 | 260 | 44445.7617
261 | 9.2075 | 43422.5547 | 0.66259 | 0.027242 | 261 | 43446.8398 | 0.66296 | 0.025208 | 261 | 43449.9805 | 0.663 | 0.024945 |  | 261 | 9.2075 | 0.66412 | 0.018826 |  | 261 | 0.66103 | 0.035822 | 261 | 0.67488 | -0.03914 | 261 | 43475.6797 | 261 | 43112.6797 | 261 | 44424.9609
262 | 9.2428 | 43462.7891 | 0.6632 | 0.023873 | 262 | 43426.7422 | 0.66265 | 0.026891 | 262 | 43440.3672 | 0.66286 | 0.02575 |  | 262 | 9.2428 | 0.66379 | 0.020614 |  | 262 | 0.66049 | 0.03881 | 262 | 0.67568 | -0.043341 | 262 | 43447.8789 | 262 | 43081.0781 | 262 | 44310.5586
263 | 9.2781 | 43518.7383 | 0.66405 | 0.019197 | 263 | 43470.7969 | 0.66332 | 0.023203 | 263 | 43489.75 | 0.66361 | 0.021618 |  | 263 | 9.2781 | 0.66433 | 0.017708 |  | 263 | 0.66085 | 0.036791 | 263 | 0.6768 | -0.049306 | 263 | 43423.8008 | 263 | 43064.2812 | 263 | 44290.7188
264 | 9.3133 | 43461.4414 | 0.66318 | 0.023985 | 264 | 43495.9961 | 0.66371 | 0.021096 | 264 | 43455.4883 | 0.66309 | 0.024484 |  | 264 | 9.3133 | 0.66457 | 0.016397 |  | 264 | 0.66085 | 0.036802 | 264 | 0.67767 | -0.05389 | 264 | 43437.4414 | 264 | 43089.7617 | 264 | 44373.7188
265 | 9.3486 | 43463.8945 | 0.66322 | 0.02378 | 265 | 43454.6641 | 0.66308 | 0.024553 | 265 | 43464.8828 | 0.66323 | 0.023697 |  | 265 | 9.3486 | 0.6637 | 0.021115 |  | 265 | 0.66095 | 0.036224 | 265 | 0.67791 | -0.055143 | 265 | 43426.1992 | 265 | 43106.8008 | 265 | 44378.7188
266 | 9.3839 | 43478.8867 | 0.66345 | 0.022526 | 266 | 43439.1406 | 0.66284 | 0.025852 | 266 | 43477.668 | 0.66343 | 0.022628 |  | 266 | 9.3839 | 0.66391 | 0.019964 |  | 266 | 0.66139 | 0.033841 | 266 | 0.67745 | -0.052704 | 266 | 43392.1602 | 266 | 43119.3594 | 266 | 44371.8008
267 | 9.4192 | 43433.2734 | 0.66275 | 0.026344 | 267 | 43450.0312 | 0.663 | 0.02494 | 267 | 43437.6328 | 0.66282 | 0.025979 |  | 267 | 9.4192 | 0.66359 | 0.02173 |  | 267 | 0.66142 | 0.033662 | 267 | 0.6772 | -0.05142 | 267 | 43412.3594 | 267 | 43134.7188 | 267 | 44394.8008
268 | 9.4544 | 43432.793 | 0.66274 | 0.026384 | 268 | 43430.7109 | 0.66271 | 0.026558 | 268 | 43455.0547 | 0.66308 | 0.02452 |  | 268 | 9.4544 | 0.66344 | 0.022528 |  | 268 | 0.66125 | 0.034613 | 268 | 0.67773 | -0.054193 | 268 | 43388 | 268 | 43122.5195 | 268 | 44399.0781
269 | 9.4897 | 43457.4414 | 0.66312 | 0.02432 | 269 | 43485.7578 | 0.66355 | 0.021952 | 269 | 43453.75 | 0.66306 | 0.024629 |  | 269 | 9.4897 | 0.66372 | 0.021028 |  | 269 | 0.66091 | 0.036465 | 269 | 0.67779 | -0.054482 | 269 | 43403.5195 | 269 | 43167.8398 | 269 | 44411.9609
270 | 9.525 | 43437.5898 | 0.66282 | 0.025982 | 270 | 43517.7773 | 0.66404 | 0.019277 | 270 | 43436.3125 | 0.6628 | 0.026089 |  | 270 | 9.525 | 0.66408 | 0.019065 |  | 270 | 0.66086 | 0.036718 | 270 | 0.67709 | -0.050839 | 270 | 43312.8008 | 270 | 43211.0781 | 270 | 44385.7617
271 | 9.5603 | 43393.2539 | 0.66214 | 0.029699 | 271 | 43475.4922 | 0.66339 | 0.02281 | 271 | 43393.3945 | 0.66214 | 0.029688 |  | 271 | 9.5603 | 0.66447 | 0.016946 |  | 271 | 0.6614 | 0.033746 | 271 | 0.67728 | -0.051797 | 271 | 43411.0781 | 271 | 43200.3203 | 271 | 44384.4414
272 | 9.5956 | 43420.2734 | 0.66255 | 0.027433 | 272 | 43434.8711 | 0.66277 | 0.02621 | 272 | 43370.1953 | 0.66179 | 0.031636 |  | 272 | 9.5956 | 0.66408 | 0.019046 |  | 272 | 0.66179 | 0.031644 | 272 | 0.67761 | -0.053551 | 272 | 43413.7188 | 272 | 43212.0391 | 272 | 44400.2383
273 | 9.6308 | 43435.1133 | 0.66278 | 0.02619 | 273 | 43448.3867 | 0.66298 | 0.025078 | 273 | 43414.0742 | 0.66246 | 0.027953 |  | 273 | 9.6308 | 0.66346 | 0.022471 |  | 273 | 0.66254 | 0.027475 | 273 | 0.67763 | -0.053679 | 273 | 43420.1992 | 273 | 43220.5195 | 273 | 44334.3594
274 | 9.6661 | 43469.1875 | 0.6633 | 0.023337 | 274 | 43444.2539 | 0.66292 | 0.025424 | 274 | 43413.582 | 0.66245 | 0.027994 |  | 274 | 9.6661 | 0.66385 | 0.020306 |  | 274 | 0.66188 | 0.031129 | 274 | 0.67736 | -0.052232 | 274 | 43432.4414 | 274 | 43143.6797 | 274 | 44267.7188
275 | 9.7014 | 43510.75 | 0.66393 | 0.019864 | 275 | 43476.2617 | 0.66341 | 0.022746 | 275 | 43435.1875 | 0.66278 | 0.026183 |  | 275 | 9.7014 | 0.66374 | 0.020902 |  | 275 | 0.66185 | 0.031278 | 275 | 0.67742 | -0.052532 | 275 | 43438.9219 | 275 | 43162.6016 | 275 | 44226.9609
276 | 9.7367 | 43481.6172 | 0.66349 | 0.022298 | 276 | 43477.1641 | 0.66342 | 0.02267 | 276 | 43467.4961 | 0.66327 | 0.023479 |  | 276 | 9.7367 | 0.6636 | 0.021654 |  | 276 | 0.66176 | 0.031789 | 276 | 0.67719 | -0.051358 | 276 | 43516.2383 | 276 | 43237.6797 | 276 | 44205.9219
277 | 9.7719 | 43477.7344 | 0.66343 | 0.022623 | 277 | 43464.043 | 0.66322 | 0.023768 | 277 | 43434.9297 | 0.66277 | 0.026205 |  | 277 | 9.7719 | 0.66432 | 0.017715 |  | 277 | 0.66247 | 0.027852 | 277 | 0.67694 | -0.05002 | 277 | 43499.3594 | 277 | 43266.3203 | 277 | 44233.9609
278 | 9.8072 | 43408.6289 | 0.66237 | 0.02841 | 278 | 43547.8594 | 0.6645 | 0.016768 | 278 | 43380.7383 | 0.66195 | 0.03075 |  | 278 | 9.8072 | 0.66525 | 0.012672 |  | 278 | 0.66209 | 0.029973 | 278 | 0.67565 | -0.043191 | 278 | 43485 | 278 | 43299.2812 | 278 | 44304.0781
279 | 9.8425 | 43454.5391 | 0.66307 | 0.024563 | 279 | 43443.9219 | 0.66291 | 0.025452 | 279 | 43518.8359 | 0.66405 | 0.019189 |  | 279 | 9.8425 | 0.66539 | 0.011878 |  | 279 | 0.66248 | 0.027803 | 279 | 0.67616 | -0.045888 | 279 | 43476 | 279 | 43348.9609 | 279 | 44346.7188
280 | 9.8778 | 43465.7734 | 0.66325 | 0.023623 | 280 | 43478.5273 | 0.66344 | 0.022556 | 280 | 43502.043 | 0.6638 | 0.020591 |  | 280 | 9.8778 | 0.66602 | 0.0084645 |  | 280 | 0.66332 | 0.023205 | 280 | 0.67612 | -0.04569 | 280 | 43458.4414 | 280 | 43326.2383 | 280 | 44357.5195
281 | 9.9131 | 43454.9922 | 0.66308 | 0.024525 | 281 | 43456.0195 | 0.6631 | 0.024439 | 281 | 43507.4453 | 0.66388 | 0.02014 |  | 281 | 9.9131 | 0.66637 | 0.0065755 |  | 281 | 0.66301 | 0.024939 | 281 | 0.67572 | -0.043577 | 281 | 43445.8008 | 281 | 43300.1602 | 281 | 44341.7188
282 | 9.9483 | 43483.4609 | 0.66352 | 0.022144 | 282 | 43459.9219 | 0.66316 | 0.024113 | 282 | 43505.0039 | 0.66384 | 0.020344 |  | 282 | 9.9483 | 0.66619 | 0.0075481 |  | 282 | 0.66384 | 0.020344 | 282 | 0.67599 | -0.045033 | 282 | 43442.6406 | 282 | 43254.0391 | 282 | 44326.7617
283 | 9.9836 | 43420.7305 | 0.66256 | 0.027395 | 283 | 43474.6953 | 0.66338 | 0.022877 | 283 | 43492.6797 | 0.66366 | 0.021373 |  | 283 | 9.9836 | 0.66588 | 0.0092269 |  | 283 | 0.66443 | 0.017162 | 283 | 0.67662 | -0.048361 | 283 | 43478 | 283 | 43246.8789 | 283 | 44335.9219
284 | 10.0189 | 43437.7852 | 0.66282 | 0.025966 | 284 | 43470.1016 | 0.66331 | 0.023261 | 284 | 43508.8555 | 0.6639 | 0.020022 |  | 284 | 10.0189 | 0.66541 | 0.011803 |  | 284 | 0.66341 | 0.022741 | 284 | 0.677 | -0.050327 | 284 | 43482.0391 | 284 | 43250.1992 | 284 | 44339.3984
285 | 10.0542 | 43412.6328 | 0.66243 | 0.028074 | 285 | 43487.0273 | 0.66357 | 0.021846 | 285 | 43445.8828 | 0.66294 | 0.025288 |  | 285 | 10.0542 | 0.66532 | 0.012275 |  | 285 | 0.6626 | 0.027186 | 285 | 0.67797 | -0.055438 | 285 | 43514.0781 | 285 | 43228.3203 | 285 | 44315.4414
286 | 10.0894 | 43444.543 | 0.66292 | 0.0254 | 286 | 43446.5781 | 0.66295 | 0.02523 | 286 | 43452.2773 | 0.66304 | 0.024752 |  | 286 | 10.0894 | 0.66553 | 0.011145 |  | 286 | 0.66225 | 0.029099 | 286 | 0.67842 | -0.057822 | 286 | 43492.1992 | 286 | 43252.1211 | 286 | 44344.8398
287 | 10.1247 | 43445.3906 | 0.66293 | 0.025329 | 287 | 43442.1992 | 0.66289 | 0.025596 | 287 | 43461.0195 | 0.66317 | 0.024021 |  | 287 | 10.1247 | 0.66582 | 0.0095704 |  | 287 | 0.66166 | 0.032343 | 287 | 0.67887 | -0.060175 | 287 | 43482.3984 | 287 | 43275.8398 | 287 | 44395.8789
288 | 10.16 | 43464.9141 | 0.66323 | 0.023695 | 288 | 43408.5586 | 0.66237 | 0.028415 | 288 | 43450.2812 | 0.66301 | 0.024919 |  | 288 | 10.16 | 0.66617 | 0.0076385 |  | 288 | 0.66178 | 0.031678 | 288 | 0.67902 | -0.060979 | 288 | 43490.5195 | 288 | 43228.7617 | 288 | 44393
289 | 10.1953 | 43442.8438 | 0.6629 | 0.025542 | 289 | 43471.7734 | 0.66334 | 0.023121 | 289 | 43419.3242 | 0.66254 | 0.027513 |  | 289 | 10.1953 | 0.66613 | 0.00788 |  | 289 | 0.66273 | 0.026462 | 289 | 0.67922 | -0.062026 | 289 | 43451.2812 | 289 | 43228.1602 | 289 | 44365.7188
290 | 10.2306 | 43442.4844 | 0.66289 | 0.025572 | 290 | 43469.3164 | 0.6633 | 0.023327 | 290 | 43413.4883 | 0.66245 | 0.028002 |  | 290 | 10.2306 | 0.66624 | 0.0072691 |  | 290 | 0.6637 | 0.021111 | 290 | 0.67954 | -0.063716 | 290 | 43438.9219 | 290 | 43223 | 290 | 44354.3594
291 | 10.2658 | 43441.543 | 0.66288 | 0.025651 | 291 | 43453.7852 | 0.66306 | 0.024626 | 291 | 43385.6055 | 0.66202 | 0.030341 |  | 291 | 10.2658 | 0.6659 | 0.0090984 |  | 291 | 0.66444 | 0.017109 | 291 | 0.67946 | -0.063295 | 291 | 43491.3203 | 291 | 43210.4805 | 291 | 44353.2383
292 | 10.3011 | 43424.2227 | 0.66261 | 0.027102 | 292 | 43403.5117 | 0.6623 | 0.028839 | 292 | 43404.4062 | 0.66231 | 0.028764 |  | 292 | 10.3011 | 0.66459 | 0.01629 |  | 292 | 0.6642 | 0.018386 | 292 | 0.67882 | -0.059924 | 292 | 43482.8008 | 292 | 43210.3594 | 292 | 44324.6406
293 | 10.3364 | 43452.8906 | 0.66305 | 0.024701 | 293 | 43375.1641 | 0.66186 | 0.031218 | 293 | 43425.5039 | 0.66263 | 0.026995 |  | 293 | 10.3364 | 0.66328 | 0.023406 |  | 293 | 0.66324 | 0.023665 | 293 | 0.67762 | -0.053599 | 293 | 43564.6797 | 293 | 43255.5195 | 293 | 44227.3203
294 | 10.3717 | 43420.9297 | 0.66256 | 0.027378 | 294 | 43435.5391 | 0.66278 | 0.026154 | 294 | 43367.0312 | 0.66174 | 0.031901 |  | 294 | 10.3717 | 0.66093 | 0.036335 |  | 294 | 0.66303 | 0.024798 | 294 | 0.67825 | -0.056919 | 294 | 43470.1992 | 294 | 43262.4414 | 294 | 44281.4805
295 | 10.4069 | 43441.3281 | 0.66287 | 0.025669 | 295 | 43423.5039 | 0.6626 | 0.027162 | 295 | 43388.4492 | 0.66207 | 0.030103 |  | 295 | 10.4069 | 0.66347 | 0.022403 |  | 295 | 0.66284 | 0.025845 | 295 | 0.67849 | -0.05818 | 295 | 43430.2383 | 295 | 43299.3203 | 295 | 44265.6016
296 | 10.4422 | 43471.2891 | 0.66333 | 0.023162 | 296 | 43399.8711 | 0.66224 | 0.029144 | 296 | 43459.332 | 0.66315 | 0.024162 |  | 296 | 10.4422 | 0.66441 | 0.01723 |  | 296 | 0.66224 | 0.029126 | 296 | 0.67872 | -0.059374 | 296 | 43460.0781 | 296 | 43302.5586 | 296 | 44262.0781
297 | 10.4775 | 43513.0898 | 0.66397 | 0.019669 | 297 | 43474.2969 | 0.66338 | 0.02291 | 297 | 43464.1211 | 0.66322 | 0.023761 |  | 297 | 10.4775 | 0.66367 | 0.021282 |  | 297 | 0.66163 | 0.032511 | 297 | 0.67846 | -0.058045 | 297 | 43457.9609 | 297 | 43366 | 297 | 44260.0391
298 | 10.5128 | 43485.6719 | 0.66355 | 0.021959 | 298 | 43523.5312 | 0.66413 | 0.018797 | 298 | 43382.75 | 0.66198 | 0.030581 |  | 298 | 10.5128 | 0.66423 | 0.018231 |  | 298 | 0.66213 | 0.029763 | 298 | 0.67832 | -0.057283 | 298 | 43490.1992 | 298 | 43355.2812 | 298 | 44288.2812
299 | 10.5481 | 43464.6094 | 0.66323 | 0.02372 | 299 | 43427.2656 | 0.66266 | 0.026847 | 299 | 43414.8672 | 0.66247 | 0.027886 |  | 299 | 10.5481 | 0.66347 | 0.022376 |  | 299 | 0.66232 | 0.028714 | 299 | 0.67741 | -0.052499 | 299 | 43521.2812 | 299 | 43333.9219 | 299 | 44352.8398
300 | 10.5833 | 43442.9414 | 0.6629 | 0.025534 | 300 | 43428.0625 | 0.66267 | 0.02678 | 300 | 43437.6172 | 0.66282 | 0.02598 |  | 300 | 10.5833 | 0.66361 | 0.021609 |  | 300 | 0.66255 | 0.027414 | 300 | 0.67863 | -0.058904 | 300 | 43577.7188 | 300 | 43353.8398 | 300 | 44378.9609
301 | 10.6186 | 43450.6484 | 0.66301 | 0.024889 | 301 | 43443.0742 | 0.6629 | 0.025523 | 301 | 43438.4219 | 0.66283 | 0.025913 |  | 301 | 10.6186 | 0.66264 | 0.02693 |  | 301 | 0.66212 | 0.02982 | 301 | 0.67884 | -0.060019 | 301 | 43553.0391 | 301 | 43368.6406 | 301 | 44409.7617
302 | 10.6539 | 43483.9922 | 0.66352 | 0.022099 | 302 | 43449.2227 | 0.66299 | 0.025008 | 302 | 43435.0625 | 0.66278 | 0.026194 |  | 302 | 10.6539 | 0.6633 | 0.023319 |  | 302 | 0.66347 | 0.02241 | 302 | 0.67872 | -0.059414 | 302 | 43582.3594 | 302 | 43379.8008 | 302 | 44450.2812
303 | 10.6892 | 43428.3008 | 0.66267 | 0.02676 | 303 | 43471.7695 | 0.66334 | 0.023121 | 303 | 43396.6992 | 0.66219 | 0.02941 |  | 303 | 10.6892 | 0.66413 | 0.018762 |  | 303 | 0.66326 | 0.023547 | 303 | 0.67793 | -0.055263 | 303 | 43633.1211 | 303 | 43474.3203 | 303 | 44437.9219
304 | 10.7244 | 43427.2539 | 0.66266 | 0.026848 | 304 | 43399.2383 | 0.66223 | 0.029197 | 304 | 43402.8281 | 0.66228 | 0.028896 |  | 304 | 10.7244 | 0.66402 | 0.019399 |  | 304 | 0.66303 | 0.024806 | 304 | 0.67755 | -0.053255 | 304 | 43636.2383 | 304 | 43479.4414 | 304 | 44439.6797
305 | 10.7597 | 43421.7266 | 0.66257 | 0.027311 | 305 | 43441.5859 | 0.66288 | 0.025648 | 305 | 43391.9453 | 0.66212 | 0.029809 |  | 305 | 10.7597 | 0.66427 | 0.018037 |  | 305 | 0.66319 | 0.02395 | 305 | 0.67636 | -0.046977 | 305 | 43604.6797 | 305 | 43418.1211 | 305 | 44457.8008
306 | 10.795 | 43450.875 | 0.66302 | 0.02487 | 306 | 43403.1797 | 0.66229 | 0.028867 | 306 | 43414.5547 | 0.66246 | 0.027913 |  | 306 | 10.795 | 0.66486 | 0.014802 |  | 306 | 0.66267 | 0.026778 | 306 | 0.67651 | -0.047772 | 306 | 43573.8789 | 306 | 43390.8398 | 306 | 44466.1602
307 | 10.8303 | 43402.4883 | 0.66228 | 0.028925 | 307 | 43454.2266 | 0.66307 | 0.024589 | 307 | 43403.1797 | 0.66229 | 0.028867 |  | 307 | 10.8303 | 0.66501 | 0.013977 |  | 307 | 0.66288 | 0.025639 | 307 | 0.67564 | -0.043165 | 307 | 43561.2812 | 307 | 43347.3203 | 307 | 44433.1992
308 | 10.8656 | 43404.3203 | 0.66231 | 0.028771 | 308 | 43399.793 | 0.66224 | 0.029151 | 308 | 43390.5469 | 0.6621 | 0.029927 |  | 308 | 10.8656 | 0.66539 | 0.011882 |  | 308 | 0.66324 | 0.023642 | 308 | 0.67597 | -0.044927 | 308 | 43557.0781 | 308 | 43310.0781 | 308 | 44371.3984
309 | 10.9008 | 43460.8281 | 0.66317 | 0.024037 | 309 | 43388.1289 | 0.66206 | 0.03013 | 309 | 43497.1328 | 0.66372 | 0.021001 |  | 309 | 10.9008 | 0.66532 | 0.012286 |  | 309 | 0.66385 | 0.020295 | 309 | 0.67592 | -0.044663 | 309 | 43578.6406 | 309 | 43340.2812 | 309 | 44388.4414
310 | 10.9361 | 43479.5078 | 0.66345 | 0.022474 | 310 | 43505.2773 | 0.66385 | 0.020321 | 310 | 43461.9844 | 0.66319 | 0.02394 |  | 310 | 10.9361 | 0.6655 | 0.011285 |  | 310 | 0.66395 | 0.019775 | 310 | 0.67632 | -0.046753 | 310 | 43607.3984 | 310 | 43416.9609 | 310 | 44409.4805
311 | 10.9714 | 43435.9023 | 0.66279 | 0.026124 | 311 | 43466.2539 | 0.66325 | 0.023583 | 311 | 43394.168 | 0.66215 | 0.029623 |  | 311 | 10.9714 | 0.66572 | 0.010114 |  | 311 | 0.66404 | 0.019263 | 311 | 0.67651 | -0.047783 | 311 | 43638.3594 | 311 | 43489.5586 | 311 | 44434.0781
312 | 11.0067 | 43370.1445 | 0.66179 | 0.03164 | 312 | 43438.7344 | 0.66283 | 0.025886 | 312 | 43352.7031 | 0.66152 | 0.033106 |  | 312 | 11.0067 | 0.6655 | 0.011311 |  | 312 | 0.66407 | 0.019118 | 312 | 0.67652 | -0.04783 | 312 | 43658.4805 | 312 | 43523.2812 | 312 | 44428.2383
313 | 11.0419 | 43373.2148 | 0.66183 | 0.031382 | 313 | 43371.7734 | 0.66181 | 0.031503 | 313 | 43387.918 | 0.66206 | 0.030147 |  | 313 | 11.0419 | 0.66516 | 0.013168 |  | 313 | 0.66443 | 0.017113 | 313 | 0.67628 | -0.04653 | 313 | 43595.5195 | 313 | 43469.6406 | 313 | 44416.4805
314 | 11.0772 | 43429.9766 | 0.6627 | 0.02662 | 314 | 43429.6289 | 0.66269 | 0.026649 | 314 | 43412.0469 | 0.66243 | 0.028123 |  | 314 | 11.0772 | 0.66566 | 0.010454 |  | 314 | 0.66462 | 0.016074 | 314 | 0.67703 | -0.050506 | 314 | 43540.0781 | 314 | 43439.3594 | 314 | 44407.3594
315 | 11.1125 | 43488.2812 | 0.66359 | 0.021741 | 315 | 43453.4453 | 0.66306 | 0.024655 | 315 | 43380.1797 | 0.66194 | 0.030797 |  | 315 | 11.1125 | 0.6659 | 0.00914 |  | 315 | 0.66484 | 0.014916 | 315 | 0.67747 | -0.052798 | 315 | 43451.3984 | 315 | 43387.4805 | 315 | 44405.8789
316 | 11.1478 | 43464.6523 | 0.66323 | 0.023717 | 316 | 43491.5352 | 0.66364 | 0.021469 | 316 | 43340.5625 | 0.66133 | 0.034126 |  | 316 | 11.1478 | 0.66551 | 0.011232 |  | 316 | 0.66511 | 0.013406 | 316 | 0.67755 | -0.053259 | 316 | 43289.1211 | 316 | 43369.3984 | 316 | 44433.7188
317 | 11.1831 | 43381.6211 | 0.66196 | 0.030676 | 317 | 43469.75 | 0.66331 | 0.02329 | 317 | 43336.3203 | 0.66127 | 0.034483 |  | 317 | 11.1831 | 0.66623 | 0.0073482 |  | 317 | 0.66571 | 0.010182 | 317 | 0.67749 | -0.052912 | 317 | 43455.6016 | 317 | 43310.6406 | 317 | 44457.7617
318 | 11.2183 | 43429.7656 | 0.6627 | 0.026638 | 318 | 43422.2656 | 0.66258 | 0.027266 | 318 | 43412.3125 | 0.66243 | 0.028101 |  | 318 | 11.2183 | 0.66649 | 0.0059158 |  | 318 | 0.66551 | 0.01124 | 318 | 0.67728 | -0.051834 | 318 | 43472.0391 | 318 | 43280.2383 | 318 | 44366.2383
319 | 11.2536 | 43409.6094 | 0.66239 | 0.028327 | 319 | 43456.4297 | 0.6631 | 0.024405 | 319 | 43432.2578 | 0.66273 | 0.026429 |  | 319 | 11.2536 | 0.66528 | 0.012506 |  | 319 | 0.66506 | 0.013705 | 319 | 0.67752 | -0.053058 | 319 | 43467.7188 | 319 | 43301.3984 | 319 | 44364.2812
320 | 11.2889 | 43344.3594 | 0.66139 | 0.033807 | 320 | 43412.918 | 0.66244 | 0.02805 | 320 | 43390.1016 | 0.66209 | 0.029964 |  | 320 | 11.2889 | 0.66613 | 0.0078686 |  | 320 | 0.66472 | 0.015529 | 320 | 0.67776 | -0.054365 | 320 | 43436.7617 | 320 | 43346.7617 | 320 | 44290
321 | 11.3242 | 43425.8477 | 0.66264 | 0.026966 | 321 | 43364.8164 | 0.6617 | 0.032087 | 321 | 43453.1523 | 0.66305 | 0.024679 |  | 321 | 11.3242 | 0.66479 | 0.015173 |  | 321 | 0.66425 | 0.018117 | 321 | 0.67779 | -0.054518 | 321 | 43450.7188 | 321 | 43345.2812 | 321 | 44256.9219
322 | 11.3594 | 43403.8789 | 0.6623 | 0.028808 | 322 | 43422.5508 | 0.66259 | 0.027242 | 322 | 43438.3164 | 0.66283 | 0.025921 |  | 322 | 11.3594 | 0.66487 | 0.014734 |  | 322 | 0.66397 | 0.019669 | 322 | 0.67796 | -0.05542 | 322 | 43457.6016 | 322 | 43388.7617 | 322 | 44242.5586
323 | 11.3947 | 43427.2422 | 0.66266 | 0.026849 | 323 | 43465.3164 | 0.66324 | 0.023661 | 323 | 43414.1602 | 0.66246 | 0.027946 |  | 323 | 11.3947 | 0.66462 | 0.016082 |  | 323 | 0.66399 | 0.019521 | 323 | 0.67794 | -0.055292 | 323 | 43422.3594 | 323 | 43407.5586 | 323 | 44258.6406
324 | 11.43 | 43456.3281 | 0.6631 | 0.024413 | 324 | 43468.9961 | 0.66329 | 0.023353 | 324 | 43404.0625 | 0.6623 | 0.028793 |  | 324 | 11.43 | 0.66366 | 0.021358 |  | 324 | 0.66414 | 0.018728 | 324 | 0.67795 | -0.055329 | 324 | 43407.1602 | 324 | 43396.8008 | 324 | 44249.5586
325 | 11.4653 | 43476.332 | 0.66341 | 0.02274 | 325 | 43494.3125 | 0.66368 | 0.021237 | 325 | 43462.957 | 0.6632 | 0.023859 |  | 325 | 11.4653 | 0.66296 | 0.02516 |  | 325 | 0.66484 | 0.014916 | 325 | 0.67742 | -0.052535 | 325 | 43459.2812 | 325 | 43380.8008 | 325 | 44278.6406
326 | 11.5006 | 43516.4141 | 0.66402 | 0.019391 | 326 | 43498.6289 | 0.66375 | 0.020876 | 326 | 43498.2891 | 0.66374 | 0.020905 |  | 326 | 11.5006 | 0.66302 | 0.024844 |  | 326 | 0.6638 | 0.020591 | 326 | 0.67775 | -0.05427 | 326 | 43481.8398 | 326 | 43393.5586 | 326 | 44294.7188
327 | 11.5358 | 43424.3594 | 0.66261 | 0.027091 | 327 | 43499.3828 | 0.66376 | 0.020813 | 327 | 43423.7695 | 0.6626 | 0.02714 |  | 327 | 11.5358 | 0.66338 | 0.022878 |  | 327 | 0.66381 | 0.020515 | 327 | 0.67801 | -0.055664 | 327 | 43471.1992 | 327 | 43373.8398 | 327 | 44315.8008
328 | 11.5711 | 43381.168 | 0.66195 | 0.030714 | 328 | 43479.957 | 0.66346 | 0.022437 | 328 | 43405.3203 | 0.66232 | 0.028687 |  | 328 | 11.5711 | 0.66392 | 0.019942 |  | 328 | 0.66381 | 0.02053 | 328 | 0.67736 | -0.052217 | 328 | 43496.0781 | 328 | 43368 | 328 | 44295.0391
329 | 11.6064 | 43482.5391 | 0.6635 | 0.022221 | 329 | 43376.4883 | 0.66188 | 0.031107 | 329 | 43439.0586 | 0.66284 | 0.025859 |  | 329 | 11.6064 | 0.66388 | 0.020124 |  | 329 | 0.66327 | 0.023494 | 329 | 0.67766 | -0.053796 | 329 | 43525.0781 | 329 | 43392 | 329 | 44295.5195
330 | 11.6417 | 43370.4414 | 0.66179 | 0.031615 | 330 | 43462.9219 | 0.6632 | 0.023862 | 330 | 43374.1172 | 0.66185 | 0.031306 |  | 330 | 11.6417 | 0.66411 | 0.018906 |  | 330 | 0.66264 | 0.026946 | 330 | 0.67752 | -0.053083 | 330 | 43554.8789 | 330 | 43413.2383 | 330 | 44320.8008
331 | 11.6769 | 43409.3281 | 0.66238 | 0.028351 | 331 | 43382.2578 | 0.66197 | 0.030622 | 331 | 43370.5586 | 0.66179 | 0.031605 |  | 331 | 11.6769 | 0.66408 | 0.01905 |  | 331 | 0.66204 | 0.030247 | 331 | 0.67704 | -0.050539 | 331 | 43572.2812 | 331 | 43427.4805 | 331 | 44347.6797
332 | 11.7122 | 43482.7656 | 0.6635 | 0.022202 | 332 | 43396.4531 | 0.66219 | 0.029431 | 332 | 43444.7188 | 0.66292 | 0.025385 |  | 332 | 11.7122 | 0.66341 | 0.022737 |  | 332 | 0.66177 | 0.031717 | 332 | 0.67584 | -0.044219 | 332 | 43568.6016 | 332 | 43448 | 332 | 44324.3594
333 | 11.7475 | 43407.2461 | 0.66235 | 0.028526 | 333 | 43488.6133 | 0.66359 | 0.021713 | 333 | 43416.1133 | 0.66249 | 0.027782 |  | 333 | 11.7475 | 0.66371 | 0.021062 |  | 333 | 0.66247 | 0.027879 | 333 | 0.6753 | -0.041339 | 333 | 43582.8789 | 333 | 43459.8789 | 333 | 44309.0391
334 | 11.7828 | 43437.3125 | 0.66281 | 0.026005 | 334 | 43443.4609 | 0.6629 | 0.025491 | 334 | 43448.1523 | 0.66298 | 0.025098 |  | 334 | 11.7828 | 0.66428 | 0.017947 |  | 334 | 0.66368 | 0.021237 | 334 | 0.67482 | -0.038805 | 334 | 43590.1211 | 334 | 43492.3984 | 334 | 44340.9219
335 | 11.8181 | 43448.582 | 0.66298 | 0.025062 | 335 | 43442.2109 | 0.66289 | 0.025595 | 335 | 43436.1953 | 0.66279 | 0.026099 |  | 335 | 11.8181 | 0.66445 | 0.017037 |  | 335 | 0.66465 | 0.015961 | 335 | 0.67479 | -0.038665 | 335 | 43559.9219 | 335 | 43526.6406 | 335 | 44345.8008
336 | 11.8533 | 43414.9922 | 0.66247 | 0.027876 | 336 | 43474.0078 | 0.66337 | 0.022934 | 336 | 43417.9102 | 0.66251 | 0.027631 |  | 336 | 11.8533 | 0.66501 | 0.013996 |  | 336 | 0.66491 | 0.014499 | 336 | 0.67491 | -0.039273 | 336 | 43618.0391 | 336 | 43550.9609 | 336 | 44343.6016
337 | 11.8886 | 43452.9492 | 0.66305 | 0.024696 | 337 | 43472.9453 | 0.66335 | 0.023023 | 337 | 43430.9102 | 0.66271 | 0.026542 |  | 337 | 11.8886 | 0.66517 | 0.013077 |  | 337 | 0.66524 | 0.012736 | 337 | 0.67484 | -0.038923 | 337 | 43632.7617 | 337 | 43560.9609 | 337 | 44354.0781
338 | 11.9239 | 43457.0195 | 0.66311 | 0.024355 | 338 | 43464.5742 | 0.66323 | 0.023723 | 338 | 43409.8164 | 0.66239 | 0.02831 |  | 338 | 11.9239 | 0.66514 | 0.013285 |  | 338 | 0.66486 | 0.014817 | 338 | 0.67462 | -0.03773 | 338 | 43560.0391 | 338 | 43556.2812 | 338 | 44353.8398
339 | 11.9592 | 43464.457 | 0.66323 | 0.023733 | 339 | 43424.082 | 0.66261 | 0.027114 | 339 | 43436.5039 | 0.6628 | 0.026073 |  | 339 | 11.9592 | 0.66525 | 0.012653 |  | 339 | 0.6652 | 0.01291 | 339 | 0.67446 | -0.036916 | 339 | 43591.8008 | 339 | 43537.8398 | 339 | 44355.5195
340 | 11.9944 | 43500.0469 | 0.66377 | 0.020758 | 340 | 43446.5703 | 0.66295 | 0.02523 | 340 | 43463.5586 | 0.66321 | 0.023808 |  | 340 | 11.9944 | 0.66591 | 0.0090647 |  | 340 | 0.66535 | 0.012116 | 340 | 0.67448 | -0.037005 | 340 | 43599.4805 | 340 | 43520.1211 | 340 | 44367
341 | 12.0297 | 43485.25 | 0.66354 | 0.021994 | 341 | 43503.2891 | 0.66382 | 0.020487 | 341 | 43455.0195 | 0.66308 | 0.024523 |  | 341 | 12.0297 | 0.6651 | 0.01347 |  | 341 | 0.66546 | 0.011496 | 341 | 0.67469 | -0.038098 | 341 | 43612.3984 | 341 | 43470.4414 | 341 | 44363.2812
342 | 12.065 | 43489.0469 | 0.6636 | 0.021677 | 342 | 43490.625 | 0.66362 | 0.021545 | 342 | 43470.8711 | 0.66332 | 0.023197 |  | 342 | 12.065 | 0.66546 | 0.0115 |  | 342 | 0.6649 | 0.014598 | 342 | 0.67438 | -0.036499 | 342 | 43597.4805 | 342 | 43482.4414 | 342 | 44381.4414
343 | 12.1003 | 43468.5859 | 0.66329 | 0.023388 | 343 | 43499.3359 | 0.66376 | 0.020817 | 343 | 43422.5586 | 0.66259 | 0.027242 |  | 343 | 12.1003 | 0.66596 | 0.0088079 |  | 343 | 0.66509 | 0.013557 | 343 | 0.67456 | -0.037406 | 343 | 43544.1211 | 343 | 43509.5586 | 343 | 44335.0781
344 | 12.1356 | 43460.8594 | 0.66317 | 0.024034 | 344 | 43439.0156 | 0.66284 | 0.025863 | 344 | 43402.5547 | 0.66228 | 0.028919 |  | 344 | 12.1356 | 0.66526 | 0.012634 |  | 344 | 0.66534 | 0.012162 | 344 | 0.67495 | -0.039493 | 344 | 43545.0781 | 344 | 43482.9219 | 344 | 44365.3984
345 | 12.1708 | 43425.7852 | 0.66264 | 0.026971 | 345 | 43494.6289 | 0.66369 | 0.02121 | 345 | 43413.4688 | 0.66245 | 0.028004 |  | 345 | 12.1708 | 0.66549 | 0.01138 |  | 345 | 0.66548 | 0.01141 | 345 | 0.67615 | -0.045844 | 345 | 43520.8398 | 345 | 43497.7617 | 345 | 44340.2812
346 | 12.2061 | 43479.1797 | 0.66345 | 0.022502 | 346 | 43463.8125 | 0.66322 | 0.023787 | 346 | 43421.2109 | 0.66257 | 0.027355 |  | 346 | 12.2061 | 0.66579 | 0.0096989 |  | 346 | 0.66581 | 0.0095967 | 346 | 0.67759 | -0.053449 | 346 | 43486.9609 | 346 | 43487.1602 | 346 | 44300.3984
347 | 12.2414 | 43474.8945 | 0.66338 | 0.02286 | 347 | 43475.8633 | 0.6634 | 0.022779 | 347 | 43420.8867 | 0.66256 | 0.027382 |  | 347 | 12.2414 | 0.66662 | 0.0051887 |  | 347 | 0.6656 | 0.010741 | 347 | 0.67829 | -0.05713 | 347 | 43419.6016 | 347 | 43487.6406 | 347 | 44231.5586
348 | 12.2767 | 43485.7031 | 0.66355 | 0.021956 | 348 | 43478.3359 | 0.66344 | 0.022572 | 348 | 43494.1172 | 0.66368 | 0.021253 |  | 348 | 12.2767 | 0.66701 | 0.0031073 |  | 348 | 0.66534 | 0.012196 | 348 | 0.67804 | -0.055807 | 348 | 43431.6016 | 348 | 43462.8398 | 348 | 44196.6016
349 | 12.3119 | 43454.4453 | 0.66307 | 0.024571 | 349 | 43512.7031 | 0.66396 | 0.019701 | 349 | 43465.9336 | 0.66325 | 0.02361 |  | 349 | 12.3119 | 0.66762 | -0.00020846 |  | 349 | 0.66675 | 0.004511 | 349 | 0.67932 | -0.062525 | 349 | 43436.5195 | 349 | 43443.0781 | 349 | 44186.5195
350 | 12.3472 | 43466.4453 | 0.66326 | 0.023567 | 350 | 43433.5547 | 0.66275 | 0.02632 | 350 | 43462.4766 | 0.66319 | 0.023899 |  | 350 | 12.3472 | 0.66765 | -0.00035854 |  | 350 | 0.66648 | 0.0059611 | 350 | 0.67892 | -0.060423 | 350 | 43462.4805 | 350 | 43421.5586 | 350 | 44179.3984
351 | 12.3825 | 43434.0078 | 0.66276 | 0.026282 | 351 | 43462.3906 | 0.66319 | 0.023906 | 351 | 43457.2109 | 0.66311 | 0.024339 |  | 351 | 12.3825 | 0.66745 | 0.00071584 |  | 351 | 0.6668 | 0.0042172 | 351 | 0.67889 | -0.06027 | 351 | 43454.5195 | 351 | 43408.4805 | 351 | 44144.8789
352 | 12.4178 | 43459.4453 | 0.66315 | 0.024152 | 352 | 43464.9492 | 0.66323 | 0.023692 | 352 | 43485.4219 | 0.66355 | 0.02198 |  | 352 | 12.4178 | 0.66682 | 0.0041231 |  | 352 | 0.66631 | 0.0068881 | 352 | 0.67816 | -0.056438 | 352 | 43475.6797 | 352 | 43384 | 352 | 44134.3594
353 | 12.4531 | 43509.418 | 0.66391 | 0.019975 | 353 | 43503.5703 | 0.66382 | 0.020464 | 353 | 43488.2344 | 0.66359 | 0.021745 |  | 353 | 12.4531 | 0.66643 | 0.0062137 |  | 353 | 0.66693 | 0.0035322 | 353 | 0.67766 | -0.053847 | 353 | 43473.3984 | 353 | 43378.6406 | 353 | 44142.5195
354 | 12.4883 | 43503.8008 | 0.66383 | 0.020444 | 354 | 43534.3906 | 0.66429 | 0.017891 | 354 | 43492.2852 | 0.66365 | 0.021406 |  | 354 | 12.4883 | 0.66555 | 0.011006 |  | 354 | 0.66673 | 0.0045977 | 354 | 0.67784 | -0.054756 | 354 | 43445.2383 | 354 | 43409.4805 | 354 | 44141.7617
355 | 12.5236 | 43409.1641 | 0.66238 | 0.028365 | 355 | 43525.5586 | 0.66416 | 0.018628 | 355 | 43459.9062 | 0.66316 | 0.024114 |  | 355 | 12.5236 | 0.6649 | 0.014556 |  | 355 | 0.66622 | 0.0073822 | 355 | 0.67737 | -0.052309 | 355 | 43440.5195 | 355 | 43452.6406 | 355 | 44127.8008
356 | 12.5589 | 43412 | 0.66242 | 0.028127 | 356 | 43417.1094 | 0.6625 | 0.027698 | 356 | 43468.4648 | 0.66329 | 0.023398 |  | 356 | 12.5589 | 0.6655 | 0.011304 |  | 356 | 0.6661 | 0.0080194 | 356 | 0.67708 | -0.050773 | 356 | 43448.8398 | 356 | 43505.7617 | 356 | 44131.4805
357 | 12.5942 | 43407.4922 | 0.66236 | 0.028505 | 357 | 43463.3008 | 0.66321 | 0.02383 | 357 | 43433.7812 | 0.66276 | 0.026301 |  | 357 | 12.5942 | 0.66548 | 0.011413 |  | 357 | 0.66568 | 0.010329 | 357 | 0.67773 | -0.054182 | 357 | 43524 | 357 | 43494.5195 | 357 | 44151.8789
358 | 12.6294 | 43421.5664 | 0.66257 | 0.027325 | 358 | 43427.3906 | 0.66266 | 0.026837 | 358 | 43449.8164 | 0.663 | 0.024958 |  | 358 | 12.6294 | 0.66495 | 0.014272 |  | 358 | 0.66546 | 0.011504 | 358 | 0.67777 | -0.054394 | 358 | 43547.4414 | 358 | 43514.5195 | 358 | 44172.1602
359 | 12.6647 | 43439.8672 | 0.66285 | 0.025791 | 359 | 43413.8594 | 0.66245 | 0.027971 | 359 | 43420.4023 | 0.66255 | 0.027422 |  | 359 | 12.6647 | 0.66509 | 0.013561 |  | 359 | 0.66485 | 0.014855 | 359 | 0.67731 | -0.051983 | 359 | 43551.9219 | 359 | 43498.0781 | 359 | 44198.5586
360 | 12.7 | 43471.4805 | 0.66333 | 0.023146 | 360 | 43417.2344 | 0.6625 | 0.027688 | 360 | 43430.4883 | 0.66271 | 0.026577 |  | 360 | 12.7 | 0.66469 | 0.015699 |  | 360 | 0.66486 | 0.014802 | 360 | 0.67733 | -0.052057 | 360 | 43586.8789 | 360 | 43518.5195 | 360 | 44282.2383
361 | 12.7353 | 43471.1367 | 0.66333 | 0.023174 | 361 | 43434.1875 | 0.66276 | 0.026267 | 361 | 43404.4727 | 0.66231 | 0.028758 |  | 361 | 12.7353 | 0.66471 | 0.015635 |  | 361 | 0.66521 | 0.012884 | 361 | 0.6773 | -0.05194 | 361 | 43567.4805 | 361 | 43561.3984 | 361 | 44361.3203
362 | 12.7706 | 43490.457 | 0.66362 | 0.021559 | 362 | 43459.668 | 0.66315 | 0.024134 | 362 | 43390.0547 | 0.66209 | 0.029968 |  | 362 | 12.7706 | 0.66548 | 0.011394 |  | 362 | 0.66548 | 0.011413 | 362 | 0.67675 | -0.049013 | 362 | 43628.8398 | 362 | 43589.6797 | 362 | 44387.1602
363 | 12.8058 | 43473.6289 | 0.66337 | 0.022966 | 363 | 43477.9609 | 0.66343 | 0.022604 | 363 | 43475.625 | 0.6634 | 0.022799 |  | 363 | 12.8058 | 0.66599 | 0.0086081 |  | 363 | 0.66518 | 0.013066 | 363 | 0.67683 | -0.049427 | 363 | 43586.2812 | 363 | 43580.0391 | 363 | 44413.4805
364 | 12.8411 | 43449.5234 | 0.663 | 0.024983 | 364 | 43498.9062 | 0.66375 | 0.020853 | 364 | 43469.0273 | 0.66329 | 0.023351 |  | 364 | 12.8411 | 0.66633 | 0.0067902 |  | 364 | 0.66464 | 0.015995 | 364 | 0.67722 | -0.051483 | 364 | 43609.8398 | 364 | 43585.4805 | 364 | 44471.3594
365 | 12.8764 | 43441.4844 | 0.66287 | 0.025656 | 365 | 43490.1836 | 0.66362 | 0.021582 | 365 | 43436.3438 | 0.6628 | 0.026087 |  | 365 | 12.8764 | 0.66652 | 0.0057464 |  | 365 | 0.66455 | 0.016461 | 365 | 0.67661 | -0.048281 | 365 | 43599.3594 | 365 | 43622.6016 | 365 | 44448.9219
366 | 12.9117 | 43441.5039 | 0.66287 | 0.025654 | 366 | 43480.5039 | 0.66347 | 0.022391 | 366 | 43452.5703 | 0.66304 | 0.024728 |  | 366 | 12.9117 | 0.66712 | 0.0024902 |  | 366 | 0.6656 | 0.01073 | 366 | 0.67562 | -0.043029 | 366 | 43621.1211 | 366 | 43577.6797 | 366 | 44456.6016
367 | 12.9469 | 43393.8984 | 0.66215 | 0.029645 | 367 | 43465.707 | 0.66324 | 0.023629 | 367 | 43471.3242 | 0.66333 | 0.023159 |  | 367 | 12.9469 | 0.66672 | 0.0046767 |  | 367 | 0.6658 | 0.0096723 | 367 | 0.67445 | -0.036839 | 367 | 43581.9219 | 367 | 43634.6406 | 367 | 44403.3984
368 | 12.9822 | 43389.9961 | 0.66209 | 0.029973 | 368 | 43414.2031 | 0.66246 | 0.027942 | 368 | 43462.707 | 0.6632 | 0.02388 |  | 368 | 12.9822 | 0.66611 | 0.0080044 |  | 368 | 0.66684 | 0.0040141 | 368 | 0.67373 | -0.03304 | 368 | 43625 | 368 | 43588.1211 | 368 | 44379.5586
369 | 13.0175 | 43430.5742 | 0.66271 | 0.02657 | 369 | 43459.3828 | 0.66315 | 0.024158 | 369 | 43435.1836 | 0.66278 | 0.026184 |  | 369 | 13.0175 | 0.66741 | 0.00092995 |  | 369 | 0.6669 | 0.003713 | 369 | 0.67395 | -0.034169 | 369 | 43690.6406 | 369 | 43580.7617 | 369 | 44363.9219
370 | 13.0528 | 43436.0352 | 0.66279 | 0.026112 | 370 | 43481.9141 | 0.66349 | 0.022273 | 370 | 43398.5352 | 0.66222 | 0.029256 |  | 370 | 13.0528 | 0.66674 | 0.0045337 |  | 370 | 0.66687 | 0.0038559 | 370 | 0.67408 | -0.034867 | 370 | 43700.9219 | 370 | 43657.3984 | 370 | 44385.3203
371 | 13.0881 | 43476.9219 | 0.66342 | 0.022691 | 371 | 43407.9414 | 0.66236 | 0.028467 | 371 | 43387.8242 | 0.66206 | 0.030155 |  | 371 | 13.0881 | 0.66669 | 0.0048158 |  | 371 | 0.66673 | 0.0045977 | 371 | 0.67487 | -0.039085 | 371 | 43707.9609 | 371 | 43640.5586 | 371 | 44336.3203
372 | 13.1233 | 43503.8242 | 0.66383 | 0.020442 | 372 | 43425.7422 | 0.66263 | 0.026975 | 372 | 43470.293 | 0.66331 | 0.023245 |  | 372 | 13.1233 | 0.66595 | 0.0088721 |  | 372 | 0.66629 | 0.0070164 | 372 | 0.67634 | -0.046838 | 372 | 43701.1602 | 372 | 43679.6016 | 372 | 44342.3203
373 | 13.1586 | 43405.2031 | 0.66232 | 0.028697 | 373 | 43502.2773 | 0.6638 | 0.020572 | 373 | 43413.5547 | 0.66245 | 0.027996 |  | 373 | 13.1586 | 0.66677 | 0.0044169 |  | 373 | 0.66613 | 0.0078498 | 373 | 0.67689 | -0.049775 | 373 | 43696.2383 | 373 | 43623.9609 | 373 | 44350.7188
374 | 13.1939 | 43444.3906 | 0.66292 | 0.025413 | 374 | 43508.6953 | 0.6639 | 0.020036 | 374 | 43376.582 | 0.66188 | 0.031099 |  | 374 | 13.1939 | 0.66678 | 0.0043528 |  | 374 | 0.66707 | 0.0027458 | 374 | 0.67673 | -0.048929 | 374 | 43653.0781 | 374 | 43659.3203 | 374 | 44345.1602
375 | 13.2292 | 43416.5117 | 0.66249 | 0.027749 | 375 | 43479.8477 | 0.66346 | 0.022446 | 375 | 43356.2617 | 0.66157 | 0.032806 |  | 375 | 13.2292 | 0.66666 | 0.0049666 |  | 375 | 0.66674 | 0.0045638 | 375 | 0.6766 | -0.048255 | 375 | 43620.7188 | 375 | 43612.1602 | 375 | 44335.8789
376 | 13.2644 | 43444.7422 | 0.66292 | 0.025383 | 376 | 43454.4883 | 0.66307 | 0.024567 | 376 | 43446.4805 | 0.66295 | 0.025238 |  | 376 | 13.2644 | 0.66685 | 0.0039652 |  | 376 | 0.66728 | 0.0016441 | 376 | 0.67651 | -0.047743 | 376 | 43586.1992 | 376 | 43609.5195 | 376 | 44335.4805
377 | 13.2997 | 43411.0352 | 0.66241 | 0.028208 | 377 | 43461.25 | 0.66318 | 0.024001 | 377 | 43425.7305 | 0.66263 | 0.026976 |  | 377 | 13.2997 | 0.66688 | 0.003811 |  | 377 | 0.6675 | 0.0004191 | 377 | 0.67651 | -0.047746 | 377 | 43567.2383 | 377 | 43605.8398 | 377 | 44281.6016
378 | 13.335 | 43396.3047 | 0.66219 | 0.029443 | 378 | 43488.2891 | 0.66359 | 0.02174 | 378 | 43435.6836 | 0.66279 | 0.026142 |  | 378 | 13.335 | 0.66633 | 0.0067828 |  | 378 | 0.66687 | 0.0038446 | 378 | 0.67713 | -0.051026 | 378 | 43563.7617 | 378 | 43598.0391 | 378 | 44278.5195
379 | 13.3703 | 43419.6055 | 0.66254 | 0.027489 | 379 | 43462.25 | 0.66319 | 0.023918 | 379 | 43392.2422 | 0.66212 | 0.029784 |  | 379 | 13.3703 | 0.66653 | 0.0056972 |  | 379 | 0.66768 | -0.00054249 | 379 | 0.67773 | -0.054208 | 379 | 43567.9219 | 379 | 43525.1602 | 379 | 44303.9609
380 | 13.4056 | 43429.8516 | 0.6627 | 0.02663 | 380 | 43424.2344 | 0.66261 | 0.027101 | 380 | 43368.7852 | 0.66177 | 0.031754 |  | 380 | 13.4056 | 0.66707 | 0.0027423 |  | 380 | 0.66808 | -0.0026931 | 380 | 0.67735 | -0.052188 | 380 | 43557.3203 | 380 | 43502.4805 | 380 | 44285.3594
381 | 13.4408 | 43402.5234 | 0.66228 | 0.028922 | 381 | 43417.0195 | 0.6625 | 0.027706 | 381 | 43390.9141 | 0.6621 | 0.029896 |  | 381 | 13.4408 | 0.66746 | 0.00065933 |  | 381 | 0.6679 | -0.0017139 | 381 | 0.6779 | -0.055081 | 381 | 43519.8789 | 381 | 43522.6406 | 381 | 44240.5586
382 | 13.4761 | 43391.3945 | 0.66211 | 0.029855 | 382 | 43461.8008 | 0.66318 | 0.023955 | 382 | 43472.7891 | 0.66335 | 0.023036 |  | 382 | 13.4761 | 0.66802 | -0.002363 |  | 382 | 0.66717 | 0.0022119 | 382 | 0.67739 | -0.052393 | 382 | 43540.3203 | 382 | 43558.1992 | 382 | 44157.4805
383 | 13.5114 | 43438.4766 | 0.66283 | 0.025908 | 383 | 43439.9414 | 0.66285 | 0.025785 | 383 | 43478.4961 | 0.66344 | 0.022559 |  | 383 | 13.5114 | 0.66683 | 0.0040555 |  | 383 | 0.66743 | 0.00082112 | 383 | 0.67715 | -0.05115 | 383 | 43564.7188 | 383 | 43567.1992 | 383 | 44086
384 | 13.5467 | 43428.8281 | 0.66268 | 0.026716 | 384 | 43470.1289 | 0.66331 | 0.023259 | 384 | 43471.9336 | 0.66334 | 0.023108 |  | 384 | 13.5467 | 0.66617 | 0.0076346 |  | 384 | 0.66773 | -0.00081677 | 384 | 0.67682 | -0.049409 | 384 | 43561.2383 | 384 | 43547.0391 | 384 | 44143.4805
385 | 13.5819 | 43411.1797 | 0.66241 | 0.028196 | 385 | 43473.1328 | 0.66336 | 0.023007 | 385 | 43426.2344 | 0.66264 | 0.026934 |  | 385 | 13.5819 | 0.66606 | 0.0082268 |  | 385 | 0.66685 | 0.0039614 | 385 | 0.67598 | -0.044964 | 385 | 43571.0391 | 385 | 43523.2812 | 385 | 44152.6797
386 | 13.6172 | 43385.2773 | 0.66202 | 0.030369 | 386 | 43462.9375 | 0.6632 | 0.02386 | 386 | 43402.5234 | 0.66228 | 0.028922 |  | 386 | 13.6172 | 0.66601 | 0.0085251 |  | 386 | 0.66594 | 0.0088948 | 386 | 0.67591 | -0.044586 | 386 | 43627.5586 | 386 | 43526.5195 | 386 | 44215.1992
387 | 13.6525 | 43362.5586 | 0.66167 | 0.032277 | 387 | 43389.3828 | 0.66208 | 0.030024 | 387 | 43376.8125 | 0.66189 | 0.03108 |  | 387 | 13.6525 | 0.66564 | 0.010545 |  | 387 | 0.66542 | 0.011712 | 387 | 0.67555 | -0.042691 | 387 | 43644.0391 | 387 | 43569.4805 | 387 | 44275.1992
388 | 13.6878 | 43363.918 | 0.66169 | 0.032163 | 388 | 43382.0117 | 0.66197 | 0.030643 | 388 | 43359.4766 | 0.66162 | 0.032536 |  | 388 | 13.6878 | 0.66593 | 0.0089778 |  | 388 | 0.66605 | 0.0083325 | 388 | 0.676 | -0.045077 | 388 | 43645.9219 | 388 | 43581.1211 | 388 | 44323.3594
389 | 13.7231 | 43432 | 0.66273 | 0.02645 | 389 | 43421.8438 | 0.66257 | 0.027302 | 389 | 43384.6523 | 0.66201 | 0.030421 |  | 389 | 13.7231 | 0.66528 | 0.012513 |  | 389 | 0.66736 | 0.0011932 | 389 | 0.67652 | -0.04779 | 389 | 43654.6406 | 389 | 43648.5195 | 389 | 44328.9219
390 | 13.7583 | 43454.9648 | 0.66308 | 0.024527 | 390 | 43463.668 | 0.66321 | 0.023799 | 390 | 43367.3008 | 0.66174 | 0.031879 |  | 390 | 13.7583 | 0.6658 | 0.0096762 |  | 390 | 0.66842 | -0.0045264 | 390 | 0.67653 | -0.047856 | 390 | 43650.1992 | 390 | 43661.5586 | 390 | 44290.8008
391 | 13.7936 | 43432.9258 | 0.66274 | 0.026373 | 391 | 43476.9688 | 0.66342 | 0.022687 | 391 | 43347.9531 | 0.66145 | 0.033505 |  | 391 | 13.7936 | 0.66607 | 0.0081967 |  | 391 | 0.66886 | -0.0069196 | 391 | 0.67637 | -0.046995 | 391 | 43656.8008 | 391 | 43666.1602 | 391 | 44276.5586
392 | 13.8289 | 43409.6484 | 0.66239 | 0.028324 | 392 | 43441.7188 | 0.66288 | 0.025636 | 392 | 43375.5781 | 0.66187 | 0.031183 |  | 392 | 13.8289 | 0.66547 | 0.011459 |  | 392 | 0.66863 | -0.0056616 | 392 | 0.67654 | -0.047915 | 392 | 43685.6406 | 392 | 43659.9219 | 392 | 44275.6797
393 | 13.8642 | 43466.8203 | 0.66326 | 0.023535 | 393 | 43414.1133 | 0.66246 | 0.02795 | 393 | 43407.875 | 0.66236 | 0.028473 |  | 393 | 13.8642 | 0.66627 | 0.0071181 |  | 393 | 0.66884 | -0.0067961 | 393 | 0.67606 | -0.045375 | 393 | 43649.8008 | 393 | 43651.6016 | 393 | 44307
394 | 13.8994 | 43418.8242 | 0.66253 | 0.027555 | 394 | 43461.3008 | 0.66318 | 0.023997 | 394 | 43406.625 | 0.66234 | 0.028578 |  | 394 | 13.8994 | 0.66687 | 0.0038672 |  | 394 | 0.66838 | -0.0043277 | 394 | 0.67602 | -0.045173 | 394 | 43598.1211 | 394 | 43647.2383 | 394 | 44351.8008
395 | 13.9347 | 43458.125 | 0.66313 | 0.024263 | 395 | 43412.6641 | 0.66243 | 0.028071 | 395 | 43439.5117 | 0.66284 | 0.025821 |  | 395 | 13.9347 | 0.66589 | 0.0091666 |  | 395 | 0.66743 | 0.00080207 | 395 | 0.67603 | -0.045243 | 395 | 43621.9609 | 395 | 43685.6406 | 395 | 44329
396 | 13.97 | 43456.6641 | 0.66311 | 0.024385 | 396 | 43476.5469 | 0.66341 | 0.022722 | 396 | 43454.3477 | 0.66307 | 0.024579 |  | 396 | 13.97 | 0.66627 | 0.0071369 |  | 396 | 0.66678 | 0.004334 | 396 | 0.67594 | -0.044744 | 396 | 43645.5586 | 396 | 43692.4414 | 396 | 44385.8789
397 | 14.0053 | 43414.5391 | 0.66246 | 0.027914 | 397 | 43431.1328 | 0.66272 | 0.026523 | 397 | 43359.5273 | 0.66162 | 0.032532 |  | 397 | 14.0053 | 0.6651 | 0.013497 |  | 397 | 0.66651 | 0.0058329 | 397 | 0.67602 | -0.045173 | 397 | 43688.3984 | 397 | 43672.5586 | 397 | 44349.7188
398 | 14.0406 | 43403.0586 | 0.66229 | 0.028877 | 398 | 43442.4453 | 0.66289 | 0.025576 | 398 | 43381.4688 | 0.66196 | 0.030689 |  | 398 | 14.0406 | 0.66469 | 0.015722 |  | 398 | 0.66704 | 0.0029417 | 398 | 0.67592 | -0.044641 | 398 | 43660.2383 | 398 | 43712.0391 | 398 | 44322.0391
399 | 14.0758 | 43428.4805 | 0.66268 | 0.026745 | 399 | 43446.6797 | 0.66295 | 0.025221 | 399 | 43438.1836 | 0.66282 | 0.025932 |  | 399 | 14.0758 | 0.66356 | 0.021905 |  | 399 | 0.6674 | 0.00095644 | 399 | 0.67505 | -0.040038 | 399 | 43642.3984 | 399 | 43694.9219 | 399 | 44303.2383
400 | 14.1111 | 43402.2539 | 0.66228 | 0.028944 | 400 | 43498.3867 | 0.66374 | 0.020897 | 400 | 43436.4336 | 0.6628 | 0.026079 |  | 400 | 14.1111 | 0.66306 | 0.02462 |  | 400 | 0.66772 | -0.00074546 | 400 | 0.67573 | -0.043635 | 400 | 43666.7188 | 400 | 43743.5195 | 400 | 44254.5195
401 | 14.1464 | 43399.7109 | 0.66224 | 0.029158 | 401 | 43464.2148 | 0.66322 | 0.023753 | 401 | 43441.6406 | 0.66288 | 0.025643 |  | 401 | 14.1464 | 0.66331 | 0.023288 |  | 401 | 0.66731 | 0.0014674 | 401 | 0.67554 | -0.042647 | 401 | 43688 | 401 | 43787.6406 | 401 | 44242.7617
402 | 14.1817 | 43454.418 | 0.66307 | 0.024573 | 402 | 43445.8789 | 0.66294 | 0.025288 | 402 | 43456.8398 | 0.66311 | 0.024371 |  | 402 | 14.1817 | 0.66305 | 0.024677 |  | 402 | 0.66717 | 0.0022158 | 402 | 0.67558 | -0.042846 | 402 | 43680.1211 | 402 | 43786.3203 | 402 | 44236.6797
403 | 14.2169 | 43374.8984 | 0.66186 | 0.03124 | 403 | 43481.3633 | 0.66348 | 0.022319 | 403 | 43374.9414 | 0.66186 | 0.031237 |  | 403 | 14.2169 | 0.66373 | 0.020982 |  | 403 | 0.66682 | 0.0041008 | 403 | 0.67509 | -0.040255 | 403 | 43632.9219 | 403 | 43733.6406 | 403 | 44230.3203
404 | 14.2522 | 43370.7461 | 0.6618 | 0.031589 | 404 | 43429.4297 | 0.66269 | 0.026666 | 404 | 43403.957 | 0.6623 | 0.028801 |  | 404 | 14.2522 | 0.66394 | 0.019828 |  | 404 | 0.66641 | 0.006353 | 404 | 0.67567 | -0.043323 | 404 | 43631.4805 | 404 | 43719.7617 | 404 | 44276.2383
405 | 14.2875 | 43430.0859 | 0.6627 | 0.026611 | 405 | 43389.4023 | 0.66208 | 0.030023 | 405 | 43429.7852 | 0.6627 | 0.026636 |  | 405 | 14.2875 | 0.66391 | 0.019968 |  | 405 | 0.66624 | 0.0072953 | 405 | 0.67533 | -0.041494 | 405 | 43647.5195 | 405 | 43697.6406 | 405 | 44302
406 | 14.3228 | 43440.1406 | 0.66285 | 0.025769 | 406 | 43447.8438 | 0.66297 | 0.025124 | 406 | 43436.0117 | 0.66279 | 0.026114 |  | 406 | 14.3228 | 0.66505 | 0.013735 |  | 406 | 0.66673 | 0.0046165 | 406 | 0.67534 | -0.041552 | 406 | 43678.1992 | 406 | 43665.3203 | 406 | 44284.4414
407 | 14.3581 | 43405.1367 | 0.66232 | 0.028702 | 407 | 43467.1211 | 0.66327 | 0.02351 | 407 | 43389.8906 | 0.66209 | 0.029982 |  | 407 | 14.3581 | 0.66501 | 0.013947 |  | 407 | 0.66719 | 0.0020991 | 407 | 0.67541 | -0.041953 | 407 | 43732.3203 | 407 | 43627.5586 | 407 | 44261.6016
408 | 14.3933 | 43407.5508 | 0.66236 | 0.0285 | 408 | 43396.1094 | 0.66218 | 0.02946 | 408 | 43391.8359 | 0.66212 | 0.029818 |  | 408 | 14.3933 | 0.66596 | 0.0087777 |  | 408 | 0.66757 | 5.0742e-05 | 408 | 0.67509 | -0.040233 | 408 | 43656.9609 | 408 | 43618.3984 | 408 | 44246.3203
409 | 14.4286 | 43405.0547 | 0.66232 | 0.028709 | 409 | 43446.6523 | 0.66295 | 0.025223 | 409 | 43402.1211 | 0.66227 | 0.028955 |  | 409 | 14.4286 | 0.66689 | 0.0037356 |  | 409 | 0.66796 | -0.0020441 | 409 | 0.6748 | -0.038687 | 409 | 43620.6406 | 409 | 43648.8398 | 409 | 44233.8398
410 | 14.4639 | 43398.6875 | 0.66222 | 0.029243 | 410 | 43474.4883 | 0.66338 | 0.022894 | 410 | 43405.0859 | 0.66232 | 0.028707 |  | 410 | 14.4639 | 0.66701 | 0.0031147 |  | 410 | 0.6682 | -0.0033719 | 410 | 0.67481 | -0.038779 | 410 | 43669 | 410 | 43723.2383 | 410 | 44248.2812
411 | 14.4992 | 43389.7461 | 0.66209 | 0.029994 | 411 | 43412.4023 | 0.66243 | 0.028093 | 411 | 43411.918 | 0.66242 | 0.028134 |  | 411 | 14.4992 | 0.66854 | -0.0052158 |  | 411 | 0.66836 | -0.0042153 | 411 | 0.67508 | -0.040166 | 411 | 43647.2383 | 411 | 43775 | 411 | 44234.5195
412 | 14.5344 | 43423.4023 | 0.6626 | 0.027171 | 412 | 43411.8398 | 0.66242 | 0.02814 | 412 | 43442.7539 | 0.66289 | 0.02555 |  | 412 | 14.5344 | 0.66803 | -0.0024568 |  | 412 | 0.6684 | -0.0044291 | 412 | 0.67524 | -0.041056 | 412 | 43657.0391 | 412 | 43780.0781 | 412 | 44242.8789
413 | 14.5697 | 43421.2656 | 0.66257 | 0.02735 | 413 | 43459.6953 | 0.66315 | 0.024132 | 413 | 43442.4062 | 0.66289 | 0.025579 |  | 413 | 14.5697 | 0.66804 | -0.0024794 |  | 413 | 0.66841 | -0.0044739 | 413 | 0.67562 | -0.043051 | 413 | 43642.7617 | 413 | 43803 | 413 | 44250.9219
414 | 14.605 | 43427.7812 | 0.66267 | 0.026804 | 414 | 43470.0391 | 0.66331 | 0.023266 | 414 | 43448.9883 | 0.66299 | 0.025028 |  | 414 | 14.605 | 0.66716 | 0.0022607 |  | 414 | 0.66859 | -0.0054404 | 414 | 0.67563 | -0.043095 | 414 | 43676.7617 | 414 | 43807.9219 | 414 | 44229.8008
415 | 14.6403 | 43352.3516 | 0.66151 | 0.033135 | 415 | 43482.0234 | 0.66349 | 0.022264 | 415 | 43357.0469 | 0.66159 | 0.03274 |  | 415 | 14.6403 | 0.66764 | -0.0003098 |  | 415 | 0.66913 | -0.0083674 | 415 | 0.67564 | -0.04318 | 415 | 43657.4414 | 415 | 43780.5195 | 415 | 44255.7617
416 | 14.6756 | 43389.3047 | 0.66208 | 0.030031 | 416 | 43416.7695 | 0.6625 | 0.027727 | 416 | 43416.7227 | 0.6625 | 0.027731 |  | 416 | 14.6756 | 0.66719 | 0.0021405 |  | 416 | 0.66944 | -0.010031 | 416 | 0.67565 | -0.043224 | 416 | 43710.3984 | 416 | 43748.1602 | 416 | 44249.1602
417 | 14.7108 | 43425.7422 | 0.66263 | 0.026975 | 417 | 43433.1016 | 0.66275 | 0.026358 | 417 | 43388.0156 | 0.66206 | 0.030139 |  | 417 | 14.7108 | 0.66678 | 0.0043567 |  | 417 | 0.66944 | -0.010042 | 417 | 0.67555 | -0.042684 | 417 | 43688.7617 | 417 | 43717.5586 | 417 | 44238.1992
418 | 14.7461 | 43466.4609 | 0.66326 | 0.023565 | 418 | 43452.7227 | 0.66305 | 0.024715 | 418 | 43426.5117 | 0.66265 | 0.02691 |  | 418 | 14.7461 | 0.6676 | -0.00011066 |  | 418 | 0.66915 | -0.0084982 | 418 | 0.67617 | -0.045947 | 418 | 43602.8008 | 418 | 43685.1992 | 418 | 44193.5586
419 | 14.7814 | 43433.8477 | 0.66276 | 0.026296 | 419 | 43442.1016 | 0.66288 | 0.025604 | 419 | 43380.2188 | 0.66194 | 0.030794 |  | 419 | 14.7814 | 0.66762 | -0.00020458 |  | 419 | 0.66849 | -0.0049237 | 419 | 0.67632 | -0.04676 | 419 | 43627.2383 | 419 | 43700.8008 | 419 | 44233.6016
420 | 14.8167 | 43461.8867 | 0.66319 | 0.023948 | 420 | 43395.8789 | 0.66218 | 0.029479 | 420 | 43411.625 | 0.66242 | 0.028158 |  | 420 | 14.8167 | 0.66754 | 0.0002425 |  | 420 | 0.66904 | -0.0078849 | 420 | 0.67719 | -0.051355 | 420 | 43552.0391 | 420 | 43736.1992 | 420 | 44225.3594
421 | 14.8519 | 43411.7734 | 0.66242 | 0.028146 | 421 | 43441.793 | 0.66288 | 0.02563 | 421 | 43394.582 | 0.66216 | 0.029588 |  | 421 | 14.8519 | 0.66763 | -0.00024202 |  | 421 | 0.66883 | -0.0067285 | 421 | 0.67834 | -0.057403 | 421 | 43533.7617 | 421 | 43756.0391 | 421 | 44213.4805
422 | 14.8872 | 43411.1172 | 0.66241 | 0.028201 | 422 | 43440.4023 | 0.66286 | 0.025747 | 422 | 43364.8789 | 0.66171 | 0.032082 |  | 422 | 14.8872 | 0.66729 | 0.0015989 |  | 422 | 0.66855 | -0.0052608 | 422 | 0.67806 | -0.055934 | 422 | 43511.3984 | 422 | 43709.7188 | 422 | 44210.5195
423 | 14.9225 | 43474.7227 | 0.66338 | 0.022874 | 423 | 43483.2266 | 0.66351 | 0.022163 | 423 | 43461.7383 | 0.66318 | 0.023961 |  | 423 | 14.9225 | 0.66762 | -0.0002233 |  | 423 | 0.66924 | -0.0089731 | 423 | 0.67764 | -0.053737 | 423 | 43455.8789 | 423 | 43706.1602 | 423 | 44208.6016
424 | 14.9578 | 43388.3047 | 0.66206 | 0.030115 | 424 | 43509.1172 | 0.66391 | 0.02 | 424 | 43382.082 | 0.66197 | 0.030637 |  | 424 | 14.9578 | 0.66708 | 0.0027045 |  | 424 | 0.67026 | -0.014465 | 424 | 0.67768 | -0.053912 | 424 | 43473.6016 | 424 | 43693.4805 | 424 | 44189.6797
425 | 14.9931 | 43391.1445 | 0.66211 | 0.029876 | 425 | 43440.6562 | 0.66286 | 0.025725 | 425 | 43358.332 | 0.66161 | 0.032632 |  | 425 | 14.9931 | 0.66685 | 0.003984 |  | 425 | 0.67064 | -0.016522 | 425 | 0.67741 | -0.052503 | 425 | 43491.7617 | 425 | 43656.3203 | 425 | 44150.8008
426 | 15.0283 | 43408.25 | 0.66237 | 0.028441 | 426 | 43435.207 | 0.66278 | 0.026182 | 426 | 43347.1914 | 0.66144 | 0.033569 |  | 426 | 15.0283 | 0.66615 | 0.0077516 |  | 426 | 0.67113 | -0.019142 | 426 | 0.67768 | -0.053923 | 426 | 43523.3594 | 426 | 43635.1211 | 426 | 44162.5195
427 | 15.0636 | 43458.8789 | 0.66314 | 0.0242 | 427 | 43432.8203 | 0.66274 | 0.026382 | 427 | 43411.6328 | 0.66242 | 0.028158 |  | 427 | 15.0636 | 0.66574 | 0.0099896 |  | 427 | 0.67086 | -0.017687 | 427 | 0.67713 | -0.051033 | 427 | 43539.0781 | 427 | 43662.0781 | 427 | 44196.0781
428 | 15.0989 | 43452.4961 | 0.66304 | 0.024734 | 428 | 43477.1836 | 0.66342 | 0.022669 | 428 | 43411.9492 | 0.66242 | 0.028131 |  | 428 | 15.0989 | 0.66524 | 0.012695 |  | 428 | 0.67036 | -0.015013 | 428 | 0.67626 | -0.046438 | 428 | 43590.7188 | 428 | 43707.3594 | 428 | 44234.0391
429 | 15.1342 | 43482.8906 | 0.66351 | 0.022192 | 429 | 43461.8633 | 0.66319 | 0.02395 | 429 | 43436.8125 | 0.6628 | 0.026047 |  | 429 | 15.1342 | 0.66581 | 0.0095853 |  | 429 | 0.67021 | -0.014185 | 429 | 0.67597 | -0.044912 | 429 | 43599.8398 | 429 | 43731.2383 | 429 | 44239.6797
430 | 15.1694 | 43478.8398 | 0.66344 | 0.02253 | 430 | 43494.5352 | 0.66368 | 0.021218 | 430 | 43445.7852 | 0.66294 | 0.025296 |  | 430 | 15.1694 | 0.66583 | 0.009525 |  | 430 | 0.66936 | -0.0096236 | 430 | 0.67563 | -0.043092 | 430 | 43629.3594 | 430 | 43751.1211 | 430 | 44210.1602
431 | 15.2047 | 43443.9258 | 0.66291 | 0.025452 | 431 | 43442.7031 | 0.66289 | 0.025554 | 431 | 43356.8945 | 0.66158 | 0.032753 |  | 431 | 15.2047 | 0.66615 | 0.0077895 |  | 431 | 0.66881 | -0.0066462 | 431 | 0.67513 | -0.04045 | 431 | 43701.8398 | 431 | 43766.9609 | 431 | 44242.8398
432 | 15.24 | 43436.9062 | 0.6628 | 0.026039 | 432 | 43426.9961 | 0.66265 | 0.02687 | 432 | 43402.4688 | 0.66228 | 0.028926 |  | 432 | 15.24 | 0.66658 | 0.0054111 |  | 432 | 0.66898 | -0.0075894 | 432 | 0.67528 | -0.041226 | 432 | 43728.7188 | 432 | 43767.8398 | 432 | 44241.2812
433 | 15.2753 | 43459.543 | 0.66315 | 0.024144 | 433 | 43424.0391 | 0.66261 | 0.027118 | 433 | 43444.1211 | 0.66291 | 0.025435 |  | 433 | 15.2753 | 0.66655 | 0.0056144 |  | 433 | 0.66885 | -0.0068746 | 433 | 0.67598 | -0.044971 | 433 | 43736.1992 | 433 | 43775.7188 | 433 | 44276.9219
434 | 15.3106 | 43439.5938 | 0.66285 | 0.025814 | 434 | 43443.9414 | 0.66291 | 0.02545 | 434 | 43424.4375 | 0.66261 | 0.027084 |  | 434 | 15.3106 | 0.66606 | 0.008246 |  | 434 | 0.66889 | -0.0070691 | 434 | 0.67653 | -0.047864 | 434 | 43779.8789 | 434 | 43813.4805 | 434 | 44292.0781
435 | 15.3458 | 43472.2617 | 0.66334 | 0.02308 | 435 | 43405.7578 | 0.66233 | 0.02865 | 435 | 43440.9531 | 0.66287 | 0.025701 |  | 435 | 15.3458 | 0.6662 | 0.0074687 |  | 435 | 0.66943 | -0.0099672 | 435 | 0.67682 | -0.049372 | 435 | 43766.2383 | 435 | 43835.0781 | 435 | 44354.0781
436 | 15.3811 | 43475.7539 | 0.6634 | 0.022788 | 436 | 43467.1055 | 0.66327 | 0.023512 | 436 | 43474.1133 | 0.66337 | 0.022925 |  | 436 | 15.3811 | 0.66623 | 0.0073255 |  | 436 | 0.66994 | -0.012749 | 436 | 0.67657 | -0.048054 | 436 | 43799.7617 | 436 | 43872.6797 | 436 | 44413.3984
437 | 15.4164 | 43431.3906 | 0.66272 | 0.026502 | 437 | 43520.2617 | 0.66408 | 0.01907 | 437 | 43456.9062 | 0.66311 | 0.024365 |  | 437 | 15.4164 | 0.66479 | 0.015166 |  | 437 | 0.67101 | -0.018502 | 437 | 0.67583 | -0.044138 | 437 | 43738.4805 | 437 | 43857.3203 | 437 | 44414
438 | 15.4517 | 43414.1406 | 0.66246 | 0.027947 | 438 | 43487.2734 | 0.66357 | 0.021825 | 438 | 43421.4453 | 0.66257 | 0.027335 |  | 438 | 15.4517 | 0.6655 | 0.011319 |  | 438 | 0.6719 | -0.023276 | 438 | 0.67565 | -0.043195 | 438 | 43705.6016 | 438 | 43842.4805 | 438 | 44362.8398
439 | 15.4869 | 43380.4102 | 0.66194 | 0.030778 | 439 | 43463.4219 | 0.66321 | 0.02382 | 439 | 43459.2891 | 0.66315 | 0.024166 |  | 439 | 15.4869 | 0.66531 | 0.012339 |  | 439 | 0.67268 | -0.027405 | 439 | 0.67532 | -0.041449 | 439 | 43703.6797 | 439 | 43861 | 439 | 44384.4805
440 | 15.5222 | 43390.1914 | 0.66209 | 0.029956 | 440 | 43477.9062 | 0.66343 | 0.022608 | 440 | 43485.2227 | 0.66354 | 0.021997 |  | 440 | 15.5222 | 0.6666 | 0.0053169 |  | 440 | 0.6737 | -0.032859 | 440 | 0.67491 | -0.039313 | 440 | 43732.9609 | 440 | 43821.4805 | 440 | 44375.4805
441 | 15.5575 | 43408.3359 | 0.66237 | 0.028434 | 441 | 43465.7266 | 0.66324 | 0.023627 | 441 | 43419.457 | 0.66254 | 0.027502 |  | 441 | 15.5575 | 0.66659 | 0.0053657 |  | 441 | 0.6735 | -0.031773 | 441 | 0.675 | -0.039747 | 441 | 43701.4414 | 441 | 43799.3203 | 441 | 44367.6406
442 | 15.5928 | 43434.3594 | 0.66277 | 0.026253 | 442 | 43451.9258 | 0.66303 | 0.024782 | 442 | 43422.6641 | 0.66259 | 0.027233 |  | 442 | 15.5928 | 0.66691 | 0.0036602 |  | 442 | 0.67259 | -0.026931 | 442 | 0.67523 | -0.040964 | 442 | 43704.3594 | 442 | 43846.0781 | 442 | 44346.6797
443 | 15.6281 | 43475.3711 | 0.66339 | 0.02282 | 443 | 43449.6523 | 0.663 | 0.024972 | 443 | 43452.4453 | 0.66304 | 0.024738 |  | 443 | 15.6281 | 0.66723 | 0.0018997 |  | 443 | 0.67239 | -0.025883 | 443 | 0.67567 | -0.043305 | 443 | 43685.6797 | 443 | 43795.7617 | 443 | 44310.8398
444 | 15.6633 | 43433.7461 | 0.66276 | 0.026304 | 444 | 43498.4258 | 0.66374 | 0.020893 | 444 | 43382.1055 | 0.66197 | 0.030635 |  | 444 | 15.6633 | 0.66718 | 0.0021666 |  | 444 | 0.67261 | -0.027068 | 444 | 0.67519 | -0.040777 | 444 | 43744.1211 | 444 | 43896.8789 | 444 | 44260.4414
445 | 15.6986 | 43393.5703 | 0.66214 | 0.029673 | 445 | 43466.1172 | 0.66325 | 0.023594 | 445 | 43365.0117 | 0.66171 | 0.032071 |  | 445 | 15.6986 | 0.66805 | -0.0025242 |  | 445 | 0.67196 | -0.023565 | 445 | 0.67505 | -0.040012 | 445 | 43723.2812 | 445 | 43905.3984 | 445 | 44231.6797
446 | 15.7339 | 43420.9453 | 0.66256 | 0.027377 | 446 | 43463.7422 | 0.66321 | 0.023793 | 446 | 43426.8906 | 0.66265 | 0.026879 |  | 446 | 15.7339 | 0.66808 | -0.0026818 |  | 446 | 0.67217 | -0.024686 | 446 | 0.67294 | -0.028793 | 446 | 43706.2383 | 446 | 43935.7617 | 446 | 44208.5586
447 | 15.7692 | 43499.9023 | 0.66377 | 0.02077 | 447 | 43454.1406 | 0.66307 | 0.024596 | 447 | 43490.1602 | 0.66362 | 0.021584 |  | 447 | 15.7692 | 0.6673 | 0.0015427 |  | 447 | 0.67133 | -0.020205 | 447 | 0.67384 | -0.033608 | 447 | 43678.6797 | 447 | 43990.8398 | 447 | 44229.3594
448 | 15.8044 | 43439.0234 | 0.66284 | 0.025862 | 448 | 43521.0664 | 0.66409 | 0.019003 | 448 | 43438.25 | 0.66283 | 0.025927 |  | 448 | 15.8044 | 0.66632 | 0.0068544 |  | 448 | 0.67153 | -0.021249 | 448 | 0.67363 | -0.032508 | 448 | 43665.1992 | 448 | 43985.8789 | 448 | 44270.2812
449 | 15.8397 | 43427.4688 | 0.66266 | 0.02683 | 449 | 43430.5273 | 0.66271 | 0.026574 | 449 | 43417.1875 | 0.6625 | 0.027692 |  | 449 | 15.8397 | 0.66637 | 0.0065528 |  | 449 | 0.67124 | -0.019703 | 449 | 0.67302 | -0.029256 | 449 | 43616 | 449 | 43952.2383 | 449 | 44305.9219
450 | 15.875 | 43430.1328 | 0.6627 | 0.026607 | 450 | 43427.4766 | 0.66266 | 0.02683 | 450 | 43406.75 | 0.66234 | 0.028567 |  | 450 | 15.875 | 0.666 | 0.0085666 |  | 450 | 0.67059 | -0.016216 | 450 | 0.67358 | -0.032228 | 450 | 43560.9609 | 450 | 43950.4805 | 450 | 44301.3984
451 | 15.9103 | 43478.1445 | 0.66343 | 0.022588 | 451 | 43386.0273 | 0.66203 | 0.030306 | 451 | 43469.2852 | 0.6633 | 0.023329 |  | 451 | 15.9103 | 0.66579 | 0.009744 |  | 451 | 0.66984 | -0.012182 | 451 | 0.67355 | -0.032061 | 451 | 43604.2812 | 451 | 43885.6406 | 451 | 44303
452 | 15.9456 | 43430.6172 | 0.66271 | 0.026566 | 452 | 43528.5938 | 0.6642 | 0.018375 | 452 | 43444.3203 | 0.66292 | 0.025419 |  | 452 | 15.9456 | 0.66572 | 0.010106 |  | 452 | 0.66945 | -0.010094 | 452 | 0.67358 | -0.032194 | 452 | 43612.2812 | 452 | 43853.5195 | 452 | 44280.5586
453 | 15.9808 | 43441.5586 | 0.66288 | 0.02565 | 453 | 43475.4648 | 0.66339 | 0.022812 | 453 | 43409.6641 | 0.66239 | 0.028323 |  | 453 | 15.9808 | 0.66594 | 0.0088909 |  | 453 | 0.66967 | -0.011297 | 453 | 0.67409 | -0.034948 | 453 | 43630.3594 | 453 | 43843.5195 | 453 | 44246
454 | 16.0161 | 43421.5234 | 0.66257 | 0.027328 | 454 | 43444.9531 | 0.66293 | 0.025366 | 454 | 43428.7383 | 0.66268 | 0.026724 |  | 454 | 16.0161 | 0.66588 | 0.009257 |  | 454 | 0.67018 | -0.014014 | 454 | 0.67422 | -0.035607 | 454 | 43669.7188 | 454 | 43827.6797 | 454 | 44231.5195
455 | 16.0514 | 43476.1562 | 0.6634 | 0.022755 | 455 | 43471.3672 | 0.66333 | 0.023155 | 455 | 43452.4102 | 0.66304 | 0.024741 |  | 455 | 16.0514 | 0.66558 | 0.010851 |  | 455 | 0.67021 | -0.014182 | 455 | 0.67435 | -0.036304 | 455 | 43673.2812 | 455 | 43833.1211 | 455 | 44200.2812
456 | 16.0867 | 43440.3477 | 0.66286 | 0.025751 | 456 | 43454.7383 | 0.66308 | 0.024546 | 456 | 43433.2812 | 0.66275 | 0.026343 |  | 456 | 16.0867 | 0.66558 | 0.010862 |  | 456 | 0.6695 | -0.010389 | 456 | 0.67416 | -0.035283 | 456 | 43651.3203 | 456 | 43861.1602 | 456 | 44213.8789
457 | 16.1219 | 43441.4648 | 0.66287 | 0.025658 | 457 | 43446.9805 | 0.66296 | 0.025196 | 457 | 43438.5078 | 0.66283 | 0.025905 |  | 457 | 16.1219 | 0.6656 | 0.01076 |  | 457 | 0.66999 | -0.013007 | 457 | 0.67451 | -0.037188 | 457 | 43659 | 457 | 43889 | 457 | 44212.2383
458 | 16.1572 | 43411.9805 | 0.66242 | 0.028128 | 458 | 43453.7695 | 0.66306 | 0.024627 | 458 | 43393.8516 | 0.66215 | 0.029649 |  | 458 | 16.1572 | 0.6655 | 0.011274 |  | 458 | 0.67046 | -0.015505 | 458 | 0.67503 | -0.039935 | 458 | 43678.9219 | 458 | 43956.1992 | 458 | 44230.1602
459 | 16.1925 | 43406.5352 | 0.66234 | 0.028585 | 459 | 43412.9141 | 0.66244 | 0.02805 | 459 | 43382.043 | 0.66197 | 0.03064 |  | 459 | 16.1925 | 0.66543 | 0.011682 |  | 459 | 0.67017 | -0.013969 | 459 | 0.67546 | -0.042185 | 459 | 43572.0781 | 459 | 43977.4805 | 459 | 44185.6406
460 | 16.2278 | 43435.7344 | 0.66279 | 0.026138 | 460 | 43411.7266 | 0.66242 | 0.02815 | 460 | 43420.0664 | 0.66255 | 0.027451 |  | 460 | 16.2278 | 0.6656 | 0.01076 |  | 460 | 0.66974 | -0.011641 | 460 | 0.67494 | -0.039431 | 460 | 43592.3594 | 460 | 44012.9609 | 460 | 44197.9219
461 | 16.2631 | 43476.168 | 0.6634 | 0.022754 | 461 | 43454.0195 | 0.66307 | 0.024607 | 461 | 43453.75 | 0.66306 | 0.024629 |  | 461 | 16.2631 | 0.66614 | 0.007797 |  | 461 | 0.67025 | -0.014413 | 461 | 0.67398 | -0.034369 | 461 | 43601.2812 | 461 | 44072.9609 | 461 | 44048.0391
462 | 16.2983 | 43476.9648 | 0.66342 | 0.022687 | 462 | 43467.5039 | 0.66327 | 0.023478 | 462 | 43437.8008 | 0.66282 | 0.025965 |  | 462 | 16.2983 | 0.66616 | 0.0077328 |  | 462 | 0.67011 | -0.013674 | 462 | 0.67342 | -0.031359 | 462 | 43604.8398 | 462 | 44060.4805 | 462 | 44106.4805
463 | 16.3336 | 43453.8867 | 0.66306 | 0.024618 | 463 | 43454.4297 | 0.66307 | 0.024572 | 463 | 43454.168 | 0.66307 | 0.024594 |  | 463 | 16.3336 | 0.66603 | 0.0083893 |  | 463 | 0.66993 | -0.012667 | 463 | 0.67242 | -0.026024 | 463 | 43643.0781 | 463 | 44046.0781 | 463 | 44107.6797
464 | 16.3689 | 43403.3906 | 0.66229 | 0.028849 | 464 | 43476.1719 | 0.6634 | 0.022753 | 464 | 43395.1914 | 0.66217 | 0.029537 |  | 464 | 16.3689 | 0.66552 | 0.011213 |  | 464 | 0.66964 | -0.011096 | 464 | 0.6722 | -0.024867 | 464 | 43707.9609 | 464 | 44021.3594 | 464 | 44055.6406
465 | 16.4042 | 43416.8555 | 0.6625 | 0.02772 | 465 | 43422.293 | 0.66258 | 0.027264 | 465 | 43425.7539 | 0.66263 | 0.026974 |  | 465 | 16.4042 | 0.66632 | 0.0068618 |  | 465 | 0.66947 | -0.010203 | 465 | 0.67232 | -0.025497 | 465 | 43676.8398 | 465 | 44030.8789 | 465 | 44068.0781
466 | 16.4394 | 43439.3555 | 0.66284 | 0.025834 | 466 | 43408.5859 | 0.66237 | 0.028413 | 466 | 43461.8633 | 0.66319 | 0.02395 |  | 466 | 16.4394 | 0.66634 | 0.0067073 |  | 466 | 0.66942 | -0.0099261 | 466 | 0.67282 | -0.028146 | 466 | 43665.1211 | 466 | 44037.6016 | 466 | 44091.4805
467 | 16.4747 | 43460.1367 | 0.66316 | 0.024095 | 467 | 43485.8164 | 0.66355 | 0.021947 | 467 | 43461.6562 | 0.66318 | 0.023967 |  | 467 | 16.4747 | 0.66617 | 0.0076725 |  | 467 | 0.66893 | -0.0072676 | 467 | 0.67325 | -0.030469 | 467 | 43689.3203 | 467 | 44039.9219 | 467 | 44097.4414
468 | 16.51 | 43472.5859 | 0.66335 | 0.023053 | 468 | 43518.2539 | 0.66405 | 0.019238 | 468 | 43431.5508 | 0.66272 | 0.026488 |  | 468 | 16.51 | 0.66661 | 0.0052379 |  | 468 | 0.66838 | -0.0043167 | 468 | 0.67349 | -0.031714 | 468 | 43715.0391 | 468 | 44022.1211 | 468 | 44104.2812
469 | 16.5453 | 43434.2539 | 0.66276 | 0.026262 | 469 | 43480.0938 | 0.66346 | 0.022425 | 469 | 43437.6797 | 0.66282 | 0.025975 |  | 469 | 16.5453 | 0.66691 | 0.0036263 |  | 469 | 0.66889 | -0.0070769 | 469 | 0.67328 | -0.030635 | 469 | 43694.0391 | 469 | 43973.6797 | 469 | 44102.7188
470 | 16.5806 | 43475.9883 | 0.6634 | 0.022769 | 470 | 43450.6797 | 0.66301 | 0.024886 | 470 | 43463.9961 | 0.66322 | 0.023772 |  | 470 | 16.5806 | 0.6672 | 0.0020729 |  | 470 | 0.66864 | -0.0057253 | 470 | 0.67355 | -0.032058 | 470 | 43675.0781 | 470 | 43925.9219 | 470 | 44133.4805
471 | 16.6158 | 43438.4375 | 0.66283 | 0.025911 | 471 | 43448.5156 | 0.66298 | 0.025067 | 471 | 43412.75 | 0.66244 | 0.028064 |  | 471 | 16.6158 | 0.66764 | -0.00029851 |  | 471 | 0.66987 | -0.012361 | 471 | 0.67413 | -0.035132 | 471 | 43676.8789 | 471 | 43901.2812 | 471 | 44126.9219
472 | 16.6511 | 43434.0312 | 0.66276 | 0.02628 | 472 | 43440.0664 | 0.66285 | 0.025775 | 472 | 43399.8867 | 0.66224 | 0.029143 |  | 472 | 16.6511 | 0.66738 | 0.0010692 |  | 472 | 0.67061 | -0.01635 | 472 | 0.67399 | -0.034405 | 472 | 43640.5195 | 472 | 43891.1992 | 472 | 44144.7617
473 | 16.6864 | 43477.3672 | 0.66342 | 0.022653 | 473 | 43407.2383 | 0.66235 | 0.028526 | 473 | 43440.7188 | 0.66286 | 0.02572 |  | 473 | 16.6864 | 0.66724 | 0.0018435 |  | 473 | 0.6713 | -0.020045 | 473 | 0.67414 | -0.035176 | 473 | 43627 | 473 | 43874.9219 | 473 | 44182.9219
474 | 16.7217 | 43455.1953 | 0.66308 | 0.024508 | 474 | 43437.2891 | 0.66281 | 0.026007 | 474 | 43404.0586 | 0.6623 | 0.028793 |  | 474 | 16.7217 | 0.6679 | -0.0017065 |  | 474 | 0.67071 | -0.016849 | 474 | 0.67513 | -0.04045 | 474 | 43603.1992 | 474 | 43884.8789 | 474 | 44192.3203
475 | 16.7569 | 43460.875 | 0.66317 | 0.024033 | 475 | 43459.1992 | 0.66314 | 0.024173 | 475 | 43452.2969 | 0.66304 | 0.024751 |  | 475 | 16.7569 | 0.66771 | -0.00068157 |  | 475 | 0.67082 | -0.01746 | 475 | 0.6751 | -0.040321 | 475 | 43624.3203 | 475 | 43894.3203 | 475 | 44175.0391
476 | 16.7922 | 43408.5156 | 0.66237 | 0.028419 | 476 | 43476.7852 | 0.66341 | 0.022702 | 476 | 43432.5039 | 0.66274 | 0.026408 |  | 476 | 16.7922 | 0.66729 | 0.0015876 |  | 476 | 0.67036 | -0.014983 | 476 | 0.67476 | -0.038514 | 476 | 43596.3984 | 476 | 43911.6016 | 476 | 44144.1992
477 | 16.8275 | 43413.2539 | 0.66244 | 0.028022 | 477 | 43462.7305 | 0.6632 | 0.023878 | 477 | 43430.7734 | 0.66271 | 0.026553 |  | 477 | 16.8275 | 0.6669 | 0.0037016 |  | 477 | 0.67015 | -0.013887 | 477 | 0.6743 | -0.036039 | 477 | 43547.9219 | 477 | 43894.5586 | 477 | 44080.7617
478 | 16.8628 | 43448.9922 | 0.66299 | 0.025027 | 478 | 43474.1211 | 0.66337 | 0.022925 | 478 | 43408.6484 | 0.66237 | 0.028408 |  | 478 | 16.8628 | 0.66621 | 0.0074425 |  | 478 | 0.67034 | -0.014871 | 478 | 0.6742 | -0.035505 | 478 | 43547.1211 | 478 | 43923.2812 | 478 | 44017.8789
479 | 16.8981 | 43479.6484 | 0.66346 | 0.022463 | 479 | 43463.3555 | 0.66321 | 0.023825 | 479 | 43459.9453 | 0.66316 | 0.024111 |  | 479 | 16.8981 | 0.66622 | 0.0073669 |  | 479 | 0.67084 | -0.017583 | 479 | 0.67384 | -0.033597 | 479 | 43576.4805 | 479 | 43933.9219 | 479 | 44005.7617
480 | 16.9333 | 43421.5273 | 0.66257 | 0.027328 | 480 | 43475.6992 | 0.6634 | 0.022793 | 480 | 43417.1562 | 0.6625 | 0.027695 |  | 480 | 16.9333 | 0.66638 | 0.0065 |  | 480 | 0.67078 | -0.017248 | 480 | 0.67371 | -0.032896 | 480 | 43584.3203 | 480 | 43923.1992 | 480 | 44025.4414
481 | 16.9686 | 43438.6523 | 0.66283 | 0.025893 | 481 | 43426.1367 | 0.66264 | 0.026942 | 481 | 43447.5781 | 0.66297 | 0.025146 |  | 481 | 16.9686 | 0.66668 | 0.0049065 |  | 481 | 0.67106 | -0.018766 | 481 | 0.67366 | -0.032667 | 481 | 43630.0781 | 481 | 43920.6797 | 481 | 44029.5195
482 | 17.0039 | 43484.8906 | 0.66354 | 0.022024 | 482 | 43451.5625 | 0.66303 | 0.024812 | 482 | 43428.8203 | 0.66268 | 0.026717 |  | 482 | 17.0039 | 0.66688 | 0.0037958 |  | 482 | 0.67178 | -0.022626 | 482 | 0.67358 | -0.032213 | 482 | 43610.7188 | 482 | 43933 | 482 | 44068.6797
483 | 17.0392 | 43498.2422 | 0.66374 | 0.020909 | 483 | 43463.2461 | 0.66321 | 0.023834 | 483 | 43402.1992 | 0.66228 | 0.028949 |  | 483 | 17.0392 | 0.66694 | 0.0034458 |  | 483 | 0.67277 | -0.027913 | 483 | 0.67393 | -0.034106 | 483 | 43635.0781 | 483 | 43891.4414 | 483 | 44080.1602
484 | 17.0744 | 43464.1992 | 0.66322 | 0.023755 | 484 | 43454.9727 | 0.66308 | 0.024527 | 484 | 43404.625 | 0.66231 | 0.028745 |  | 484 | 17.0744 | 0.66632 | 0.0068165 |  | 484 | 0.67297 | -0.028941 | 484 | 0.67409 | -0.034929 | 484 | 43596.7617 | 484 | 43899.1602 | 484 | 44067.3203
485 | 17.1097 | 43390.3945 | 0.66209 | 0.029939 | 485 | 43439.8594 | 0.66285 | 0.025792 | 485 | 43391.9336 | 0.66212 | 0.02981 |  | 485 | 17.1097 | 0.66546 | 0.011538 |  | 485 | 0.67292 | -0.028723 | 485 | 0.67459 | -0.037601 | 485 | 43559.6016 | 485 | 43842.2812 | 485 | 44104.8008
486 | 17.145 | 43401.6211 | 0.66227 | 0.028997 | 486 | 43450.7383 | 0.66302 | 0.024881 | 486 | 43401.5117 | 0.66226 | 0.029007 |  | 486 | 17.145 | 0.66657 | 0.0054564 |  | 486 | 0.67333 | -0.030908 | 486 | 0.67482 | -0.038798 | 486 | 43578.3984 | 486 | 43857.0781 | 486 | 44123.5586
487 | 17.1803 | 43433.0742 | 0.66275 | 0.02636 | 487 | 43445.0117 | 0.66293 | 0.025361 | 487 | 43395.6641 | 0.66218 | 0.029497 |  | 487 | 17.1803 | 0.66711 | 0.0025729 |  | 487 | 0.67296 | -0.02889 | 487 | 0.67492 | -0.039361 | 487 | 43654.9609 | 487 | 43859.1211 | 487 | 44130
488 | 17.2156 | 43431.2461 | 0.66272 | 0.026514 | 488 | 43453.1523 | 0.66305 | 0.024679 | 488 | 43414.7266 | 0.66247 | 0.027898 |  | 488 | 17.2156 | 0.66822 | -0.0034843 |  | 488 | 0.6726 | -0.026998 | 488 | 0.67464 | -0.037829 | 488 | 43631.6797 | 488 | 43808.1211 | 488 | 44132.3203
489 | 17.2508 | 43421.0586 | 0.66256 | 0.027367 | 489 | 43438.7422 | 0.66283 | 0.025886 | 489 | 43400.7109 | 0.66225 | 0.029074 |  | 489 | 17.2508 | 0.66806 | -0.0025954 |  | 489 | 0.67242 | -0.026053 | 489 | 0.67414 | -0.035198 | 489 | 43666.1211 | 489 | 43785.7617 | 489 | 44188.3594
490 | 17.2861 | 43474.3477 | 0.66338 | 0.022906 | 490 | 43434.1133 | 0.66276 | 0.026273 | 490 | 43439.6953 | 0.66285 | 0.025806 |  | 490 | 17.2861 | 0.66797 | -0.0020893 |  | 490 | 0.67314 | -0.02987 | 490 | 0.67443 | -0.036724 | 490 | 43673.0391 | 490 | 43818.6406 | 490 | 44219.6016
491 | 17.3214 | 43449.2383 | 0.66299 | 0.025007 | 491 | 43476.0586 | 0.6634 | 0.022763 | 491 | 43420.293 | 0.66255 | 0.027432 |  | 491 | 17.3214 | 0.6679 | -0.0017365 |  | 491 | 0.67347 | -0.031659 | 491 | 0.67441 | -0.036647 | 491 | 43698.8398 | 491 | 43784.4805 | 491 | 44196.2383
492 | 17.3567 | 43431.7266 | 0.66273 | 0.026473 | 492 | 43449.5547 | 0.663 | 0.02498 | 492 | 43466.6914 | 0.66326 | 0.023546 |  | 492 | 17.3567 | 0.66775 | -0.00094422 |  | 492 | 0.67363 | -0.032471 | 492 | 0.67481 | -0.038772 | 492 | 43722 | 492 | 43866.6406 | 492 | 44156.3203
493 | 17.3919 | 43377.3633 | 0.6619 | 0.031033 | 493 | 43460.7461 | 0.66317 | 0.024044 | 493 | 43425.75 | 0.66263 | 0.026974 |  | 493 | 17.3919 | 0.66723 | 0.001911 |  | 493 | 0.67381 | -0.033442 | 493 | 0.67468 | -0.038043 | 493 | 43719.1602 | 493 | 43949.1211 | 493 | 44148.9219
494 | 17.4272 | 43510.2344 | 0.66392 | 0.019907 | 494 | 43433.7617 | 0.66276 | 0.026303 | 494 | 43461.3945 | 0.66318 | 0.023989 |  | 494 | 17.4272 | 0.66709 | 0.002667 |  | 494 | 0.673 | -0.029115 | 494 | 0.67402 | -0.034557 | 494 | 43712.6406 | 494 | 43960.3984 | 494 | 44140.1211
495 | 17.4625 | 43474.9258 | 0.66338 | 0.022857 | 495 | 43483.1953 | 0.66351 | 0.022166 | 495 | 43512.7852 | 0.66396 | 0.019694 |  | 495 | 17.4625 | 0.66731 | 0.00149 |  | 495 | 0.67227 | -0.025223 | 495 | 0.67402 | -0.034579 | 495 | 43705.1602 | 495 | 43944.5195 | 495 | 44116.8789
496 | 17.4978 | 43429.3398 | 0.66269 | 0.026673 | 496 | 43432.4648 | 0.66274 | 0.026412 | 496 | 43469.0352 | 0.66329 | 0.02335 |  | 496 | 17.4978 | 0.66815 | -0.0030718 |  | 496 | 0.6725 | -0.026453 | 496 | 0.67432 | -0.036168 | 496 | 43720.4414 | 496 | 43946.9609 | 496 | 44107.3594
497 | 17.5331 | 43455.2578 | 0.66308 | 0.024503 | 497 | 43432.4023 | 0.66274 | 0.026417 | 497 | 43451.9961 | 0.66303 | 0.024776 |  | 497 | 17.5331 | 0.66818 | -0.0032256 |  | 497 | 0.67233 | -0.025564 | 497 | 0.67422 | -0.035622 | 497 | 43724.8398 | 497 | 43901.7617 | 497 | 44107.2812
498 | 17.5683 | 43470.8047 | 0.66332 | 0.023202 | 498 | 43443.6016 | 0.66291 | 0.025479 | 498 | 43414.1328 | 0.66246 | 0.027948 |  | 498 | 17.5683 | 0.66871 | -0.0061221 |  | 498 | 0.67191 | -0.023328 | 498 | 0.67455 | -0.037402 | 498 | 43703.1211 | 498 | 43892.9609 | 498 | 44118.6797
499 | 17.6036 | 43411 | 0.66241 | 0.028211 | 499 | 43449.8945 | 0.663 | 0.024952 | 499 | 43357.7148 | 0.6616 | 0.032684 |  | 499 | 17.6036 | 0.66915 | -0.0084795 |  | 499 | 0.67181 | -0.022753 | 499 | 0.67479 | -0.038654 | 499 | 43674.3594 | 499 | 43898.8398 | 499 | 44127.7188
500 | 17.6389 | 43453.2305 | 0.66305 | 0.024673 | 500 | 43437.5586 | 0.66281 | 0.025985 | 500 | 43378.6602 | 0.66192 | 0.030925 |  | 500 | 17.6389 | 0.66972 | -0.011562 |  | 500 | 0.672 | -0.023788 | 500 | 0.67515 | -0.040534 | 500 | 43653.4414 | 500 | 43908.8789 | 500 | 44147
501 | 17.6742 | 43513.4844 | 0.66397 | 0.019636 | 501 | 43473.8555 | 0.66337 | 0.022947 | 501 | 43421.875 | 0.66258 | 0.027299 |  | 501 | 17.6742 | 0.66995 | -0.012813 |  | 501 | 0.67224 | -0.025064 | 501 | 0.67571 | -0.043532 | 501 | 43634.8008 | 501 | 43938.1211 | 501 | 44159.6016
502 | 17.7094 | 43483.7578 | 0.66352 | 0.022119 | 502 | 43480.8203 | 0.66347 | 0.022365 | 502 | 43478.4961 | 0.66344 | 0.022559 |  | 502 | 17.7094 | 0.67008 | -0.013507 |  | 502 | 0.67193 | -0.023406 | 502 | 0.67522 | -0.040942 | 502 | 43648.3984 | 502 | 43971.2383 | 502 | 44154.1992
503 | 17.7447 | 43432.3477 | 0.66274 | 0.026421 | 503 | 43483.1328 | 0.66351 | 0.022171 | 503 | 43424.4219 | 0.66261 | 0.027085 |  | 503 | 17.7447 | 0.66984 | -0.012201 |  | 503 | 0.67189 | -0.023198 | 503 | 0.67533 | -0.041538 | 503 | 43662.9609 | 503 | 44019.8008 | 503 | 44171.1211
504 | 17.78 | 43406.7656 | 0.66234 | 0.028566 | 504 | 43450.9414 | 0.66302 | 0.024864 | 504 | 43421.6289 | 0.66257 | 0.02732 |  | 504 | 17.78 | 0.66977 | -0.011838 |  | 504 | 0.67181 | -0.022782 | 504 | 0.67543 | -0.042063 | 504 | 43687.7188 | 504 | 44064.5586 | 504 | 44138.1602
505 | 17.8153 | 43436.1914 | 0.66279 | 0.026099 | 505 | 43437.5742 | 0.66281 | 0.025984 | 505 | 43494.1094 | 0.66368 | 0.021254 |  | 505 | 17.8153 | 0.66979 | -0.011928 |  | 505 | 0.67167 | -0.022044 | 505 | 0.67558 | -0.042864 | 505 | 43678.7188 | 505 | 44082.6797 | 505 | 44141.9219
506 | 17.8506 | 43465.6719 | 0.66324 | 0.023631 | 506 | 43460.1133 | 0.66316 | 0.024097 | 506 | 43491.7305 | 0.66364 | 0.021453 |  | 506 | 17.8506 | 0.66981 | -0.01204 |  | 506 | 0.67148 | -0.021019 | 506 | 0.67558 | -0.042838 | 506 | 43673.9609 | 506 | 44088.0391 | 506 | 44175.0781
507 | 17.8858 | 43514.0352 | 0.66398 | 0.01959 | 507 | 43468.5312 | 0.66329 | 0.023392 | 507 | 43505.4805 | 0.66385 | 0.020304 |  | 507 | 17.8858 | 0.66904 | -0.0078662 |  | 507 | 0.67178 | -0.022589 | 507 | 0.67552 | -0.042537 | 507 | 43622.8398 | 507 | 44101.4805 | 507 | 44151.9609
508 | 17.9211 | 43492.3281 | 0.66365 | 0.021403 | 508 | 43496.8281 | 0.66372 | 0.021027 | 508 | 43474.4414 | 0.66338 | 0.022898 |  | 508 | 17.9211 | 0.66906 | -0.0080196 |  | 508 | 0.67258 | -0.026887 | 508 | 0.67542 | -0.042001 | 508 | 43677.6406 | 508 | 44087.7617 | 508 | 44143.3984
509 | 17.9564 | 43444.2539 | 0.66292 | 0.025424 | 509 | 43515.8945 | 0.66401 | 0.019434 | 509 | 43427.2266 | 0.66266 | 0.02685 |  | 509 | 17.9564 | 0.66821 | -0.0034057 |  | 509 | 0.67342 | -0.031344 | 509 | 0.67498 | -0.03964 | 509 | 43726.4414 | 509 | 44088.2812 | 509 | 44129.0391
510 | 17.9917 | 43425.293 | 0.66263 | 0.027012 | 510 | 43475.8594 | 0.6634 | 0.022779 | 510 | 43403.6367 | 0.6623 | 0.028828 |  | 510 | 17.9917 | 0.66786 | -0.0015336 |  | 510 | 0.67361 | -0.03239 | 510 | 0.6754 | -0.041887 | 510 | 43757.3984 | 510 | 44075.6016 | 510 | 44123.7617
511 | 18.0269 | 43435.5703 | 0.66278 | 0.026151 | 511 | 43445.4023 | 0.66293 | 0.025328 | 511 | 43365.6992 | 0.66172 | 0.032013 |  | 511 | 18.0269 | 0.6673 | 0.0015201 |  | 511 | 0.67328 | -0.030609 | 511 | 0.67563 | -0.043099 | 511 | 43749.6406 | 511 | 44067.6797 | 511 | 44180.9219
512 | 18.0622 | 43394.1719 | 0.66215 | 0.029622 | 512 | 43451.2617 | 0.66302 | 0.024837 | 512 | 43343.8281 | 0.66138 | 0.033852 |  | 512 | 18.0622 | 0.66783 | -0.0013649 |  | 512 | 0.67259 | -0.026931 | 512 | 0.67517 | -0.040656 | 512 | 43741.7188 | 512 | 44113.8008 | 512 | 44172.8789
513 | 18.0975 | 43427.9883 | 0.66267 | 0.026787 | 513 | 43418.1133 | 0.66252 | 0.027614 | 513 | 43374.0039 | 0.66184 | 0.031316 |  | 513 | 18.0975 | 0.66732 | 0.0014112 |  | 513 | 0.67243 | -0.026098 | 513 | 0.67509 | -0.040218 | 513 | 43737.9219 | 513 | 44136.8398 | 513 | 44184.1992
514 | 18.1328 | 43445.3047 | 0.66293 | 0.025336 | 514 | 43410.3828 | 0.6624 | 0.028262 | 514 | 43396.3828 | 0.66219 | 0.029437 |  | 514 | 18.1328 | 0.66754 | 0.0002425 |  | 514 | 0.67204 | -0.024022 | 514 | 0.67448 | -0.036979 | 514 | 43714.8008 | 514 | 44136.3203 | 514 | 44201.0391
515 | 18.1681 | 43460.1055 | 0.66316 | 0.024097 | 515 | 43449.1094 | 0.66299 | 0.025018 | 515 | 43421.7773 | 0.66257 | 0.027307 |  | 515 | 18.1681 | 0.66747 | 0.00058797 |  | 515 | 0.67173 | -0.022348 | 515 | 0.67405 | -0.034737 | 515 | 43689.6797 | 515 | 44087.9219 | 515 | 44194.5586
516 | 18.2033 | 43450.0352 | 0.66301 | 0.02494 | 516 | 43450.5781 | 0.66301 | 0.024895 | 516 | 43393.8125 | 0.66215 | 0.029653 |  | 516 | 18.2033 | 0.66727 | 0.0016593 |  | 516 | 0.67145 | -0.020833 | 516 | 0.67384 | -0.033608 | 516 | 43684.2812 | 516 | 44070.3984 | 516 | 44241.7188
517 | 18.2386 | 43465.7188 | 0.66324 | 0.023628 | 517 | 43400.8633 | 0.66225 | 0.029061 | 517 | 43437 | 0.66281 | 0.026032 |  | 517 | 18.2386 | 0.6671 | 0.0025994 |  | 517 | 0.67156 | -0.021442 | 517 | 0.67393 | -0.034088 | 517 | 43705.4414 | 517 | 44078.6406 | 517 | 44246.3203
518 | 18.2739 | 43467.7695 | 0.66328 | 0.023456 | 518 | 43459.2266 | 0.66315 | 0.024171 | 518 | 43421.332 | 0.66257 | 0.027344 |  | 518 | 18.2739 | 0.6671 | 0.0026256 |  | 518 | 0.67121 | -0.019536 | 518 | 0.67323 | -0.03035 | 518 | 43747.4414 | 518 | 44052.6016 | 518 | 44254.8789
519 | 18.3092 | 43466.9258 | 0.66326 | 0.023527 | 519 | 43442.7031 | 0.66289 | 0.025554 | 519 | 43431.4062 | 0.66272 | 0.0265 |  | 519 | 18.3092 | 0.66726 | 0.0017269 |  | 519 | 0.67126 | -0.019826 | 519 | 0.67272 | -0.027624 | 519 | 43781.6016 | 519 | 44051.8398 | 519 | 44255
520 | 18.3444 | 43445.8828 | 0.66294 | 0.025288 | 520 | 43433.1094 | 0.66275 | 0.026358 | 520 | 43427.3359 | 0.66266 | 0.026841 |  | 520 | 18.3444 | 0.66728 | 0.001648 |  | 520 | 0.67141 | -0.020624 | 520 | 0.67257 | -0.026824 | 520 | 43810.9219 | 520 | 44034.6406 | 520 | 44235.6406
521 | 18.3797 | 43463.3047 | 0.66321 | 0.02383 | 521 | 43441.3125 | 0.66287 | 0.02567 | 521 | 43425.75 | 0.66263 | 0.026974 |  | 521 | 18.3797 | 0.66646 | 0.0061042 |  | 521 | 0.67142 | -0.020684 | 521 | 0.67267 | -0.027376 | 521 | 43853.9219 | 521 | 44051.3203 | 521 | 44237.9219
522 | 18.415 | 43429.793 | 0.6627 | 0.026635 | 522 | 43487.2617 | 0.66357 | 0.021826 | 522 | 43438.7617 | 0.66283 | 0.025884 |  | 522 | 18.415 | 0.66559 | 0.010805 |  | 522 | 0.67182 | -0.022812 | 522 | 0.67126 | -0.019837 | 522 | 43891.3203 | 522 | 44046.0391 | 522 | 44241.6016
523 | 18.4503 | 43540.5195 | 0.66439 | 0.01738 | 523 | 43458.8828 | 0.66314 | 0.0242 | 523 | 43464.5508 | 0.66323 | 0.023725 |  | 523 | 18.4503 | 0.66551 | 0.01124 |  | 523 | 0.67225 | -0.025127 | 523 | 0.67192 | -0.023372 | 523 | 43907.0391 | 523 | 44049.5195 | 523 | 44224.7617
524 | 18.4856 | 43518.7695 | 0.66405 | 0.019195 | 524 | 43503.8906 | 0.66383 | 0.020437 | 524 | 43462.7188 | 0.6632 | 0.023879 |  | 524 | 18.4856 | 0.66587 | 0.0092684 |  | 524 | 0.67214 | -0.024541 | 524 | 0.67279 | -0.028005 | 524 | 43895.8008 | 524 | 44049.8789 | 524 | 44195.1992
525 | 18.5208 | 43472.2148 | 0.66334 | 0.023084 | 525 | 43483.9258 | 0.66352 | 0.022105 | 525 | 43406.1133 | 0.66233 | 0.028621 |  | 525 | 18.5208 | 0.6663 | 0.0069522 |  | 525 | 0.6729 | -0.028579 | 525 | 0.6736 | -0.032327 | 525 | 43892.7188 | 525 | 44060 | 525 | 44186.3594
526 | 18.5561 | 43509.7422 | 0.66392 | 0.019948 | 526 | 43436.293 | 0.6628 | 0.026091 | 526 | 43457.9531 | 0.66313 | 0.024277 |  | 526 | 18.5561 | 0.66675 | 0.0045223 |  | 526 | 0.67286 | -0.028375 | 526 | 0.67364 | -0.032523 | 526 | 43867.6016 | 526 | 44048.8398 | 526 | 44205.3594
527 | 18.5914 | 43493.125 | 0.66366 | 0.021336 | 527 | 43476.5508 | 0.66341 | 0.022722 | 527 | 43481.8203 | 0.66349 | 0.022281 |  | 527 | 18.5914 | 0.6675 | 0.00042653 |  | 527 | 0.67292 | -0.028686 | 527 | 0.67363 | -0.032464 | 527 | 43879.1602 | 527 | 44039.8008 | 527 | 44200.5195
528 | 18.6267 | 43488.8398 | 0.6636 | 0.021694 | 528 | 43445.3281 | 0.66293 | 0.025334 | 528 | 43341.1953 | 0.66134 | 0.034073 |  | 528 | 18.6267 | 0.66796 | -0.0020328 |  | 528 | 0.67339 | -0.031204 | 528 | 0.67372 | -0.032944 | 528 | 43868.5195 | 528 | 44055.7188 | 528 | 44206.6797
529 | 18.6619 | 43483.4648 | 0.66352 | 0.022144 | 529 | 43435.1602 | 0.66278 | 0.026186 | 529 | 43255.8555 | 0.66004 | 0.041265 |  | 529 | 18.6619 | 0.6684 | -0.0044066 |  | 529 | 0.67346 | -0.031588 | 529 | 0.67375 | -0.033125 | 529 | 43814.2383 | 529 | 44081.8008 | 529 | 44180.9219
530 | 18.6972 | 43435.75 | 0.66279 | 0.026136 | 530 | 43462.8242 | 0.6632 | 0.02387 | 530 | 43450.5156 | 0.66301 | 0.0249 |  | 530 | 18.6972 | 0.66834 | -0.0041215 |  | 530 | 0.67399 | -0.034398 | 530 | 0.67348 | -0.031707 | 530 | 43817.5586 | 530 | 44122.5195 | 530 | 44144.8789
531 | 18.7325 | 43428.4141 | 0.66268 | 0.026751 | 531 | 43418.9453 | 0.66253 | 0.027545 | 531 | 43464.7695 | 0.66323 | 0.023707 |  | 531 | 18.7325 | 0.66757 | 5.8166e-05 |  | 531 | 0.67401 | -0.034498 | 531 | 0.67323 | -0.030361 | 531 | 43761.6016 | 531 | 44149.3984 | 531 | 44148.1211
532 | 18.7678 | 43450.3945 | 0.66301 | 0.02491 | 532 | 43408.8633 | 0.66238 | 0.02839 | 532 | 43462.75 | 0.6632 | 0.023876 |  | 532 | 18.7678 | 0.66784 | -0.0013875 |  | 532 | 0.67367 | -0.032671 | 532 | 0.67315 | -0.029903 | 532 | 43732.6406 | 532 | 44147.5586 | 532 | 44130.8008
[14 empty] | 533 | 18.8031 | 0.66724 | 0.0018396 |  | 533 | 0.67371 | -0.032896 | 533 | 0.67309 | -0.029626 | 533 | 43717.7617 | 533 | 44108.8789 | 533 | 44061.7617
[14 empty] | 534 | 18.8383 | 0.66705 | 0.0028777 |  | 534 | 0.67354 | -0.031995 | 534 | 0.67294 | -0.028804 | 534 | 43717.1992 | 534 | 44098.2812 | 534 | 44054
[14 empty] | 535 | 18.8736 | 0.66726 | 0.0017117 |  | 535 | 0.67356 | -0.032113 | 535 | 0.67249 | -0.026424 | 535 | 43671.7617 | 535 | 44051 | 535 | 44049.7188
[14 empty] | 536 | 18.9089 | 0.66694 | 0.0034798 |  | 536 | 0.67319 | -0.030147 | 536 | 0.67294 | -0.028823 | 536 | 43698.8008 | 536 | 44036.8789 | 536 | 44029.4414
[14 empty] | 537 | 18.9442 | 0.66657 | 0.0054826 |  | 537 | 0.67231 | -0.025449 | 537 | 0.67276 | -0.027842 | 537 | 43708.5586 | 537 | 44003.8398 | 537 | 43985.9609
[14 empty] | 538 | 18.9794 | 0.66586 | 0.0093628 |  | 538 | 0.67206 | -0.024133 | 538 | 0.67234 | -0.025616 | 538 | 43672.7617 | 538 | 44003.7617 | 538 | 44008.1992
[14 empty] | 539 | 19.0147 | 0.66605 | 0.0083098 |  | 539 | 0.67147 | -0.020941 | 539 | 0.67271 | -0.027601 | 539 | 43653.4414 | 539 | 43963.4414 | 539 | 44039.1602
[14 empty] | 540 | 19.05 | 0.66595 | 0.0088344 |  | 540 | 0.67166 | -0.021951 | 540 | 0.67235 | -0.025664 | 540 | 43674.1211 | 540 | 43937.9609 | 540 | 44097.3203
[14 empty] | 541 | 19.0853 | 0.66609 | 0.0080988 |  | 541 | 0.67177 | -0.022575 | 541 | 0.67199 | -0.023755 | 541 | 43704.7188 | 541 | 43985.6406 | 541 | 44122.9609
[14 empty] | 542 | 19.1206 | 0.66744 | 0.000742 |  | 542 | 0.67242 | -0.026005 | 542 | 0.67284 | -0.028257 | 542 | 43695 | 542 | 44000.2812 | 542 | 44114.4805
[14 empty] | 543 | 19.1558 | 0.66779 | -0.0011433 |  | 543 | 0.67309 | -0.029592 | 543 | 0.67315 | -0.029947 | 543 | 43682.3984 | 543 | 44043.3984 | 543 | 44108.5195
[14 empty] | 544 | 19.1911 | 0.66794 | -0.001939 |  | 544 | 0.67358 | -0.032217 | 544 | 0.67308 | -0.029555 | 544 | 43647.1602 | 544 | 44041.8398 | 544 | 44100.2812
[14 empty] | 545 | 19.2264 | 0.66806 | -0.0025845 |  | 545 | 0.67427 | -0.035914 | 545 | 0.67328 | -0.030635 | 545 | 43627.3203 | 545 | 44033.2383 | 545 | 44098.8008
[14 empty] | 546 | 19.2617 | 0.6684 | -0.0044591 |  | 546 | 0.67464 | -0.037859 | 546 | 0.67348 | -0.031662 | 546 | 43652.8789 | 546 | 44104.3203 | 546 | 44098.3594
[14 empty] | 547 | 19.2969 | 0.6689 | -0.007118 |  | 547 | 0.67541 | -0.041935 | 547 | 0.67259 | -0.026935 | 547 | 43687.1211 | 547 | 44124 | 547 | 44091.2812
[14 empty] | 548 | 19.3322 | 0.66931 | -0.0093506 |  | 548 | 0.67579 | -0.043925 | 548 | 0.67282 | -0.028186 | 548 | 43711.5195 | 548 | 44120.8008 | 548 | 44096.0781
[14 empty] | 549 | 19.3675 | 0.66945 | -0.010094 |  | 549 | 0.67527 | -0.041196 | 549 | 0.67287 | -0.028431 | 549 | 43709.0781 | 549 | 44133.0391 | 549 | 44073.4805
[14 empty] | 550 | 19.4028 | 0.66941 | -0.0098963 |  | 550 | 0.67553 | -0.042581 | 550 | 0.67269 | -0.027479 | 550 | 43752.8789 | 550 | 44122.7617 | 550 | 44044.3594
[14 empty] | 551 | 19.4381 | 0.66976 | -0.011749 |  | 551 | 0.6762 | -0.046101 | 551 | 0.67256 | -0.026805 | 551 | 43776.8789 | 551 | 44146.9219 | 551 | 44056.3984
[14 empty] | 552 | 19.4733 | 0.66989 | -0.012443 |  | 552 | 0.67745 | -0.0527 | 552 | 0.67265 | -0.027283 | 552 | 43792.6797 | 552 | 44172.0781 | 552 | 44039.9609
[14 empty] | 553 | 19.5086 | 0.66998 | -0.012965 |  | 553 | 0.67769 | -0.053971 | 553 | 0.67265 | -0.027257 | 553 | 43760.5195 | 553 | 44176.2812 | 553 | 44005.8008
[14 empty] | 554 | 19.5439 | 0.67058 | -0.016157 |  | 554 | 0.67766 | -0.053843 | 554 | 0.67279 | -0.027983 | 554 | 43728.0781 | 554 | 44175 | 554 | 44038.4805
[14 empty] | 555 | 19.5792 | 0.67097 | -0.018267 |  | 555 | 0.67798 | -0.055482 | 555 | 0.67268 | -0.027394 | 555 | 43740.2383 | 555 | 44168.3203 | 555 | 44051.2383
[14 empty] | 556 | 19.6144 | 0.67082 | -0.017456 |  | 556 | 0.67775 | -0.054285 | 556 | 0.67314 | -0.029884 | 556 | 43711.0781 | 556 | 44157.4805 | 556 | 44029.5586
[14 empty] | 557 | 19.6497 | 0.6705 | -0.015721 |  | 557 | 0.67752 | -0.053061 | 557 | 0.67344 | -0.031463 | 557 | 43699.9609 | 557 | 44137.4414 | 557 | 44039.2812
[14 empty] | 558 | 19.685 | 0.67031 | -0.014741 |  | 558 | 0.67761 | -0.053551 | 558 | 0.67346 | -0.031563 | 558 | 43673.9609 | 558 | 44106.7188 | 558 | 44052.7188
[14 empty] | 559 | 19.7203 | 0.67074 | -0.017028 |  | 559 | 0.67739 | -0.0524 | 559 | 0.67419 | -0.035486 | 559 | 43672.8789 | 559 | 44133.5586 | 559 | 44065.6016
[14 empty] | 560 | 19.7556 | 0.67089 | -0.017839 |  | 560 | 0.67708 | -0.05074 | 560 | 0.67488 | -0.03914 | 560 | 43642.0781 | 560 | 44097.2812 | 560 | 44098.0391
[14 empty] | 561 | 19.7908 | 0.6705 | -0.015743 |  | 561 | 0.6766 | -0.048233 | 561 | 0.67502 | -0.039887 | 561 | 43615.5195 | 561 | 44068.8008 | 561 | 44076.2383
[14 empty] | 562 | 19.8261 | 0.67025 | -0.014394 |  | 562 | 0.67655 | -0.047988 | 562 | 0.67474 | -0.038374 | 562 | 43641.3984 | 562 | 44049.2812 | 562 | 44058.6016
[14 empty] | 563 | 19.8614 | 0.66987 | -0.012335 |  | 563 | 0.67615 | -0.045859 | 563 | 0.67408 | -0.034881 | 563 | 43644.6797 | 563 | 44109.9219 | 563 | 44049.9609
[14 empty] | 564 | 19.8967 | 0.66884 | -0.0067848 |  | 564 | 0.67627 | -0.046471 | 564 | 0.67422 | -0.035648 | 564 | 43716.1992 | 564 | 44136.8789 | 564 | 44035.1602
[14 empty] | 565 | 19.9319 | 0.66891 | -0.0071926 |  | 565 | 0.67621 | -0.046189 | 565 | 0.67396 | -0.034243 | 565 | 43739.1211 | 565 | 44164.8008 | 565 | 44032.3594
[14 empty] | 566 | 19.9672 | 0.66888 | -0.0070466 |  | 566 | 0.67655 | -0.047955 | 566 | 0.67368 | -0.032734 | 566 | 43752 | 566 | 44193.0781 | 566 | 44043.3984
[14 empty] | 567 | 20.0025 | 0.66849 | -0.0049199 |  | 567 | 0.67579 | -0.043943 | 567 | 0.67349 | -0.031762 | 567 | 43756.4805 | 567 | 44228.1992 | 567 | 44047.0781
[14 empty] | 568 | 20.0378 | 0.66739 | 0.0010465 |  | 568 | 0.67612 | -0.045694 | 568 | 0.67423 | -0.035674 | 568 | 43789.0781 | 568 | 44264.1211 | 568 | 44053
[14 empty] | 569 | 20.0731 | 0.66649 | 0.0059009 |  | 569 | 0.67602 | -0.04518 | 569 | 0.67427 | -0.035902 | 569 | 43848.7617 | 569 | 44318.6016 | 569 | 44058.8789
[14 empty] | 570 | 20.1083 | 0.66713 | 0.0024452 |  | 570 | 0.67574 | -0.043705 | 570 | 0.67422 | -0.035648 | 570 | 43867.6797 | 570 | 44325.8789 | 570 | 44064.8398
[14 empty] | 571 | 20.1436 | 0.66787 | -0.00156 |  | 571 | 0.67565 | -0.043191 | 571 | 0.67464 | -0.037874 | 571 | 43864.5586 | 571 | 44329.9219 | 571 | 44070.3203
[14 empty] | 572 | 20.1789 | 0.66805 | -0.0025468 |  | 572 | 0.67588 | -0.044435 | 572 | 0.67488 | -0.039122 | 572 | 43874.4805 | 572 | 44350.6016 | 572 | 44096.1211
[14 empty] | 573 | 20.2142 | 0.66817 | -0.0032033 |  | 573 | 0.67743 | -0.052601 | 573 | 0.67424 | -0.035737 | 573 | 43872.7188 | 573 | 44388.4414 | 573 | 44115.8789
[14 empty] | 574 | 20.2494 | 0.66868 | -0.0059573 |  | 574 | 0.67704 | -0.050554 | 574 | 0.67263 | -0.027135 | 574 | 43909.8789 | 574 | 44426.0391 | 574 | 44140.5586
[14 empty] | 575 | 20.2847 | 0.66921 | -0.0087787 |  | 575 | 0.67663 | -0.048402 | 575 | 0.67182 | -0.022808 | 575 | 43908.1992 | 575 | 44434.3203 | 575 | 44182.1211
[14 empty] | 576 | 20.32 | 0.66984 | -0.012178 |  | 576 | 0.67749 | -0.052912 | 576 | 0.67158 | -0.021557 | 576 | 43928.6016 | 576 | 44441.1992 | 576 | 44197.7188
[14 empty] | 577 | 20.3553 | 0.66998 | -0.012965 |  | 577 | 0.67785 | -0.054821 | 577 | 0.671 | -0.018427 | 577 | 43939.4414 | 577 | 44437.6016 | 577 | 44192.9609
[14 empty] | 578 | 20.3906 | 0.66986 | -0.012324 |  | 578 | 0.67834 | -0.057382 | 578 | 0.67121 | -0.019558 | 578 | 43947.7617 | 578 | 44440.5195 | 578 | 44155.8008
[14 empty] | 579 | 20.4258 | 0.66959 | -0.010838 |  | 579 | 0.67861 | -0.058795 | 579 | 0.67107 | -0.018814 | 579 | 43915.8789 | 579 | 44435.7188 | 579 | 44144.9609
[14 empty] | 580 | 20.4611 | 0.66974 | -0.011678 |  | 580 | 0.67851 | -0.058303 | 580 | 0.67118 | -0.019376 | 580 | 43881.3203 | 580 | 44462.5586 | 580 | 44104.4414
[14 empty] | 581 | 20.4964 | 0.66956 | -0.010666 |  | 581 | 0.67814 | -0.056372 | 581 | 0.67108 | -0.018851 | 581 | 43893.2383 | 581 | 44421.7617 | 581 | 44069.8398
[14 empty] | 582 | 20.5317 | 0.66845 | -0.0047025 |  | 582 | 0.67736 | -0.052261 | 582 | 0.67104 | -0.018654 | 582 | 43913.5195 | 582 | 44399.9219 | 582 | 44077.9609
[14 empty] | 583 | 20.5669 | 0.66815 | -0.0030796 |  | 583 | 0.67746 | -0.052762 | 583 | 0.6714 | -0.020591 | 583 | 43921.7617 | 583 | 44370.2812 | 583 | 44099
[14 empty] | 584 | 20.6022 | 0.66805 | -0.0025168 |  | 584 | 0.67725 | -0.051669 | 584 | 0.67179 | -0.022649 | 584 | 43868.6016 | 584 | 44342.1992 | 584 | 44091.2812
[14 empty] | 585 | 20.6375 | 0.66707 | 0.0027837 |  | 585 | 0.67743 | -0.05259 | 585 | 0.67182 | -0.022805 | 585 | 43882.1211 | 585 | 44327.4805 | 585 | 44131.6797
[14 empty] | 586 | 20.6728 | 0.6679 | -0.0017026 |  | 586 | 0.67758 | -0.053387 | 586 | 0.67183 | -0.022897 | 586 | 43792.0781 | 586 | 44325.0781 | 586 | 44169.3203
[14 empty] | 587 | 20.7081 | 0.66823 | -0.003522 |  | 587 | 0.6773 | -0.051911 | 587 | 0.67179 | -0.022656 | 587 | 43779.6797 | 587 | 44309 | 587 | 44208.9219
[14 empty] | 588 | 20.7433 | 0.66891 | -0.0071701 |  | 588 | 0.67758 | -0.05339 | 588 | 0.67187 | -0.023098 | 588 | 43813.4414 | 588 | 44295.7617 | 588 | 44124.4414
[14 empty] | 589 | 20.7786 | 0.66867 | -0.0059087 |  | 589 | 0.67687 | -0.049661 | 589 | 0.6716 | -0.021635 | 589 | 43775.3594 | 589 | 44288.6797 | 589 | 44049.9609
[14 empty] | 590 | 20.8139 | 0.6688 | -0.006605 |  | 590 | 0.67734 | -0.052126 | 590 | 0.67222 | -0.024938 | 590 | 43710.4805 | 590 | 44241 | 590 | 43992.1602
[14 empty] | 591 | 20.8492 | 0.66907 | -0.0080234 |  | 591 | 0.67734 | -0.052133 | 591 | 0.67251 | -0.026487 | 591 | 43658.2383 | 591 | 44268.7188 | 591 | 43956.7617
[14 empty] | 592 | 20.8844 | 0.66949 | -0.010289 |  | 592 | 0.67788 | -0.054975 | 592 | 0.67228 | -0.025271 | 592 | 43684.8789 | 592 | 44265.9219 | 592 | 43918.9219
[14 empty] | 593 | 20.9197 | 0.66945 | -0.010124 |  | 593 | 0.67808 | -0.056051 | 593 | 0.67151 | -0.021137 | 593 | 43750.6406 | 593 | 44294.8008 | 593 | 43931.8789
[14 empty] | 594 | 20.955 | 0.66919 | -0.0087189 |  | 594 | 0.67781 | -0.054613 | 594 | 0.67176 | -0.022511 | 594 | 43762.2383 | 594 | 44385.5195 | 594 | 43923.6406
[14 empty] | 595 | 20.9903 | 0.66884 | -0.0068298 |  | 595 | 0.67734 | -0.052126 | 595 | 0.67162 | -0.021758 | 595 | 43766.1602 | 595 | 44389.1602 | 595 | 43921.9219
[14 empty] | 596 | 21.0256 | 0.66866 | -0.0058302 |  | 596 | 0.67806 | -0.055923 | 596 | 0.67164 | -0.021869 | 596 | 43803.2383 | 596 | 44361.1992 | 596 | 43933.2383
[14 empty] | 597 | 21.0608 | 0.66857 | -0.005332 |  | 597 | 0.6783 | -0.057203 | 597 | 0.67191 | -0.023306 | 597 | 43852.8398 | 597 | 44381.6016 | 597 | 43940.6797
[14 empty] | 598 | 21.0961 | 0.66749 | 0.00047528 |  | 598 | 0.67862 | -0.058843 | 598 | 0.67189 | -0.023209 | 598 | 43899.9609 | 598 | 44391.2383 | 598 | 43970.8398
[14 empty] | 599 | 21.1314 | 0.66706 | 0.0028137 |  | 599 | 0.67853 | -0.058391 | 599 | 0.67245 | -0.026213 | 599 | 43841.3594 | 599 | 44448.5195 | 599 | 43973.6016
[14 empty] | 600 | 21.1667 | 0.66748 | 0.00055407 |  | 600 | 0.67793 | -0.055252 | 600 | 0.67264 | -0.027227 | 600 | 43876.1602 | 600 | 44500.3984 | 600 | 43999
[14 empty] | 601 | 21.2019 | 0.66837 | -0.0042604 |  | 601 | 0.67821 | -0.056722 | 601 | 0.67235 | -0.025675 | 601 | 43872.0781 | 601 | 44505.3594 | 601 | 43992.4414
[14 empty] | 602 | 21.2372 | 0.66937 | -0.009676 |  | 602 | 0.67857 | -0.058602 | 602 | 0.67216 | -0.024615 | 602 | 43872.0391 | 602 | 44505.0391 | 602 | 43982.1992
[14 empty] | 603 | 21.2725 | 0.67012 | -0.013678 |  | 603 | 0.67792 | -0.055205 | 603 | 0.67214 | -0.02453 | 603 | 43857.6797 | 603 | 44442.9219 | 603 | 44004.8789
[14 empty] | 604 | 21.3078 | 0.67055 | -0.01603 |  | 604 | 0.67724 | -0.051614 | 604 | 0.67127 | -0.019889 | 604 | 43790.3203 | 604 | 44368.2383 | 604 | 43973.8398
[14 empty] | 605 | 21.3431 | 0.67013 | -0.013745 |  | 605 | 0.67733 | -0.052079 | 605 | 0.67079 | -0.017285 | 605 | 43740.8008 | 605 | 44373.7188 | 605 | 44016.4414
[14 empty] | 606 | 21.3783 | 0.66985 | -0.012227 |  | 606 | 0.67727 | -0.051786 | 606 | 0.6707 | -0.016823 | 606 | 43728.6406 | 606 | 44407.4414 | 606 | 44025.9219
[14 empty] | 607 | 21.4136 | 0.67028 | -0.014573 |  | 607 | 0.67731 | -0.051983 | 607 | 0.67062 | -0.016406 | 607 | 43671.2383 | 607 | 44408.6797 | 607 | 44019.7617
[14 empty] | 608 | 21.4489 | 0.66871 | -0.0061034 |  | 608 | 0.67774 | -0.054267 | 608 | 0.67093 | -0.018078 | 608 | 43730.0781 | 608 | 44378.9609 | 608 | 43979.8398
[14 empty] | 609 | 21.4842 | 0.66895 | -0.0074171 |  | 609 | 0.67805 | -0.055894 | 609 | 0.6722 | -0.024838 | 609 | 43795.7188 | 609 | 44406.8008 | 609 | 43995.0391
[14 empty] | 610 | 21.5194 | 0.66942 | -0.0099114 |  | 610 | 0.67794 | -0.05531 | 610 | 0.67179 | -0.022675 | 610 | 43818.4414 | 610 | 44382.9219 | 610 | 43985.4414
[14 empty] | 611 | 21.5547 | 0.66891 | -0.0071852 |  | 611 | 0.67836 | -0.057487 | 611 | 0.67188 | -0.023142 | 611 | 43802.6797 | 611 | 44365.3203 | 611 | 43979.1992
[14 empty] | 612 | 21.59 | 0.66853 | -0.0051221 |  | 612 | 0.67964 | -0.064218 | 612 | 0.67104 | -0.018651 | 612 | 43834.6016 | 612 | 44394.1211 | 612 | 43993.0781
[14 empty] | 613 | 21.6253 | 0.66859 | -0.0054594 |  | 613 | 0.67935 | -0.062688 | 613 | 0.67088 | -0.017802 | 613 | 43851.1211 | 613 | 44415.5195 | 613 | 44001.4414
[14 empty] | 614 | 21.6606 | 0.66816 | -0.0031508 |  | 614 | 0.67974 | -0.06473 | 614 | 0.67067 | -0.016648 | 614 | 43868.1602 | 614 | 44413.0781 | 614 | 44021.3984
[14 empty] | 615 | 21.6958 | 0.66817 | -0.0032069 |  | 615 | 0.68037 | -0.068012 | 615 | 0.67091 | -0.017973 | 615 | 43854.6797 | 615 | 44420.2812 | 615 | 44022.0391
[14 empty] | 616 | 21.7311 | 0.66799 | -0.0021979 |  | 616 | 0.68014 | -0.066852 | 616 | 0.67062 | -0.016395 | 616 | 43848.1992 | 616 | 44413.2383 | 616 | 43987.2812
[14 empty] | 617 | 21.7664 | 0.66777 | -0.0010007 |  | 617 | 0.68008 | -0.066522 | 617 | 0.67094 | -0.018122 | 617 | 43845.5586 | 617 | 44423.1992 | 617 | 43990.2383
[14 empty] | 618 | 21.8017 | 0.66779 | -0.001132 |  | 618 | 0.68046 | -0.068531 | 618 | 0.67106 | -0.018732 | 618 | 43843.4805 | 618 | 44457.1211 | 618 | 43979.6406
[14 empty] | 619 | 21.8369 | 0.66833 | -0.0040317 |  | 619 | 0.68039 | -0.068146 | 619 | 0.67148 | -0.020978 | 619 | 43818.8008 | 619 | 44500.2812 | 619 | 43938.5586
[14 empty] | 620 | 21.8722 | 0.66855 | -0.0052682 |  | 620 | 0.6809 | -0.070801 | 620 | 0.6715 | -0.021085 | 620 | 43767.3984 | 620 | 44479.8008 | 620 | 43916.1602
[14 empty] | 621 | 21.9075 | 0.66886 | -0.0069308 |  | 621 | 0.6812 | -0.072346 | 621 | 0.67132 | -0.020142 | 621 | 43715.2383 | 621 | 44473.4805 | 621 | 43922.0781
[14 empty] | 622 | 21.9428 | 0.669 | -0.0076492 |  | 622 | 0.6801 | -0.066606 | 622 | 0.67182 | -0.022845 | 622 | 43750.9219 | 622 | 44496.4805 | 622 | 43911.9219
[14 empty] | 623 | 21.9781 | 0.66962 | -0.010987 |  | 623 | 0.6792 | -0.061896 | 623 | 0.67164 | -0.021836 | 623 | 43799.9609 | 623 | 44487.2812 | 623 | 43915.4414
[14 empty] | 624 | 22.0133 | 0.66919 | -0.0086852 |  | 624 | 0.67862 | -0.058861 | 624 | 0.67145 | -0.020855 | 624 | 43839.2383 | 624 | 44441.9609 | 624 | 43964.2383
[14 empty] | 625 | 22.0486 | 0.6696 | -0.010916 |  | 625 | 0.67828 | -0.057057 | 625 | 0.67155 | -0.021379 | 625 | 43878.3594 | 625 | 44442.3594 | 625 | 43968.3594
[14 empty] | 626 | 22.0839 | 0.66943 | -0.0099749 |  | 626 | 0.67836 | -0.057498 | 626 | 0.67147 | -0.020944 | 626 | 43898.4805 | 626 | 44423.0781 | 626 | 43981.0781
[14 empty] | 627 | 22.1192 | 0.66939 | -0.0097842 |  | 627 | 0.67834 | -0.057414 | 627 | 0.67156 | -0.021412 | 627 | 43896.6016 | 627 | 44391.8789 | 627 | 43936.3984
[14 empty] | 628 | 22.1544 | 0.66983 | -0.012148 |  | 628 | 0.67909 | -0.06131 | 628 | 0.67182 | -0.022849 | 628 | 43903.6016 | 628 | 44445.5195 | 628 | 43901.0391
[14 empty] | 629 | 22.1897 | 0.67033 | -0.014841 |  | 629 | 0.67941 | -0.063008 | 629 | 0.67207 | -0.024155 | 629 | 43889.1211 | 629 | 44465.4414 | 629 | 43956.4805
[14 empty] | 630 | 22.225 | 0.6704 | -0.015221 |  | 630 | 0.68001 | -0.066167 | 630 | 0.67195 | -0.023532 | 630 | 43784.4414 | 630 | 44492.9219 | 630 | 43954.3594
[14 empty] | 631 | 22.2603 | 0.67027 | -0.014513 |  | 631 | 0.68046 | -0.068494 | 631 | 0.67171 | -0.022248 | 631 | 43850.3984 | 631 | 44528.5586 | 631 | 43917.5586
[14 empty] | 632 | 22.2956 | 0.67003 | -0.013223 |  | 632 | 0.68079 | -0.070251 | 632 | 0.67128 | -0.019948 | 632 | 43831 | 632 | 44548.8789 | 632 | 43948.7617
[14 empty] | 633 | 22.3308 | 0.67085 | -0.017642 |  | 633 | 0.68078 | -0.070149 | 633 | 0.67135 | -0.020331 | 633 | 43812.8008 | 633 | 44553.0781 | 633 | 43953.7188
[14 empty] | 634 | 22.3661 | 0.67065 | -0.016551 |  | 634 | 0.68068 | -0.069675 | 634 | 0.67047 | -0.015609 | 634 | 43792.7188 | 634 | 44562.8008 | 634 | 43921.9219
[14 empty] | 635 | 22.4014 | 0.67059 | -0.01622 |  | 635 | 0.68103 | -0.071495 | 635 | 0.67077 | -0.017169 | 635 | 43757.8008 | 635 | 44550.3984 | 635 | 43947.9609
[14 empty] | 636 | 22.4367 | 0.67014 | -0.01382 |  | 636 | 0.68168 | -0.074864 | 636 | 0.67109 | -0.018892 | 636 | 43748.6406 | 636 | 44555.9609 | 636 | 43940.2383
[14 empty] | 637 | 22.4719 | 0.67106 | -0.018751 |  | 637 | 0.68127 | -0.072707 | 637 | 0.67132 | -0.020134 | 637 | 43773.5195 | 637 | 44565.5195 | 637 | 43971.2812
[14 empty] | 638 | 22.5072 | 0.67156 | -0.021434 |  | 638 | 0.68109 | -0.071814 | 638 | 0.6719 | -0.02325 | 638 | 43773.1602 | 638 | 44596.0781 | 638 | 43969.3594
[14 empty] | 639 | 22.5425 | 0.67184 | -0.022931 |  | 639 | 0.68076 | -0.070066 | 639 | 0.6723 | -0.025401 | 639 | 43767.7617 | 639 | 44616.5586 | 639 | 43958.5195
[14 empty] | 640 | 22.5778 | 0.67204 | -0.024022 |  | 640 | 0.68124 | -0.072552 | 640 | 0.67263 | -0.027176 | 640 | 43784.6797 | 640 | 44633.1992 | 640 | 43950.1602
[14 empty] | 641 | 22.6131 | 0.67169 | -0.022151 |  | 641 | 0.68203 | -0.076665 | 641 | 0.67284 | -0.028261 | 641 | 43804.8789 | 641 | 44666.6797 | 641 | 43941.7617
[14 empty] | 642 | 22.6483 | 0.67162 | -0.021728 |  | 642 | 0.68256 | -0.079433 | 642 | 0.67204 | -0.024003 | 642 | 43816.0391 | 642 | 44669.5586 | 642 | 43931.2812
[14 empty] | 643 | 22.6836 | 0.67192 | -0.023383 |  | 643 | 0.68255 | -0.079376 | 643 | 0.67194 | -0.023473 | 643 | 43839.6016 | 643 | 44627.7617 | 643 | 43955.6016
[14 empty] | 644 | 22.7189 | 0.67191 | -0.023317 |  | 644 | 0.6826 | -0.07966 | 644 | 0.67184 | -0.022909 | 644 | 43845.6406 | 644 | 44554.8789 | 644 | 43978.3594
[14 empty] | 645 | 22.7542 | 0.67172 | -0.022289 |  | 645 | 0.6835 | -0.084323 | 645 | 0.67176 | -0.022482 | 645 | 43869.5586 | 645 | 44530.0781 | 645 | 43982.3984
[14 empty] | 646 | 22.7894 | 0.67169 | -0.022125 |  | 646 | 0.68418 | -0.087854 | 646 | 0.67202 | -0.023888 | 646 | 43861.8789 | 646 | 44527.2383 | 646 | 43941.4805
[14 empty] | 647 | 22.8247 | 0.67137 | -0.020394 |  | 647 | 0.68423 | -0.088109 | 647 | 0.67259 | -0.026924 | 647 | 43854.7188 | 647 | 44537.9219 | 647 | 43915.8789
[14 empty] | 648 | 22.86 | 0.67166 | -0.021992 |  | 648 | 0.68395 | -0.086628 | 648 | 0.67325 | -0.030447 | 648 | 43838.6797 | 648 | 44539.8398 | 648 | 43939.7188
[14 empty] | 649 | 22.8953 | 0.672 | -0.023803 |  | 649 | 0.68417 | -0.087776 | 649 | 0.67333 | -0.030912 | 649 | 43878.8008 | 649 | 44558.7617 | 649 | 43887.1602
[14 empty] | 650 | 22.9306 | 0.67259 | -0.026931 |  | 650 | 0.68431 | -0.088486 | 650 | 0.67321 | -0.030247 | 650 | 43880.8008 | 650 | 44572.8008 | 650 | 43881.1602
[14 empty] | 651 | 22.9658 | 0.67298 | -0.029041 |  | 651 | 0.68377 | -0.085731 | 651 | 0.67317 | -0.030018 | 651 | 43895.9219 | 651 | 44604.5586 | 651 | 43914.0391
[14 empty] | 652 | 23.0011 | 0.67357 | -0.032154 |  | 652 | 0.68369 | -0.085292 | 652 | 0.67346 | -0.031566 | 652 | 43949.8789 | 652 | 44619.0781 | 652 | 43948.7188
[14 empty] | 653 | 23.0364 | 0.67331 | -0.030805 |  | 653 | 0.68388 | -0.086283 | 653 | 0.67365 | -0.032597 | 653 | 43943.0391 | 653 | 44612.8008 | 653 | 43977.3594
[14 empty] | 654 | 23.0717 | 0.67357 | -0.032139 |  | 654 | 0.68441 | -0.089009 | 654 | 0.67353 | -0.031943 | 654 | 43908.9219 | 654 | 44628.5195 | 654 | 43987.9609
[25 empty] | 655 | 43962.9609 | 655 | 44637.6016 | 655 | 44015.2383
[25 empty] | 656 | 43955.5586 | 656 | 44644.6406 | 656 | 43997.8789
[25 empty] | 657 | 43924.3594 | 657 | 44681.2812 | 657 | 43998.4805
[25 empty] | 658 | 43916.0391 | 658 | 44654.5586 | 658 | 44006.1992
[25 empty] | 659 | 43939.2383 | 659 | 44604.9219 | 659 | 43986.6797
[25 empty] | 660 | 44003.8789 | 660 | 44617 | 660 | 43998.6406
[25 empty] | 661 | 44026.7617 | 661 | 44657.3594 | 661 | 43990.4805
[25 empty] | 662 | 44049.8008 | 662 | 44663.8008 | 662 | 43997.1602
[25 empty] | 663 | 44029.1211 | 663 | 44714.6016 | 663 | 44030.0781
[25 empty] | 664 | 43989.5586 | 664 | 44735.1992 | 664 | 44042.1992
[25 empty] | 665 | 44012.4805 | 665 | 44755.8008 | 665 | 44045.6406
[25 empty] | 666 | 44025.2383 | 666 | 44797.7617 | 666 | 44042.8398
[25 empty] | 667 | 43999.5586 | 667 | 44813.1992 | 667 | 44049.7617
[25 empty] | 668 | 44005.1211 | 668 | 44840.0391 | 668 | 44076.8789
[25 empty] | 669 | 44013.2383 | 669 | 44840.5586 | 669 | 44088.7617
[25 empty] | 670 | 44026.7188 | 670 | 44863.4414 | 670 | 44087.5195
[25 empty] | 671 | 44039.7188 | 671 | 44885.6406 | 671 | 44089.1992
[25 empty] | 672 | 44068.6016 | 672 | 44870.4414 | 672 | 44058.8789
[25 empty] | 673 | 44085 | 673 | 44839.6797 | 673 | 44056.6406
[25 empty] | 674 | 44123.4805 | 674 | 44857.1992 | 674 | 43991.6797
[25 empty] | 675 | 44082.8008 | 675 | 44867.6016 | 675 | 43947.5195
[25 empty] | 676 | 44137.2383 | 676 | 44856.9219 | 676 | 43926.3203
[25 empty] | 677 | 44096 | 677 | 44878.7617 | 677 | 43928.7188
[25 empty] | 678 | 44089.3203 | 678 | 44891.2383 | 678 | 43946.5586
[25 empty] | 679 | 44025.1992 | 679 | 44897.0391 | 679 | 43904.5195
[25 empty] | 680 | 44028.5195 | 680 | 44916.7188 | 680 | 43931.6797
[25 empty] | 681 | 43991.7617 | 681 | 44860.9609 | 681 | 43950.5195
[25 empty] | 682 | 44024.3203 | 682 | 44859.9609 | 682 | 43975.6797
[25 empty] | 683 | 44066.9219 | 683 | 44915.8398 | 683 | 44025.3594
[25 empty] | 684 | 44066.5195 | 684 | 44902.8398 | 684 | 44039.1602
[25 empty] | 685 | 44045.6797 | 685 | 44913.0391 | 685 | 44061.1992
[25 empty] | 686 | 44071.5586 | 686 | 44921.7188 | 686 | 43997.5586
[25 empty] | 687 | 44071.6406 | 687 | 44957.5195 | 687 | 43987.4414
[25 empty] | 688 | 44077.1602 | 688 | 44959.5195 | 688 | 44013.0391
[25 empty] | 689 | 44055.9609 | 689 | 45006.9219 | 689 | 44002.3984

## Profiler vs epson unflat
|  |  |  | dose=c+b/(X-a) | [7 empty] | pixels to cm conversion | 1px = |  | 1/ | 1 inch = 2.54 centimetres
a |  |  |  | -0.040473 | [11 empty] | dpi | ppcm | cm/px
b |  |  |  | 2.7151 | [11 empty] | 72 | 28.3465 | 0.035278
c |  |  |  | -3.8345
centre across | dist cm |  |  | normalized film response | Gy
0 | 0 | -9.5779 | 42527.668 | 0.64893 | 0.10373 | 4.7575 | -11.2 | 1.3544
1 | 0.035278 | -9.5426 | 42548.2227 | 0.64924 | 0.10194 | 4.6753 | -10.8 | 1.4443
2 | 0.070556 | -9.5074 | 42472 | 0.64808 | 0.10859 | 4.9803 | -10.4 | 1.5476
3 | 0.10583 | -9.4721 | 42531 | 0.64898 | 0.10344 | 4.7442 | -10 | 1.6799
4 | 0.14111 | -9.4368 | 42565.5547 | 0.64951 | 0.10043 | 4.6061 | -9.6 | 1.8313
5 | 0.17639 | -9.4015 | 42536.4453 | 0.64906 | 0.10297 | 4.7224 | -9.2 | 2.0226
6 | 0.21167 | -9.3663 | 42515.668 | 0.64875 | 0.10478 | 4.8055 | -8.8 | 2.205
7 | 0.24694 | -9.331 | 42477.8906 | 0.64817 | 0.10808 | 4.9567 | -8.4 | 2.379
8 | 0.28222 | -9.2957 | 42488.668 | 0.64834 | 0.10714 | 4.9135 | -8 | 2.8574
9 | 0.3175 | -9.2604 | 42537.1094 | 0.64907 | 0.10291 | 4.7197 | -7.6 | 3.2205
10 | 0.35278 | -9.2251 | 42552 | 0.6493 | 0.10161 | 4.6603 | -7.2 | 3.6975
11 | 0.38806 | -9.1899 | 42513.4453 | 0.64871 | 0.10497 | 4.8144 | -6.8 | 4.333
12 | 0.42333 | -9.1546 | 42552.5547 | 0.64931 | 0.10156 | 4.658 | -6.4 | 5.0721
13 | 0.45861 | -9.1193 | 42514.8906 | 0.64874 | 0.10485 | 4.8086 | -6 | 6.4392
14 | 0.49389 | -9.084 | 42549.668 | 0.64927 | 0.10182 | 4.6696 | -5.6 | 9.2202
15 | 0.52917 | -9.0488 | 42511.332 | 0.64868 | 0.10516 | 4.8228 | -5.2 | 20.0409
16 | 0.56444 | -9.0135 | 42523.4453 | 0.64887 | 0.1041 | 4.7744 | -4.8 | 84.8158
17 | 0.59972 | -8.9782 | 42584.2227 | 0.64979 | 0.098809 | 4.5317 | -4.4 | 98.2231
18 | 0.635 | -8.9429 | 42541.332 | 0.64914 | 0.10254 | 4.7029 | -4 | 100.3047
19 | 0.67028 | -8.9076 | 42544.5547 | 0.64919 | 0.10226 | 4.69 | -3.6 | 101.0664
20 | 0.70556 | -8.8724 | 42494.7773 | 0.64843 | 0.1066 | 4.8891 | -3.2 | 101.0972
21 | 0.74083 | -8.8371 | 42515.4453 | 0.64874 | 0.1048 | 4.8064 | -2.8 | 100.9568
22 | 0.77611 | -8.8018 | 42528.4453 | 0.64894 | 0.10367 | 4.7544 | -2.4 | 100.497
23 | 0.81139 | -8.7665 | 42542.332 | 0.64915 | 0.10246 | 4.6989 | -2 | 100.0592
24 | 0.84667 | -8.7312 | 42537.7773 | 0.64908 | 0.10285 | 4.7171 | -1.6 | 99.4802
25 | 0.88194 | -8.696 | 42479.4453 | 0.64819 | 0.10794 | 4.9505 | -1.2 | 99.6109
26 | 0.91722 | -8.6607 | 42466.1094 | 0.64799 | 0.10911 | 5.0039 | -0.8 | 99.742
27 | 0.9525 | -8.6254 | 42475.1094 | 0.64813 | 0.10832 | 4.9678 | -0.4 | 99.9903 |  | -11.2 | 1.3544 |  | 0 | 42527.668
28 | 0.98778 | -8.5901 | 42421 | 0.6473 | 0.11305 | 5.1849 | 0 | 100 |  | -10.8 | 1.4443 |  | 1 | 42548.2227
29 | 1.0231 | -8.5549 | 42484.2227 | 0.64827 | 0.10752 | 4.9313 | 0.4 | 100.1212 |  | -10.4 | 1.5476 |  | 2 | 42472
30 | 1.0583 | -8.5196 | 42585.2227 | 0.64981 | 0.098722 | 4.5277 | 0.8 | 100.0051 |  | -10 | 1.6799 |  | 3 | 42531
31 | 1.0936 | -8.4843 | 42585.8906 | 0.64982 | 0.098664 | 4.525 | 1.2 | 99.834 |  | -9.6 | 1.8313 |  | 4 | 42565.5547
32 | 1.1289 | -8.449 | 42570.332 | 0.64958 | 0.10002 | 4.5871 | 1.6 | 99.7528 |  | -9.2 | 2.0226 |  | 5 | 42536.4453
33 | 1.1642 | -8.4138 | 42522.668 | 0.64885 | 0.10417 | 4.7775 | 2 | 100.0306 |  | -8.8 | 2.205 |  | 6 | 42515.668
34 | 1.1994 | -8.3785 | 42447.5547 | 0.64771 | 0.11073 | 5.0783 | 2.4 | 100.5114 |  | -8.4 | 2.379 |  | 7 | 42477.8906
35 | 1.2347 | -8.3432 | 42538.8906 | 0.6491 | 0.10275 | 4.7126 | 2.8 | 100.9674 |  | -8 | 2.8574 |  | 8 | 42488.668
36 | 1.27 | -8.3079 | 42598.5547 | 0.65001 | 0.097563 | 4.4745 | 3.2 | 101.1327 |  | -7.6 | 3.2205 |  | 9 | 42537.1094
37 | 1.3053 | -8.2726 | 42533.2227 | 0.64902 | 0.10325 | 4.7353 | 3.6 | 100.9958 |  | -7.2 | 3.6975 |  | 10 | 42552
38 | 1.3406 | -8.2374 | 42493.7773 | 0.64841 | 0.10669 | 4.8931 | 4 | 100.2152 |  | -6.8 | 4.333 |  | 11 | 42513.4453
39 | 1.3758 | -8.2021 | 42488.1094 | 0.64833 | 0.10718 | 4.9158 | 4.4 | 96.7088 |  | -6.4 | 5.0721 |  | 12 | 42552.5547
40 | 1.4111 | -8.1668 | 42492.4453 | 0.64839 | 0.10681 | 4.8984 | 4.8 | 75.9462 |  | -6 | 6.4392 |  | 13 | 42514.8906
41 | 1.4464 | -8.1315 | 42477.332 | 0.64816 | 0.10813 | 4.9589 | 5.2 | 15.3962 |  | -5.6 | 9.2202 |  | 14 | 42549.668
42 | 1.4817 | -8.0962 | 42502.8906 | 0.64855 | 0.10589 | 4.8566 | 5.6 | 8.0949 |  | -5.2 | 20.0409 |  | 15 | 42511.332
43 | 1.5169 | -8.061 | 42533 | 0.64901 | 0.10327 | 4.7362 | 6 | 5.9523 |  | -4.8 | 84.8158 |  | 16 | 42523.4453
44 | 1.5522 | -8.0257 | 42545.5547 | 0.6492 | 0.10217 | 4.686 | 6.4 | 4.9166 |  | -4.4 | 98.2231 |  | 17 | 42584.2227
45 | 1.5875 | -7.9904 | 42514.7773 | 0.64873 | 0.10486 | 4.809 | 6.8 | 4.1063 |  | -4 | 100.3047 |  | 18 | 42541.332
46 | 1.6228 | -7.9551 | 42481.2227 | 0.64822 | 0.10779 | 4.9433 | 7.2 | 3.5706 |  | -3.6 | 101.0664 |  | 19 | 42544.5547
47 | 1.6581 | -7.9199 | 42528.1094 | 0.64894 | 0.10369 | 4.7557 | 7.6 | 3.0823 |  | -3.2 | 101.0972 |  | 20 | 42494.7773
48 | 1.6933 | -7.8846 | 42462.332 | 0.64793 | 0.10944 | 5.0191 | 8 | 2.639 |  | -2.8 | 100.9568 |  | 21 | 42515.4453
49 | 1.7286 | -7.8493 | 42413.668 | 0.64719 | 0.1137 | 5.2144 | 8.4 | 2.3103 |  | -2.4 | 100.497 |  | 22 | 42528.4453
50 | 1.7639 | -7.814 | 42447.1094 | 0.6477 | 0.11077 | 5.0801 | 8.8 | 2.0938 |  | -2 | 100.0592 |  | 23 | 42542.332
51 | 1.7992 | -7.7788 | 42446.7773 | 0.6477 | 0.1108 | 5.0814 | 9.2 | 1.915 |  | -1.6 | 99.4802 |  | 24 | 42537.7773
52 | 1.8344 | -7.7435 | 42414.7773 | 0.64721 | 0.1136 | 5.2099 | 9.6 | 1.6851 |  | -1.2 | 99.6109 |  | 25 | 42479.4453
53 | 1.8697 | -7.7082 | 42406 | 0.64707 | 0.11437 | 5.2452 | 10 | 1.5497 |  | -0.8 | 99.742 |  | 26 | 42466.1094
54 | 1.905 | -7.6729 | 42410.4453 | 0.64714 | 0.11398 | 5.2273 | 10.4 | 1.4568 |  | -0.4 | 99.9903 |  | 27 | 42475.1094
55 | 1.9403 | -7.6376 | 42447.1094 | 0.6477 | 0.11077 | 5.0801 | 10.8 | 1.3746 |  | 0 | 100 |  | 28 | 42421
56 | 1.9756 | -7.6024 | 42450.8906 | 0.64776 | 0.11044 | 5.0649 | 11.2 | 1.2596 |  | 0.4 | 100.1212 |  | 29 | 42484.2227
57 | 2.0108 | -7.5671 | 42462.4453 | 0.64794 | 0.10943 | 5.0186 |  |  |  | 0.8 | 100.0051 |  | 30 | 42585.2227
58 | 2.0461 | -7.5318 | 42508.2227 | 0.64863 | 0.10543 | 4.8352 | Y Axis Position(cm) | Set 1 |  | 1.2 | 99.834 |  | 31 | 42585.8906
59 | 2.0814 | -7.4965 | 42526.668 | 0.64892 | 0.10382 | 4.7615 | -16.4 | 2.4625 |  | 1.6 | 99.7528 |  | 32 | 42570.332
60 | 2.1167 | -7.4612 | 42550.332 | 0.64928 | 0.10176 | 4.6669 | -16 | 2.6876 |  | 2 | 100.0306 |  | 33 | 42522.668
61 | 2.1519 | -7.426 | 42425.7773 | 0.64738 | 0.11263 | 5.1657 | -15.6 | 2.8697 |  | 2.4 | 100.5114 |  | 34 | 42447.5547
62 | 2.1872 | -7.3907 | 42463.4453 | 0.64795 | 0.10934 | 5.0146 | -15.2 | 3.1415 |  | 2.8 | 100.9674 |  | 35 | 42538.8906
63 | 2.2225 | -7.3554 | 42351.1094 | 0.64624 | 0.11918 | 5.4661 | -14.8 | 3.348 |  | 3.2 | 101.1327 |  | 36 | 42598.5547
64 | 2.2578 | -7.3201 | 42377.332 | 0.64664 | 0.11688 | 5.3605 | -14.4 | 3.7106 |  | 3.6 | 100.9958 |  | 37 | 42533.2227
65 | 2.2931 | -7.2849 | 42352.8906 | 0.64626 | 0.11903 | 5.4589 | -14 | 4.1405 |  | 4 | 100.2152 |  | 38 | 42493.7773
66 | 2.3283 | -7.2496 | 42396.332 | 0.64693 | 0.11521 | 5.2841 | -13.6 | 4.5166 |  | 4.4 | 96.7088 |  | 39 | 42488.1094
67 | 2.3636 | -7.2143 | 42270.668 | 0.64501 | 0.12626 | 5.7908 | -13.2 | 5.0184 |  | 4.8 | 75.9462 |  | 40 | 42492.4453
68 | 2.3989 | -7.179 | 42213.668 | 0.64414 | 0.1313 | 6.0216 | -12.8 | 5.5321 |  | 5.2 | 15.3962 |  | 41 | 42477.332
69 | 2.4342 | -7.1438 | 42327.668 | 0.64588 | 0.12124 | 5.5606 | -12.4 | 6.1006 |  | 5.6 | 8.0949 |  | 42 | 42502.8906
70 | 2.4694 | -7.1085 | 42296.7773 | 0.64541 | 0.12396 | 5.6853 | -12 | 6.7372 |  | 6 | 5.9523 |  | 43 | 42533
71 | 2.5047 | -7.0732 | 42393 | 0.64688 | 0.11551 | 5.2975 | -11.6 | 7.6521 |  | 6.4 | 4.9166 |  | 44 | 42545.5547
72 | 2.54 | -7.0379 | 42420 | 0.64729 | 0.11314 | 5.1889 | -11.2 | 9.0169 |  | 6.8 | 4.1063 |  | 45 | 42514.7773
73 | 2.5753 | -7.0026 | 42336.7773 | 0.64602 | 0.12044 | 5.5239 | -10.8 | 11.9192 |  | 7.2 | 3.5706 |  | 46 | 42481.2227
74 | 2.6106 | -6.9674 | 42234.2227 | 0.64445 | 0.12948 | 5.9383 | -10.4 | 23.1964 |  | 7.6 | 3.0823 |  | 47 | 42528.1094
75 | 2.6458 | -6.9321 | 42130.4453 | 0.64287 | 0.13867 | 6.3596 | -10 | 70.0223 |  | 8 | 2.639 |  | 48 | 42462.332
76 | 2.6811 | -6.8968 | 42169.332 | 0.64346 | 0.13522 | 6.2015 | -9.6 | 96.3873 |  | 8.4 | 2.3103 |  | 49 | 42413.668
77 | 2.7164 | -6.8615 | 42105.668 | 0.64249 | 0.14086 | 6.4604 | -9.2 | 100.5738 |  | 8.8 | 2.0938 |  | 50 | 42447.1094
78 | 2.7517 | -6.8262 | 42124.1094 | 0.64277 | 0.13923 | 6.3854 | -8.8 | 101.6498 |  | 9.2 | 1.915 |  | 51 | 42446.7773
79 | 2.7869 | -6.791 | 42122.668 | 0.64275 | 0.13936 | 6.3912 | -8.4 | 102.3072 |  | 9.6 | 1.6851 |  | 52 | 42414.7773
80 | 2.8222 | -6.7557 | 42076.332 | 0.64204 | 0.14347 | 6.58 | -8 | 102.7132 |  | 10 | 1.5497 |  | 53 | 42406
81 | 2.8575 | -6.7204 | 42067.4453 | 0.64191 | 0.14426 | 6.6163 | -7.6 | 103.0261 |  | 10.4 | 1.4568 |  | 54 | 42410.4453
82 | 2.8928 | -6.6851 | 42142.332 | 0.64305 | 0.13761 | 6.3112 | -7.2 | 103.1994 |  | 10.8 | 1.3746 |  | 55 | 42447.1094
83 | 2.9281 | -6.6499 | 42118.5547 | 0.64269 | 0.13972 | 6.408 | -6.8 | 103.3055 |  | 11.2 | 1.2596 |  | 56 | 42450.8906
84 | 2.9633 | -6.6146 | 42126.5547 | 0.64281 | 0.13901 | 6.3754 | -6.4 | 103.506 | Detector ID | Y Axis Position(cm) | Set 1 |  | 57 | 42462.4453
85 | 2.9986 | -6.5793 | 42003.5547 | 0.64093 | 0.14996 | 6.8774 | -6 | 103.4357 |  | -16.4 | 0.67405 |  | 58 | 42508.2227
86 | 3.0339 | -6.544 | 42022.4453 | 0.64122 | 0.14827 | 6.8001 | -5.6 | 103.3058 |  | -16 | 0.75086 |  | 59 | 42526.668
87 | 3.0692 | -6.5088 | 42050.5547 | 0.64165 | 0.14577 | 6.6852 | -5.2 | 103.2712 |  | -15.6 | 0.77318 |  | 60 | 42550.332
88 | 3.1044 | -6.4735 | 42029.668 | 0.64133 | 0.14763 | 6.7706 | -4.8 | 103.2441 |  | -15.2 | 0.85447 |  | 61 | 42425.7773
89 | 3.1397 | -6.4382 | 42020.8906 | 0.6412 | 0.14841 | 6.8064 | -4.4 | 103.0501 |  | -14.8 | 0.92808 |  | 62 | 42463.4453
90 | 3.175 | -6.4029 | 42074.668 | 0.64202 | 0.14362 | 6.5868 | -4 | 102.5345 |  | -14.4 | 0.99125 |  | 63 | 42351.1094
91 | 3.2103 | -6.3676 | 41960.1094 | 0.64027 | 0.15384 | 7.0553 | -3.6 | 101.8602 |  | -14 | 1.0613 |  | 64 | 42377.332
92 | 3.2456 | -6.3324 | 41969.5547 | 0.64041 | 0.15299 | 7.0166 | -3.2 | 101.437 |  | -13.6 | 1.1297 |  | 65 | 42352.8906
93 | 3.2808 | -6.2971 | 41922.332 | 0.63969 | 0.15722 | 7.2103 | -2.8 | 100.637 |  | -13.2 | 1.1934 |  | 66 | 42396.332
94 | 3.3161 | -6.2618 | 41880.8906 | 0.63906 | 0.16093 | 7.3807 | -2.4 | 100.2303 |  | -12.8 | 1.2675 |  | 67 | 42270.668
95 | 3.3514 | -6.2265 | 41789.2227 | 0.63766 | 0.16917 | 7.7587 | -2 | 99.61 |  | -12.4 | 1.3777 |  | 68 | 42213.668
96 | 3.3867 | -6.1913 | 41477.2227 | 0.6329 | 0.19748 | 9.0569 | -1.6 | 99.4825 |  | -12 | 1.4477 |  | 69 | 42327.668
97 | 3.4219 | -6.156 | 41330.8906 | 0.63067 | 0.21089 | 9.6721 | -1.2 | 99.6622 |  | -11.6 | 1.5635 |  | 70 | 42296.7773
98 | 3.4572 | -6.1207 | 41604.332 | 0.63484 | 0.1859 | 8.5258 | -0.8 | 99.8668 |  | -11.2 | 1.6924 |  | 71 | 42393
99 | 3.4925 | -6.0854 | 41774.332 | 0.63744 | 0.17051 | 7.8202 | -0.4 | 99.8628 |  | -10.8 | 1.8483 |  | 72 | 42420
100 | 3.5278 | -6.0501 | 41760 | 0.63722 | 0.17181 | 7.8795 | 0 | 100 |  | -10.4 | 1.9765 |  | 73 | 42336.7773
101 | 3.5631 | -6.0149 | 41675.5547 | 0.63593 | 0.17944 | 8.2295 | 0.4 | 100.2499 |  | -10 | 2.1579 |  | 74 | 42234.2227
102 | 3.5983 | -5.9796 | 41697.2227 | 0.63626 | 0.17748 | 8.1396 | 0.8 | 100.3314 |  | -9.6 | 2.3587 |  | 75 | 42130.4453
103 | 3.6336 | -5.9443 | 41711.332 | 0.63647 | 0.1762 | 8.0811 | 1.2 | 100.5899 |  | -9.2 | 2.5998 |  | 76 | 42169.332
104 | 3.6689 | -5.909 | 41635.5547 | 0.63532 | 0.18306 | 8.3958 | 1.6 | 101.2232 |  | -8.8 | 2.8183 |  | 77 | 42105.668
105 | 3.7042 | -5.8738 | 41632.5547 | 0.63527 | 0.18334 | 8.4083 | 2 | 101.6594 |  | -8.4 | 3.1722 |  | 78 | 42124.1094
106 | 3.7394 | -5.8385 | 41553.668 | 0.63407 | 0.19051 | 8.7371 | 2.4 | 102.6837 |  | -8 | 3.5408 |  | 79 | 42122.668
107 | 3.7747 | -5.8032 | 41584.5547 | 0.63454 | 0.1877 | 8.6082 | 2.8 | 103.2703 |  | -7.6 | 3.9315 |  | 80 | 42076.332
108 | 3.81 | -5.7679 | 41538.2227 | 0.63383 | 0.19191 | 8.8016 | 3.2 | 104.0326 |  | -7.2 | 4.3838 |  | 81 | 42067.4453
109 | 3.8453 | -5.7326 | 41476.8906 | 0.6329 | 0.19751 | 9.0583 | 3.6 | 104.4274 |  | -6.8 | 5.0034 |  | 82 | 42142.332
110 | 3.8806 | -5.6974 | 41367.8906 | 0.63123 | 0.20749 | 9.5162 | 4 | 104.8436 |  | -6.4 | 5.6755 |  | 83 | 42118.5547
111 | 3.9158 | -5.6621 | 41391.668 | 0.6316 | 0.20531 | 9.4161 | 4.4 | 104.9309 |  | -6 | 6.8761 |  | 84 | 42126.5547
112 | 3.9511 | -5.6268 | 41281.332 | 0.62991 | 0.21546 | 9.8814 | 4.8 | 104.9451 |  | -5.6 | 9.9146 |  | 85 | 42003.5547
113 | 3.9864 | -5.5915 | 41298.4453 | 0.63017 | 0.21388 | 9.8091 | 5.2 | 105.0835 |  | -5.2 | 22.9712 |  | 86 | 42022.4453
114 | 4.0217 | -5.5563 | 41190.668 | 0.62853 | 0.22383 | 10.2655 | 5.6 | 105.0255 |  | -4.8 | 67.2426 |  | 87 | 42050.5547
115 | 4.0569 | -5.521 | 41089.1094 | 0.62698 | 0.23325 | 10.6977 | 6 | 104.9896 |  | -4.4 | 91.2343 |  | 88 | 42029.668
116 | 4.0922 | -5.4857 | 41029.1094 | 0.62606 | 0.23884 | 10.9539 | 6.4 | 104.7757 |  | -4 | 98.0957 |  | 89 | 42020.8906
117 | 4.1275 | -5.4504 | 40932.5547 | 0.62459 | 0.24787 | 11.3678 | 6.8 | 104.8769 |  | -3.6 | 99.4905 |  | 90 | 42074.668
118 | 4.1628 | -5.4151 | 40715.1094 | 0.62127 | 0.26833 | 12.3066 | 7.2 | 104.5792 |  | -3.2 | 100.0978 |  | 91 | 41960.1094
119 | 4.1981 | -5.3799 | 40565.332 | 0.61899 | 0.28255 | 12.9587 | 7.6 | 104.2662 |  | -2.8 | 99.8854 |  | 92 | 41969.5547
120 | 4.2333 | -5.3446 | 40401.5547 | 0.61649 | 0.29822 | 13.677 | 8 | 104.0358 |  | -2.4 | 99.8753 |  | 93 | 41922.332
121 | 4.2686 | -5.3093 | 40218.8906 | 0.6137 | 0.31582 | 14.4846 | 8.4 | 103.5482 |  | -2 | 99.4692 |  | 94 | 41880.8906
122 | 4.3039 | -5.274 | 40016.332 | 0.61061 | 0.33553 | 15.3882 | 8.8 | 102.1359 |  | -1.6 | 99.4543 |  | 95 | 41789.2227
123 | 4.3392 | -5.2388 | 39645.4453 | 0.60495 | 0.37209 | 17.0652 | 9.2 | 99.2299 |  | -1.2 | 99.5821 |  | 96 | 41477.2227
124 | 4.3744 | -5.2035 | 39222.668 | 0.5985 | 0.41456 | 19.013 | 9.6 | 83.9414 |  | -0.8 | 99.8154 |  | 97 | 41330.8906
125 | 4.4097 | -5.1682 | 38733.7773 | 0.59104 | 0.46476 | 21.315 | 10 | 32.5274 |  | -0.4 | 99.8557 |  | 98 | 41604.332
126 | 4.445 | -5.1329 | 38095.668 | 0.5813 | 0.53208 | 24.4028 | 10.4 | 13.2724 |  | 0 | 100 |  | 99 | 41774.332
127 | 4.4803 | -5.0976 | 37395.332 | 0.57062 | 0.60845 | 27.905 | 10.8 | 9.6188 |  | 0.4 | 100.114 |  | 100 | 41760
128 | 4.5156 | -5.0624 | 36695.5547 | 0.55994 | 0.68746 | 31.5288 | 11.2 | 8.3009 |  | 0.8 | 100.0654 |  | 101 | 41675.5547
129 | 4.5508 | -5.0271 | 35769.332 | 0.54581 | 0.79647 | 36.5284 | 11.6 | 7.21 |  | 1.2 | 100.0297 |  | 102 | 41697.2227
130 | 4.5861 | -4.9918 | 34736.4453 | 0.53004 | 0.92441 | 42.3958 | 12 | 6.4134 |  | 1.6 | 99.9109 |  | 103 | 41711.332
131 | 4.6214 | -4.9565 | 33612.2227 | 0.51289 | 1.0719 | 49.1619 | 12.4 | 5.7384 |  | 2 | 100.0918 |  | 104 | 41635.5547
132 | 4.6567 | -4.9213 | 32413.666 | 0.4946 | 1.2396 | 56.8533 | 12.8 | 5.2678 |  | 2.4 | 100.4326 |  | 105 | 41632.5547
133 | 4.6919 | -4.886 | 31253.2227 | 0.47689 | 1.4133 | 64.8182 | 13.2 | 4.7581 |  | 2.8 | 100.8204 |  | 106 | 41553.668
134 | 4.7272 | -4.8507 | 30224.7773 | 0.4612 | 1.5775 | 72.347 | 13.6 | 4.3577 |  | 3.2 | 101.0339 |  | 107 | 41584.5547
135 | 4.7625 | -4.8154 | 29501.666 | 0.45017 | 1.6992 | 77.929 | 14 | 3.9579 |  | 3.6 | 100.9378 |  | 108 | 41538.2227
136 | 4.7978 | -4.7801 | 28880.666 | 0.44069 | 1.8082 | 82.9271 | 14.4 | 3.656 |  | 4 | 100.0377 |  | 109 | 41476.8906
137 | 4.8331 | -4.7449 | 28435.1113 | 0.43389 | 1.889 | 86.6361 | 14.8 | 3.3305 |  | 4.4 | 95.9061 |  | 110 | 41367.8906
138 | 4.8683 | -4.7096 | 28135.7773 | 0.42932 | 1.9447 | 89.1882 | 15.2 | 3.1023 |  | 4.8 | 77.3493 |  | 111 | 41391.668
139 | 4.9036 | -4.6743 | 27919 | 0.42602 | 1.9857 | 91.0677 | 15.6 | 2.8256 |  | 5.2 | 33.8011 |  | 112 | 41281.332
140 | 4.9389 | -4.639 | 27680 | 0.42237 | 2.0315 | 93.1709 | 16 | 2.5966 |  | 5.6 | 11.7398 |  | 113 | 41298.4453
141 | 4.9742 | -4.6038 | 27465.334 | 0.41909 | 2.0733 | 95.0885 | 16.4 | 2.4349 |  | 6 | 7.8283 |  | 114 | 41190.668
142 | 5.0094 | -4.5685 | 27304.8887 | 0.41665 | 2.105 | 96.5396 |  |  |  | 6.4 | 6.2151 |  | 115 | 41089.1094
143 | 5.0447 | -4.5332 | 27209.5547 | 0.41519 | 2.1239 | 97.4093 |  |  |  | 6.8 | 5.2689 |  | 116 | 41029.1094
144 | 5.08 | -4.4979 | 27084.4453 | 0.41328 | 2.149 | 98.559 |  |  |  | 7.2 | 4.6441 |  | 117 | 40932.5547
145 | 5.1153 | -4.4626 | 26980.1113 | 0.41169 | 2.1701 | 99.5252 |  |  |  | 7.6 | 4.0928 |  | 118 | 40715.1094
146 | 5.1506 | -4.4274 | 26927.8887 | 0.41089 | 2.1807 | 100.0114 |  |  |  | 8 | 3.6586 |  | 119 | 40565.332
147 | 5.1858 | -4.3921 | 26780 | 0.40864 | 2.2109 | 101.3976 |  |  |  | 8.4 | 3.2977 |  | 120 | 40401.5547
148 | 5.2211 | -4.3568 | 26719.666 | 0.40772 | 2.2233 | 101.9671 |  |  |  | 8.8 | 2.9718 |  | 121 | 40218.8906
149 | 5.2564 | -4.3215 | 26714.666 | 0.40764 | 2.2243 | 102.0144 |  |  |  | 9.2 | 2.6762 |  | 122 | 40016.332
150 | 5.2917 | -4.2863 | 26713 | 0.40761 | 2.2247 | 102.0302 |  |  |  | 9.6 | 2.4238 |  | 123 | 39645.4453
151 | 5.3269 | -4.251 | 26706.666 | 0.40752 | 2.226 | 102.0901 |  |  |  | 10 | 2.2345 |  | 124 | 39222.668
152 | 5.3622 | -4.2157 | 26668.8887 | 0.40694 | 2.2338 | 102.4482 |  |  |  | 10.4 | 2.0553 |  | 125 | 38733.7773
153 | 5.3975 | -4.1804 | 26596.2227 | 0.40583 | 2.2489 | 103.1397 |  |  |  | 10.8 | 1.8666 |  | 126 | 38095.668
154 | 5.4328 | -4.1451 | 26638.4453 | 0.40648 | 2.2401 | 102.7375 |  |  |  | 11.2 | 1.7822 |  | 127 | 37395.332
155 | 5.4681 | -4.1099 | 26627.5547 | 0.40631 | 2.2424 | 102.8411 |  |  |  | 11.6 | 1.6177 |  | 128 | 36695.5547
156 | 5.5033 | -4.0746 | 26709.334 | 0.40756 | 2.2254 | 102.0649 |  |  |  | 12 | 1.4997 |  | 129 | 35769.332
157 | 5.5386 | -4.0393 | 26694.2227 | 0.40733 | 2.2286 | 102.208 |  |  |  | 12.4 | 1.3855 |  | 130 | 34736.4453
158 | 5.5739 | -4.004 | 26786.5547 | 0.40874 | 2.2095 | 101.3359 |  |  |  | 12.8 | 1.3324 |  | 131 | 33612.2227
159 | 5.6092 | -3.9688 | 26687.334 | 0.40722 | 2.23 | 102.2733 |  |  |  | 13.2 | 1.2369 |  | 132 | 32413.666
160 | 5.6444 | -3.9335 | 26681.1113 | 0.40713 | 2.2313 | 102.3323 |  |  |  | 13.6 | 1.1886 |  | 133 | 31253.2227
161 | 5.6797 | -3.8982 | 26656.2227 | 0.40675 | 2.2364 | 102.5685 |  |  |  | 14 | 1.0981 |  | 134 | 30224.7773
162 | 5.715 | -3.8629 | 26640.4453 | 0.40651 | 2.2397 | 102.7185 |  |  |  | 14.4 | 1.047 |  | 135 | 29501.666
163 | 5.7503 | -3.8276 | 26649.1113 | 0.40664 | 2.2379 | 102.6361 |  |  |  | 14.8 | 0.98912 |  | 136 | 28880.666
164 | 5.7856 | -3.7924 | 26647.8887 | 0.40662 | 2.2382 | 102.6477 |  |  |  | 15.2 | 0.93402 |  | 137 | 28435.1113
165 | 5.8208 | -3.7571 | 26665.1113 | 0.40688 | 2.2346 | 102.4841 |  |  |  | 15.6 | 0.87866 |  | 138 | 28135.7773
166 | 5.8561 | -3.7218 | 26682.4453 | 0.40715 | 2.231 | 102.3196 |  |  |  | 16 | 0.82943 |  | 139 | 27919
167 | 5.8914 | -3.6865 | 26699.8887 | 0.40741 | 2.2274 | 102.1543 |  |  |  | 16.4 | 0.77833 |  | 140 | 27680
168 | 5.9267 | -3.6513 | 26665 | 0.40688 | 2.2346 | 102.4852 | [6 empty] | 141 | 27465.334
169 | 5.9619 | -3.616 | 26625.5547 | 0.40628 | 2.2428 | 102.8602 | [6 empty] | 142 | 27304.8887
170 | 5.9972 | -3.5807 | 26591.5547 | 0.40576 | 2.2498 | 103.1842 | [6 empty] | 143 | 27209.5547
171 | 6.0325 | -3.5454 | 26576.8887 | 0.40554 | 2.2529 | 103.3242 | [6 empty] | 144 | 27084.4453
172 | 6.0678 | -3.5101 | 26566.5547 | 0.40538 | 2.2551 | 103.423 | [6 empty] | 145 | 26980.1113
173 | 6.1031 | -3.4749 | 26604.666 | 0.40596 | 2.2471 | 103.0592 | [6 empty] | 146 | 26927.8887
174 | 6.1383 | -3.4396 | 26584.334 | 0.40565 | 2.2514 | 103.2531 | [6 empty] | 147 | 26780
175 | 6.1736 | -3.4043 | 26564.5547 | 0.40535 | 2.2555 | 103.4421 | [6 empty] | 148 | 26719.666
176 | 6.2089 | -3.369 | 26549.2227 | 0.40512 | 2.2587 | 103.5887 | [6 empty] | 149 | 26714.666
177 | 6.2442 | -3.3338 | 26565.666 | 0.40537 | 2.2552 | 103.4315 | [6 empty] | 150 | 26713
178 | 6.2794 | -3.2985 | 26618.666 | 0.40617 | 2.2442 | 102.9258 | [6 empty] | 151 | 26706.666
179 | 6.3147 | -3.2632 | 26583.4453 | 0.40564 | 2.2515 | 103.2616 | [6 empty] | 152 | 26668.8887
180 | 6.35 | -3.2279 | 26532.5547 | 0.40486 | 2.2621 | 103.7483 | [6 empty] | 153 | 26596.2227
181 | 6.3853 | -3.1926 | 26622.334 | 0.40623 | 2.2435 | 102.8908 | [6 empty] | 154 | 26638.4453
182 | 6.4206 | -3.1574 | 26531.4453 | 0.40484 | 2.2624 | 103.759 | [6 empty] | 155 | 26627.5547
183 | 6.4558 | -3.1221 | 26559 | 0.40526 | 2.2566 | 103.4952 | [6 empty] | 156 | 26709.334
184 | 6.4911 | -3.0868 | 26538.5547 | 0.40495 | 2.2609 | 103.6909 | [6 empty] | 157 | 26694.2227
185 | 6.5264 | -3.0515 | 26433.7773 | 0.40335 | 2.2829 | 104.6979 | [6 empty] | 158 | 26786.5547
186 | 6.5617 | -3.0163 | 26462.666 | 0.40379 | 2.2768 | 104.4195 | [6 empty] | 159 | 26687.334
187 | 6.5969 | -2.981 | 26589.2227 | 0.40573 | 2.2503 | 103.2065 | [6 empty] | 160 | 26681.1113
188 | 6.6322 | -2.9457 | 26577.8887 | 0.40555 | 2.2527 | 103.3147 | [6 empty] | 161 | 26656.2227
189 | 6.6675 | -2.9104 | 26572.5547 | 0.40547 | 2.2538 | 103.3656 | [6 empty] | 162 | 26640.4453
190 | 6.7028 | -2.8751 | 26589.2227 | 0.40573 | 2.2503 | 103.2065 | [6 empty] | 163 | 26649.1113
191 | 6.7381 | -2.8399 | 26549.334 | 0.40512 | 2.2586 | 103.5877 | [6 empty] | 164 | 26647.8887
192 | 6.7733 | -2.8046 | 26529 | 0.40481 | 2.2629 | 103.7824 | [6 empty] | 165 | 26665.1113
193 | 6.8086 | -2.7693 | 26351.5547 | 0.4021 | 2.3002 | 105.4933 | [6 empty] | 166 | 26682.4453
194 | 6.8439 | -2.734 | 26295.334 | 0.40124 | 2.3121 | 106.0397 | [6 empty] | 167 | 26699.8887
195 | 6.8792 | -2.6988 | 26454.334 | 0.40367 | 2.2785 | 104.4998 | [6 empty] | 168 | 26665
196 | 6.9144 | -2.6635 | 26483.7773 | 0.40412 | 2.2724 | 104.2164 | [6 empty] | 169 | 26625.5547
197 | 6.9497 | -2.6282 | 26495.7773 | 0.4043 | 2.2698 | 104.1011 | [6 empty] | 170 | 26591.5547
198 | 6.985 | -2.5929 | 26535.2227 | 0.4049 | 2.2616 | 103.7228 | [6 empty] | 171 | 26576.8887
199 | 7.0203 | -2.5576 | 26574 | 0.40549 | 2.2535 | 103.3518 | [6 empty] | 172 | 26566.5547
200 | 7.0556 | -2.5224 | 26624.5547 | 0.40626 | 2.243 | 102.8697 | [6 empty] | 173 | 26604.666
201 | 7.0908 | -2.4871 | 26603.4453 | 0.40594 | 2.2474 | 103.0708 | [6 empty] | 174 | 26584.334
202 | 7.1261 | -2.4518 | 26555 | 0.4052 | 2.2575 | 103.5335 | [6 empty] | 175 | 26564.5547
203 | 7.1614 | -2.4165 | 26472.4453 | 0.40394 | 2.2747 | 104.3254 | [6 empty] | 176 | 26549.2227
204 | 7.1967 | -2.3813 | 26576.7773 | 0.40554 | 2.2529 | 103.3253 | [6 empty] | 177 | 26565.666
205 | 7.2319 | -2.346 | 26695.334 | 0.40734 | 2.2283 | 102.1975 | [6 empty] | 178 | 26618.666
206 | 7.2672 | -2.3107 | 26718.334 | 0.4077 | 2.2236 | 101.9797 | [6 empty] | 179 | 26583.4453
207 | 7.3025 | -2.2754 | 26645.4453 | 0.40658 | 2.2387 | 102.6709 | [6 empty] | 180 | 26532.5547
208 | 7.3378 | -2.2401 | 26764.8887 | 0.40841 | 2.214 | 101.54 | [6 empty] | 181 | 26622.334
209 | 7.3731 | -2.2049 | 26767.2227 | 0.40844 | 2.2135 | 101.518 | [6 empty] | 182 | 26531.4453
210 | 7.4083 | -2.1696 | 26769.666 | 0.40848 | 2.213 | 101.495 | [6 empty] | 183 | 26559
211 | 7.4436 | -2.1343 | 26756.334 | 0.40828 | 2.2158 | 101.6207 | [6 empty] | 184 | 26538.5547
212 | 7.4789 | -2.099 | 26765.1113 | 0.40841 | 2.214 | 101.5379 | [6 empty] | 185 | 26433.7773
213 | 7.5142 | -2.0638 | 26823.8887 | 0.40931 | 2.2019 | 100.9848 | [6 empty] | 186 | 26462.666
214 | 7.5494 | -2.0285 | 26846.7773 | 0.40966 | 2.1972 | 100.77 | [6 empty] | 187 | 26589.2227
215 | 7.5847 | -1.9932 | 26858.334 | 0.40983 | 2.1948 | 100.6616 | [6 empty] | 188 | 26577.8887
216 | 7.62 | -1.9579 | 26895.334 | 0.4104 | 2.1873 | 100.3154 | [6 empty] | 189 | 26572.5547
217 | 7.6553 | -1.9226 | 26902.8887 | 0.41051 | 2.1858 | 100.2448 | [6 empty] | 190 | 26589.2227
218 | 7.6906 | -1.8874 | 26882.7773 | 0.4102 | 2.1899 | 100.4328 | [6 empty] | 191 | 26549.334
219 | 7.7258 | -1.8521 | 26870 | 0.41001 | 2.1925 | 100.5523 | [6 empty] | 192 | 26529
220 | 7.7611 | -1.8168 | 26854.8887 | 0.40978 | 2.1955 | 100.6939 | [6 empty] | 193 | 26351.5547
221 | 7.7964 | -1.7815 | 26875.666 | 0.4101 | 2.1913 | 100.4993 | [6 empty] | 194 | 26295.334
222 | 7.8317 | -1.7463 | 26898.5547 | 0.41045 | 2.1866 | 100.2853 | [6 empty] | 195 | 26454.334
223 | 7.8669 | -1.711 | 26898.1113 | 0.41044 | 2.1867 | 100.2894 | [6 empty] | 196 | 26483.7773
224 | 7.9022 | -1.6757 | 26921.334 | 0.41079 | 2.182 | 100.0725 | [6 empty] | 197 | 26495.7773
225 | 7.9375 | -1.6404 | 26967.666 | 0.4115 | 2.1726 | 99.6409 | [6 empty] | 198 | 26535.2227
226 | 7.9728 | -1.6051 | 26985.1113 | 0.41177 | 2.1691 | 99.4788 | [6 empty] | 199 | 26574
227 | 8.0081 | -1.5699 | 27013.7773 | 0.4122 | 2.1633 | 99.2127 | [6 empty] | 200 | 26624.5547
228 | 8.0433 | -1.5346 | 26882.8887 | 0.41021 | 2.1898 | 100.4317 | [6 empty] | 201 | 26603.4453
229 | 8.0786 | -1.4993 | 26812.7773 | 0.40914 | 2.2042 | 101.0892 | [6 empty] | 202 | 26555
230 | 8.1139 | -1.464 | 26834.8887 | 0.40947 | 2.1996 | 100.8815 | [6 empty] | 203 | 26472.4453
231 | 8.1492 | -1.4288 | 26837.5547 | 0.40951 | 2.1991 | 100.8565 | [6 empty] | 204 | 26576.7773
232 | 8.1844 | -1.3935 | 26844.2227 | 0.40962 | 2.1977 | 100.7939 | [6 empty] | 205 | 26695.334
233 | 8.2197 | -1.3582 | 26907.8887 | 0.41059 | 2.1847 | 100.1981 | [6 empty] | 206 | 26718.334
234 | 8.255 | -1.3229 | 26870.2227 | 0.41001 | 2.1924 | 100.5503 | [6 empty] | 207 | 26645.4453
235 | 8.2903 | -1.2876 | 26823.666 | 0.4093 | 2.2019 | 100.9869 | [6 empty] | 208 | 26764.8887
236 | 8.3256 | -1.2524 | 26867.1113 | 0.40997 | 2.1931 | 100.5794 | [6 empty] | 209 | 26767.2227
237 | 8.3608 | -1.2171 | 26800.334 | 0.40895 | 2.2067 | 101.2062 | [6 empty] | 210 | 26769.666
238 | 8.3961 | -1.1818 | 26777.334 | 0.4086 | 2.2114 | 101.4227 | [6 empty] | 211 | 26756.334
239 | 8.4314 | -1.1465 | 26801.8887 | 0.40897 | 2.2064 | 101.1915 | [6 empty] | 212 | 26765.1113
240 | 8.4667 | -1.1113 | 26821.1113 | 0.40926 | 2.2025 | 101.0108 | [6 empty] | 213 | 26823.8887
241 | 8.5019 | -1.076 | 26825.4453 | 0.40933 | 2.2016 | 100.9701 | [6 empty] | 214 | 26846.7773
242 | 8.5372 | -1.0407 | 26803 | 0.40899 | 2.2062 | 101.1811 | [6 empty] | 215 | 26858.334
243 | 8.5725 | -1.0054 | 26797 | 0.4089 | 2.2074 | 101.2375 | [6 empty] | 216 | 26895.334
244 | 8.6078 | -0.97014 | 26846.666 | 0.40965 | 2.1972 | 100.771 | [6 empty] | 217 | 26902.8887
245 | 8.6431 | -0.93486 | 26917.7773 | 0.41074 | 2.1827 | 100.1057 | [6 empty] | 218 | 26882.7773
246 | 8.6783 | -0.89958 | 26940.2227 | 0.41108 | 2.1782 | 99.8964 | [6 empty] | 219 | 26870
247 | 8.7136 | -0.86431 | 26965 | 0.41146 | 2.1731 | 99.6657 | [6 empty] | 220 | 26854.8887
248 | 8.7489 | -0.82903 | 26842.4453 | 0.40959 | 2.1981 | 100.8106 | [6 empty] | 221 | 26875.666
249 | 8.7842 | -0.79375 | 26762.2227 | 0.40837 | 2.2145 | 101.5652 | [6 empty] | 222 | 26898.5547
250 | 8.8194 | -0.75847 | 26832 | 0.40943 | 2.2002 | 100.9086 | [6 empty] | 223 | 26898.1113
251 | 8.8547 | -0.72319 | 26797.666 | 0.40891 | 2.2073 | 101.2313 | [6 empty] | 224 | 26921.334
252 | 8.89 | -0.68792 | 26840.1113 | 0.40955 | 2.1986 | 100.8325 | [6 empty] | 225 | 26967.666
253 | 8.9253 | -0.65264 | 26806 | 0.40903 | 2.2056 | 101.1529 | [6 empty] | 226 | 26985.1113
254 | 8.9606 | -0.61736 | 26799.2227 | 0.40893 | 2.2069 | 101.2166 | [6 empty] | 227 | 27013.7773
255 | 8.9958 | -0.58208 | 26887.334 | 0.41027 | 2.1889 | 100.3901 | [6 empty] | 228 | 26882.8887
256 | 9.0311 | -0.54681 | 26914.7773 | 0.41069 | 2.1833 | 100.1337 | [6 empty] | 229 | 26812.7773
257 | 9.0664 | -0.51153 | 26950.666 | 0.41124 | 2.176 | 99.7991 | [6 empty] | 230 | 26834.8887
258 | 9.1017 | -0.47625 | 26874 | 0.41007 | 2.1916 | 100.5149 | [6 empty] | 231 | 26837.5547
259 | 9.1369 | -0.44097 | 26903 | 0.41051 | 2.1857 | 100.2437 | [6 empty] | 232 | 26844.2227
260 | 9.1722 | -0.40569 | 26936.2227 | 0.41102 | 2.179 | 99.9337 | [6 empty] | 233 | 26907.8887
261 | 9.2075 | -0.37042 | 26895.1113 | 0.41039 | 2.1873 | 100.3174 | [6 empty] | 234 | 26870.2227
262 | 9.2428 | -0.33514 | 26917 | 0.41073 | 2.1829 | 100.113 | [6 empty] | 235 | 26823.666
263 | 9.2781 | -0.29986 | 26890.4453 | 0.41032 | 2.1883 | 100.3611 | [6 empty] | 236 | 26867.1113
264 | 9.3133 | -0.26458 | 26911.5547 | 0.41064 | 2.184 | 100.1638 | [6 empty] | 237 | 26800.334
265 | 9.3486 | -0.22931 | 26959.334 | 0.41137 | 2.1743 | 99.7184 | [6 empty] | 238 | 26777.334
266 | 9.3839 | -0.19403 | 26946.5547 | 0.41118 | 2.1769 | 99.8374 | [6 empty] | 239 | 26801.8887
267 | 9.4192 | -0.15875 | 26897.2227 | 0.41043 | 2.1869 | 100.2977 | [6 empty] | 240 | 26821.1113
268 | 9.4544 | -0.12347 | 26858.334 | 0.40983 | 2.1948 | 100.6616 | [6 empty] | 241 | 26825.4453
269 | 9.4897 | -0.088194 | 26880.5547 | 0.41017 | 2.1903 | 100.4536 | [6 empty] | 242 | 26803
270 | 9.525 | -0.052917 | 26939 | 0.41106 | 2.1784 | 99.9078 | [6 empty] | 243 | 26797
271 | 9.5603 | -0.017639 | 26929.1113 | 0.41091 | 2.1804 | 100 | [6 empty] | 244 | 26846.666
272 | 9.5956 | 0.017639 | 26900.2227 | 0.41047 | 2.1863 | 100.2697 | [6 empty] | 245 | 26917.7773
273 | 9.6308 | 0.052917 | 26875 | 0.41009 | 2.1914 | 100.5055 | [6 empty] | 246 | 26940.2227
274 | 9.6661 | 0.088194 | 26894.4453 | 0.41038 | 2.1875 | 100.3237 | [6 empty] | 247 | 26965
275 | 9.7014 | 0.12347 | 26882.7773 | 0.4102 | 2.1899 | 100.4328 | [6 empty] | 248 | 26842.4453
276 | 9.7367 | 0.15875 | 26909.1113 | 0.41061 | 2.1845 | 100.1866 | [6 empty] | 249 | 26762.2227
277 | 9.7719 | 0.19403 | 27016.666 | 0.41225 | 2.1627 | 99.1859 | [6 empty] | 250 | 26832
278 | 9.8072 | 0.22931 | 26987.8887 | 0.41181 | 2.1685 | 99.453 | [6 empty] | 251 | 26797.666
279 | 9.8425 | 0.26458 | 27019.8887 | 0.4123 | 2.162 | 99.156 | [6 empty] | 252 | 26840.1113
280 | 9.8778 | 0.29986 | 26994.2227 | 0.41191 | 2.1672 | 99.3941 | [6 empty] | 253 | 26806
281 | 9.9131 | 0.33514 | 27053.5547 | 0.41281 | 2.1552 | 98.8444 | [6 empty] | 254 | 26799.2227
282 | 9.9483 | 0.37042 | 27044.5547 | 0.41267 | 2.157 | 98.9276 | [6 empty] | 255 | 26887.334
283 | 9.9836 | 0.40569 | 27049.334 | 0.41275 | 2.1561 | 98.8834 | [6 empty] | 256 | 26914.7773
284 | 10.0189 | 0.44097 | 27075.1113 | 0.41314 | 2.1509 | 98.6452 | [6 empty] | 257 | 26950.666
285 | 10.0542 | 0.47625 | 27108.5547 | 0.41365 | 2.1442 | 98.3367 | [6 empty] | 258 | 26874
286 | 10.0894 | 0.51153 | 27078.8887 | 0.4132 | 2.1501 | 98.6103 | [6 empty] | 259 | 26903
287 | 10.1247 | 0.54681 | 27019.1113 | 0.41229 | 2.1622 | 99.1632 | [6 empty] | 260 | 26936.2227
288 | 10.16 | 0.58208 | 27030.5547 | 0.41246 | 2.1599 | 99.0572 | [6 empty] | 261 | 26895.1113
289 | 10.1953 | 0.61736 | 27082.334 | 0.41325 | 2.1494 | 98.5785 | [6 empty] | 262 | 26917
290 | 10.2306 | 0.65264 | 27075.2227 | 0.41314 | 2.1509 | 98.6441 | [6 empty] | 263 | 26890.4453
291 | 10.2658 | 0.68792 | 27037 | 0.41256 | 2.1586 | 98.9975 | [6 empty] | 264 | 26911.5547
292 | 10.3011 | 0.72319 | 27021.5547 | 0.41232 | 2.1617 | 99.1406 | [6 empty] | 265 | 26959.334
293 | 10.3364 | 0.75847 | 27096 | 0.41346 | 2.1467 | 98.4524 | [6 empty] | 266 | 26946.5547
294 | 10.3717 | 0.79375 | 27070.334 | 0.41307 | 2.1518 | 98.6893 | [6 empty] | 267 | 26897.2227
295 | 10.4069 | 0.82903 | 27031.334 | 0.41247 | 2.1597 | 99.05 | [6 empty] | 268 | 26858.334
296 | 10.4422 | 0.86431 | 27132.666 | 0.41402 | 2.1393 | 98.1147 | [6 empty] | 269 | 26880.5547
297 | 10.4775 | 0.89958 | 27183.666 | 0.4148 | 2.1291 | 97.6464 | [6 empty] | 270 | 26939
298 | 10.5128 | 0.93486 | 27183.1113 | 0.41479 | 2.1292 | 97.6515 | [6 empty] | 271 | 26929.1113
299 | 10.5481 | 0.97014 | 27152.5547 | 0.41432 | 2.1353 | 97.9319 | [6 empty] | 272 | 26900.2227
300 | 10.5833 | 1.0054 | 27090.7773 | 0.41338 | 2.1477 | 98.5006 | [6 empty] | 273 | 26875
301 | 10.6186 | 1.0407 | 27158.5547 | 0.41441 | 2.1341 | 97.8768 | [6 empty] | 274 | 26894.4453
302 | 10.6539 | 1.076 | 27063.666 | 0.41297 | 2.1532 | 98.7509 | [6 empty] | 275 | 26882.7773
303 | 10.6892 | 1.1113 | 27092.2227 | 0.4134 | 2.1474 | 98.4872 | [6 empty] | 276 | 26909.1113
304 | 10.7244 | 1.1465 | 27087.5547 | 0.41333 | 2.1484 | 98.5303 | [6 empty] | 277 | 27016.666
305 | 10.7597 | 1.1818 | 27026.4453 | 0.4124 | 2.1607 | 99.0953 | [6 empty] | 278 | 26987.8887
306 | 10.795 | 1.2171 | 27065.2227 | 0.41299 | 2.1529 | 98.7365 | [6 empty] | 279 | 27019.8887
307 | 10.8303 | 1.2524 | 27044.7773 | 0.41268 | 2.157 | 98.9256 | [6 empty] | 280 | 26994.2227
308 | 10.8656 | 1.2876 | 27013 | 0.41219 | 2.1634 | 99.2199 | [6 empty] | 281 | 27053.5547
309 | 10.9008 | 1.3229 | 26993.7773 | 0.4119 | 2.1673 | 99.3983 | [6 empty] | 282 | 27044.5547
310 | 10.9361 | 1.3582 | 27054 | 0.41282 | 2.1551 | 98.8402 | [6 empty] | 283 | 27049.334
311 | 10.9714 | 1.3935 | 27039.666 | 0.4126 | 2.158 | 98.9729 | [6 empty] | 284 | 27075.1113
312 | 11.0067 | 1.4287 | 27003.5547 | 0.41205 | 2.1653 | 99.3075 | [6 empty] | 285 | 27108.5547
313 | 11.0419 | 1.464 | 26934.334 | 0.41099 | 2.1794 | 99.9513 | [6 empty] | 286 | 27078.8887
314 | 11.0772 | 1.4993 | 26941.2227 | 0.4111 | 2.178 | 99.8871 | [6 empty] | 287 | 27019.1113
315 | 11.1125 | 1.5346 | 26922.8887 | 0.41082 | 2.1817 | 100.058 | [6 empty] | 288 | 27030.5547
316 | 11.1478 | 1.5699 | 26862.2227 | 0.40989 | 2.1941 | 100.6252 | [6 empty] | 289 | 27082.334
317 | 11.1831 | 1.6051 | 26905 | 0.41054 | 2.1853 | 100.225 | [6 empty] | 290 | 27075.2227
318 | 11.2183 | 1.6404 | 26943.8887 | 0.41114 | 2.1774 | 99.8623 | [6 empty] | 291 | 27037
319 | 11.2536 | 1.6757 | 26938.7773 | 0.41106 | 2.1785 | 99.9099 | [6 empty] | 292 | 27021.5547
320 | 11.2889 | 1.711 | 26952 | 0.41126 | 2.1758 | 99.7867 | [6 empty] | 293 | 27096
321 | 11.3242 | 1.7462 | 26950.5547 | 0.41124 | 2.1761 | 99.8002 | [6 empty] | 294 | 27070.334
322 | 11.3594 | 1.7815 | 26906.2227 | 0.41056 | 2.1851 | 100.2136 | [6 empty] | 295 | 27031.334
323 | 11.3947 | 1.8168 | 26864.666 | 0.40993 | 2.1936 | 100.6023 | [6 empty] | 296 | 27132.666
324 | 11.43 | 1.8521 | 26925.666 | 0.41086 | 2.1811 | 100.0321 | [6 empty] | 297 | 27183.666
325 | 11.4653 | 1.8874 | 26922.334 | 0.41081 | 2.1818 | 100.0632 | [6 empty] | 298 | 27183.1113
326 | 11.5006 | 1.9226 | 26891.7773 | 0.41034 | 2.188 | 100.3486 | [6 empty] | 299 | 27152.5547
327 | 11.5358 | 1.9579 | 26919.2227 | 0.41076 | 2.1824 | 100.0922 | [6 empty] | 300 | 27090.7773
328 | 11.5711 | 1.9932 | 26890.666 | 0.41033 | 2.1882 | 100.359 | [6 empty] | 301 | 27158.5547
329 | 11.6064 | 2.0285 | 26897 | 0.41042 | 2.187 | 100.2998 | [6 empty] | 302 | 27063.666
330 | 11.6417 | 2.0637 | 26926 | 0.41086 | 2.1811 | 100.029 | [6 empty] | 303 | 27092.2227
331 | 11.6769 | 2.099 | 26902.334 | 0.4105 | 2.1859 | 100.2499 | [6 empty] | 304 | 27087.5547
332 | 11.7122 | 2.1343 | 26951.1113 | 0.41125 | 2.1759 | 99.795 | [6 empty] | 305 | 27026.4453
333 | 11.7475 | 2.1696 | 26948.4453 | 0.41121 | 2.1765 | 99.8198 | [6 empty] | 306 | 27065.2227
334 | 11.7828 | 2.2049 | 26986.8887 | 0.41179 | 2.1687 | 99.4622 | [6 empty] | 307 | 27044.7773
335 | 11.8181 | 2.2401 | 26982 | 0.41172 | 2.1697 | 99.5077 | [6 empty] | 308 | 27013
336 | 11.8533 | 2.2754 | 27003.4453 | 0.41205 | 2.1653 | 99.3085 | [6 empty] | 309 | 26993.7773
337 | 11.8886 | 2.3107 | 27002.1113 | 0.41203 | 2.1656 | 99.3209 | [6 empty] | 310 | 27054
338 | 11.9239 | 2.346 | 26977.7773 | 0.41165 | 2.1705 | 99.5469 | [6 empty] | 311 | 27039.666
339 | 11.9592 | 2.3812 | 26967.8887 | 0.4115 | 2.1725 | 99.6389 | [6 empty] | 312 | 27003.5547
340 | 11.9944 | 2.4165 | 26993.7773 | 0.4119 | 2.1673 | 99.3983 | [6 empty] | 313 | 26934.334
341 | 12.0297 | 2.4518 | 27018.666 | 0.41228 | 2.1623 | 99.1674 | [6 empty] | 314 | 26941.2227
342 | 12.065 | 2.4871 | 26999 | 0.41198 | 2.1662 | 99.3498 | [6 empty] | 315 | 26922.8887
343 | 12.1003 | 2.5224 | 26984.8887 | 0.41176 | 2.1691 | 99.4808 | [6 empty] | 316 | 26862.2227
344 | 12.1356 | 2.5576 | 26995.334 | 0.41192 | 2.167 | 99.3838 | [6 empty] | 317 | 26905
345 | 12.1708 | 2.5929 | 26990.666 | 0.41185 | 2.1679 | 99.4272 | [6 empty] | 318 | 26943.8887
346 | 12.2061 | 2.6282 | 27059.5547 | 0.4129 | 2.154 | 98.7889 | [6 empty] | 319 | 26938.7773
347 | 12.2414 | 2.6635 | 27067.666 | 0.41303 | 2.1524 | 98.7139 | [6 empty] | 320 | 26952
348 | 12.2767 | 2.6987 | 27123.8887 | 0.41388 | 2.1411 | 98.1955 | [6 empty] | 321 | 26950.5547
349 | 12.3119 | 2.734 | 27086.5547 | 0.41331 | 2.1486 | 98.5395 | [6 empty] | 322 | 26906.2227
350 | 12.3472 | 2.7693 | 27118.7773 | 0.41381 | 2.1421 | 98.2425 | [6 empty] | 323 | 26864.666
351 | 12.3825 | 2.8046 | 27201 | 0.41506 | 2.1256 | 97.4876 | [6 empty] | 324 | 26925.666
352 | 12.4178 | 2.8399 | 27178.4453 | 0.41472 | 2.1301 | 97.6943 | [6 empty] | 325 | 26922.334
353 | 12.4531 | 2.8751 | 27080.1113 | 0.41322 | 2.1499 | 98.599 | [6 empty] | 326 | 26891.7773
354 | 12.4883 | 2.9104 | 27121.1113 | 0.41384 | 2.1416 | 98.221 | [6 empty] | 327 | 26919.2227
355 | 12.5236 | 2.9457 | 27095.7773 | 0.41346 | 2.1467 | 98.4545 | [6 empty] | 328 | 26890.666
356 | 12.5589 | 2.981 | 27136.7773 | 0.41408 | 2.1385 | 98.0769 | [6 empty] | 329 | 26897
357 | 12.5942 | 3.0162 | 27106.666 | 0.41362 | 2.1445 | 98.3541 | [6 empty] | 330 | 26926
358 | 12.6294 | 3.0515 | 27037.666 | 0.41257 | 2.1584 | 98.9914 | [6 empty] | 331 | 26902.334
359 | 12.6647 | 3.0868 | 26986.2227 | 0.41178 | 2.1688 | 99.4684 | [6 empty] | 332 | 26951.1113
360 | 12.7 | 3.1221 | 26984.7773 | 0.41176 | 2.1691 | 99.4819 | [6 empty] | 333 | 26948.4453
361 | 12.7353 | 3.1574 | 26950.334 | 0.41124 | 2.1761 | 99.8022 | [6 empty] | 334 | 26986.8887
362 | 12.7706 | 3.1926 | 26960.5547 | 0.41139 | 2.174 | 99.7071 | [6 empty] | 335 | 26982
363 | 12.8058 | 3.2279 | 26962.2227 | 0.41142 | 2.1737 | 99.6916 | [6 empty] | 336 | 27003.4453
364 | 12.8411 | 3.2632 | 26976.4453 | 0.41163 | 2.1708 | 99.5593 | [6 empty] | 337 | 27002.1113
365 | 12.8764 | 3.2985 | 26993.8887 | 0.4119 | 2.1673 | 99.3972 | [6 empty] | 338 | 26977.7773
366 | 12.9117 | 3.3337 | 27019.666 | 0.41229 | 2.1621 | 99.1581 | [6 empty] | 339 | 26967.8887
367 | 12.9469 | 3.369 | 26990.8887 | 0.41185 | 2.1679 | 99.4251 | [6 empty] | 340 | 26993.7773
368 | 12.9822 | 3.4043 | 27015.5547 | 0.41223 | 2.1629 | 99.1962 | [6 empty] | 341 | 27018.666
369 | 13.0175 | 3.4396 | 26949.7773 | 0.41123 | 2.1762 | 99.8074 | [6 empty] | 342 | 26999
370 | 13.0528 | 3.4749 | 26930.7773 | 0.41094 | 2.1801 | 99.9845 | [6 empty] | 343 | 26984.8887
371 | 13.0881 | 3.5101 | 26958.5547 | 0.41136 | 2.1744 | 99.7257 | [6 empty] | 344 | 26995.334
372 | 13.1233 | 3.5454 | 26918.4453 | 0.41075 | 2.1826 | 100.0995 | [6 empty] | 345 | 26990.666
373 | 13.1586 | 3.5807 | 26799.8887 | 0.40894 | 2.2068 | 101.2104 | [6 empty] | 346 | 27059.5547
374 | 13.1939 | 3.616 | 26856.334 | 0.4098 | 2.1953 | 100.6804 | [6 empty] | 347 | 27067.666
375 | 13.2292 | 3.6512 | 26903.5547 | 0.41052 | 2.1856 | 100.2385 | [6 empty] | 348 | 27123.8887
376 | 13.2644 | 3.6865 | 26996.7773 | 0.41194 | 2.1667 | 99.3704 | [6 empty] | 349 | 27086.5547
377 | 13.2997 | 3.7218 | 26976 | 0.41163 | 2.1709 | 99.5634 | [6 empty] | 350 | 27118.7773
378 | 13.335 | 3.7571 | 26922 | 0.4108 | 2.1819 | 100.0663 | [6 empty] | 351 | 27201
379 | 13.3703 | 3.7924 | 26898.4453 | 0.41044 | 2.1867 | 100.2863 | [6 empty] | 352 | 27178.4453
380 | 13.4056 | 3.8276 | 26842.8887 | 0.4096 | 2.198 | 100.8064 | [6 empty] | 353 | 27080.1113
381 | 13.4408 | 3.8629 | 26854.334 | 0.40977 | 2.1957 | 100.6991 | [6 empty] | 354 | 27121.1113
382 | 13.4761 | 3.8982 | 26832 | 0.40943 | 2.2002 | 100.9086 | [6 empty] | 355 | 27095.7773
383 | 13.5114 | 3.9335 | 26834.7773 | 0.40947 | 2.1997 | 100.8825 | [6 empty] | 356 | 27136.7773
384 | 13.5467 | 3.9688 | 26870.5547 | 0.41002 | 2.1923 | 100.5472 | [6 empty] | 357 | 27106.666
385 | 13.5819 | 4.004 | 26882.5547 | 0.4102 | 2.1899 | 100.4348 | [6 empty] | 358 | 27037.666
386 | 13.6172 | 4.0393 | 26856.666 | 0.40981 | 2.1952 | 100.6772 | [6 empty] | 359 | 26986.2227
387 | 13.6525 | 4.0746 | 26859.8887 | 0.40986 | 2.1945 | 100.6471 | [6 empty] | 360 | 26984.7773
388 | 13.6878 | 4.1099 | 26885.7773 | 0.41025 | 2.1892 | 100.4047 | [6 empty] | 361 | 26950.334
389 | 13.7231 | 4.1451 | 26901 | 0.41048 | 2.1861 | 100.2624 | [6 empty] | 362 | 26960.5547
390 | 13.7583 | 4.1804 | 26919 | 0.41076 | 2.1825 | 100.0943 | [6 empty] | 363 | 26962.2227
391 | 13.7936 | 4.2157 | 26932.666 | 0.41097 | 2.1797 | 99.9669 | [6 empty] | 364 | 26976.4453
392 | 13.8289 | 4.251 | 26986.7773 | 0.41179 | 2.1687 | 99.4633 | [6 empty] | 365 | 26993.8887
393 | 13.8642 | 4.2862 | 26991.8887 | 0.41187 | 2.1677 | 99.4158 | [6 empty] | 366 | 27019.666
394 | 13.8994 | 4.3215 | 27028.666 | 0.41243 | 2.1602 | 99.0747 | [6 empty] | 367 | 26990.8887
395 | 13.9347 | 4.3568 | 27136.4453 | 0.41408 | 2.1386 | 98.08 | [6 empty] | 368 | 27015.5547
396 | 13.97 | 4.3921 | 27092.4453 | 0.4134 | 2.1474 | 98.4852 | [6 empty] | 369 | 26949.7773
397 | 14.0053 | 4.4274 | 27097.2227 | 0.41348 | 2.1464 | 98.4411 | [6 empty] | 370 | 26930.7773
398 | 14.0406 | 4.4626 | 27111.334 | 0.41369 | 2.1436 | 98.3111 | [6 empty] | 371 | 26958.5547
399 | 14.0758 | 4.4979 | 27165 | 0.41451 | 2.1328 | 97.8176 | [6 empty] | 372 | 26918.4453
400 | 14.1111 | 4.5332 | 27201.1113 | 0.41506 | 2.1256 | 97.4866 | [6 empty] | 373 | 26799.8887
401 | 14.1464 | 4.5685 | 27258.666 | 0.41594 | 2.1141 | 96.9606 | [6 empty] | 374 | 26856.334
402 | 14.1817 | 4.6037 | 27320.334 | 0.41688 | 2.1019 | 96.3993 | [6 empty] | 375 | 26903.5547
403 | 14.2169 | 4.639 | 27480.1113 | 0.41932 | 2.0704 | 94.9556 | [6 empty] | 376 | 26996.7773
404 | 14.2522 | 4.6743 | 27624.4453 | 0.42152 | 2.0423 | 93.6646 | [6 empty] | 377 | 26976
405 | 14.2875 | 4.7096 | 27783 | 0.42394 | 2.0117 | 92.2605 | [6 empty] | 378 | 26922
406 | 14.3228 | 4.7449 | 27955.7773 | 0.42658 | 1.9787 | 90.7469 | [6 empty] | 379 | 26898.4453
407 | 14.3581 | 4.7801 | 28221.5547 | 0.43063 | 1.9286 | 88.4518 | [6 empty] | 380 | 26842.8887
408 | 14.3933 | 4.8154 | 28587.666 | 0.43622 | 1.8611 | 85.3543 | [6 empty] | 381 | 26854.334
409 | 14.4286 | 4.8507 | 29023.334 | 0.44287 | 1.7827 | 81.7615 | [6 empty] | 382 | 26832
410 | 14.4639 | 4.886 | 29663.1113 | 0.45263 | 1.6715 | 76.6611 | [6 empty] | 383 | 26834.7773
411 | 14.4992 | 4.9212 | 30507.4453 | 0.46551 | 1.5313 | 70.2312 | [6 empty] | 384 | 26870.5547
412 | 14.5344 | 4.9565 | 31471.8887 | 0.48023 | 1.3797 | 63.2759 | [6 empty] | 385 | 26882.5547
413 | 14.5697 | 4.9918 | 32605.4453 | 0.49753 | 1.212 | 55.5875 | [6 empty] | 386 | 26856.666
414 | 14.605 | 5.0271 | 33905.8906 | 0.51737 | 1.0325 | 47.3544 | [6 empty] | 387 | 26859.8887
415 | 14.6403 | 5.0624 | 35077.8906 | 0.53525 | 0.88134 | 40.4206 | [6 empty] | 388 | 26885.7773
416 | 14.6756 | 5.0976 | 36038.1094 | 0.54991 | 0.7643 | 35.0529 | [6 empty] | 389 | 26901
417 | 14.7108 | 5.1329 | 36858.8906 | 0.56243 | 0.66877 | 30.6715 | [6 empty] | 390 | 26919
418 | 14.7461 | 5.1682 | 37666.668 | 0.57476 | 0.57855 | 26.5337 | [6 empty] | 391 | 26932.666
419 | 14.7814 | 5.2035 | 38346.7773 | 0.58513 | 0.50534 | 23.1762 | [6 empty] | 392 | 26986.7773
420 | 14.8167 | 5.2387 | 38896.5547 | 0.59352 | 0.44791 | 20.5425 | [6 empty] | 393 | 26991.8887
421 | 14.8519 | 5.274 | 39354.8906 | 0.60052 | 0.40119 | 18.3996 | [6 empty] | 394 | 27028.666
422 | 14.8872 | 5.3093 | 39767.5547 | 0.60681 | 0.35998 | 16.5098 | [6 empty] | 395 | 27136.4453
423 | 14.9225 | 5.3446 | 40116.5547 | 0.61214 | 0.32575 | 14.94 | [6 empty] | 396 | 27092.4453
424 | 14.9578 | 5.3799 | 40397 | 0.61642 | 0.29865 | 13.697 | [6 empty] | 397 | 27097.2227
425 | 14.9931 | 5.4151 | 40668.1094 | 0.62056 | 0.27279 | 12.5107 | [6 empty] | 398 | 27111.334
426 | 15.0283 | 5.4504 | 41030.332 | 0.62608 | 0.23873 | 10.9487 | [6 empty] | 399 | 27165
427 | 15.0636 | 5.4857 | 41176.2227 | 0.62831 | 0.22517 | 10.3269 | [6 empty] | 400 | 27201.1113
428 | 15.0989 | 5.521 | 41210.4453 | 0.62883 | 0.222 | 10.1816 | [6 empty] | 401 | 27258.666
429 | 15.1342 | 5.5562 | 41336.668 | 0.63076 | 0.21036 | 9.6478 | [6 empty] | 402 | 27320.334
430 | 15.1694 | 5.5915 | 41413.8906 | 0.63194 | 0.20327 | 9.3227 | [6 empty] | 403 | 27480.1113
431 | 15.2047 | 5.6268 | 41495.2227 | 0.63318 | 0.19583 | 8.9815 | [6 empty] | 404 | 27624.4453
432 | 15.24 | 5.6621 | 41586.4453 | 0.63457 | 0.18752 | 8.6003 | [6 empty] | 405 | 27783
433 | 15.2753 | 5.6974 | 41668.8906 | 0.63583 | 0.18004 | 8.2572 | [6 empty] | 406 | 27955.7773
434 | 15.3106 | 5.7326 | 41726.668 | 0.63671 | 0.17482 | 8.0175 | [6 empty] | 407 | 28221.5547
435 | 15.3458 | 5.7679 | 41822.1094 | 0.63816 | 0.16621 | 7.6229 | [6 empty] | 408 | 28587.666
436 | 15.3811 | 5.8032 | 41887.7773 | 0.63917 | 0.16031 | 7.3524 | [6 empty] | 409 | 29023.334
437 | 15.4164 | 5.8385 | 41941 | 0.63998 | 0.15554 | 7.1337 | [6 empty] | 410 | 29663.1113
438 | 15.4517 | 5.8737 | 42020.332 | 0.64119 | 0.14846 | 6.8087 | [6 empty] | 411 | 30507.4453
439 | 15.4869 | 5.909 | 42095.7773 | 0.64234 | 0.14174 | 6.5007 | [6 empty] | 412 | 31471.8887
440 | 15.5222 | 5.9443 | 42017.1094 | 0.64114 | 0.14875 | 6.8219 | [6 empty] | 413 | 32605.4453
441 | 15.5575 | 5.9796 | 41993.332 | 0.64078 | 0.15087 | 6.9192 | [6 empty] | 414 | 33905.8906
442 | 15.5928 | 6.0149 | 42177.2227 | 0.64358 | 0.13452 | 6.1694 | [6 empty] | 415 | 35077.8906
443 | 15.6281 | 6.0501 | 42188.332 | 0.64375 | 0.13354 | 6.1243 | [6 empty] | 416 | 36038.1094
444 | 15.6633 | 6.0854 | 42192.332 | 0.64381 | 0.13318 | 6.1081 | [6 empty] | 417 | 36858.8906
445 | 15.6986 | 6.1207 | 42222.5547 | 0.64427 | 0.13051 | 5.9855 | [6 empty] | 418 | 37666.668
446 | 15.7339 | 6.156 | 42326.1094 | 0.64586 | 0.12138 | 5.5669 | [6 empty] | 419 | 38346.7773
447 | 15.7692 | 6.1912 | 42295.7773 | 0.64539 | 0.12405 | 5.6893 | [6 empty] | 420 | 38896.5547
448 | 15.8044 | 6.2265 | 42290.4453 | 0.64531 | 0.12452 | 5.7108 | [6 empty] | 421 | 39354.8906
449 | 15.8397 | 6.2618 | 42318 | 0.64573 | 0.12209 | 5.5996 | [6 empty] | 422 | 39767.5547
450 | 15.875 | 6.2971 | 42366.5547 | 0.64647 | 0.11783 | 5.4039 | [6 empty] | 423 | 40116.5547
451 | 15.9103 | 6.3324 | 42448.5547 | 0.64772 | 0.11064 | 5.0743 | [6 empty] | 424 | 40397
452 | 15.9456 | 6.3676 | 42429.8906 | 0.64744 | 0.11227 | 5.1492 | [6 empty] | 425 | 40668.1094
453 | 15.9808 | 6.4029 | 42362.332 | 0.64641 | 0.1182 | 5.4209 | [6 empty] | 426 | 41030.332
454 | 16.0161 | 6.4382 | 42361.7773 | 0.6464 | 0.11825 | 5.4231 | [6 empty] | 427 | 41176.2227
455 | 16.0514 | 6.4735 | 42415.7773 | 0.64722 | 0.11351 | 5.2059 | [6 empty] | 428 | 41210.4453
456 | 16.0867 | 6.5087 | 42484 | 0.64826 | 0.10754 | 4.9322 | [6 empty] | 429 | 41336.668
457 | 16.1219 | 6.544 | 42486.5547 | 0.6483 | 0.10732 | 4.922 | [6 empty] | 430 | 41413.8906
458 | 16.1572 | 6.5793 | 42488.5547 | 0.64833 | 0.10715 | 4.914 | [6 empty] | 431 | 41495.2227
459 | 16.1925 | 6.6146 | 42491.668 | 0.64838 | 0.10687 | 4.9015 | [6 empty] | 432 | 41586.4453
460 | 16.2278 | 6.6499 | 42486.2227 | 0.6483 | 0.10735 | 4.9233 | [6 empty] | 433 | 41668.8906
461 | 16.2631 | 6.6851 | 42471.1094 | 0.64807 | 0.10867 | 4.9839 | [6 empty] | 434 | 41726.668
462 | 16.2983 | 6.7204 | 42394.4453 | 0.6469 | 0.11538 | 5.2916 | [6 empty] | 435 | 41822.1094
463 | 16.3336 | 6.7557 | 42441.7773 | 0.64762 | 0.11123 | 5.1015 | [6 empty] | 436 | 41887.7773
464 | 16.3689 | 6.791 | 42597.8906 | 0.65 | 0.097621 | 4.4772 | [6 empty] | 437 | 41941
465 | 16.4042 | 6.8262 | 42593.332 | 0.64993 | 0.098017 | 4.4953 | [6 empty] | 438 | 42020.332
466 | 16.4394 | 6.8615 | 42528 | 0.64894 | 0.1037 | 4.7561 | [6 empty] | 439 | 42095.7773
467 | 16.4747 | 6.8968 | 42602.7773 | 0.65008 | 0.097197 | 4.4577 | [6 empty] | 440 | 42017.1094
468 | 16.51 | 6.9321 | 42541.8906 | 0.64915 | 0.10249 | 4.7006 | [6 empty] | 441 | 41993.332
469 | 16.5453 | 6.9674 | 42578.668 | 0.64971 | 0.099292 | 4.5538 | [6 empty] | 442 | 42177.2227
470 | 16.5806 | 7.0026 | 42708.8906 | 0.6517 | 0.087999 | 4.0359 | [6 empty] | 443 | 42188.332
471 | 16.6158 | 7.0379 | 42667.7773 | 0.65107 | 0.091558 | 4.1991 | [6 empty] | 444 | 42192.332
472 | 16.6511 | 7.0732 | 42694.668 | 0.65148 | 0.089229 | 4.0923 | [6 empty] | 445 | 42222.5547
473 | 16.6864 | 7.1085 | 42717.7773 | 0.65183 | 0.087231 | 4.0006 | [6 empty] | 446 | 42326.1094
474 | 16.7217 | 7.1438 | 42728.4453 | 0.65199 | 0.086309 | 3.9584 | [6 empty] | 447 | 42295.7773
475 | 16.7569 | 7.179 | 42773.1094 | 0.65268 | 0.082454 | 3.7816 | [6 empty] | 448 | 42290.4453
476 | 16.7922 | 7.2143 | 42728.8906 | 0.652 | 0.08627 | 3.9566 | [6 empty] | 449 | 42318
477 | 16.8275 | 7.2496 | 42761.8906 | 0.6525 | 0.083421 | 3.8259 | [6 empty] | 450 | 42366.5547
478 | 16.8628 | 7.2849 | 42801.2227 | 0.6531 | 0.080031 | 3.6704 | [6 empty] | 451 | 42448.5547
479 | 16.8981 | 7.3201 | 42835.332 | 0.65363 | 0.077096 | 3.5358 | [6 empty] | 452 | 42429.8906
480 | 16.9333 | 7.3554 | 42793 | 0.65298 | 0.080739 | 3.7029 | [6 empty] | 453 | 42362.332
481 | 16.9686 | 7.3907 | 42791 | 0.65295 | 0.080912 | 3.7108 | [6 empty] | 454 | 42361.7773
482 | 17.0039 | 7.426 | 42783.4453 | 0.65283 | 0.081563 | 3.7407 | [6 empty] | 455 | 42415.7773
483 | 17.0392 | 7.4612 | 42755.5547 | 0.65241 | 0.083968 | 3.851 | [6 empty] | 456 | 42484
484 | 17.0744 | 7.4965 | 42773.668 | 0.65268 | 0.082406 | 3.7793 | [6 empty] | 457 | 42486.5547
485 | 17.1097 | 7.5318 | 42854.4453 | 0.65392 | 0.075453 | 3.4605 | [6 empty] | 458 | 42488.5547
486 | 17.145 | 7.5671 | 42851.5547 | 0.65387 | 0.075701 | 3.4719 | [6 empty] | 459 | 42491.668
487 | 17.1803 | 7.6024 | 42869.1094 | 0.65414 | 0.074193 | 3.4027 | [6 empty] | 460 | 42486.2227
488 | 17.2156 | 7.6376 | 42815.8906 | 0.65333 | 0.078768 | 3.6125 | [6 empty] | 461 | 42471.1094
489 | 17.2508 | 7.6729 | 42881.8906 | 0.65434 | 0.073096 | 3.3524 | [6 empty] | 462 | 42394.4453
490 | 17.2861 | 7.7082 | 42874.2227 | 0.65422 | 0.073754 | 3.3826 | [6 empty] | 463 | 42441.7773
491 | 17.3214 | 7.7435 | 42881.7773 | 0.65433 | 0.073106 | 3.3528 | [6 empty] | 464 | 42597.8906
492 | 17.3567 | 7.7787 | 42947.332 | 0.65533 | 0.067488 | 3.0952 | [6 empty] | 465 | 42593.332
493 | 17.3919 | 7.814 | 42937 | 0.65518 | 0.068372 | 3.1357 | [6 empty] | 466 | 42528
494 | 17.4272 | 7.8493 | 42823.4453 | 0.65344 | 0.078118 | 3.5827 | [6 empty] | 467 | 42602.7773
495 | 17.4625 | 7.8846 | 42828.2227 | 0.65352 | 0.077707 | 3.5639 | [6 empty] | 468 | 42541.8906
496 | 17.4978 | 7.9199 | 42817.5547 | 0.65335 | 0.078625 | 3.606 | [6 empty] | 469 | 42578.668
497 | 17.5331 | 7.9551 | 42882.7773 | 0.65435 | 0.07302 | 3.3489 | [6 empty] | 470 | 42708.8906
498 | 17.5683 | 7.9904 | 42889.8906 | 0.65446 | 0.07241 | 3.3209 | [6 empty] | 471 | 42667.7773
499 | 17.6036 | 8.0257 | 42878.2227 | 0.65428 | 0.073411 | 3.3668 | [6 empty] | 472 | 42694.668
500 | 17.6389 | 8.061 | 42904.668 | 0.65468 | 0.071142 | 3.2628 | [6 empty] | 473 | 42717.7773
501 | 17.6742 | 8.0962 | 42917.1094 | 0.65487 | 0.070076 | 3.2139 | [6 empty] | 474 | 42728.4453
502 | 17.7094 | 8.1315 | 42937.668 | 0.65519 | 0.068315 | 3.1331 | [6 empty] | 475 | 42773.1094
503 | 17.7447 | 8.1668 | 42947.332 | 0.65533 | 0.067488 | 3.0952 | [6 empty] | 476 | 42728.8906
504 | 17.78 | 8.2021 | 43020.668 | 0.65645 | 0.061223 | 2.8078 | [6 empty] | 477 | 42761.8906
505 | 17.8153 | 8.2374 | 43030.8906 | 0.65661 | 0.060351 | 2.7679 | [6 empty] | 478 | 42801.2227
506 | 17.8506 | 8.2726 | 43023.668 | 0.6565 | 0.060967 | 2.7961 | [6 empty] | 479 | 42835.332
507 | 17.8858 | 8.3079 | 43022.5547 | 0.65648 | 0.061062 | 2.8005 | [6 empty] | 480 | 42793
508 | 17.9211 | 8.3432 | 43053.332 | 0.65695 | 0.058439 | 2.6802 | [6 empty] | 481 | 42791
509 | 17.9564 | 8.3785 | 43090.5547 | 0.65752 | 0.055271 | 2.5349 | [6 empty] | 482 | 42783.4453
510 | 17.9917 | 8.4138 | 43066.8906 | 0.65716 | 0.057284 | 2.6272 | [6 empty] | 483 | 42755.5547
511 | 18.0269 | 8.449 | 43079.2227 | 0.65735 | 0.056235 | 2.5791 | [6 empty] | 484 | 42773.668
512 | 18.0622 | 8.4843 | 43049.1094 | 0.65689 | 0.058798 | 2.6966 | [6 empty] | 485 | 42854.4453
513 | 18.0975 | 8.5196 | 43125.8906 | 0.65806 | 0.052268 | 2.3972 | [6 empty] | 486 | 42851.5547
514 | 18.1328 | 8.5549 | 43061.1094 | 0.65707 | 0.057776 | 2.6498 | [6 empty] | 487 | 42869.1094
515 | 18.1681 | 8.5901 | 43064.332 | 0.65712 | 0.057502 | 2.6372 | [6 empty] | 488 | 42815.8906
516 | 18.2033 | 8.6254 | 43066.4453 | 0.65715 | 0.057322 | 2.6289 | [6 empty] | 489 | 42881.8906
517 | 18.2386 | 8.6607 | 43084.8906 | 0.65743 | 0.055752 | 2.557 | [6 empty] | 490 | 42874.2227
518 | 18.2739 | 8.696 | 43173 | 0.65878 | 0.048272 | 2.2139 | [6 empty] | 491 | 42881.7773
519 | 18.3092 | 8.7312 | 43154.4453 | 0.65849 | 0.049845 | 2.286 | [6 empty] | 492 | 42947.332
520 | 18.3444 | 8.7665 | 43151.7773 | 0.65845 | 0.050071 | 2.2964 | [6 empty] | 493 | 42937
521 | 18.3797 | 8.8018 | 43186.8906 | 0.65899 | 0.047096 | 2.1599 | [6 empty] | 494 | 42823.4453
522 | 18.415 | 8.8371 | 43097.332 | 0.65762 | 0.054694 | 2.5084 | [6 empty] | 495 | 42828.2227
523 | 18.4503 | 8.8724 | 43061.4453 | 0.65708 | 0.057748 | 2.6485 | [6 empty] | 496 | 42817.5547
524 | 18.4856 | 8.9076 | 43022.5547 | 0.65648 | 0.061062 | 2.8005 | [6 empty] | 497 | 42882.7773
525 | 18.5208 | 8.9429 | 42979 | 0.65582 | 0.06478 | 2.971 | [6 empty] | 498 | 42889.8906
526 | 18.5561 | 8.9782 | 42948 | 0.65534 | 0.067431 | 3.0926 | [6 empty] | 499 | 42878.2227
527 | 18.5914 | 9.0135 | 42994.668 | 0.65606 | 0.063442 | 2.9096 | [6 empty] | 500 | 42904.668
528 | 18.6267 | 9.0487 | 42993.8906 | 0.65604 | 0.063508 | 2.9127 | [6 empty] | 501 | 42917.1094
529 | 18.6619 | 9.084 | 42985.1094 | 0.65591 | 0.064258 | 2.9471 | [6 empty] | 502 | 42937.668
530 | 18.6972 | 9.1193 | 42974.2227 | 0.65574 | 0.065188 | 2.9897 | [6 empty] | 503 | 42947.332
531 | 18.7325 | 9.1546 | 42901 | 0.65463 | 0.071457 | 3.2772 | [6 empty] | 504 | 43020.668
532 | 18.7678 | 9.1899 | 42936.2227 | 0.65516 | 0.068439 | 3.1388 | [6 empty] | 505 | 43030.8906
533 | 18.8031 | 9.2251 | 42952 | [9 empty] | 506 | 43023.668
534 | 18.8383 | 9.2604 | 42908.2227 | [9 empty] | 507 | 43022.5547
535 | 18.8736 | 9.2957 | 42927.7773 | [9 empty] | 508 | 43053.332
536 | 18.9089 | 9.331 | 42912.5547 | [9 empty] | 509 | 43090.5547
537 | 18.9442 | 9.3662 | 42929.2227 | [9 empty] | 510 | 43066.8906
538 | 18.9794 | 9.4015 | 42916 | [9 empty] | 511 | 43079.2227
539 | 19.0147 | 9.4368 | 43048.2227 | [9 empty] | 512 | 43049.1094
540 | 19.05 | 9.4721 | 42891.7773 | [9 empty] | 513 | 43125.8906
541 | 19.0853 | 9.5074 | 42976.5547 | [9 empty] | 514 | 43061.1094
542 | 19.1206 | 9.5426 | 42976 | [9 empty] | 515 | 43064.332
543 | 19.1558 | 9.5779 | 43146.3555 | [9 empty] | 516 | 43066.4453
[13 empty] | 517 | 43084.8906
[13 empty] | 518 | 43173
[13 empty] | 519 | 43154.4453
[13 empty] | 520 | 43151.7773
[13 empty] | 521 | 43186.8906
[13 empty] | 522 | 43097.332
[13 empty] | 523 | 43061.4453
[13 empty] | 524 | 43022.5547
[13 empty] | 525 | 42979
[13 empty] | 526 | 42948
[13 empty] | 527 | 42994.668
[13 empty] | 528 | 42993.8906
[13 empty] | 529 | 42985.1094
[13 empty] | 530 | 42974.2227
[13 empty] | 531 | 42901
[13 empty] | 532 | 42936.2227
[13 empty] | 533 | 42952
[13 empty] | 534 | 42908.2227
[13 empty] | 535 | 42927.7773
[13 empty] | 536 | 42912.5547
[13 empty] | 537 | 42929.2227
[13 empty] | 538 | 42916
[13 empty] | 539 | 43048.2227
[13 empty] | 540 | 42891.7773
[13 empty] | 541 | 42976.5547
[13 empty] | 542 | 42976

## Profiler vs epson
|  |  |  | dose=c+b/(X-a) | [7 empty] | pixels to cm conversion | 1px = | 1/ | 1 inch = 2.54 centimetres
a |  |  |  | -0.040473 | [10 empty] | dpi | ppcm | cm/px
b |  |  |  | 2.7151 | [10 empty] | 72 | 28.3465 | 0.035278
c |  |  |  | -3.8345
centre across | dist cm |  |  | normalized film response | Gy
0 | 0 | -9.5779 | 43060.7539 | 0.65706 | 0.057806 | 2.6572 | -11.2 | 1.3544
1 | 0.035278 | -9.5426 | 43098.0781 | 0.65763 | 0.054631 | 2.5112 | -10.8 | 1.4443
2 | 0.070556 | -9.5074 | 43096.3789 | 0.65761 | 0.054775 | 2.5179 | -10.4 | 1.5476
3 | 0.10583 | -9.4721 | 43013.6406 | 0.65635 | 0.061822 | 2.8418 | -10 | 1.6799
4 | 0.14111 | -9.4368 | 43058.8867 | 0.65704 | 0.057965 | 2.6645 | -9.6 | 1.8313
5 | 0.17639 | -9.4015 | 43083.8984 | 0.65742 | 0.055837 | 2.5667 | -9.2 | 2.0226
6 | 0.21167 | -9.3663 | 43073.7109 | 0.65726 | 0.056704 | 2.6065 | -8.8 | 2.205
7 | 0.24694 | -9.331 | 43059.6836 | 0.65705 | 0.057898 | 2.6614 | -8.4 | 2.379
8 | 0.28222 | -9.2957 | 43037.6406 | 0.65671 | 0.059775 | 2.7477 | -8 | 2.8574
9 | 0.3175 | -9.2604 | 43053.9883 | 0.65696 | 0.058383 | 2.6837 | -7.6 | 3.2205
10 | 0.35278 | -9.2251 | 43108.2109 | 0.65779 | 0.05377 | 2.4717 | -7.2 | 3.6975
11 | 0.38806 | -9.1899 | 43129.5156 | 0.65811 | 0.05196 | 2.3885 | -6.8 | 4.333
12 | 0.42333 | -9.1546 | 43071.375 | 0.65723 | 0.056902 | 2.6156 | -6.4 | 5.0721
13 | 0.45861 | -9.1193 | 43125.5352 | 0.65805 | 0.052298 | 2.404 | -6 | 6.4392
14 | 0.49389 | -9.084 | 43090.2969 | 0.65752 | 0.055293 | 2.5417 | -5.6 | 9.2202
15 | 0.52917 | -9.0488 | 43126.9062 | 0.65807 | 0.052182 | 2.3987 | -5.2 | 20.0409
16 | 0.56444 | -9.0135 | 43078.7227 | 0.65734 | 0.056277 | 2.5869 | -4.8 | 84.8158
17 | 0.59972 | -8.9782 | 43103.5703 | 0.65772 | 0.054164 | 2.4898 | -4.4 | 98.2231
18 | 0.635 | -8.9429 | 43142.1445 | 0.65831 | 0.050889 | 2.3392 | -4 | 100.3047
19 | 0.67028 | -8.9076 | 43107.9375 | 0.65778 | 0.053793 | 2.4727 | -3.6 | 101.0664
20 | 0.70556 | -8.8724 | 43122.9062 | 0.65801 | 0.052522 | 2.4143 | -3.2 | 101.0972
21 | 0.74083 | -8.8371 | 43048.625 | 0.65688 | 0.058839 | 2.7047 | -2.8 | 100.9568
22 | 0.77611 | -8.8018 | 43069.7188 | 0.6572 | 0.057043 | 2.6221 | -2.4 | 100.497
23 | 0.81139 | -8.7665 | 43074.543 | 0.65728 | 0.056633 | 2.6033 | -2 | 100.0592
24 | 0.84667 | -8.7312 | 43085.9648 | 0.65745 | 0.055661 | 2.5586 | -1.6 | 99.4802
25 | 0.88194 | -8.696 | 43069.1875 | 0.65719 | 0.057089 | 2.6242 | -1.2 | 99.6109
26 | 0.91722 | -8.6607 | 43008.9023 | 0.65627 | 0.062227 | 2.8604 | -0.8 | 99.742
27 | 0.9525 | -8.6254 | 42981.0625 | 0.65585 | 0.064604 | 2.9697 | -0.4 | 99.9903 |  | -11.2 | 1.3544
28 | 0.98778 | -8.5901 | 42981.0352 | 0.65585 | 0.064606 | 2.9698 | 0 | 100 |  | -10.8 | 1.4443
29 | 1.0231 | -8.5549 | 42923.0586 | 0.65496 | 0.069566 | 3.1978 | 0.4 | 100.1212 |  | -10.4 | 1.5476
30 | 1.0583 | -8.5196 | 42985.918 | 0.65592 | 0.064189 | 2.9506 | 0.8 | 100.0051 |  | -10 | 1.6799
31 | 1.0936 | -8.4843 | 43081.4375 | 0.65738 | 0.056046 | 2.5763 | 1.2 | 99.834 |  | -9.6 | 1.8313
32 | 1.1289 | -8.449 | 43086.0391 | 0.65745 | 0.055655 | 2.5583 | 1.6 | 99.7528 |  | -9.2 | 2.0226
33 | 1.1642 | -8.4138 | 43065.8516 | 0.65714 | 0.057372 | 2.6373 | 2 | 100.0306 |  | -8.8 | 2.205
34 | 1.1994 | -8.3785 | 43005.3633 | 0.65622 | 0.062529 | 2.8743 | 2.4 | 100.5114 |  | -8.4 | 2.379
35 | 1.2347 | -8.3432 | 42922.832 | 0.65496 | 0.069586 | 3.1987 | 2.8 | 100.9674 |  | -8 | 2.8574
36 | 1.27 | -8.3079 | 43044.1562 | 0.65681 | 0.05922 | 2.7222 | 3.2 | 101.1327 |  | -7.6 | 3.2205
37 | 1.3053 | -8.2726 | 43081.8555 | 0.65739 | 0.056011 | 2.5747 | 3.6 | 100.9958 |  | -7.2 | 3.6975
38 | 1.3406 | -8.2374 | 43018.0859 | 0.65641 | 0.061443 | 2.8244 | 4 | 100.2152 |  | -6.8 | 4.333
39 | 1.3758 | -8.2021 | 42983.8477 | 0.65589 | 0.064366 | 2.9587 | 4.4 | 96.7088 |  | -6.4 | 5.0721
40 | 1.4111 | -8.1668 | 42972.1172 | 0.65571 | 0.065368 | 3.0048 | 4.8 | 75.9462 |  | -6 | 6.4392
41 | 1.4464 | -8.1315 | 42977.2266 | 0.65579 | 0.064932 | 2.9847 | 5.2 | 15.3962 |  | -5.6 | 9.2202
42 | 1.4817 | -8.0962 | 42946.6836 | 0.65532 | 0.067544 | 3.1048 | 5.6 | 8.0949 |  | -5.2 | 20.0409
43 | 1.5169 | -8.061 | 42939.4492 | 0.65521 | 0.068163 | 3.1333 | 6 | 5.9523 |  | -4.8 | 84.8158
44 | 1.5522 | -8.0257 | 42939.668 | 0.65522 | 0.068144 | 3.1324 | 6.4 | 4.9166 |  | -4.4 | 98.2231
45 | 1.5875 | -7.9904 | 42935.2344 | 0.65515 | 0.068524 | 3.1498 | 6.8 | 4.1063 |  | -4 | 100.3047
46 | 1.6228 | -7.9551 | 42892.7656 | 0.6545 | 0.072163 | 3.3171 | 7.2 | 3.5706 |  | -3.6 | 101.0664
47 | 1.6581 | -7.9199 | 42859.8281 | 0.654 | 0.07499 | 3.4471 | 7.6 | 3.0823 |  | -3.2 | 101.0972
48 | 1.6933 | -7.8846 | 42912.6094 | 0.6548 | 0.070462 | 3.2389 | 8 | 2.639 |  | -2.8 | 100.9568
49 | 1.7286 | -7.8493 | 42864.9648 | 0.65408 | 0.074549 | 3.4268 | 8.4 | 2.3103 |  | -2.4 | 100.497
50 | 1.7639 | -7.814 | 42828.3594 | 0.65352 | 0.077695 | 3.5714 | 8.8 | 2.0938 |  | -2 | 100.0592
51 | 1.7992 | -7.7788 | 42863.8828 | 0.65406 | 0.074642 | 3.4311 | 9.2 | 1.915 |  | -1.6 | 99.4802
52 | 1.8344 | -7.7435 | 42865.2578 | 0.65408 | 0.074524 | 3.4257 | 9.6 | 1.6851 |  | -1.2 | 99.6109
53 | 1.8697 | -7.7082 | 42842.25 | 0.65373 | 0.076501 | 3.5165 | 10 | 1.5497 |  | -0.8 | 99.742
54 | 1.905 | -7.6729 | 42843.0117 | 0.65374 | 0.076435 | 3.5135 | 10.4 | 1.4568 |  | -0.4 | 99.9903
55 | 1.9403 | -7.6376 | 42829.5859 | 0.65354 | 0.07759 | 3.5666 | 10.8 | 1.3746 |  | 0 | 100
56 | 1.9756 | -7.6024 | 42862.0547 | 0.65403 | 0.074799 | 3.4383 | 11.2 | 1.2596 |  | 0.4 | 100.1212
57 | 2.0108 | -7.5671 | 42865.8867 | 0.65409 | 0.07447 | 3.4232 |  |  |  | 0.8 | 100.0051
58 | 2.0461 | -7.5318 | 42877.5391 | 0.65427 | 0.07347 | 3.3772 | Y Axis Position(cm) | Set 1 |  | 1.2 | 99.834
59 | 2.0814 | -7.4965 | 42928.1953 | 0.65504 | 0.069126 | 3.1776 | -16.4 | 2.4625 |  | 1.6 | 99.7528
60 | 2.1167 | -7.4612 | 42936.2422 | 0.65517 | 0.068437 | 3.1459 | -16 | 2.6876 |  | 2 | 100.0306
61 | 2.1519 | -7.426 | 42970.4336 | 0.65569 | 0.065512 | 3.0114 | -15.6 | 2.8697 |  | 2.4 | 100.5114
62 | 2.1872 | -7.3907 | 42869.5273 | 0.65415 | 0.074157 | 3.4088 | -15.2 | 3.1415 |  | 2.8 | 100.9674
63 | 2.2225 | -7.3554 | 42926.2656 | 0.65501 | 0.069292 | 3.1851 | -14.8 | 3.348 |  | 3.2 | 101.1327
64 | 2.2578 | -7.3201 | 42818.3047 | 0.65337 | 0.07856 | 3.6112 | -14.4 | 3.7106 |  | 3.6 | 100.9958
65 | 2.2931 | -7.2849 | 42855.6484 | 0.65394 | 0.075349 | 3.4636 | -14 | 4.1405 |  | 4 | 100.2152
66 | 2.3283 | -7.2496 | 42837.0703 | 0.65365 | 0.076946 | 3.537 | -13.6 | 4.5166 |  | 4.4 | 96.7088
67 | 2.3636 | -7.2143 | 42865.75 | 0.65409 | 0.074482 | 3.4237 | -13.2 | 5.0184 |  | 4.8 | 75.9462
68 | 2.3989 | -7.179 | 42747.1953 | 0.65228 | 0.08469 | 3.893 | -12.8 | 5.5321 |  | 5.2 | 15.3962
69 | 2.4342 | -7.1438 | 42672.3438 | 0.65114 | 0.091162 | 4.1905 | -12.4 | 6.1006 |  | 5.6 | 8.0949
70 | 2.4694 | -7.1085 | 42774.0312 | 0.65269 | 0.082374 | 3.7865 | -12 | 6.7372 |  | 6 | 5.9523
71 | 2.5047 | -7.0732 | 42742.2695 | 0.65221 | 0.085115 | 3.9125 | -11.6 | 7.6521 |  | 6.4 | 4.9166
72 | 2.54 | -7.0379 | 42831.5117 | 0.65357 | 0.077424 | 3.559 | -11.2 | 9.0169 |  | 6.8 | 4.1063
73 | 2.5753 | -7.0026 | 42863.8867 | 0.65406 | 0.074642 | 3.4311 | -10.8 | 11.9192 |  | 7.2 | 3.5706
74 | 2.6106 | -6.9674 | 42775.0234 | 0.65271 | 0.082289 | 3.7826 | -10.4 | 23.1964 |  | 7.6 | 3.0823
75 | 2.6458 | -6.9321 | 42676.2734 | 0.6512 | 0.090822 | 4.1748 | -10 | 70.0223 |  | 8 | 2.639
76 | 2.6811 | -6.8968 | 42591.9609 | 0.64991 | 0.098137 | 4.5111 | -9.6 | 96.3873 |  | 8.4 | 2.3103
77 | 2.7164 | -6.8615 | 42633.6719 | 0.65055 | 0.094514 | 4.3446 | -9.2 | 100.5738 |  | 8.8 | 2.0938
78 | 2.7517 | -6.8262 | 42577.543 | 0.64969 | 0.09939 | 4.5687 | -8.8 | 101.6498 |  | 9.2 | 1.915
79 | 2.7869 | -6.791 | 42582.6914 | 0.64977 | 0.098942 | 4.5481 | -8.4 | 102.3072 |  | 9.6 | 1.6851
80 | 2.8222 | -6.7557 | 42585.6367 | 0.64982 | 0.098686 | 4.5363 | -8 | 102.7132 |  | 10 | 1.5497
81 | 2.8575 | -6.7204 | 42537.8086 | 0.64909 | 0.10285 | 4.7277 | -7.6 | 103.0261 |  | 10.4 | 1.4568
82 | 2.8928 | -6.6851 | 42519.1016 | 0.6488 | 0.10448 | 4.8027 | -7.2 | 103.1994 |  | 10.8 | 1.3746
83 | 2.9281 | -6.6499 | 42607.3242 | 0.65015 | 0.096802 | 4.4497 | -6.8 | 103.3055 |  | 11.2 | 1.2596
84 | 2.9633 | -6.6146 | 42572.5586 | 0.64962 | 0.099824 | 4.5886 | -6.4 | 103.506 | Detector ID | Y Axis Position(cm) | Set 1
85 | 2.9986 | -6.5793 | 42580.0781 | 0.64973 | 0.09917 | 4.5586 | -6 | 103.4357 |  | -16.4 | 0.67405
86 | 3.0339 | -6.544 | 42466.8359 | 0.648 | 0.10904 | 5.0124 | -5.6 | 103.3058 |  | -16 | 0.75086
87 | 3.0692 | -6.5088 | 42477.2227 | 0.64816 | 0.10814 | 4.9707 | -5.2 | 103.2712 |  | -15.6 | 0.77318
88 | 3.1044 | -6.4735 | 42497.4453 | 0.64847 | 0.10637 | 4.8895 | -4.8 | 103.2441 |  | -15.2 | 0.85447
89 | 3.1397 | -6.4382 | 42481.9258 | 0.64823 | 0.10772 | 4.9518 | -4.4 | 103.0501 |  | -14.8 | 0.92808
90 | 3.175 | -6.4029 | 42464.125 | 0.64796 | 0.10928 | 5.0233 | -4 | 102.5345 |  | -14.4 | 0.99125
91 | 3.2103 | -6.3676 | 42506.4375 | 0.64861 | 0.10558 | 4.8534 | -3.6 | 101.8602 |  | -14 | 1.0613
92 | 3.2456 | -6.3324 | 42406.875 | 0.64709 | 0.11429 | 5.2536 | -3.2 | 101.437 |  | -13.6 | 1.1297
93 | 3.2808 | -6.2971 | 42415.7891 | 0.64722 | 0.11351 | 5.2177 | -2.8 | 100.637 |  | -13.2 | 1.1934
94 | 3.3161 | -6.2618 | 42362.5859 | 0.64641 | 0.11818 | 5.4322 | -2.4 | 100.2303 |  | -12.8 | 1.2675
95 | 3.3514 | -6.2265 | 42306.8594 | 0.64556 | 0.12308 | 5.6574 | -2 | 99.61 |  | -12.4 | 1.3777
96 | 3.3867 | -6.1913 | 42215.2344 | 0.64416 | 0.13116 | 6.0289 | -1.6 | 99.4825 |  | -12 | 1.4477
97 | 3.4219 | -6.156 | 41926.8359 | 0.63976 | 0.15681 | 7.2083 | -1.2 | 99.6622 |  | -11.6 | 1.5635
98 | 3.4572 | -6.1207 | 41722.6914 | 0.63665 | 0.17517 | 8.0523 | -0.8 | 99.8668 |  | -11.2 | 1.6924
99 | 3.4925 | -6.0854 | 41938.1445 | 0.63994 | 0.1558 | 7.1617 | -0.4 | 99.8628 |  | -10.8 | 1.8483
100 | 3.5278 | -6.0501 | 42150.3086 | 0.64317 | 0.1369 | 6.2931 | 0 | 100 |  | -10.4 | 1.9765
101 | 3.5631 | -6.0149 | 42146.8008 | 0.64312 | 0.13721 | 6.3074 | 0.4 | 100.2499 |  | -10 | 2.1579
102 | 3.5983 | -5.9796 | 42053.7617 | 0.6417 | 0.14548 | 6.6874 | 0.8 | 100.3314 |  | -9.6 | 2.3587
103 | 3.6336 | -5.9443 | 42069.6719 | 0.64194 | 0.14406 | 6.6223 | 1.2 | 100.5899 |  | -9.2 | 2.5998
104 | 3.6689 | -5.909 | 42077.8086 | 0.64207 | 0.14334 | 6.589 | 1.6 | 101.2232 |  | -8.8 | 2.8183
105 | 3.7042 | -5.8738 | 41983.4023 | 0.64063 | 0.15175 | 6.9757 | 2 | 101.6594 |  | -8.4 | 3.1722
106 | 3.7394 | -5.8385 | 41991.0664 | 0.64074 | 0.15107 | 6.9443 | 2.4 | 102.6837 |  | -8 | 3.5408
107 | 3.7747 | -5.8032 | 41907.5664 | 0.63947 | 0.15854 | 7.2876 | 2.8 | 103.2703 |  | -7.6 | 3.9315
108 | 3.81 | -5.7679 | 41927.2383 | 0.63977 | 0.15678 | 7.2066 | 3.2 | 104.0326 |  | -7.2 | 4.3838
109 | 3.8453 | -5.7326 | 41892.8516 | 0.63924 | 0.15986 | 7.3482 | 3.6 | 104.4274 |  | -6.8 | 5.0034
110 | 3.8806 | -5.6974 | 41835.5312 | 0.63837 | 0.165 | 7.5848 | 4 | 104.8436 |  | -6.4 | 5.6755
111 | 3.9158 | -5.6621 | 41745.543 | 0.637 | 0.17311 | 7.9574 | 4.4 | 104.9309 |  | -6 | 6.8761
112 | 3.9511 | -5.6268 | 41741.7695 | 0.63694 | 0.17345 | 7.9731 | 4.8 | 104.9451 |  | -5.6 | 9.9146
113 | 3.9864 | -5.5915 | 41639.2969 | 0.63537 | 0.18272 | 8.3993 | 5.2 | 105.0835 |  | -5.2 | 22.9712
114 | 4.0217 | -5.5563 | 41641.7695 | 0.63541 | 0.1825 | 8.389 | 5.6 | 105.0255 |  | -4.8 | 67.2426
115 | 4.0569 | -5.521 | 41520.7109 | 0.63357 | 0.19351 | 8.8951 | 6 | 104.9896 |  | -4.4 | 91.2343
116 | 4.0922 | -5.4857 | 41418.8711 | 0.63201 | 0.20282 | 9.3229 | 6.4 | 104.7757 |  | -4 | 98.0957
117 | 4.1275 | -5.4504 | 41332.8281 | 0.6307 | 0.21071 | 9.686 | 6.8 | 104.8769 |  | -3.6 | 99.4905
118 | 4.1628 | -5.4151 | 41226.5391 | 0.62908 | 0.22051 | 10.1364 | 7.2 | 104.5792 |  | -3.2 | 100.0978
119 | 4.1981 | -5.3799 | 41014.5391 | 0.62584 | 0.2402 | 11.0414 | 7.6 | 104.2662 |  | -2.8 | 99.8854
120 | 4.2333 | -5.3446 | 40856.1562 | 0.62342 | 0.25503 | 11.7232 | 8 | 104.0358 |  | -2.4 | 99.8753
121 | 4.2686 | -5.3093 | 40677.4492 | 0.6207 | 0.2719 | 12.4985 | 8.4 | 103.5482 |  | -2 | 99.4692
122 | 4.3039 | -5.274 | 40475.0703 | 0.61761 | 0.29117 | 13.3843 | 8.8 | 102.1359 |  | -1.6 | 99.4543
123 | 4.3392 | -5.2388 | 40267.6562 | 0.61445 | 0.31111 | 14.3008 | 9.2 | 99.2299 |  | -1.2 | 99.5821
124 | 4.3744 | -5.2035 | 39921.3633 | 0.60916 | 0.34483 | 15.8509 | 9.6 | 83.9414 |  | -0.8 | 99.8154
125 | 4.4097 | -5.1682 | 39504.6328 | 0.6028 | 0.38614 | 17.75 | 10 | 32.5274 |  | -0.4 | 99.8557
126 | 4.445 | -5.1329 | 38999.5391 | 0.59509 | 0.43733 | 20.1027 | 10.4 | 13.2724 |  | 0 | 100
127 | 4.4803 | -5.0976 | 38369.8203 | 0.58549 | 0.5029 | 23.117 | 10.8 | 9.6188 |  | 0.4 | 100.114
128 | 4.5156 | -5.0624 | 37656.3438 | 0.5746 | 0.57968 | 26.6461 | 11.2 | 8.3009 |  | 0.8 | 100.0654
129 | 4.5508 | -5.0271 | 36915.9531 | 0.5633 | 0.66227 | 30.4429 | 11.6 | 7.21 |  | 1.2 | 100.0297
130 | 4.5861 | -4.9918 | 36021.5312 | 0.54965 | 0.76627 | 35.2235 | 12 | 6.4134 |  | 1.6 | 99.9109
131 | 4.6214 | -4.9565 | 35002.4453 | 0.5341 | 0.89079 | 40.9472 | 12.4 | 5.7384 |  | 2 | 100.0918
132 | 4.6567 | -4.9213 | 33843.6953 | 0.51642 | 1.0408 | 47.8436 | 12.8 | 5.2678 |  | 2.4 | 100.4326
133 | 4.6919 | -4.886 | 32631.5742 | 0.49793 | 1.2083 | 55.5425 | 13.2 | 4.7581 |  | 2.8 | 100.8204
134 | 4.7272 | -4.8507 | 31458.6699 | 0.48003 | 1.3817 | 63.5131 | 13.6 | 4.3577 |  | 3.2 | 101.0339
135 | 4.7625 | -4.8154 | 30432.8418 | 0.46438 | 1.5434 | 70.9475 | 14 | 3.9579 |  | 3.6 | 100.9378
136 | 4.7978 | -4.7801 | 29706.0312 | 0.45328 | 1.6642 | 76.5002 | 14.4 | 3.656 |  | 4 | 100.0377
137 | 4.8331 | -4.7449 | 29090.3867 | 0.44389 | 1.7709 | 81.4025 | 14.8 | 3.3305 |  | 4.4 | 95.9061
138 | 4.8683 | -4.7096 | 28651.0352 | 0.43719 | 1.8496 | 85.0189 | 15.2 | 3.1023 |  | 4.8 | 77.3493
139 | 4.9036 | -4.6743 | 28371.1328 | 0.43292 | 1.9008 | 87.3763 | 15.6 | 2.8256 |  | 5.2 | 33.8011
140 | 4.9389 | -4.639 | 28152.8496 | 0.42958 | 1.9415 | 89.2444 | 16 | 2.5966 |  | 5.6 | 11.7398
141 | 4.9742 | -4.6038 | 27926.8105 | 0.42614 | 1.9842 | 91.207 | 16.4 | 2.4349 |  | 6 | 7.8283
142 | 5.0094 | -4.5685 | 27710.2227 | 0.42283 | 2.0257 | 93.115 |  |  |  | 6.4 | 6.2151
143 | 5.0447 | -4.5332 | 27549.8574 | 0.42038 | 2.0568 | 94.5453 |  |  |  | 6.8 | 5.2689
144 | 5.08 | -4.4979 | 27434.957 | 0.41863 | 2.0793 | 95.5795 |  |  |  | 7.2 | 4.6441
145 | 5.1153 | -4.4626 | 27305.5312 | 0.41666 | 2.1048 | 96.7539 |  |  |  | 7.6 | 4.0928
146 | 5.1506 | -4.4274 | 27193.123 | 0.41494 | 2.1272 | 97.7822 |  |  |  | 8 | 3.6586
147 | 5.1858 | -4.3921 | 27124.7012 | 0.4139 | 2.1409 | 98.4119 |  |  |  | 8.4 | 3.2977
148 | 5.2211 | -4.3568 | 26985.3633 | 0.41177 | 2.169 | 99.7033 |  |  |  | 8.8 | 2.9718
149 | 5.2564 | -4.3215 | 26915.2734 | 0.4107 | 2.1832 | 100.3575 |  |  |  | 9.2 | 2.6762
150 | 5.2917 | -4.2863 | 26894.3574 | 0.41038 | 2.1875 | 100.5533 |  |  |  | 9.6 | 2.4238
151 | 5.3269 | -4.251 | 26884.0098 | 0.41022 | 2.1896 | 100.6503 |  |  |  | 10 | 2.2345
152 | 5.3622 | -4.2157 | 26887.25 | 0.41027 | 2.1889 | 100.6199 |  |  |  | 10.4 | 2.0553
153 | 5.3975 | -4.1804 | 26841.3652 | 0.40957 | 2.1983 | 101.0506 |  |  |  | 10.8 | 1.8666
154 | 5.4328 | -4.1451 | 26753.873 | 0.40824 | 2.2163 | 101.8757 |  |  |  | 11.2 | 1.7822
155 | 5.4681 | -4.1099 | 26795.8047 | 0.40888 | 2.2076 | 101.4797 |  |  |  | 11.6 | 1.6177
156 | 5.5033 | -4.0746 | 26786.668 | 0.40874 | 2.2095 | 101.5659 |  |  |  | 12 | 1.4997
157 | 5.5386 | -4.0393 | 26867.8926 | 0.40998 | 2.1929 | 100.8014 |  |  |  | 12.4 | 1.3855
158 | 5.5739 | -4.004 | 26849.3965 | 0.4097 | 2.1967 | 100.9752 |  |  |  | 12.8 | 1.3324
159 | 5.6092 | -3.9688 | 26924.7578 | 0.41085 | 2.1813 | 100.2688 |  |  |  | 13.2 | 1.2369
160 | 5.6444 | -3.9335 | 26848.834 | 0.40969 | 2.1968 | 100.9804 |  |  |  | 13.6 | 1.1886
161 | 5.6797 | -3.8982 | 26830.6406 | 0.40941 | 2.2005 | 101.1515 |  |  |  | 14 | 1.0981
162 | 5.715 | -3.8629 | 26804.9062 | 0.40902 | 2.2058 | 101.3939 |  |  |  | 14.4 | 1.047
163 | 5.7503 | -3.8276 | 26781.1406 | 0.40865 | 2.2107 | 101.6181 |  |  |  | 14.8 | 0.98912
164 | 5.7856 | -3.7924 | 26796.5898 | 0.40889 | 2.2075 | 101.4723 |  |  |  | 15.2 | 0.93402
165 | 5.8208 | -3.7571 | 26802.125 | 0.40897 | 2.2064 | 101.4201 |  |  |  | 15.6 | 0.87866
166 | 5.8561 | -3.7218 | 26827.582 | 0.40936 | 2.2011 | 101.1803 |  |  |  | 16 | 0.82943
167 | 5.8914 | -3.6865 | 26845.1426 | 0.40963 | 2.1975 | 101.0151 |  |  |  | 16.4 | 0.77833
168 | 5.9267 | -3.6513 | 26848.6094 | 0.40968 | 2.1968 | 100.9825
169 | 5.9619 | -3.616 | 26820.082 | 0.40925 | 2.2027 | 101.2509
170 | 5.9972 | -3.5807 | 26791.3086 | 0.40881 | 2.2086 | 101.5221
171 | 6.0325 | -3.5454 | 26755.9727 | 0.40827 | 2.2158 | 101.8559
172 | 6.0678 | -3.5101 | 26731.1191 | 0.40789 | 2.2209 | 102.0911
173 | 6.1031 | -3.4749 | 26708.4062 | 0.40754 | 2.2256 | 102.3064
174 | 6.1383 | -3.4396 | 26730.5879 | 0.40788 | 2.2211 | 102.0962
175 | 6.1736 | -3.4043 | 26710.3047 | 0.40757 | 2.2252 | 102.2884
176 | 6.2089 | -3.369 | 26691.3828 | 0.40728 | 2.2291 | 102.4681
177 | 6.2442 | -3.3338 | 26660.1504 | 0.40681 | 2.2356 | 102.765
178 | 6.2794 | -3.2985 | 26694.0195 | 0.40732 | 2.2286 | 102.443
179 | 6.3147 | -3.2632 | 26734.877 | 0.40795 | 2.2202 | 102.0555
180 | 6.35 | -3.2279 | 26701.5664 | 0.40744 | 2.227 | 102.3714
181 | 6.3853 | -3.1926 | 26622.8672 | 0.40624 | 2.2433 | 103.1204
182 | 6.4206 | -3.1574 | 26702.1211 | 0.40745 | 2.2269 | 102.3661
183 | 6.4558 | -3.1221 | 26609.791 | 0.40604 | 2.2461 | 103.2452
184 | 6.4911 | -3.0868 | 26625.6309 | 0.40628 | 2.2428 | 103.094
185 | 6.5264 | -3.0515 | 26603.2012 | 0.40594 | 2.2474 | 103.3082
186 | 6.5617 | -3.0163 | 26513.7734 | 0.40457 | 2.2661 | 104.1654
187 | 6.5969 | -2.981 | 26554.3848 | 0.40519 | 2.2576 | 103.7755
188 | 6.6322 | -2.9457 | 26682.7461 | 0.40715 | 2.2309 | 102.5501
189 | 6.6675 | -2.9104 | 26677.9863 | 0.40708 | 2.2319 | 102.5954
190 | 6.7028 | -2.8751 | 26680.9648 | 0.40713 | 2.2313 | 102.5671
191 | 6.7381 | -2.8399 | 26711.6172 | 0.40759 | 2.225 | 102.276
192 | 6.7733 | -2.8046 | 26654.6641 | 0.40672 | 2.2367 | 102.8173
193 | 6.8086 | -2.7693 | 26595.4199 | 0.40582 | 2.249 | 103.3826
194 | 6.8439 | -2.734 | 26444.5703 | 0.40352 | 2.2806 | 104.8324
195 | 6.8792 | -2.6988 | 26424.1445 | 0.40321 | 2.2849 | 105.0298
196 | 6.9144 | -2.6635 | 26570.2559 | 0.40544 | 2.2543 | 103.6234
197 | 6.9497 | -2.6282 | 26606.041 | 0.40598 | 2.2468 | 103.2811
198 | 6.985 | -2.5929 | 26627.0859 | 0.4063 | 2.2425 | 103.0801
199 | 7.0203 | -2.5576 | 26655.2949 | 0.40673 | 2.2366 | 102.8113
200 | 7.0556 | -2.5224 | 26710.5137 | 0.40758 | 2.2252 | 102.2865
201 | 7.0908 | -2.4871 | 26784.5996 | 0.40871 | 2.2099 | 101.5854
202 | 7.1261 | -2.4518 | 26754.1582 | 0.40824 | 2.2162 | 101.873
203 | 7.1614 | -2.4165 | 26696.3066 | 0.40736 | 2.2281 | 102.4213
204 | 7.1967 | -2.3813 | 26613.5898 | 0.4061 | 2.2453 | 103.209
205 | 7.2319 | -2.346 | 26728.3066 | 0.40785 | 2.2215 | 102.1178
206 | 7.2672 | -2.3107 | 26849.123 | 0.40969 | 2.1967 | 100.9777
207 | 7.3025 | -2.2754 | 26876.043 | 0.4101 | 2.1912 | 100.725
208 | 7.3378 | -2.2401 | 26796.6816 | 0.40889 | 2.2075 | 101.4714
209 | 7.3731 | -2.2049 | 26901.5234 | 0.41049 | 2.186 | 100.4862
210 | 7.4083 | -2.1696 | 26894.6699 | 0.41039 | 2.1874 | 100.5504
211 | 7.4436 | -2.1343 | 26895.8672 | 0.4104 | 2.1872 | 100.5391
212 | 7.4789 | -2.099 | 26895.3633 | 0.4104 | 2.1873 | 100.5439
213 | 7.5142 | -2.0638 | 26905.8945 | 0.41056 | 2.1851 | 100.4452
214 | 7.5494 | -2.0285 | 26957.1406 | 0.41134 | 2.1747 | 99.9663
215 | 7.5847 | -1.9932 | 26982.0762 | 0.41172 | 2.1697 | 99.7339
216 | 7.62 | -1.9579 | 26996.0391 | 0.41193 | 2.1668 | 99.6039
217 | 7.6553 | -1.9226 | 27045.1543 | 0.41268 | 2.1569 | 99.1477
218 | 7.6906 | -1.8874 | 27054.8125 | 0.41283 | 2.155 | 99.0581
219 | 7.7258 | -1.8521 | 27019.5645 | 0.41229 | 2.1621 | 99.3852
220 | 7.7611 | -1.8168 | 26996.25 | 0.41194 | 2.1668 | 99.6019
221 | 7.7964 | -1.7815 | 26988.2832 | 0.41181 | 2.1684 | 99.6761
222 | 7.8317 | -1.7463 | 27004.9785 | 0.41207 | 2.165 | 99.5208
223 | 7.8669 | -1.711 | 27043.625 | 0.41266 | 2.1572 | 99.1619
224 | 7.9022 | -1.6757 | 27048.375 | 0.41273 | 2.1563 | 99.1178
225 | 7.9375 | -1.6404 | 27053.9824 | 0.41282 | 2.1551 | 99.0658
226 | 7.9728 | -1.6051 | 27086.209 | 0.41331 | 2.1486 | 98.7675
227 | 8.0081 | -1.5699 | 27094.5664 | 0.41344 | 2.147 | 98.6902
228 | 8.0433 | -1.5346 | 27115.1934 | 0.41375 | 2.1428 | 98.4997
229 | 8.0786 | -1.4993 | 26974.2539 | 0.4116 | 2.1713 | 99.8068
230 | 8.1139 | -1.464 | 26914.5391 | 0.41069 | 2.1834 | 100.3643
231 | 8.1492 | -1.4288 | 26937.5078 | 0.41104 | 2.1787 | 100.1496
232 | 8.1844 | -1.3935 | 26938.0586 | 0.41105 | 2.1786 | 100.1445
233 | 8.2197 | -1.3582 | 26947.0918 | 0.41119 | 2.1768 | 100.0601
234 | 8.255 | -1.3229 | 27008.0703 | 0.41212 | 2.1644 | 99.492
235 | 8.2903 | -1.2876 | 26964.2715 | 0.41145 | 2.1733 | 99.8998
236 | 8.3256 | -1.2524 | 26910.9375 | 0.41063 | 2.1841 | 100.398
237 | 8.3608 | -1.2171 | 26936.3223 | 0.41102 | 2.179 | 100.1607
238 | 8.3961 | -1.1818 | 26878.0312 | 0.41013 | 2.1908 | 100.7063
239 | 8.4314 | -1.1465 | 26846.5996 | 0.40965 | 2.1972 | 101.0014
240 | 8.4667 | -1.1113 | 26866.209 | 0.40995 | 2.1932 | 100.8172
241 | 8.5019 | -1.076 | 26878.3008 | 0.41014 | 2.1908 | 100.7038
242 | 8.5372 | -1.0407 | 26878.4316 | 0.41014 | 2.1907 | 100.7026
243 | 8.5725 | -1.0054 | 26858.5996 | 0.40984 | 2.1948 | 100.8887
244 | 8.6078 | -0.97014 | 26865.9902 | 0.40995 | 2.1933 | 100.8193
245 | 8.6431 | -0.93486 | 26886.6016 | 0.41026 | 2.1891 | 100.626
246 | 8.6783 | -0.89958 | 26955.3223 | 0.41131 | 2.1751 | 99.9833
247 | 8.7136 | -0.86431 | 26987.7891 | 0.41181 | 2.1685 | 99.6807
248 | 8.7489 | -0.82903 | 27000.0137 | 0.41199 | 2.166 | 99.5669
249 | 8.7842 | -0.79375 | 26902.793 | 0.41051 | 2.1858 | 100.4743
250 | 8.8194 | -0.75847 | 26831.2148 | 0.40942 | 2.2004 | 101.1461
251 | 8.8547 | -0.72319 | 26889.7812 | 0.41031 | 2.1884 | 100.5962
252 | 8.89 | -0.68792 | 26841.4746 | 0.40957 | 2.1983 | 101.0496
253 | 8.9253 | -0.65264 | 26877.2617 | 0.41012 | 2.191 | 100.7135
254 | 8.9606 | -0.61736 | 26863.166 | 0.40991 | 2.1939 | 100.8458
255 | 8.9958 | -0.58208 | 26859.2324 | 0.40985 | 2.1947 | 100.8827
256 | 9.0311 | -0.54681 | 26919.9961 | 0.41077 | 2.1823 | 100.3133
257 | 9.0664 | -0.51153 | 26942.834 | 0.41112 | 2.1776 | 100.0999
258 | 9.1017 | -0.47625 | 26985.3164 | 0.41177 | 2.169 | 99.7037
259 | 9.1369 | -0.44097 | 26920.5254 | 0.41078 | 2.1822 | 100.3083
260 | 9.1722 | -0.40569 | 26947.2129 | 0.41119 | 2.1767 | 100.059
261 | 9.2075 | -0.37042 | 26965.9844 | 0.41147 | 2.1729 | 99.8838
262 | 9.2428 | -0.33514 | 26928.8066 | 0.41091 | 2.1805 | 100.2309
263 | 9.2781 | -0.29986 | 26942.0938 | 0.41111 | 2.1778 | 100.1068
264 | 9.3133 | -0.26458 | 26911.4355 | 0.41064 | 2.184 | 100.3934
265 | 9.3486 | -0.22931 | 26938.3145 | 0.41105 | 2.1785 | 100.1421
266 | 9.3839 | -0.19403 | 26973.0371 | 0.41158 | 2.1715 | 99.8181
267 | 9.4192 | -0.15875 | 26978.041 | 0.41166 | 2.1705 | 99.7715
268 | 9.4544 | -0.12347 | 26934.2656 | 0.41099 | 2.1794 | 100.1799
269 | 9.4897 | -0.088194 | 26892.0137 | 0.41035 | 2.188 | 100.5752
270 | 9.525 | -0.052917 | 26914.3281 | 0.41069 | 2.1834 | 100.3663
271 | 9.5603 | -0.017639 | 26953.5312 | 0.41128 | 2.1755 | 100
272 | 9.5956 | 0.017639 | 26961.7734 | 0.41141 | 2.1738 | 99.9231
273 | 9.6308 | 0.052917 | 26928.8242 | 0.41091 | 2.1805 | 100.2307
274 | 9.6661 | 0.088194 | 26898.0488 | 0.41044 | 2.1867 | 100.5187
275 | 9.7014 | 0.12347 | 26910.082 | 0.41062 | 2.1843 | 100.406
276 | 9.7367 | 0.15875 | 26897.1934 | 0.41042 | 2.1869 | 100.5267
277 | 9.7719 | 0.19403 | 26919.9395 | 0.41077 | 2.1823 | 100.3138
278 | 9.8072 | 0.22931 | 27000.2969 | 0.412 | 2.166 | 99.5643
279 | 9.8425 | 0.26458 | 26973.5508 | 0.41159 | 2.1714 | 99.8133
280 | 9.8778 | 0.29986 | 26999.7422 | 0.41199 | 2.1661 | 99.5695
281 | 9.9131 | 0.33514 | 26981.457 | 0.41171 | 2.1698 | 99.7397
282 | 9.9483 | 0.37042 | 27031.9102 | 0.41248 | 2.1596 | 99.2705
283 | 9.9836 | 0.40569 | 27027.6719 | 0.41242 | 2.1604 | 99.3099
284 | 10.0189 | 0.44097 | 27031.75 | 0.41248 | 2.1596 | 99.272
285 | 10.0542 | 0.47625 | 27078.4277 | 0.41319 | 2.1502 | 98.8394
286 | 10.0894 | 0.51153 | 27101.0117 | 0.41353 | 2.1457 | 98.6306
287 | 10.1247 | 0.54681 | 27063.0195 | 0.41296 | 2.1533 | 98.9821
288 | 10.16 | 0.58208 | 27013.1387 | 0.41219 | 2.1634 | 99.4449
289 | 10.1953 | 0.61736 | 27024.2773 | 0.41236 | 2.1611 | 99.3414
290 | 10.2306 | 0.65264 | 27066.7246 | 0.41301 | 2.1526 | 98.9478
291 | 10.2658 | 0.68792 | 27056.5312 | 0.41286 | 2.1546 | 99.0422
292 | 10.3011 | 0.72319 | 27027.375 | 0.41241 | 2.1605 | 99.3127
293 | 10.3364 | 0.75847 | 27004.8926 | 0.41207 | 2.165 | 99.5216
294 | 10.3717 | 0.79375 | 27061.3184 | 0.41293 | 2.1537 | 98.9978
295 | 10.4069 | 0.82903 | 27045.6797 | 0.41269 | 2.1568 | 99.1428
296 | 10.4422 | 0.86431 | 27010.4395 | 0.41215 | 2.1639 | 99.47
297 | 10.4775 | 0.89958 | 27097.0547 | 0.41347 | 2.1465 | 98.6672
298 | 10.5128 | 0.93486 | 27143.2734 | 0.41418 | 2.1372 | 98.2407
299 | 10.5481 | 0.97014 | 27143.2969 | 0.41418 | 2.1372 | 98.2405
300 | 10.5833 | 1.0054 | 27111.1602 | 0.41369 | 2.1436 | 98.5369
301 | 10.6186 | 1.0407 | 27045.0547 | 0.41268 | 2.1569 | 99.1486
302 | 10.6539 | 1.076 | 27109.1777 | 0.41366 | 2.144 | 98.5552
303 | 10.6892 | 1.1113 | 27039.9746 | 0.4126 | 2.158 | 99.1957
304 | 10.7244 | 1.1465 | 27083.0996 | 0.41326 | 2.1493 | 98.7962
305 | 10.7597 | 1.1818 | 27077.207 | 0.41317 | 2.1505 | 98.8507
306 | 10.795 | 1.2171 | 27032.7598 | 0.41249 | 2.1594 | 99.2627
307 | 10.8303 | 1.2524 | 27081.5527 | 0.41324 | 2.1496 | 98.8105
308 | 10.8656 | 1.2876 | 27063.2559 | 0.41296 | 2.1533 | 98.9799
309 | 10.9008 | 1.3229 | 27038.2188 | 0.41258 | 2.1583 | 99.212
310 | 10.9361 | 1.3582 | 27018.0918 | 0.41227 | 2.1624 | 99.3989
311 | 10.9714 | 1.3935 | 27061.6582 | 0.41293 | 2.1536 | 98.9947
312 | 11.0067 | 1.4287 | 27050.1016 | 0.41276 | 2.1559 | 99.1018
313 | 11.0419 | 1.464 | 27015.6074 | 0.41223 | 2.1629 | 99.422
314 | 11.0772 | 1.4993 | 26948.791 | 0.41121 | 2.1764 | 100.0442
315 | 11.1125 | 1.5346 | 26961.6738 | 0.41141 | 2.1738 | 99.924
316 | 11.1478 | 1.5699 | 26955.6367 | 0.41132 | 2.175 | 99.9804
317 | 11.1831 | 1.6051 | 26902.7969 | 0.41051 | 2.1858 | 100.4742
318 | 11.2183 | 1.6404 | 26933.1348 | 0.41097 | 2.1796 | 100.1905
319 | 11.2536 | 1.6757 | 26968.1113 | 0.41151 | 2.1725 | 99.864
320 | 11.2889 | 1.711 | 26979.0664 | 0.41167 | 2.1703 | 99.7619
321 | 11.3242 | 1.7462 | 26987.1348 | 0.4118 | 2.1686 | 99.6868
322 | 11.3594 | 1.7815 | 26993.3203 | 0.41189 | 2.1674 | 99.6292
323 | 11.3947 | 1.8168 | 26960.959 | 0.4114 | 2.174 | 99.9307
324 | 11.43 | 1.8521 | 26896.9297 | 0.41042 | 2.187 | 100.5292
325 | 11.4653 | 1.8874 | 26930.6035 | 0.41093 | 2.1801 | 100.2141
326 | 11.5006 | 1.9226 | 26947.9102 | 0.4112 | 2.1766 | 100.0525
327 | 11.5358 | 1.9579 | 26899.6758 | 0.41046 | 2.1864 | 100.5035
328 | 11.5711 | 1.9932 | 26922.3125 | 0.41081 | 2.1818 | 100.2916
329 | 11.6064 | 2.0285 | 26894.4941 | 0.41038 | 2.1875 | 100.552
330 | 11.6417 | 2.0637 | 26892.377 | 0.41035 | 2.1879 | 100.5718
331 | 11.6769 | 2.099 | 26916.3398 | 0.41072 | 2.183 | 100.3475
332 | 11.7122 | 2.1343 | 26895.3652 | 0.4104 | 2.1873 | 100.5438
333 | 11.7475 | 2.1696 | 26939.0391 | 0.41106 | 2.1784 | 100.1353
334 | 11.7828 | 2.2049 | 26929.2891 | 0.41091 | 2.1804 | 100.2264
335 | 11.8181 | 2.2401 | 26958.2188 | 0.41136 | 2.1745 | 99.9563
336 | 11.8533 | 2.2754 | 26951 | 0.41125 | 2.176 | 100.0236
337 | 11.8886 | 2.3107 | 26966.4551 | 0.41148 | 2.1728 | 99.8795
338 | 11.9239 | 2.346 | 26976.8555 | 0.41164 | 2.1707 | 99.7825
339 | 11.9592 | 2.3812 | 26969.3965 | 0.41153 | 2.1722 | 99.852
340 | 11.9944 | 2.4165 | 26953.2344 | 0.41128 | 2.1755 | 100.0028
341 | 12.0297 | 2.4518 | 26975.2344 | 0.41162 | 2.1711 | 99.7976
342 | 12.065 | 2.4871 | 27001.9727 | 0.41202 | 2.1656 | 99.5487
343 | 12.1003 | 2.5224 | 26984.4531 | 0.41176 | 2.1692 | 99.7118
344 | 12.1356 | 2.5576 | 26957.7539 | 0.41135 | 2.1746 | 99.9606
345 | 12.1708 | 2.5929 | 26974.2734 | 0.4116 | 2.1712 | 99.8066
346 | 12.2061 | 2.6282 | 26966.1152 | 0.41148 | 2.1729 | 99.8826
347 | 12.2414 | 2.6635 | 27026.8809 | 0.4124 | 2.1606 | 99.3172
348 | 12.2767 | 2.6987 | 27031.6016 | 0.41248 | 2.1597 | 99.2734
349 | 12.3119 | 2.734 | 27089.5312 | 0.41336 | 2.148 | 98.7367
350 | 12.3472 | 2.7693 | 27068.8301 | 0.41304 | 2.1521 | 98.9283
351 | 12.3825 | 2.8046 | 27101.2812 | 0.41354 | 2.1456 | 98.6281
352 | 12.4178 | 2.8399 | 27176.5996 | 0.41469 | 2.1305 | 97.934
353 | 12.4531 | 2.8751 | 27132.041 | 0.41401 | 2.1394 | 98.3442
354 | 12.4883 | 2.9104 | 27041.3184 | 0.41262 | 2.1577 | 99.1832
355 | 12.5236 | 2.9457 | 27075.8594 | 0.41315 | 2.1507 | 98.8632
356 | 12.5589 | 2.981 | 27056.5684 | 0.41286 | 2.1546 | 99.0419
357 | 12.5942 | 3.0162 | 27090.1426 | 0.41337 | 2.1479 | 98.7311
358 | 12.6294 | 3.0515 | 27055.084 | 0.41283 | 2.1549 | 99.0556
359 | 12.6647 | 3.0868 | 26990.0488 | 0.41184 | 2.1681 | 99.6597
360 | 12.7 | 3.1221 | 26944.6152 | 0.41115 | 2.1773 | 100.0832
361 | 12.7353 | 3.1574 | 26923.4668 | 0.41083 | 2.1816 | 100.2808
362 | 12.7706 | 3.1926 | 26877.3457 | 0.41012 | 2.191 | 100.7128
363 | 12.8058 | 3.2279 | 26887.0898 | 0.41027 | 2.189 | 100.6214
364 | 12.8411 | 3.2632 | 26905.1934 | 0.41055 | 2.1853 | 100.4518
365 | 12.8764 | 3.2985 | 26924.6777 | 0.41084 | 2.1813 | 100.2695
366 | 12.9117 | 3.3337 | 26957.2383 | 0.41134 | 2.1747 | 99.9654
367 | 12.9469 | 3.369 | 26976.5723 | 0.41164 | 2.1708 | 99.7852
368 | 12.9822 | 3.4043 | 26967.5195 | 0.4115 | 2.1726 | 99.8695
369 | 13.0175 | 3.4396 | 27014.9219 | 0.41222 | 2.163 | 99.4283
370 | 13.0528 | 3.4749 | 26948.5195 | 0.41121 | 2.1765 | 100.0468
371 | 13.0881 | 3.5101 | 26942.0586 | 0.41111 | 2.1778 | 100.1071
372 | 13.1233 | 3.5454 | 26953.5742 | 0.41129 | 2.1754 | 99.9996
373 | 13.1586 | 3.5807 | 26905.498 | 0.41055 | 2.1852 | 100.4489
374 | 13.1939 | 3.616 | 26816.2656 | 0.40919 | 2.2035 | 101.2868
375 | 13.2292 | 3.6512 | 26883.4043 | 0.41021 | 2.1897 | 100.6559
376 | 13.2644 | 3.6865 | 26918.5645 | 0.41075 | 2.1826 | 100.3267
377 | 13.2997 | 3.7218 | 27003.0254 | 0.41204 | 2.1654 | 99.5389
378 | 13.335 | 3.7571 | 26977.8477 | 0.41166 | 2.1705 | 99.7733
379 | 13.3703 | 3.7924 | 26926.875 | 0.41088 | 2.1809 | 100.249
380 | 13.4056 | 3.8276 | 26898.7695 | 0.41045 | 2.1866 | 100.5119
381 | 13.4408 | 3.8629 | 26843.8789 | 0.40961 | 2.1978 | 101.027
382 | 13.4761 | 3.8982 | 26863.1699 | 0.40991 | 2.1939 | 100.8458
383 | 13.5114 | 3.9335 | 26833.5996 | 0.40945 | 2.1999 | 101.1237
384 | 13.5467 | 3.9688 | 26833.3125 | 0.40945 | 2.2 | 101.1264
385 | 13.5819 | 4.004 | 26891.1934 | 0.41033 | 2.1881 | 100.5829
386 | 13.6172 | 4.0393 | 26920.873 | 0.41079 | 2.1821 | 100.3051
387 | 13.6525 | 4.0746 | 26893.5352 | 0.41037 | 2.1877 | 100.561
388 | 13.6878 | 4.1099 | 26892.6348 | 0.41036 | 2.1878 | 100.5694
389 | 13.7231 | 4.1451 | 26917.8926 | 0.41074 | 2.1827 | 100.333
390 | 13.7583 | 4.1804 | 26926.3535 | 0.41087 | 2.181 | 100.2538
391 | 13.7936 | 4.2157 | 26951.7754 | 0.41126 | 2.1758 | 100.0164
392 | 13.8289 | 4.251 | 26962.6641 | 0.41142 | 2.1736 | 99.9148
393 | 13.8642 | 4.2862 | 27004.1836 | 0.41206 | 2.1652 | 99.5282
394 | 13.8994 | 4.3215 | 26993.75 | 0.4119 | 2.1673 | 99.6252
395 | 13.9347 | 4.3568 | 27032.8184 | 0.41249 | 2.1594 | 99.2621
396 | 13.97 | 4.3921 | 27152.1816 | 0.41432 | 2.1354 | 98.1587
397 | 14.0053 | 4.4274 | 27096.752 | 0.41347 | 2.1465 | 98.67
398 | 14.0406 | 4.4626 | 27094.8965 | 0.41344 | 2.1469 | 98.6871
399 | 14.0758 | 4.4979 | 27120.7266 | 0.41384 | 2.1417 | 98.4486
400 | 14.1111 | 4.5332 | 27180.2754 | 0.41474 | 2.1298 | 97.9002
401 | 14.1464 | 4.5685 | 27208.8145 | 0.41518 | 2.1241 | 97.6382
402 | 14.1817 | 4.6037 | 27252.1445 | 0.41584 | 2.1154 | 97.2414
403 | 14.2169 | 4.639 | 27309.5 | 0.41672 | 2.1041 | 96.7178
404 | 14.2522 | 4.6743 | 27457.1855 | 0.41897 | 2.0749 | 95.3788
405 | 14.2875 | 4.7096 | 27601.498 | 0.42117 | 2.0467 | 94.0831
406 | 14.3228 | 4.7449 | 27763.8047 | 0.42365 | 2.0154 | 92.6405
407 | 14.3581 | 4.7801 | 27943.0312 | 0.42638 | 1.9811 | 91.0652
408 | 14.3933 | 4.8154 | 28208.8555 | 0.43044 | 1.931 | 88.7626
409 | 14.4286 | 4.8507 | 28573.7539 | 0.43601 | 1.8636 | 85.6656
410 | 14.4639 | 4.886 | 29018.498 | 0.44279 | 1.7836 | 81.9874
411 | 14.4992 | 4.9212 | 29651.0859 | 0.45245 | 1.6736 | 76.9301
412 | 14.5344 | 4.9565 | 30487.6309 | 0.46521 | 1.5345 | 70.5388
413 | 14.5697 | 4.9918 | 31451.9844 | 0.47993 | 1.3827 | 63.5601
414 | 14.605 | 5.0271 | 32580.0566 | 0.49714 | 1.2157 | 55.8814
415 | 14.6403 | 5.0624 | 33882.1875 | 0.51701 | 1.0357 | 47.6075
416 | 14.6756 | 5.0976 | 35068.043 | 0.5351 | 0.88257 | 40.5694
417 | 14.7108 | 5.1329 | 36045.7969 | 0.55002 | 0.76339 | 35.0909
418 | 14.7461 | 5.1682 | 36892.8633 | 0.56295 | 0.6649 | 30.5636
419 | 14.7814 | 5.2035 | 37680.5469 | 0.57497 | 0.57703 | 26.5244
420 | 14.8167 | 5.2387 | 38367.6016 | 0.58545 | 0.50314 | 23.1278
421 | 14.8519 | 5.274 | 38934.4844 | 0.5941 | 0.44401 | 20.4099
422 | 14.8872 | 5.3093 | 39389.3906 | 0.60104 | 0.39771 | 18.2818
423 | 14.9225 | 5.3446 | 39807.2031 | 0.60742 | 0.35607 | 16.3674
424 | 14.9578 | 5.3799 | 40169.2227 | 0.61294 | 0.32064 | 14.7389
425 | 14.9931 | 5.4151 | 40430.6523 | 0.61693 | 0.29542 | 13.5799
426 | 15.0283 | 5.4504 | 40707.0156 | 0.62115 | 0.2691 | 12.3698
427 | 15.0636 | 5.4857 | 41045.4414 | 0.62631 | 0.23732 | 10.9089
428 | 15.0989 | 5.521 | 41188.3906 | 0.62849 | 0.22404 | 10.2986
429 | 15.1342 | 5.5562 | 41227.918 | 0.6291 | 0.22039 | 10.1306
430 | 15.1694 | 5.5915 | 41345.9766 | 0.6309 | 0.20951 | 9.6304
431 | 15.2047 | 5.6268 | 41424.5625 | 0.6321 | 0.2023 | 9.299
432 | 15.24 | 5.6621 | 41509.9297 | 0.6334 | 0.19449 | 8.9403
433 | 15.2753 | 5.6974 | 41593.3555 | 0.63467 | 0.1869 | 8.5911
434 | 15.3106 | 5.7326 | 41674.3359 | 0.63591 | 0.17955 | 8.2534
435 | 15.3458 | 5.7679 | 41723.6992 | 0.63666 | 0.17508 | 8.0481
436 | 15.3811 | 5.8032 | 41813.8945 | 0.63804 | 0.16695 | 7.6743
437 | 15.4164 | 5.8385 | 41885.4258 | 0.63913 | 0.16052 | 7.3788
438 | 15.4517 | 5.8737 | 41939.7773 | 0.63996 | 0.15565 | 7.155
439 | 15.4869 | 5.909 | 41985.1562 | 0.64065 | 0.1516 | 6.9685
440 | 15.5222 | 5.9443 | 42063.1719 | 0.64184 | 0.14464 | 6.6489
441 | 15.5575 | 5.9796 | 41995.4453 | 0.64081 | 0.15068 | 6.9263
442 | 15.5928 | 6.0149 | 41949.6797 | 0.64011 | 0.15477 | 7.1143
443 | 15.6281 | 6.0501 | 42160.5781 | 0.64333 | 0.13599 | 6.2513
444 | 15.6633 | 6.0854 | 42197.7617 | 0.6439 | 0.1327 | 6.1
445 | 15.6986 | 6.1207 | 42185.4414 | 0.64371 | 0.13379 | 6.1501
446 | 15.7339 | 6.156 | 42268.7695 | 0.64498 | 0.12643 | 5.8117
447 | 15.7692 | 6.1912 | 42370.0898 | 0.64653 | 0.11752 | 5.4019
448 | 15.8044 | 6.2265 | 42351.457 | 0.64624 | 0.11915 | 5.4772
449 | 15.8397 | 6.2618 | 42350.6445 | 0.64623 | 0.11922 | 5.4804
450 | 15.875 | 6.2971 | 42367.8164 | 0.64649 | 0.11772 | 5.4111
451 | 15.9103 | 6.3324 | 42407.5352 | 0.6471 | 0.11423 | 5.251
452 | 15.9456 | 6.3676 | 42479.5938 | 0.6482 | 0.10793 | 4.9612
453 | 15.9808 | 6.4029 | 42436.207 | 0.64754 | 0.11172 | 5.1355
454 | 16.0161 | 6.4382 | 42349.7656 | 0.64622 | 0.1193 | 5.484
455 | 16.0514 | 6.4735 | 42345.8906 | 0.64616 | 0.11964 | 5.4996
456 | 16.0867 | 6.5087 | 42392.5 | 0.64687 | 0.11555 | 5.3116
457 | 16.1219 | 6.544 | 42453.7188 | 0.6478 | 0.11019 | 5.0651
458 | 16.1572 | 6.5793 | 42460.3008 | 0.6479 | 0.10961 | 5.0387
459 | 16.1925 | 6.6146 | 42477.8906 | 0.64817 | 0.10808 | 4.968
460 | 16.2278 | 6.6499 | 42496.9141 | 0.64846 | 0.10642 | 4.8916
461 | 16.2631 | 6.6851 | 42498.0625 | 0.64848 | 0.10632 | 4.887
462 | 16.2983 | 6.7204 | 42494.625 | 0.64843 | 0.10662 | 4.9008
463 | 16.3336 | 6.7557 | 42430.0703 | 0.64744 | 0.11226 | 5.1602
464 | 16.3689 | 6.791 | 42490.5117 | 0.64836 | 0.10697 | 4.9173
465 | 16.4042 | 6.8262 | 42626.6758 | 0.65044 | 0.095121 | 4.3725
466 | 16.4394 | 6.8615 | 42643.0312 | 0.65069 | 0.093702 | 4.3073
467 | 16.4747 | 6.8968 | 42584.4102 | 0.6498 | 0.098793 | 4.5413
468 | 16.51 | 6.9321 | 42656.1328 | 0.65089 | 0.092567 | 4.255
469 | 16.5453 | 6.9674 | 42617.2891 | 0.6503 | 0.095936 | 4.4099
470 | 16.5806 | 7.0026 | 42651.3516 | 0.65082 | 0.092981 | 4.2741
471 | 16.6158 | 7.0379 | 42764.0117 | 0.65254 | 0.083238 | 3.8262
472 | 16.6511 | 7.0732 | 42746.0977 | 0.65226 | 0.084784 | 3.8973
473 | 16.6864 | 7.1085 | 42773.5625 | 0.65268 | 0.082415 | 3.7884
474 | 16.7217 | 7.1438 | 42785.0273 | 0.65286 | 0.081426 | 3.743
475 | 16.7569 | 7.179 | 42808.2383 | 0.65321 | 0.079427 | 3.651
476 | 16.7922 | 7.2143 | 42856.9492 | 0.65396 | 0.075238 | 3.4585
477 | 16.8275 | 7.2496 | 42819.2812 | 0.65338 | 0.078476 | 3.6074
478 | 16.8628 | 7.2849 | 42835.3594 | 0.65363 | 0.077093 | 3.5438
479 | 16.8981 | 7.3201 | 42857.5508 | 0.65396 | 0.075186 | 3.4561
480 | 16.9333 | 7.3554 | 42883.8633 | 0.65437 | 0.072927 | 3.3522
481 | 16.9686 | 7.3907 | 42847.4023 | 0.65381 | 0.076058 | 3.4962
482 | 17.0039 | 7.426 | 42848.0508 | 0.65382 | 0.076002 | 3.4936
483 | 17.0392 | 7.4612 | 42840.9648 | 0.65371 | 0.076611 | 3.5216
484 | 17.0744 | 7.4965 | 42822.1445 | 0.65342 | 0.07823 | 3.596
485 | 17.1097 | 7.5318 | 42847.8906 | 0.65382 | 0.076016 | 3.4943
486 | 17.145 | 7.5671 | 42921.4609 | 0.65494 | 0.069703 | 3.2041
487 | 17.1803 | 7.6024 | 42924.0508 | 0.65498 | 0.069481 | 3.1939
488 | 17.2156 | 7.6376 | 42955.9805 | 0.65547 | 0.066748 | 3.0682
489 | 17.2508 | 7.6729 | 42921.6289 | 0.65494 | 0.069689 | 3.2034
490 | 17.2861 | 7.7082 | 42969.5039 | 0.65567 | 0.065592 | 3.0151
491 | 17.3214 | 7.7435 | 42953.5391 | 0.65543 | 0.066957 | 3.0778
492 | 17.3567 | 7.7787 | 42955.5234 | 0.65546 | 0.066787 | 3.07
493 | 17.3919 | 7.814 | 43026.6875 | 0.65655 | 0.060709 | 2.7906
494 | 17.4272 | 7.8493 | 43012.2383 | 0.65632 | 0.061942 | 2.8473
495 | 17.4625 | 7.8846 | 42891.3711 | 0.65448 | 0.072283 | 3.3226
496 | 17.4978 | 7.9199 | 42894.5703 | 0.65453 | 0.072008 | 3.31
497 | 17.5331 | 7.9551 | 42869.0469 | 0.65414 | 0.074199 | 3.4107
498 | 17.5683 | 7.9904 | 42941.9453 | 0.65525 | 0.067949 | 3.1234
499 | 17.6036 | 8.0257 | 42948.418 | 0.65535 | 0.067395 | 3.098
500 | 17.6389 | 8.061 | 42954.6523 | 0.65545 | 0.066862 | 3.0735
501 | 17.6742 | 8.0962 | 42987.5 | 0.65595 | 0.064054 | 2.9444
502 | 17.7094 | 8.1315 | 43003.8164 | 0.6562 | 0.062661 | 2.8803
503 | 17.7447 | 8.1668 | 43019.6367 | 0.65644 | 0.061311 | 2.8183
504 | 17.78 | 8.2021 | 43026.1992 | 0.65654 | 0.060751 | 2.7926
505 | 17.8153 | 8.2374 | 43078.6836 | 0.65734 | 0.05628 | 2.5871
506 | 17.8506 | 8.2726 | 43077.6797 | 0.65732 | 0.056366 | 2.591
507 | 17.8858 | 8.3079 | 43075.4297 | 0.65729 | 0.056557 | 2.5998
508 | 17.9211 | 8.3432 | 43089.2188 | 0.6575 | 0.055384 | 2.5459
509 | 17.9564 | 8.3785 | 43113.0664 | 0.65786 | 0.053357 | 2.4527
510 | 17.9917 | 8.4138 | 43144.7031 | 0.65835 | 0.050672 | 2.3292
511 | 18.0269 | 8.449 | 43127.5586 | 0.65808 | 0.052127 | 2.3961
512 | 18.0622 | 8.4843 | 43151.8984 | 0.65846 | 0.050061 | 2.3012
513 | 18.0975 | 8.5196 | 43113.8789 | 0.65788 | 0.053288 | 2.4495
514 | 18.1328 | 8.5549 | 43197.4492 | 0.65915 | 0.046202 | 2.1238
515 | 18.1681 | 8.5901 | 43123.082 | 0.65802 | 0.052507 | 2.4136
516 | 18.2033 | 8.6254 | 43135.9414 | 0.65821 | 0.051415 | 2.3634
517 | 18.2386 | 8.6607 | 43121.3867 | 0.65799 | 0.052651 | 2.4202
518 | 18.2739 | 8.696 | 43132.2148 | 0.65816 | 0.051731 | 2.3779
519 | 18.3092 | 8.7312 | 43223.25 | 0.65954 | 0.044019 | 2.0235
520 | 18.3444 | 8.7665 | 43220 | 0.65949 | 0.044294 | 2.0361
521 | 18.3797 | 8.8018 | 43226.957 | 0.6596 | 0.043706 | 2.009
522 | 18.415 | 8.8371 | 43263.3906 | 0.66016 | 0.040629 | 1.8676
523 | 18.4503 | 8.8724 | 43187.6094 | 0.659 | 0.047035 | 2.1621
524 | 18.4856 | 8.9076 | 43162.625 | 0.65862 | 0.049152 | 2.2594
525 | 18.5208 | 8.9429 | 43125.043 | 0.65805 | 0.05234 | 2.4059
526 | 18.5561 | 8.9782 | 43090.8594 | 0.65752 | 0.055245 | 2.5395
527 | 18.5914 | 9.0135 | 43050.1641 | 0.6569 | 0.058708 | 2.6987
528 | 18.6267 | 9.0487 | 43081.9961 | 0.65739 | 0.055999 | 2.5741
529 | 18.6619 | 9.084 | 43073.6914 | 0.65726 | 0.056705 | 2.6066
530 | 18.6972 | 9.1193 | 43069.2461 | 0.65719 | 0.057084 | 2.624
531 | 18.7325 | 9.1546 | 43072.918 | 0.65725 | 0.056771 | 2.6096
532 | 18.7678 | 9.1899 | 43008.8164 | 0.65627 | 0.062234 | 2.8607
533 | 18.8031 | 9.2251 | 43039.5625
534 | 18.8383 | 9.2604 | 43064.5117
535 | 18.8736 | 9.2957 | 43037.4141
536 | 18.9089 | 9.331 | 43064.6094
537 | 18.9442 | 9.3662 | 43067.8984
538 | 18.9794 | 9.4015 | 43094.1328
539 | 19.0147 | 9.4368 | 43090.7969
540 | 19.05 | 9.4721 | 43202.5469
541 | 19.0853 | 9.5074 | 43069.6719
542 | 19.1206 | 9.5426 | 43151.3828
543 | 19.1558 | 9.5779 | 43146.3555

## Sheet1
0 | 0 | 42281.41
5 | 0.5 | 38866.02
10 | 1 | 36454.16
15 | 2 | 31724.05
20 | 4 | 25538.42
25 | 7 | 19556.78
30 | 9 | 17246.94
35 | 10.5 | 15622.84
40 | 12 | 14459.44
45 | 13.5 | 13257.07
50 | 15 | 12040.34
55 | 19.5 | 9853.722

## uniformity over various heights
The scan area, which is defined by the template, the inner borders of the template to be more precise, is 24cm long, 12 cm to the centre. So my plan is to measure in 5 areas across the scanner plate.
The five areas are actually planes across the scanner bed in the lateral direction; the planes are defined by the labels thusly
The idea is then to scan the step wedge film at the corresponding heights with the following particular steps that correspond to the approximate doses of 0Gy, 1Gy and 14Gy.
The heights were marked on to the scan plate template.
The steps that correspond to the three doses of 0 1 and 14 Gy are
1Gy; the fourth strip after the large black strip at the top of the film.
14Gy is the tenth strip after the large black strip at the top of the film.
it is not possible to obtain a 0Gy strip using this step wedge; a blank film can be scanned for this purpose.
height | 1Gy (4th strip) | 14Gy (tenth strip)
middle | y
4.5cm up
11.5cm up
4.5cm down
11.5cm down
| middle 1Gy |  | 4.5cm up 1Gy |  | 4.5cm down 1Gy |  | 11.5cm up 1Gy |  | 11.5cm down 1Gy |  | middle 14Gy |  | 4.5cm up 14Gy |  | 10cm up 14Gy |  | 10cm up 14Gy; scanner covered |  | 4.5cm down 14Gy |  | 4.5cm down 14Gy; covered |  | 10cm down 14Gy
| cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy
| -9.4735 | 123.951 | -9.579 | 130.1937 | -9.5087 | 127.4692 | -9.5087 | 127.41 | -9.5614 | 131.3782 | -9.4911 | 1562.2456 | -9.579 | 1552.5978 | -9.579 | 1549.2523 | -9.6141 | 1554.7835 | -9.6141 | 1562.9289 | -9.6317 | 1572.9293 | -9.6317 | 1554.7835
| -9.4384 | 123.951 | -9.5438 | 129.3252 | -9.4735 | 127.4692 | -9.4735 | 126.9557 | -9.5262 | 131.2006 | -9.4559 | 1562.2456 | -9.5438 | 1552.5978 | -9.5438 | 1548.2527 | -9.579 | 1554.7835 | -9.579 | 1560.0305 | -9.5966 | 1569.9802 | -9.5966 | 1557.147
| -9.4032 | 124.0103 | -9.5087 | 129.3054 | -9.4384 | 127.6272 | -9.4384 | 126.7385 | -9.4911 | 130.9044 | -9.4208 | 1560.8814 | -9.5087 | 1549.4191 | -9.5087 | 1547.4211 | -9.5438 | 1554.4466 | -9.5438 | 1559.0111 | -9.5614 | 1568.5976 | -9.5614 | 1552.4301
| -9.3681 | 124.0103 | -9.4735 | 129.3054 | -9.4032 | 127.8839 | -9.4032 | 126.3433 | -9.4559 | 129.8384 | -9.3856 | 1559.1808 | -9.4735 | 1546.5908 | -9.4735 | 1543.2825 | -9.5087 | 1550.7553 | -9.5087 | 1557.9935 | -9.5262 | 1564.6408 | -9.5262 | 1548.5857
| -9.3329 | 124.3467 | -9.4384 | 129.2659 | -9.3681 | 127.9629 | -9.3681 | 125.9481 | -9.4208 | 128.9698 | -9.3505 | 1555.1205 | -9.4384 | 1551.0898 | -9.4384 | 1542.6233 | -9.4735 | 1547.2549 | -9.4735 | 1556.9778 | -9.4911 | 1561.2221 | -9.4911 | 1547.7536
| -9.2978 | 124.6434 | -9.4032 | 129.3054 | -9.3329 | 128.5947 | -9.3329 | 125.6517 | -9.3856 | 128.0222 | -9.3153 | 1551.7595 | -9.4032 | 1551.0898 | -9.4032 | 1543.6124 | -9.4384 | 1546.259 | -9.4384 | 1556.9778 | -9.4559 | 1559.6905 | -9.4559 | 1545.265
| -9.2626 | 124.5445 | -9.3681 | 129.6608 | -9.2978 | 128.1011 | -9.2978 | 126.1458 | -9.3505 | 127.5285 | -9.2802 | 1546.9228 | -9.3681 | 1553.1014 | -9.3681 | 1543.2825 | -9.4032 | 1544.1076 | -9.4032 | 1555.2891 | -9.4208 | 1556.3017 | -9.4208 | 1544.9341
| -9.2275 | 124.228 | -9.3329 | 128.9501 | -9.2626 | 127.5482 | -9.2626 | 126.1458 | -9.3153 | 127.41 | -9.245 | 1545.4305 | -9.3329 | 1552.0947 | -9.3329 | 1543.2825 | -9.3681 | 1543.9425 | -9.3681 | 1552.5978 | -9.3856 | 1555.1205 | -9.3856 | 1544.9341
| -9.1923 | 124.228 | -9.2978 | 127.9827 | -9.2275 | 126.5014 | -9.2275 | 125.9086 | -9.2802 | 127.41 | -9.2099 | 1545.7618 | -9.2978 | 1551.0898 | -9.2978 | 1542.9528 | -9.3329 | 1542.9528 | -9.3329 | 1551.4246 | -9.3505 | 1553.7736 | -9.3505 | 1543.9425
| -9.1571 | 124.0697 | -9.2626 | 127.8642 | -9.1923 | 126.5014 | -9.1923 | 125.9086 | -9.245 | 127.0545 | -9.1747 | 1547.7536 | -9.2626 | 1547.2549 | -9.2626 | 1541.3074 | -9.2978 | 1542.9528 | -9.2978 | 1550.0868 | -9.3153 | 1552.0947 | -9.3153 | 1541.965
| -9.122 | 124.0103 | -9.2275 | 127.5285 | -9.1571 | 126.5014 | -9.1571 | 125.8098 | -9.2099 | 127.1335 | -9.1396 | 1548.0863 | -9.2275 | 1546.4249 | -9.2275 | 1537.869 | -9.2626 | 1542.6233 | -9.2626 | 1549.7528 | -9.2802 | 1551.592 | -9.2802 | 1540.979
| -9.0868 | 123.5155 | -9.1923 | 128.0024 | -9.122 | 127.568 | -9.122 | 125.0191 | -9.1747 | 127.6667 | -9.1044 | 1548.9189 | -9.1923 | 1545.9275 | -9.1923 | 1536.2397 | -9.2275 | 1540.979 | -9.2275 | 1550.0868 | -9.245 | 1550.5881 | -9.245 | 1538.8491
| -9.0517 | 122.9017 | -9.1571 | 128.0024 | -9.0868 | 126.9557 | -9.0868 | 124.7225 | -9.1396 | 127.7655 | -9.0693 | 1551.592 | -9.1571 | 1545.9275 | -9.1571 | 1537.2167 | -9.1923 | 1541.3074 | -9.1923 | 1549.7528 | -9.2099 | 1550.0868 | -9.2099 | 1537.0537
| -9.0165 | 122.0298 | -9.122 | 128.4565 | -9.0517 | 126.9557 | -9.0517 | 124.7422 | -9.1044 | 127.5087 | -9.0341 | 1553.7736 | -9.122 | 1545.5961 | -9.122 | 1538.8491 | -9.1571 | 1541.3074 | -9.1571 | 1549.9198 | -9.1747 | 1550.0868 | -9.1747 | 1536.5652
| -8.9814 | 122.0298 | -9.0868 | 128.3578 | -9.0165 | 125.6912 | -9.0165 | 124.7225 | -9.0693 | 127.489 | -8.999 | 1553.7736 | -9.0868 | 1544.1076 | -9.0868 | 1539.6673 | -9.122 | 1543.2825 | -9.122 | 1550.9225 | -9.1396 | 1551.9271 | -9.1396 | 1535.2646
| -8.9462 | 122.1091 | -9.0517 | 127.5087 | -8.9814 | 125.5924 | -8.9814 | 124.7225 | -9.0341 | 126.6397 | -8.9638 | 1553.7736 | -9.0517 | 1545.4305 | -9.0517 | 1541.4717 | -9.0868 | 1543.6124 | -9.0868 | 1550.4209 | -9.1044 | 1550.5881 | -9.1044 | 1536.5652
| -8.9111 | 122.8423 | -9.0165 | 127.173 | -8.9462 | 125.0982 | -8.9462 | 124.7225 | -8.999 | 126.5607 | -8.9287 | 1553.4374 | -9.0165 | 1542.4587 | -9.0165 | 1541.4717 | -9.0517 | 1546.0932 | -9.0517 | 1550.4209 | -9.0693 | 1553.6055 | -9.0693 | 1535.9145
| -8.8759 | 122.8423 | -8.9814 | 126.7582 | -8.9111 | 125.3157 | -8.9111 | 124.5247 | -8.9638 | 126.4619 | -8.8935 | 1549.4191 | -8.9814 | 1542.1295 | -8.9814 | 1540.4867 | -9.0165 | 1545.7618 | -9.0165 | 1549.7528 | -9.0341 | 1554.11 | -9.0341 | 1535.7519
| -8.8408 | 122.8026 | -8.9462 | 126.5211 | -8.8759 | 125.3157 | -8.8759 | 124.0499 | -8.9287 | 126.5607 | -8.8584 | 1545.9275 | -8.9462 | 1539.5035 | -8.9462 | 1537.0537 | -8.9814 | 1544.1076 | -8.9814 | 1546.9228 | -8.999 | 1554.11 | -8.999 | 1535.2646
| -8.8056 | 122.2875 | -8.9111 | 126.3829 | -8.8408 | 125.3157 | -8.8408 | 123.2581 | -8.8935 | 126.778 | -8.8232 | 1541.3074 | -8.9111 | 1539.5035 | -8.9111 | 1536.728 | -8.9462 | 1543.2825 | -8.9462 | 1542.788 | -8.9638 | 1550.9225 | -8.9638 | 1534.94
| -8.7705 | 122.01 | -8.8759 | 126.3829 | -8.8056 | 124.9795 | -8.8056 | 123.1195 | -8.8584 | 126.8372 | -8.7881 | 1536.5652 | -8.8759 | 1537.7059 | -8.8759 | 1532.9967 | -8.9111 | 1541.1432 | -8.9111 | 1541.8005 | -8.9287 | 1547.4211 | -8.9287 | 1535.5894
| -8.7353 | 122.01 | -8.8408 | 126.7582 | -8.7705 | 124.3467 | -8.7705 | 123.1195 | -8.8232 | 126.7187 | -8.7529 | 1535.2646 | -8.8408 | 1535.9145 | -8.8408 | 1530.578 | -8.8759 | 1537.0537 | -8.8759 | 1541.8005 | -8.8935 | 1543.6124 | -8.8935 | 1535.7519
| -8.7002 | 122.1091 | -8.8056 | 125.8493 | -8.7353 | 123.9707 | -8.7353 | 123.1393 | -8.7881 | 126.3038 | -8.7177 | 1532.9967 | -8.8056 | 1533.9674 | -8.8056 | 1528.0109 | -8.8408 | 1533.4818 | -8.8408 | 1539.5035 | -8.8584 | 1543.4474 | -8.8584 | 1535.5894
| -8.665 | 122.0893 | -8.7705 | 125.8493 | -8.7002 | 123.8916 | -8.7002 | 122.9809 | -8.7529 | 125.9877 | -8.6826 | 1531.3829 | -8.7705 | 1527.691 | -8.7705 | 1524.9799 | -8.8056 | 1530.2564 | -8.8056 | 1538.5222 | -8.8232 | 1540.3227 | -8.8232 | 1534.94
| -8.6299 | 121.9901 | -8.7353 | 126.1458 | -8.665 | 124.2082 | -8.665 | 122.8819 | -8.7177 | 125.9086 | -8.6474 | 1529.7743 | -8.7353 | 1525.6164 | -8.7353 | 1521.4936 | -8.7705 | 1529.1324 | -8.7705 | 1535.427 | -8.7881 | 1539.831 | -8.7881 | 1531.8665
| -8.5947 | 121.772 | -8.7002 | 126.0865 | -8.6299 | 124.584 | -8.6299 | 122.6442 | -8.6826 | 125.6122 | -8.6123 | 1524.8208 | -8.7002 | 1524.0266 | -8.7002 | 1519.6027 | -8.7353 | 1528.0109 | -8.7353 | 1531.8665 | -8.7529 | 1539.3398 | -8.7529 | 1524.0266
| -8.5596 | 121.415 | -8.665 | 125.5133 | -8.5947 | 124.7027 | -8.5947 | 122.5649 | -8.6474 | 125.5529 | -8.5771 | 1523.3921 | -8.665 | 1523.3921 | -8.665 | 1517.4061 | -8.7002 | 1526.892 | -8.7002 | 1529.6138 | -8.7177 | 1534.1294 | -8.7177 | 1522.4419
| -8.5244 | 121.296 | -8.6299 | 125.8493 | -8.5596 | 124.7027 | -8.5596 | 122.5649 | -8.6123 | 125.454 | -8.542 | 1521.4936 | -8.6299 | 1521.178 | -8.6299 | 1513.1986 | -8.665 | 1522.7584 | -8.665 | 1529.2928 | -8.6826 | 1532.3505 | -8.6826 | 1522.4419
| -8.4893 | 121.1372 | -8.5947 | 125.9877 | -8.5244 | 124.2082 | -8.5244 | 122.6442 | -8.5771 | 124.7027 | -8.5068 | 1516.7804 | -8.5947 | 1518.9741 | -8.5947 | 1510.4148 | -8.6299 | 1517.8759 | -8.6299 | 1527.8509 | -8.6474 | 1526.892 | -8.6474 | 1520.7048
| -8.4541 | 120.8197 | -8.5596 | 126.0074 | -8.4893 | 123.7926 | -8.4893 | 122.5649 | -8.542 | 123.6739 | -8.4717 | 1514.4413 | -8.5596 | 1516.9367 | -8.5596 | 1503.9852 | -8.5947 | 1514.7525 | -8.5947 | 1519.4455 | -8.6123 | 1524.8208 | -8.6123 | 1515.9994
| -8.419 | 120.4821 | -8.5244 | 125.9877 | -8.4541 | 123.8322 | -8.4541 | 122.5451 | -8.5068 | 123.3175 | -8.4365 | 1512.7335 | -8.5244 | 1514.5969 | -8.5244 | 1502.0138 | -8.5596 | 1510.569 | -8.5596 | 1517.2496 | -8.5771 | 1523.2336 | -8.5771 | 1514.1303
| -8.3838 | 120.1444 | -8.4893 | 125.9877 | -8.419 | 123.4561 | -8.419 | 122.2082 | -8.4717 | 123.3175 | -8.4014 | 1511.3409 | -8.4893 | 1508.7219 | -8.4893 | 1501.409 | -8.5244 | 1508.8756 | -8.5244 | 1515.8433 | -8.542 | 1518.5031 | -8.542 | 1510.1066
| -8.3486 | 119.4486 | -8.4541 | 125.5924 | -8.3838 | 123.2977 | -8.3838 | 122.1091 | -8.4365 | 124.0895 | -8.3662 | 1508.5684 | -8.4541 | 1497.8018 | -8.4541 | 1500.0512 | -8.4893 | 1506.7294 | -8.4893 | 1512.2688 | -8.5068 | 1516.3116 | -8.5068 | 1508.5684
| -8.3135 | 118.7323 | -8.419 | 124.8016 | -8.3486 | 123.2977 | -8.3486 | 121.9505 | -8.4014 | 124.1686 | -8.3311 | 1507.3416 | -8.419 | 1497.3544 | -8.419 | 1497.9512 | -8.4541 | 1503.5294 | -8.4541 | 1508.2614 | -8.4717 | 1515.8433 | -8.4717 | 1508.1079
| -8.2783 | 118.1547 | -8.3838 | 124.5049 | -8.3135 | 123.0799 | -8.3135 | 121.8117 | -8.3662 | 124.0895 | -8.2959 | 1499.4494 | -8.3838 | 1495.8686 | -8.3838 | 1494.3917 | -8.419 | 1499.4494 | -8.419 | 1507.0354 | -8.4365 | 1514.1303 | -8.4365 | 1501.409
| -8.2432 | 117.6563 | -8.3486 | 124.1488 | -8.2783 | 122.5055 | -8.2783 | 121.6134 | -8.3311 | 124.3862 | -8.2608 | 1496.4619 | -8.3486 | 1492.4846 | -8.3486 | 1491.9008 | -8.3838 | 1494.6864 | -8.3838 | 1502.6194 | -8.4014 | 1509.7985 | -8.4014 | 1499.7501
| -8.208 | 117.6563 | -8.3135 | 124.1488 | -8.2432 | 122.3073 | -8.2432 | 120.5218 | -8.2959 | 124.3862 | -8.2256 | 1493.6565 | -8.3135 | 1489.4343 | -8.3135 | 1489.29 | -8.3486 | 1493.6565 | -8.3486 | 1497.9512 | -8.3662 | 1503.2259 | -8.3662 | 1492.9234
| -8.1729 | 118.2742 | -8.2783 | 124.0103 | -8.208 | 122.5649 | -8.208 | 120.6608 | -8.2608 | 124.3862 | -8.1905 | 1490.447 | -8.2783 | 1489.0016 | -8.2783 | 1485.1399 | -8.3135 | 1492.0466 | -8.3135 | 1497.8018 | -8.3311 | 1502.1651 | -8.3311 | 1491.9008
| -8.1377 | 118.2742 | -8.2432 | 124.0103 | -8.1729 | 121.7919 | -8.1729 | 120.4225 | -8.2256 | 124.4653 | -8.1553 | 1488.282 | -8.2432 | 1489.0016 | -8.2432 | 1482.4572 | -8.2783 | 1488.1383 | -8.2783 | 1494.9814 | -8.2959 | 1499.9006 | -8.2959 | 1490.1572
| -8.1026 | 118.1945 | -8.208 | 123.8916 | -8.1377 | 121.6531 | -8.1377 | 120.1047 | -8.1905 | 124.4258 | -8.1202 | 1487.5644 | -8.208 | 1480.4985 | -8.208 | 1475.2557 | -8.2432 | 1485.7082 | -8.2432 | 1488.4257 | -8.2608 | 1496.1651 | -8.2608 | 1488.8575
| -8.0674 | 117.8955 | -8.1729 | 122.9215 | -8.1026 | 120.8594 | -8.1026 | 119.9854 | -8.1553 | 123.0601 | -8.085 | 1485.8505 | -8.1729 | 1476.8999 | -8.1729 | 1474.4375 | -8.208 | 1485.566 | -8.208 | 1485.2818 | -8.2256 | 1492.3385 | -8.2256 | 1486.706
| -8.0323 | 117.736 | -8.1377 | 121.9703 | -8.0674 | 119.7071 | -8.0674 | 119.9854 | -8.1202 | 122.2677 | -8.0499 | 1480.7774 | -8.1377 | 1475.9395 | -8.1377 | 1473.0794 | -8.1729 | 1478.8313 | -8.1729 | 1482.4572 | -8.1905 | 1489.0016 | -8.1905 | 1482.7383
| -7.9971 | 117.8955 | -8.1026 | 121.5539 | -8.0323 | 119.4685 | -8.0323 | 119.6674 | -8.085 | 122.2677 | -8.0147 | 1478.4162 | -8.1026 | 1474.8463 | -8.1026 | 1471.3244 | -8.1377 | 1476.8999 | -8.1377 | 1480.0807 | -8.1553 | 1483.1604 | -8.1553 | 1475.2557
| -7.962 | 117.5167 | -8.0674 | 121.2563 | -7.9971 | 119.389 | -7.9971 | 119.2298 | -8.0499 | 122.1686 | -7.9795 | 1478.278 | -8.0674 | 1472.944 | -8.0674 | 1470.1161 | -8.1026 | 1474.5737 | -8.1026 | 1478.278 | -8.1202 | 1481.476 | -8.1202 | 1473.3505
| -7.9268 | 117.5964 | -8.0323 | 120.7005 | -7.962 | 119.4685 | -7.962 | 119.0906 | -8.0147 | 122.0893 | -7.9444 | 1477.588 | -8.0323 | 1472.944 | -8.0323 | 1470.6524 | -8.0674 | 1474.5737 | -8.0674 | 1477.3125 | -8.085 | 1478.8313 | -8.085 | 1472.1329
| -7.8917 | 117.7959 | -7.9971 | 120.6013 | -7.9268 | 119.5481 | -7.9268 | 118.2144 | -7.9795 | 122.1686 | -7.9092 | 1477.4502 | -7.9971 | 1472.8087 | -7.9971 | 1468.5134 | -8.0323 | 1471.8631 | -8.0323 | 1476.3507 | -8.0499 | 1475.8026 | -8.0499 | 1469.5808
| -7.8565 | 117.7959 | -7.962 | 120.6013 | -7.8917 | 120.0053 | -7.8917 | 117.756 | -7.9444 | 122.0893 | -7.8741 | 1478.4162 | -7.962 | 1469.4471 | -7.962 | 1463.8924 | -7.9971 | 1471.7283 | -7.9971 | 1473.2149 | -8.0147 | 1473.2149 | -8.0147 | 1465.2048
| -7.8214 | 117.7959 | -7.9268 | 120.7005 | -7.8565 | 120.0053 | -7.8565 | 116.9978 | -7.9092 | 121.5737 | -7.8389 | 1479.5247 | -7.9268 | 1465.3364 | -7.9268 | 1461.6755 | -7.962 | 1470.6524 | -7.962 | 1470.25 | -7.9795 | 1473.7576 | -7.9795 | 1464.8104
| -7.7862 | 118.2543 | -7.8917 | 120.8395 | -7.8214 | 119.8861 | -7.8214 | 116.9978 | -7.8741 | 120.7204 | -7.8038 | 1477.588 | -7.8917 | 1466.7879 | -7.8917 | 1461.0268 | -7.9268 | 1464.6791 | -7.9268 | 1470.921 | -7.9444 | 1473.0794 | -7.9444 | 1464.5478
| -7.7511 | 118.2543 | -7.8565 | 120.8395 | -7.7862 | 119.4287 | -7.7862 | 117.1974 | -7.8389 | 119.4884 | -7.7686 | 1479.1084 | -7.8565 | 1465.0733 | -7.8565 | 1460.6383 | -7.8917 | 1462.7166 | -7.8917 | 1469.4471 | -7.9092 | 1472.8087 | -7.9092 | 1463.6307
| -7.7159 | 118.7124 | -7.8214 | 120.4622 | -7.7511 | 118.8119 | -7.7511 | 116.9778 | -7.8038 | 118.9512 | -7.7335 | 1479.1084 | -7.8214 | 1465.0733 | -7.8214 | 1462.5863 | -7.8565 | 1462.456 | -7.8565 | 1468.2473 | -7.8741 | 1472.944 | -7.8741 | 1461.4159
| -7.6808 | 119.0508 | -7.7862 | 119.9854 | -7.7159 | 118.1547 | -7.7159 | 116.9778 | -7.7686 | 118.3539 | -7.6983 | 1477.3125 | -7.7862 | 1469.18 | -7.7862 | 1464.9418 | -7.8214 | 1462.456 | -7.8214 | 1469.4471 | -7.8389 | 1472.944 | -7.8389 | 1461.1565
| -7.6456 | 118.9911 | -7.7511 | 119.9854 | -7.6808 | 117.8955 | -7.6808 | 117.0177 | -7.7335 | 118.2742 | -7.6632 | 1477.0374 | -7.7511 | 1469.18 | -7.7511 | 1465.0733 | -7.7862 | 1462.847 | -7.7862 | 1472.8087 | -7.8038 | 1471.4589 | -7.8038 | 1461.4159
| -7.6105 | 118.8517 | -7.7159 | 119.9656 | -7.6456 | 117.8955 | -7.6456 | 116.9179 | -7.6983 | 118.4336 | -7.628 | 1476.8999 | -7.7159 | 1469.18 | -7.7159 | 1465.468 | -7.7511 | 1463.108 | -7.7511 | 1470.5182 | -7.7686 | 1471.4589 | -7.7686 | 1461.4159
| -7.5753 | 118.9911 | -7.6808 | 120.045 | -7.6105 | 118.1547 | -7.6105 | 116.9778 | -7.6632 | 118.7124 | -7.5929 | 1475.2557 | -7.6808 | 1469.9821 | -7.6808 | 1465.8633 | -7.7159 | 1463.7615 | -7.7159 | 1470.5182 | -7.7335 | 1470.5182 | -7.7335 | 1463.4999
| -7.5401 | 118.6128 | -7.6456 | 120.4225 | -7.5753 | 118.334 | -7.5753 | 117.0177 | -7.628 | 118.2941 | -7.5577 | 1473.8934 | -7.6456 | 1467.0527 | -7.6456 | 1465.0733 | -7.6808 | 1464.4166 | -7.6808 | 1470.1161 | -7.6983 | 1470.25 | -7.6983 | 1464.6791
| -7.505 | 118.1148 | -7.6105 | 120.502 | -7.5401 | 119.1105 | -7.5401 | 117.2174 | -7.5929 | 118.1347 | -7.5226 | 1470.5182 | -7.6105 | 1467.0527 | -7.6105 | 1463.3692 | -7.6456 | 1463.108 | -7.6456 | 1469.7145 | -7.6632 | 1470.25 | -7.6632 | 1465.468
| -7.4698 | 117.9155 | -7.5753 | 120.641 | -7.505 | 119.1105 | -7.505 | 117.5565 | -7.5577 | 118.2941 | -7.4874 | 1470.1161 | -7.5753 | 1466.7879 | -7.5753 | 1460.2503 | -7.6105 | 1464.0234 | -7.6105 | 1468.2473 | -7.628 | 1470.25 | -7.628 | 1465.468
| -7.4347 | 118.0151 | -7.5401 | 120.7998 | -7.4698 | 119.3293 | -7.4698 | 117.2772 | -7.5226 | 118.573 | -7.4523 | 1468.6466 | -7.5401 | 1464.6791 | -7.5401 | 1459.7338 | -7.5753 | 1464.2855 | -7.5753 | 1466.5234 | -7.5929 | 1467.7157 | -7.5929 | 1463.8924
| -7.3995 | 118.1945 | -7.505 | 120.8594 | -7.4347 | 119.9457 | -7.4347 | 117.2772 | -7.4874 | 118.1746 | -7.4171 | 1468.6466 | -7.505 | 1464.5478 | -7.505 | 1458.3182 | -7.5401 | 1462.7166 | -7.5401 | 1466.5234 | -7.5577 | 1468.2473 | -7.5577 | 1464.5478
| -7.3644 | 118.2543 | -7.4698 | 121.1769 | -7.3995 | 120.4225 | -7.3995 | 117.9553 | -7.4523 | 118.5332 | -7.382 | 1470.5182 | -7.4698 | 1464.9418 | -7.4698 | 1459.7338 | -7.505 | 1460.7678 | -7.505 | 1469.0466 | -7.5226 | 1467.8485 | -7.5226 | 1462.847
| -7.3292 | 118.2742 | -7.4347 | 121.7324 | -7.3644 | 120.502 | -7.3644 | 118.1547 | -7.4171 | 118.5332 | -7.3468 | 1474.4375 | -7.4347 | 1467.9813 | -7.4347 | 1462.3258 | -7.4698 | 1460.7678 | -7.4698 | 1468.2473 | -7.4874 | 1467.7157 | -7.4874 | 1461.0268
| -7.2941 | 118.4535 | -7.3995 | 121.7324 | -7.3292 | 120.3629 | -7.3292 | 118.4336 | -7.382 | 118.4734 | -7.3117 | 1475.9395 | -7.3995 | 1468.1143 | -7.3995 | 1464.5478 | -7.4347 | 1460.7678 | -7.4347 | 1469.5808 | -7.4523 | 1467.8485 | -7.4523 | 1456.6541
| -7.2589 | 118.4535 | -7.3644 | 121.7522 | -7.2941 | 119.9854 | -7.2941 | 119.0906 | -7.3468 | 118.4734 | -7.2765 | 1476.2135 | -7.3644 | 1469.8483 | -7.3644 | 1464.8104 | -7.3995 | 1460.3796 | -7.3995 | 1469.9821 | -7.4171 | 1468.5134 | -7.4171 | 1455.8893
| -7.2238 | 118.3738 | -7.3292 | 121.4745 | -7.2589 | 119.6674 | -7.2589 | 119.2895 | -7.3117 | 118.9313 | -7.2414 | 1476.3507 | -7.3292 | 1469.7145 | -7.3292 | 1465.7315 | -7.3644 | 1461.4159 | -7.3644 | 1470.5182 | -7.382 | 1468.9132 | -7.382 | 1456.6541
| -7.1886 | 118.055 | -7.2941 | 121.3158 | -7.2238 | 119.4089 | -7.2238 | 119.4486 | -7.2765 | 119.3094 | -7.2062 | 1478.0018 | -7.2941 | 1469.7145 | -7.2941 | 1465.7315 | -7.3292 | 1462.0655 | -7.3292 | 1470.5182 | -7.3468 | 1469.3135 | -7.3468 | 1462.7166
| -7.1535 | 118.334 | -7.2589 | 121.3158 | -7.1886 | 119.4089 | -7.1886 | 119.4486 | -7.2414 | 120.2239 | -7.171 | 1479.9415 | -7.2589 | 1469.8483 | -7.2589 | 1466.5234 | -7.2941 | 1462.7166 | -7.2941 | 1471.5936 | -7.3117 | 1470.5182 | -7.3117 | 1465.2048
| -7.1183 | 118.3539 | -7.2238 | 121.296 | -7.1535 | 119.4287 | -7.1535 | 119.2895 | -7.2062 | 120.2835 | -7.1359 | 1483.5833 | -7.2238 | 1472.944 | -7.2238 | 1469.18 | -7.2589 | 1463.6307 | -7.2589 | 1471.5936 | -7.2765 | 1471.0554 | -7.2765 | 1465.2048
| -7.0832 | 118.4535 | -7.1886 | 121.296 | -7.1183 | 119.4486 | -7.1183 | 118.9313 | -7.171 | 120.3431 | -7.1007 | 1485.9929 | -7.1886 | 1474.0293 | -7.1886 | 1473.3505 | -7.2238 | 1464.6791 | -7.2238 | 1476.0765 | -7.2414 | 1472.944 | -7.2414 | 1465.0733
| -7.048 | 118.7124 | -7.1535 | 121.296 | -7.0832 | 119.6077 | -7.0832 | 118.2742 | -7.1359 | 120.2437 | -7.0656 | 1489.29 | -7.1535 | 1478.9698 | -7.1535 | 1475.3923 | -7.1886 | 1471.3244 | -7.1886 | 1478.5545 | -7.2062 | 1475.6658 | -7.2062 | 1465.5997
| -7.0129 | 118.3937 | -7.1183 | 120.7403 | -7.048 | 120.1841 | -7.048 | 117.5765 | -7.1007 | 119.7071 | -7.0304 | 1492.0466 | -7.1183 | 1481.476 | -7.1183 | 1479.5247 | -7.1535 | 1473.8934 | -7.1535 | 1481.1963 | -7.171 | 1477.4502 | -7.171 | 1466.1272
| -6.9777 | 118.7124 | -7.0832 | 120.7403 | -7.0129 | 120.3431 | -7.0129 | 117.1774 | -7.0656 | 119.2696 | -6.9953 | 1494.2445 | -7.0832 | 1484.5727 | -7.0832 | 1480.2198 | -7.1183 | 1476.4879 | -7.1183 | 1481.476 | -7.1359 | 1478.1399 | -7.1359 | 1466.7879
| -6.9426 | 118.3937 | -7.048 | 120.7403 | -6.9777 | 120.3431 | -6.9777 | 117.357 | -7.0304 | 119.1502 | -6.9601 | 1497.8018 | -7.048 | 1485.2818 | -7.048 | 1481.896 | -7.0832 | 1478.1399 | -7.0832 | 1482.1765 | -7.1007 | 1479.5247 | -7.1007 | 1470.921
| -6.9074 | 118.1148 | -7.0129 | 120.7403 | -6.9426 | 119.9258 | -6.9426 | 117.6164 | -6.9953 | 119.2298 | -6.925 | 1497.9512 | -7.0129 | 1487.4211 | -7.0129 | 1484.7144 | -7.048 | 1479.5247 | -7.048 | 1483.0196 | -7.0656 | 1480.6379 | -7.0656 | 1473.7576
| -6.8723 | 118.0949 | -6.9777 | 120.8395 | -6.9074 | 119.6276 | -6.9074 | 118.3539 | -6.9601 | 119.2696 | -6.8898 | 1500.3526 | -6.9777 | 1494.0974 | -6.9777 | 1486.5632 | -7.0129 | 1481.0566 | -7.0129 | 1485.8505 | -7.0304 | 1481.476 | -7.0304 | 1475.9395
| -6.8371 | 118.1148 | -6.9426 | 120.4821 | -6.8723 | 119.2696 | -6.8723 | 118.2543 | -6.925 | 119.4685 | -6.8547 | 1499.9006 | -6.9426 | 1494.9814 | -6.9426 | 1488.4257 | -6.9777 | 1484.5727 | -6.9777 | 1487.9947 | -6.9953 | 1482.3168 | -6.9953 | 1478.0018
| -6.802 | 118.055 | -6.9074 | 120.2835 | -6.8371 | 119.2696 | -6.8371 | 118.4535 | -6.8898 | 119.7071 | -6.8195 | 1497.9512 | -6.9074 | 1494.9814 | -6.9074 | 1487.8512 | -6.9426 | 1486.4205 | -6.9426 | 1492.4846 | -6.9601 | 1485.9929 | -6.9601 | 1479.1084
| -6.7668 | 116.9179 | -6.8723 | 120.1245 | -6.802 | 119.2099 | -6.802 | 118.2543 | -6.8547 | 120.1047 | -6.7844 | 1497.2054 | -6.8723 | 1492.6308 | -6.8723 | 1487.8512 | -6.9074 | 1486.8489 | -6.9074 | 1492.4846 | -6.925 | 1487.2779 | -6.925 | 1481.3361
| -6.7316 | 115.8786 | -6.8371 | 120.2835 | -6.7668 | 119.5878 | -6.7668 | 117.9155 | -6.8195 | 119.6674 | -6.7492 | 1497.6526 | -6.8371 | 1490.1572 | -6.8371 | 1487.8512 | -6.8723 | 1486.9918 | -6.8723 | 1492.6308 | -6.8898 | 1489.0016 | -6.8898 | 1481.896
| -6.6965 | 115.3782 | -6.802 | 120.4027 | -6.7316 | 119.5282 | -6.7316 | 117.9155 | -6.7844 | 120.045 | -6.7141 | 1494.0974 | -6.802 | 1488.8575 | -6.802 | 1484.8561 | -6.8371 | 1484.4311 | -6.8371 | 1489.0016 | -6.8547 | 1488.8575 | -6.8547 | 1483.3013
| -6.6613 | 115.0377 | -6.7668 | 120.4027 | -6.6965 | 119.5282 | -6.6965 | 118.1148 | -6.7492 | 120.0053 | -6.6789 | 1496.0168 | -6.7668 | 1487.7077 | -6.7668 | 1482.5977 | -6.802 | 1480.9169 | -6.802 | 1489.0016 | -6.8195 | 1489.4343 | -6.8195 | 1484.1482
| -6.6262 | 114.9775 | -6.7316 | 120.5615 | -6.6613 | 119.2099 | -6.6613 | 118.1148 | -6.7141 | 119.3691 | -6.6438 | 1495.2768 | -6.7316 | 1489.5787 | -6.7316 | 1488.1383 | -6.7668 | 1480.9169 | -6.7668 | 1487.9947 | -6.7844 | 1486.1354 | -6.7844 | 1484.1482
| -6.591 | 114.9575 | -6.6965 | 120.4225 | -6.6262 | 119.4685 | -6.6262 | 117.6164 | -6.6789 | 119.3691 | -6.6086 | 1496.9077 | -6.6965 | 1490.1572 | -6.6965 | 1489.7232 | -6.7316 | 1480.4985 | -6.7316 | 1488.1383 | -6.7492 | 1483.8656 | -6.7492 | 1483.0196
| -6.5559 | 114.396 | -6.6613 | 120.3629 | -6.591 | 118.4535 | -6.591 | 117.6563 | -6.6438 | 119.4685 | -6.5735 | 1493.6565 | -6.6613 | 1488.7135 | -6.6613 | 1488.7135 | -6.6965 | 1480.4985 | -6.6965 | 1488.4257 | -6.7141 | 1483.3013 | -6.7141 | 1480.0807
| -6.5207 | 114.2756 | -6.6262 | 120.204 | -6.5559 | 118.4734 | -6.5559 | 117.5565 | -6.6086 | 119.4486 | -6.5383 | 1493.6565 | -6.6262 | 1489.8678 | -6.6262 | 1489.8678 | -6.6613 | 1481.6159 | -6.6613 | 1488.4257 | -6.6789 | 1486.2779 | -6.6789 | 1482.0362
| -6.4856 | 113.4923 | -6.591 | 119.9059 | -6.5207 | 118.5132 | -6.5207 | 117.0577 | -6.5735 | 119.2497 | -6.5032 | 1493.6565 | -6.591 | 1489.8678 | -6.591 | 1490.3021 | -6.6262 | 1484.0069 | -6.6262 | 1488.4257 | -6.6438 | 1490.0125 | -6.6438 | 1480.0807
| -6.4504 | 113.2309 | -6.5559 | 119.7668 | -6.4856 | 118.5132 | -6.4856 | 117.0577 | -6.5383 | 119.1303 | -6.468 | 1493.363 | -6.5559 | 1489.0016 | -6.5559 | 1488.282 | -6.591 | 1484.8561 | -6.591 | 1489.0016 | -6.6086 | 1491.755 | -6.6086 | 1480.0807
| -6.4153 | 113.251 | -6.5207 | 119.4884 | -6.4504 | 117.736 | -6.4504 | 117.0377 | -6.5032 | 119.0508 | -6.4329 | 1492.4846 | -6.5207 | 1489.0016 | -6.5207 | 1485.4239 | -6.5559 | 1486.9918 | -6.5559 | 1490.8822 | -6.5735 | 1492.3385 | -6.5735 | 1480.6379
| -6.3801 | 113.2912 | -6.4856 | 118.9712 | -6.4153 | 117.1974 | -6.4153 | 116.7182 | -6.468 | 118.9512 | -6.3977 | 1491.9008 | -6.4856 | 1489.0016 | -6.4856 | 1487.5644 | -6.5207 | 1487.4211 | -6.5207 | 1493.0699 | -6.5383 | 1493.2164 | -6.5383 | 1481.3361
| -6.345 | 113.2711 | -6.4504 | 118.3738 | -6.3801 | 117.1176 | -6.3801 | 116.0586 | -6.4329 | 118.6526 | -6.3625 | 1491.755 | -6.4504 | 1487.2779 | -6.4504 | 1487.4211 | -6.4856 | 1486.9918 | -6.4856 | 1491.6093 | -6.5032 | 1491.3182 | -6.5032 | 1480.3591
| -6.3098 | 113.7937 | -6.4153 | 117.9553 | -6.345 | 116.858 | -6.345 | 115.6985 | -6.3977 | 118.9512 | -6.3274 | 1489.7232 | -6.4153 | 1488.7135 | -6.4153 | 1485.7082 | -6.4504 | 1487.2779 | -6.4504 | 1490.7371 | -6.468 | 1489.8678 | -6.468 | 1479.9415
| -6.2747 | 114.396 | -6.3801 | 117.5765 | -6.3098 | 116.3385 | -6.3098 | 115.4783 | -6.3625 | 118.9712 | -6.2922 | 1489.7232 | -6.3801 | 1488.7135 | -6.3801 | 1485.7082 | -6.4153 | 1486.4205 | -6.4153 | 1490.1572 | -6.4329 | 1489.1457 | -6.4329 | 1479.5247
| -6.2395 | 114.717 | -6.345 | 117.0377 | -6.2747 | 115.4583 | -6.2747 | 115.0777 | -6.3274 | 118.8318 | -6.2571 | 1489.7232 | -6.345 | 1488.282 | -6.345 | 1484.8561 | -6.3801 | 1485.9929 | -6.3801 | 1489.0016 | -6.3977 | 1488.4257 | -6.3977 | 1479.9415
| -6.2044 | 114.717 | -6.3098 | 116.4984 | -6.2395 | 114.5164 | -6.2395 | 114.9575 | -6.2922 | 118.3539 | -6.2219 | 1486.706 | -6.3098 | 1488.282 | -6.3098 | 1478.0018 | -6.345 | 1484.8561 | -6.345 | 1486.4205 | -6.3625 | 1486.9918 | -6.3625 | 1479.3858
| -6.1692 | 114.717 | -6.2747 | 116.2786 | -6.2044 | 114.3559 | -6.2044 | 114.115 | -6.2571 | 118.314 | -6.1868 | 1483.8656 | -6.2747 | 1488.282 | -6.2747 | 1474.8463 | -6.3098 | 1481.476 | -6.3098 | 1485.9929 | -6.3274 | 1485.7082 | -6.3274 | 1479.1084
| -6.1341 | 113.9544 | -6.2395 | 115.5384 | -6.1692 | 114.4562 | -6.1692 | 113.9544 | -6.2219 | 117.9753 | -6.1516 | 1483.7244 | -6.2395 | 1485.4239 | -6.2395 | 1473.0794 | -6.2747 | 1477.8638 | -6.2747 | 1483.4423 | -6.2922 | 1485.7082 | -6.2922 | 1478.278
| -6.0989 | 113.5325 | -6.2044 | 115.4183 | -6.1341 | 114.8974 | -6.1341 | 113.9142 | -6.1868 | 117.6563 | -6.1165 | 1482.4572 | -6.2044 | 1478.8313 | -6.2044 | 1472.268 | -6.2395 | 1476.3507 | -6.2395 | 1481.1963 | -6.2571 | 1483.0196 | -6.2571 | 1476.2135
| -6.0638 | 113.854 | -6.1692 | 115.6585 | -6.0989 | 115.7185 | -6.0989 | 113.0095 | -6.1516 | 117.1575 | -6.0813 | 1480.2198 | -6.1692 | 1471.5936 | -6.1692 | 1471.998 | -6.2044 | 1473.6218 | -6.2044 | 1479.5247 | -6.2219 | 1481.7559 | -6.2219 | 1475.9395
| -6.0286 | 113.9544 | -6.1341 | 115.4383 | -6.0638 | 115.7385 | -6.0638 | 112.7881 | -6.1165 | 116.6383 | -6.0462 | 1479.1084 | -6.1341 | 1470.1161 | -6.1341 | 1470.3841 | -6.1692 | 1471.998 | -6.1692 | 1479.5247 | -6.1868 | 1479.2471 | -6.1868 | 1475.529
| -5.9935 | 113.6531 | -6.0989 | 115.4383 | -6.0286 | 115.7385 | -6.0286 | 112.7277 | -6.0813 | 115.9186 | -6.011 | 1477.3125 | -6.0989 | 1468.6466 | -6.0989 | 1468.9132 | -6.1341 | 1471.7283 | -6.1341 | 1477.3125 | -6.1516 | 1476.8999 | -6.1516 | 1475.2557
| -5.9583 | 113.251 | -6.0638 | 115.8386 | -5.9935 | 115.6185 | -5.9935 | 112.3853 | -6.0462 | 115.0377 | -5.9759 | 1477.3125 | -6.0638 | 1468.5134 | -6.0638 | 1467.5829 | -6.0989 | 1471.7283 | -6.0989 | 1473.4861 | -6.1165 | 1474.5737 | -6.1165 | 1474.4375
| -5.9231 | 113.5928 | -6.0286 | 115.8386 | -5.9583 | 115.2981 | -5.9583 | 112.3853 | -6.011 | 114.4963 | -5.9407 | 1478.6929 | -6.0286 | 1468.6466 | -6.0286 | 1466.9203 | -6.0638 | 1470.6524 | -6.0638 | 1473.4861 | -6.0813 | 1474.5737 | -6.0813 | 1469.18
| -5.888 | 112.6069 | -5.9935 | 115.4383 | -5.9231 | 114.7971 | -5.9231 | 113.1906 | -5.9759 | 113.9946 | -5.9056 | 1478.278 | -5.9935 | 1469.0466 | -5.9935 | 1466.6556 | -6.0286 | 1469.9821 | -6.0286 | 1472.5382 | -6.0462 | 1473.3505 | -6.0462 | 1467.5829
| -5.8528 | 112.345 | -5.9583 | 115.8186 | -5.888 | 114.7571 | -5.888 | 112.6472 | -5.9407 | 114.5365 | -5.8704 | 1479.5247 | -5.9583 | 1469.4471 | -5.9583 | 1466.9203 | -5.9935 | 1470.1161 | -5.9935 | 1471.5936 | -6.011 | 1473.7576 | -6.011 | 1465.9952
| -5.8177 | 112.103 | -5.9231 | 115.3782 | -5.8528 | 113.8339 | -5.8528 | 112.7277 | -5.9056 | 114.4963 | -5.8353 | 1479.8025 | -5.9231 | 1470.7867 | -5.9231 | 1468.7799 | -5.9583 | 1470.5182 | -5.9583 | 1471.5936 | -5.9759 | 1473.7576 | -5.9759 | 1466.9203
| -5.7825 | 111.6793 | -5.888 | 115.3782 | -5.8177 | 113.5727 | -5.8177 | 112.768 | -5.8704 | 114.5365 | -5.8001 | 1479.5247 | -5.888 | 1471.998 | -5.888 | 1467.5829 | -5.9231 | 1466.9203 | -5.9231 | 1471.998 | -5.9407 | 1470.7867 | -5.9407 | 1466.2592
| -5.7474 | 111.4572 | -5.8528 | 115.3782 | -5.7825 | 114.8773 | -5.7825 | 112.6472 | -5.8353 | 114.6768 | -5.765 | 1473.6218 | -5.8528 | 1469.4471 | -5.8528 | 1466.6556 | -5.888 | 1468.3803 | -5.888 | 1471.998 | -5.9056 | 1472.8087 | -5.9056 | 1466.2592
| -5.7122 | 111.0531 | -5.8177 | 115.3782 | -5.7474 | 114.8773 | -5.7474 | 112.5666 | -5.8001 | 115.218 | -5.7298 | 1472.944 | -5.8177 | 1467.5829 | -5.8177 | 1466.6556 | -5.8528 | 1468.6466 | -5.8528 | 1471.998 | -5.8704 | 1472.8087 | -5.8704 | 1466.2592
| -5.6771 | 110.8306 | -5.7825 | 115.3782 | -5.7122 | 114.3358 | -5.7122 | 112.4256 | -5.765 | 115.1579 | -5.6947 | 1471.4589 | -5.7825 | 1466.1272 | -5.7825 | 1465.9952 | -5.8177 | 1468.3803 | -5.8177 | 1472.8087 | -5.8353 | 1472.8087 | -5.8353 | 1466.9203
| -5.6419 | 110.6283 | -5.7474 | 115.0577 | -5.6771 | 114.5565 | -5.6771 | 112.7076 | -5.7298 | 115.1579 | -5.6595 | 1471.3244 | -5.7474 | 1462.9775 | -5.7474 | 1463.2386 | -5.7825 | 1467.8485 | -5.7825 | 1472.1329 | -5.8001 | 1472.1329 | -5.8001 | 1467.9813
| -5.6068 | 110.5474 | -5.7122 | 114.9375 | -5.6419 | 114.5565 | -5.6419 | 112.8888 | -5.6947 | 114.2957 | -5.6244 | 1471.3244 | -5.7122 | 1463.3692 | -5.7122 | 1459.7338 | -5.7474 | 1467.8485 | -5.7474 | 1470.25 | -5.765 | 1470.3841 | -5.765 | 1466.7879
| -5.5716 | 109.9599 | -5.6771 | 114.9375 | -5.6068 | 113.7535 | -5.6068 | 112.8485 | -5.6595 | 114.8372 | -5.5892 | 1470.6524 | -5.6771 | 1463.3692 | -5.6771 | 1457.8052 | -5.7122 | 1464.5478 | -5.7122 | 1468.2473 | -5.7298 | 1469.0466 | -5.7298 | 1466.5234
| -5.5365 | 109.9599 | -5.6419 | 114.6768 | -5.5716 | 113.251 | -5.5716 | 112.7277 | -5.6244 | 114.717 | -5.554 | 1470.5182 | -5.6419 | 1462.847 | -5.6419 | 1457.421 | -5.6771 | 1463.3692 | -5.6771 | 1466.1272 | -5.6947 | 1466.3913 | -5.6947 | 1466.1272
| -5.5013 | 109.6963 | -5.6068 | 114.396 | -5.5365 | 113.0498 | -5.5365 | 112.8485 | -5.5892 | 114.6768 | -5.5189 | 1468.2473 | -5.6068 | 1462.5863 | -5.6068 | 1457.421 | -5.6419 | 1462.1956 | -5.6419 | 1464.1544 | -5.6595 | 1463.8924 | -5.6595 | 1460.2503
| -5.4662 | 109.2497 | -5.5716 | 114.3559 | -5.5013 | 112.7277 | -5.5013 | 112.2442 | -5.554 | 114.6768 | -5.4837 | 1465.3364 | -5.5716 | 1463.7615 | -5.5716 | 1456.5265 | -5.6068 | 1459.7338 | -5.6068 | 1463.2386 | -5.6244 | 1463.108 | -5.6244 | 1458.0616
| -5.431 | 109.2497 | -5.5365 | 113.9343 | -5.4662 | 112.7076 | -5.4662 | 111.6793 | -5.5189 | 114.8773 | -5.4486 | 1463.108 | -5.5365 | 1462.9775 | -5.5365 | 1455.2534 | -5.5716 | 1458.5751 | -5.5716 | 1462.5863 | -5.5892 | 1463.108 | -5.5892 | 1458.3182
| -5.3959 | 109.0465 | -5.5013 | 113.9343 | -5.431 | 112.5062 | -5.431 | 111.0531 | -5.4837 | 114.6768 | -5.4134 | 1461.6755 | -5.5013 | 1462.3258 | -5.5013 | 1454.7456 | -5.5365 | 1457.0373 | -5.5365 | 1459.3471 | -5.554 | 1462.7166 | -5.554 | 1458.0616
| -5.3607 | 109.0465 | -5.4662 | 113.5928 | -5.3959 | 112.4256 | -5.3959 | 110.77 | -5.4486 | 114.396 | -5.3783 | 1461.5457 | -5.4662 | 1460.8973 | -5.4662 | 1453.2275 | -5.5013 | 1456.2714 | -5.5013 | 1458.8322 | -5.5189 | 1462.7166 | -5.5189 | 1458.0616
| -5.3256 | 109.2091 | -5.431 | 113.7133 | -5.3607 | 112.2442 | -5.3607 | 110.77 | -5.4134 | 114.095 | -5.3431 | 1461.5457 | -5.431 | 1458.1899 | -5.431 | 1450.3387 | -5.4662 | 1454.9994 | -5.4662 | 1458.8322 | -5.4837 | 1462.0655 | -5.4837 | 1457.9334
| -5.2904 | 109.3715 | -5.3959 | 113.6129 | -5.3256 | 111.5178 | -5.3256 | 110.5271 | -5.3783 | 113.5928 | -5.308 | 1462.3258 | -5.3959 | 1456.0166 | -5.3959 | 1449.091 | -5.431 | 1454.8725 | -5.431 | 1458.1899 | -5.4486 | 1461.5457 | -5.4486 | 1456.6541
| -5.2553 | 109.8585 | -5.3607 | 113.3113 | -5.2904 | 110.8913 | -5.2904 | 110.3246 | -5.3431 | 113.251 | -5.2728 | 1462.3258 | -5.3607 | 1454.8725 | -5.3607 | 1448.4689 | -5.3959 | 1454.8725 | -5.3959 | 1458.1899 | -5.4134 | 1460.8973 | -5.4134 | 1455.762
| -5.2201 | 110.2841 | -5.3256 | 112.0829 | -5.2553 | 110.5878 | -5.2553 | 109.7165 | -5.308 | 113.1906 | -5.2377 | 1458.7036 | -5.3256 | 1452.9752 | -5.3256 | 1447.4762 | -5.3607 | 1454.3654 | -5.3607 | 1456.144 | -5.3783 | 1459.7338 | -5.3783 | 1454.7456
| -5.1849 | 110.1626 | -5.2904 | 112.1434 | -5.2201 | 110.1423 | -5.2201 | 109.0872 | -5.2728 | 113.0297 | -5.2025 | 1458.0616 | -5.2904 | 1451.9682 | -5.2904 | 1446.3629 | -5.3256 | 1453.8591 | -5.3256 | 1454.4921 | -5.3431 | 1456.9095 | -5.3431 | 1453.3537
| -5.1498 | 109.818 | -5.2553 | 111.7803 | -5.1849 | 109.5136 | -5.1849 | 108.7619 | -5.2377 | 111.982 | -5.1674 | 1455.2534 | -5.2553 | 1449.2155 | -5.2553 | 1447.4762 | -5.2904 | 1451.8425 | -5.2904 | 1454.4921 | -5.308 | 1455.6347 | -5.308 | 1448.7176
| -5.1146 | 109.5745 | -5.2201 | 111.6793 | -5.1498 | 110.1423 | -5.1498 | 108.5992 | -5.2025 | 111.336 | -5.1322 | 1453.6063 | -5.2201 | 1447.6001 | -5.2201 | 1445.8693 | -5.2553 | 1448.4689 | -5.2553 | 1454.4921 | -5.2728 | 1454.6188 | -5.2728 | 1445.9926
| -5.0795 | 109.5745 | -5.1849 | 111.639 | -5.1146 | 109.7977 | -5.1146 | 108.5992 | -5.1674 | 111.336 | -5.0971 | 1450.8392 | -5.1849 | 1445.2533 | -5.1849 | 1444.2702 | -5.2201 | 1445.8693 | -5.2201 | 1452.4713 | -5.2377 | 1453.1013 | -5.2377 | 1443.5346
| -5.0443 | 108.7823 | -5.1498 | 111.639 | -5.0795 | 109.5136 | -5.0795 | 108.4364 | -5.1322 | 111.1137 | -5.0619 | 1450.8392 | -5.1498 | 1445.2533 | -5.1498 | 1443.2898 | -5.1849 | 1444.1475 | -5.1849 | 1449.2155 | -5.2025 | 1451.5914 | -5.2025 | 1442.1902
| -5.0092 | 108.192 | -5.1146 | 111.2956 | -5.0443 | 110.2436 | -5.0443 | 108.4364 | -5.0971 | 110.952 | -5.0268 | 1450.0887 | -5.1146 | 1444.3929 | -5.1146 | 1440.6078 | -5.1498 | 1439.8799 | -5.1498 | 1445.746 | -5.1674 | 1448.2204 | -5.1674 | 1442.5563
| -4.974 | 107.9271 | -5.0795 | 111.861 | -5.0092 | 110.4056 | -5.0092 | 108.3346 | -5.0619 | 110.9924 | -4.9916 | 1449.091 | -5.0795 | 1442.0682 | -5.0795 | 1436.7417 | -5.1146 | 1439.8799 | -5.1146 | 1443.9022 | -5.1322 | 1447.3523 | -5.1322 | 1441.3373
| -4.9389 | 107.8864 | -5.0443 | 112.103 | -4.974 | 110.9924 | -4.974 | 108.3346 | -5.0268 | 110.8104 | -4.9565 | 1448.2204 | -5.0443 | 1439.8799 | -5.0443 | 1436.3812 | -5.0795 | 1437.2228 | -5.0795 | 1442.4342 | -5.0971 | 1446.4864 | -5.0971 | 1439.5164
| -4.9037 | 107.7028 | -5.0092 | 112.1232 | -4.9389 | 111.336 | -4.9389 | 108.4974 | -4.9916 | 110.7093 | -4.9213 | 1446.4864 | -5.0092 | 1439.5164 | -5.0092 | 1435.6613 | -5.0443 | 1436.9822 | -5.0443 | 1442.1902 | -5.0619 | 1444.8843 | -5.0619 | 1438.4282
| -4.8686 | 107.7028 | -4.974 | 112.345 | -4.9037 | 111.3158 | -4.9037 | 108.5992 | -4.9565 | 110.77 | -4.8862 | 1444.8843 | -4.974 | 1438.1868 | -4.974 | 1434.9427 | -5.0092 | 1434.7035 | -5.0092 | 1440.2437 | -5.0268 | 1442.0682 | -5.0268 | 1437.2228
| -4.8334 | 108.355 | -4.9389 | 112.3046 | -4.8686 | 110.9924 | -4.8686 | 108.9043 | -4.9213 | 110.7497 | -4.851 | 1443.5346 | -4.9389 | 1436.8619 | -4.9389 | 1434.2254 | -4.974 | 1434.4643 | -4.974 | 1438.9115 | -4.9916 | 1439.3954 | -4.9916 | 1433.9865
| -4.7983 | 110.122 | -4.9037 | 112.0022 | -4.8334 | 110.4056 | -4.8334 | 109.0465 | -4.8862 | 110.6688 | -4.8159 | 1441.9463 | -4.9037 | 1436.8619 | -4.9037 | 1433.9865 | -4.9389 | 1433.5093 | -4.9389 | 1438.9115 | -4.9565 | 1439.3954 | -4.9565 | 1432.1996
| -4.7631 | 111.7197 | -4.8686 | 111.8408 | -4.7983 | 109.4324 | -4.7983 | 109.1278 | -4.851 | 110.6688 | -4.7807 | 1442.0682 | -4.8686 | 1436.8619 | -4.8686 | 1432.0807 | -4.9037 | 1433.6285 | -4.9037 | 1438.9115 | -4.9213 | 1437.825 | -4.9213 | 1430.8939
| -4.728 | 112.345 | -4.8334 | 111.5582 | -4.7631 | 109.1888 | -4.7631 | 108.9449 | -4.8159 | 110.6688 | -4.7455 | 1441.3373 | -4.8334 | 1436.8619 | -4.8334 | 1432.0807 | -4.8686 | 1433.2708 | -4.8686 | 1437.825 | -4.8862 | 1437.825 | -4.8862 | 1430.42
| -4.6928 | 112.345 | -4.7983 | 110.6081 | -4.728 | 109.1888 | -4.728 | 108.9246 | -4.7807 | 110.6688 | -4.7104 | 1439.0324 | -4.7983 | 1435.1821 | -4.7983 | 1431.9619 | -4.8334 | 1433.9865 | -4.8334 | 1437.825 | -4.851 | 1437.7045 | -4.851 | 1428.6474
| -4.6577 | 112.1635 | -4.7631 | 110.6688 | -4.6928 | 108.8839 | -4.6928 | 108.8839 | -4.7455 | 110.8104 | -4.6752 | 1437.3432 | -4.7631 | 1433.1517 | -4.7631 | 1430.065 | -4.7983 | 1433.2708 | -4.7983 | 1436.9822 | -4.8159 | 1436.5013 | -4.8159 | 1428.6474
| -4.6225 | 111.8004 | -4.728 | 111.2552 | -4.6577 | 108.7619 | -4.6577 | 108.0902 | -4.7104 | 111.5784 | -4.6401 | 1436.8619 | -4.728 | 1430.1833 | -4.728 | 1429.0014 | -4.7631 | 1432.9135 | -4.7631 | 1433.6285 | -4.7807 | 1436.5013 | -4.7807 | 1428.6474
| -4.5874 | 111.3764 | -4.6928 | 110.3651 | -4.6225 | 108.9652 | -4.6225 | 107.4988 | -4.6752 | 110.5676 | -4.6049 | 1436.2612 | -4.6928 | 1429.3557 | -4.6928 | 1424.5364 | -4.728 | 1428.4116 | -4.728 | 1431.2496 | -4.7455 | 1435.5415 | -4.7455 | 1427.1166
| -4.5522 | 111.2956 | -4.6577 | 109.9599 | -4.5874 | 109.27 | -4.5874 | 107.4988 | -4.6401 | 110.2639 | -4.5698 | 1435.1821 | -4.6577 | 1427.2341 | -4.6577 | 1421.6193 | -4.6928 | 1427.4694 | -4.6928 | 1429.9467 | -4.7104 | 1432.7944 | -4.7104 | 1425.0046
| -4.5171 | 111.3966 | -4.6225 | 110.9116 | -4.5522 | 109.818 | -4.5522 | 107.4988 | -4.6049 | 109.6963 | -4.5346 | 1435.6613 | -4.6225 | 1426.2942 | -4.6225 | 1420.6888 | -4.6577 | 1425.2388 | -4.6577 | 1428.0581 | -4.6752 | 1427.9403 | -4.6752 | 1422.901
| -4.4819 | 111.2956 | -4.5874 | 110.9116 | -4.5171 | 112.103 | -4.5171 | 107.4988 | -4.5698 | 109.1684 | -4.4995 | 1435.9011 | -4.5874 | 1425.356 | -4.5874 | 1420.9213 | -4.6225 | 1422.4346 | -4.6225 | 1425.5904 | -4.6401 | 1424.7704 | -4.6401 | 1420.9213
| -4.4468 | 111.6591 | -4.5522 | 112.6673 | -4.4819 | 112.2038 | -4.4819 | 107.7844 | -4.5346 | 108.7619 | -4.4643 | 1435.9011 | -4.5522 | 1423.8349 | -4.5522 | 1421.2702 | -4.5874 | 1422.4346 | -4.5874 | 1424.6534 | -4.6049 | 1422.085 | -4.6049 | 1418.6003
| -4.4116 | 111.6591 | -4.5171 | 114.396 | -4.4468 | 112.627 | -4.4468 | 108.3346 | -4.4995 | 108.9449 | -4.4292 | 1437.3432 | -4.5171 | 1426.2942 | -4.5171 | 1423.3677 | -4.5522 | 1424.4194 | -4.5522 | 1424.6534 | -4.5698 | 1422.085 | -4.5698 | 1416.8647
| -4.3764 | 112.0829 | -4.4819 | 114.9575 | -4.4116 | 112.8888 | -4.4116 | 109.7165 | -4.4643 | 109.1684 | -4.394 | 1438.9115 | -4.4819 | 1426.999 | -4.4819 | 1424.1856 | -4.5171 | 1424.6534 | -4.5171 | 1426.999 | -4.5346 | 1423.8349 | -4.5346 | 1418.6003
| -4.3413 | 112.3248 | -4.4468 | 115.0176 | -4.3764 | 112.8687 | -4.3764 | 110.4866 | -4.4292 | 108.8636 | -4.3589 | 1440.0011 | -4.4468 | 1430.065 | -4.4468 | 1425.8249 | -4.4819 | 1426.1768 | -4.4819 | 1428.1759 | -4.4995 | 1425.4732 | -4.4995 | 1418.9479
| -4.3061 | 112.5867 | -4.4116 | 114.8573 | -4.3413 | 112.627 | -4.3413 | 111.0733 | -4.394 | 109.7977 | -4.3237 | 1440.0011 | -4.4116 | 1428.6474 | -4.4116 | 1428.4116 | -4.4468 | 1427.2341 | -4.4468 | 1429.2376 | -4.4643 | 1427.4694 | -4.4643 | 1420.2241
| -4.271 | 112.7076 | -4.3764 | 114.717 | -4.3061 | 112.5867 | -4.3061 | 111.1137 | -4.3589 | 110.041 | -4.2886 | 1442.4342 | -4.3764 | 1430.065 | -4.3764 | 1428.8834 | -4.4116 | 1428.5295 | -4.4116 | 1431.8431 | -4.4292 | 1429.8284 | -4.4292 | 1422.318
| -4.2358 | 113.1906 | -4.3413 | 114.3559 | -4.271 | 112.5867 | -4.271 | 111.134 | -4.3237 | 110.7497 | -4.2534 | 1444.3929 | -4.3413 | 1431.4869 | -4.3413 | 1429.3557 | -4.3764 | 1429.8284 | -4.3764 | 1430.1833 | -4.394 | 1430.42 | -4.394 | 1425.2388
| -4.2007 | 113.5124 | -4.3061 | 114.0147 | -4.2358 | 112.7479 | -4.2358 | 111.0733 | -4.2886 | 111.982 | -4.2183 | 1444.0248 | -4.3061 | 1430.065 | -4.3061 | 1430.8939 | -4.3413 | 1429.8284 | -4.3413 | 1431.8431 | -4.3589 | 1430.42 | -4.3589 | 1426.2942
| -4.1655 | 113.5526 | -4.271 | 114.0147 | -4.2007 | 113.2107 | -4.2007 | 110.5676 | -4.2534 | 112.2643 | -4.1831 | 1443.6571 | -4.271 | 1431.4869 | -4.271 | 1430.1833 | -4.3061 | 1430.42 | -4.3061 | 1434.3448 | -4.3237 | 1431.8431 | -4.3237 | 1424.8875
| -4.1304 | 113.2309 | -4.2358 | 113.9343 | -4.1655 | 113.4521 | -4.1655 | 110.1423 | -4.2183 | 112.4457 | -4.148 | 1443.2898 | -4.2358 | 1434.3448 | -4.2358 | 1431.9619 | -4.271 | 1430.42 | -4.271 | 1430.5384 | -4.2886 | 1431.8431 | -4.2886 | 1426.2942
| -4.0952 | 112.8284 | -4.2007 | 114.2756 | -4.1304 | 113.4521 | -4.1304 | 110.1423 | -4.1831 | 112.6069 | -4.1128 | 1441.8244 | -4.2007 | 1434.1059 | -4.2007 | 1430.8939 | -4.2358 | 1429.2376 | -4.2358 | 1431.3682 | -4.2534 | 1432.0807 | -4.2534 | 1427.7048
| -4.0601 | 112.3046 | -4.1655 | 114.4161 | -4.0952 | 113.6129 | -4.0952 | 110.3854 | -4.148 | 112.4659 | -4.0777 | 1440.6078 | -4.1655 | 1433.6285 | -4.1655 | 1429.9467 | -4.2007 | 1428.2937 | -4.2007 | 1431.8431 | -4.2183 | 1432.0807 | -4.2183 | 1424.8875
| -4.0249 | 111.336 | -4.1304 | 114.4161 | -4.0601 | 113.6732 | -4.0601 | 111.1137 | -4.1128 | 112.486 | -4.0425 | 1438.549 | -4.1304 | 1431.7243 | -4.1304 | 1429.4738 | -4.1655 | 1428.2937 | -4.1655 | 1431.8431 | -4.1831 | 1432.5564 | -4.1831 | 1425.8249
| -3.9898 | 110.4866 | -4.0952 | 114.4161 | -4.0249 | 113.6732 | -4.0249 | 111.336 | -4.0777 | 112.4659 | -4.0074 | 1436.5013 | -4.0952 | 1430.8939 | -4.0952 | 1427.1166 | -4.1304 | 1429.2376 | -4.1304 | 1430.5384 | -4.148 | 1432.5564 | -4.148 | 1426.529
| -3.9546 | 110.0004 | -4.0601 | 115.1178 | -3.9898 | 114.0347 | -3.9898 | 111.6591 | -4.0425 | 112.486 | -3.9722 | 1436.2612 | -4.0601 | 1430.3016 | -4.0601 | 1424.5364 | -4.0952 | 1428.6474 | -4.0952 | 1430.8939 | -4.1128 | 1432.7944 | -4.1128 | 1426.529
| -3.9195 | 109.7774 | -4.0249 | 115.3181 | -3.9546 | 114.0347 | -3.9546 | 111.639 | -4.0074 | 112.768 | -3.937 | 1434.4643 | -4.0249 | 1430.6569 | -4.0249 | 1425.356 | -4.0601 | 1428.1759 | -4.0601 | 1429.3557 | -4.0777 | 1431.4869 | -4.0777 | 1423.8349
| -3.8843 | 109.4527 | -3.9898 | 115.5184 | -3.9195 | 114.0347 | -3.9195 | 112.0627 | -3.9722 | 112.8687 | -3.9019 | 1432.7944 | -3.9898 | 1429.4738 | -3.9898 | 1424.8875 | -4.0249 | 1427.8225 | -4.0249 | 1429.2376 | -4.0425 | 1432.7944 | -4.0425 | 1424.1856
| -3.8492 | 109.0872 | -3.9546 | 115.9786 | -3.8843 | 114.0347 | -3.8843 | 112.2038 | -3.937 | 113.7937 | -3.8667 | 1431.131 | -3.9546 | 1429.2376 | -3.9546 | 1422.4346 | -3.9898 | 1427.3518 | -3.9898 | 1426.8815 | -4.0074 | 1430.1833 | -4.0074 | 1422.7843
| -3.814 | 108.8229 | -3.9195 | 115.9786 | -3.8492 | 113.8339 | -3.8492 | 112.2038 | -3.9019 | 114.2957 | -3.8316 | 1432.7944 | -3.9195 | 1425.0046 | -3.9195 | 1421.2702 | -3.9546 | 1425.9422 | -3.9546 | 1425.8249 | -3.9722 | 1429.3557 | -3.9722 | 1422.4346
| -3.7789 | 108.8229 | -3.8843 | 115.9786 | -3.814 | 113.8138 | -3.814 | 112.2038 | -3.8667 | 114.095 | -3.7964 | 1433.6285 | -3.8843 | 1422.4346 | -3.8843 | 1421.1539 | -3.9195 | 1423.251 | -3.9195 | 1423.0176 | -3.937 | 1428.0581 | -3.937 | 1420.3403
| -3.7437 | 109.0465 | -3.8492 | 115.7385 | -3.7789 | 113.8138 | -3.7789 | 112.1837 | -3.8316 | 113.5124 | -3.7613 | 1432.5564 | -3.8492 | 1419.1798 | -3.8492 | 1418.4845 | -3.8843 | 1421.1539 | -3.8843 | 1421.0376 | -3.9019 | 1424.6534 | -3.9019 | 1419.0638
| -3.7086 | 109.473 | -3.814 | 115.3181 | -3.7437 | 112.2643 | -3.7437 | 112.103 | -3.7964 | 113.5124 | -3.7261 | 1431.2496 | -3.814 | 1422.4346 | -3.814 | 1417.4428 | -3.8492 | 1419.4117 | -3.8492 | 1420.2241 | -3.8667 | 1421.2702 | -3.8667 | 1416.9803
| -3.6734 | 109.676 | -3.7789 | 114.0347 | -3.7086 | 111.9618 | -3.7086 | 111.7399 | -3.7613 | 113.5526 | -3.691 | 1429.9467 | -3.7789 | 1423.251 | -3.7789 | 1417.4428 | -3.814 | 1419.4117 | -3.814 | 1419.5277 | -3.8316 | 1420.108 | -3.8316 | 1416.2871
| -3.6383 | 109.7165 | -3.7437 | 113.7736 | -3.6734 | 111.4168 | -3.6734 | 111.4572 | -3.7261 | 113.5928 | -3.6558 | 1425.5904 | -3.7437 | 1423.7181 | -3.7437 | 1417.0959 | -3.7789 | 1419.4117 | -3.7789 | 1419.2957 | -3.7964 | 1419.5277 | -3.7964 | 1414.096
| -3.6031 | 109.5745 | -3.7086 | 112.9894 | -3.6383 | 111.336 | -3.6383 | 110.8509 | -3.691 | 113.5928 | -3.6207 | 1420.805 | -3.7086 | 1429.4738 | -3.7086 | 1417.0959 | -3.7437 | 1418.832 | -3.7437 | 1419.1798 | -3.7613 | 1419.5277 | -3.7613 | 1413.2902
| -3.5679 | 109.0669 | -3.6734 | 112.8284 | -3.6031 | 111.336 | -3.6031 | 110.0004 | -3.6558 | 113.5928 | -3.5855 | 1418.1371 | -3.6734 | 1429.8284 | -3.6734 | 1417.5585 | -3.7086 | 1412.715 | -3.7086 | 1421.0376 | -3.7261 | 1421.2702 | -3.7261 | 1412.1402
| -3.5328 | 108.9856 | -3.6383 | 112.345 | -3.5679 | 111.336 | -3.5679 | 109.7571 | -3.6207 | 113.3716 | -3.5504 | 1417.6741 | -3.6383 | 1421.2702 | -3.6383 | 1417.5585 | -3.6734 | 1415.9408 | -3.6734 | 1422.2015 | -3.691 | 1421.2702 | -3.691 | 1411.7955
| -3.4976 | 108.2939 | -3.6031 | 111.982 | -3.5328 | 111.2148 | -3.5328 | 109.6557 | -3.5855 | 113.3113 | -3.5152 | 1417.6741 | -3.6031 | 1418.7162 | -3.6031 | 1414.096 | -3.6383 | 1415.9408 | -3.6383 | 1422.2015 | -3.6558 | 1421.9685 | -3.6558 | 1411.7955
| -3.4625 | 107.866 | -3.5679 | 111.538 | -3.4976 | 110.77 | -3.4976 | 109.473 | -3.5504 | 113.2309 | -3.4801 | 1419.0638 | -3.5679 | 1410.5324 | -3.5679 | 1416.0562 | -3.6031 | 1415.9408 | -3.6031 | 1417.9056 | -3.6207 | 1421.9685 | -3.6207 | 1411.6806
| -3.4273 | 107.866 | -3.5328 | 111.5986 | -3.4625 | 110.77 | -3.4625 | 108.9856 | -3.5152 | 113.0297 | -3.4449 | 1425.8249 | -3.5328 | 1410.5324 | -3.5328 | 1415.8254 | -3.5679 | 1415.9408 | -3.5679 | 1413.2902 | -3.5855 | 1419.1798 | -3.5855 | 1410.5324
| -3.3922 | 107.866 | -3.4976 | 111.9417 | -3.4273 | 110.6486 | -3.4273 | 108.7416 | -3.4801 | 112.3853 | -3.4098 | 1428.1759 | -3.4976 | 1406.98 | -3.4976 | 1414.096 | -3.5328 | 1416.8647 | -3.5328 | 1411.2212 | -3.5504 | 1415.3638 | -3.5504 | 1409.8441
| -3.357 | 107.4376 | -3.4625 | 112.4457 | -3.3922 | 110.3044 | -3.3922 | 108.1513 | -3.4449 | 111.6995 | -3.3746 | 1431.7243 | -3.4625 | 1412.1402 | -3.4625 | 1414.096 | -3.4976 | 1419.5277 | -3.4976 | 1411.2212 | -3.5152 | 1413.9809 | -3.5152 | 1409.8441
| -3.3219 | 108.2328 | -3.4273 | 112.9492 | -3.357 | 110.0815 | -3.357 | 108.1513 | -3.4098 | 111.6995 | -3.3395 | 1430.42 | -3.4273 | 1416.2871 | -3.4273 | 1415.8254 | -3.4625 | 1413.8657 | -3.4625 | 1418.3687 | -3.4801 | 1413.9809 | -3.4801 | 1407.7814
| -3.2867 | 108.4364 | -3.3922 | 112.6874 | -3.3219 | 110.1423 | -3.3219 | 108.1513 | -3.3746 | 111.0935 | -3.3043 | 1430.42 | -3.3922 | 1416.2871 | -3.3922 | 1411.4509 | -3.4273 | 1412.3701 | -3.4273 | 1423.9518 | -3.4449 | 1418.1371 | -3.4449 | 1408.1249
| -3.2516 | 108.5992 | -3.357 | 112.7277 | -3.2867 | 110.0612 | -3.2867 | 108.6602 | -3.3395 | 110.7902 | -3.2692 | 1426.8815 | -3.357 | 1416.2871 | -3.357 | 1412.1402 | -3.3922 | 1412.3701 | -3.3922 | 1425.356 | -3.4098 | 1421.3865 | -3.4098 | 1408.1249
| -3.2164 | 108.5992 | -3.3219 | 111.7803 | -3.2516 | 109.6557 | -3.2516 | 108.7416 | -3.3043 | 110.2233 | -3.234 | 1426.6465 | -3.3219 | 1416.2871 | -3.3219 | 1415.9408 | -3.357 | 1414.7873 | -3.357 | 1423.9518 | -3.3746 | 1428.5295 | -3.3746 | 1409.9588
| -3.1813 | 108.5992 | -3.2867 | 111.4572 | -3.2164 | 109.4527 | -3.2164 | 108.9043 | -3.2692 | 109.8788 | -3.1989 | 1423.251 | -3.2867 | 1415.8254 | -3.2867 | 1417.0959 | -3.3219 | 1414.9026 | -3.3219 | 1413.5204 | -3.3395 | 1428.5295 | -3.3395 | 1409.9588
| -3.1461 | 108.2939 | -3.2516 | 111.3562 | -3.1813 | 109.2497 | -3.1813 | 108.9043 | -3.234 | 109.9802 | -3.1637 | 1422.318 | -3.2516 | 1415.71 | -3.2516 | 1413.6355 | -3.2867 | 1414.9026 | -3.2867 | 1412.2552 | -3.3043 | 1421.3865 | -3.3043 | 1410.5324
| -3.111 | 108.0698 | -3.2164 | 111.0329 | -3.1461 | 108.8229 | -3.1461 | 108.7416 | -3.1989 | 110.2233 | -3.1285 | 1416.8647 | -3.2164 | 1413.7506 | -3.2164 | 1412.1402 | -3.2516 | 1412.9451 | -3.2516 | 1409.6147 | -3.2692 | 1415.1332 | -3.2692 | 1408.1249
| -3.0758 | 107.5396 | -3.1813 | 110.7497 | -3.111 | 108.8229 | -3.111 | 107.7436 | -3.1637 | 110.1626 | -3.0934 | 1414.096 | -3.1813 | 1413.6355 | -3.1813 | 1406.6367 | -3.2164 | 1411.336 | -3.2164 | 1413.4053 | -3.234 | 1410.1882 | -3.234 | 1406.6367
| -3.0407 | 106.9269 | -3.1461 | 110.689 | -3.0758 | 109.2497 | -3.0758 | 107.3151 | -3.1285 | 110.0612 | -3.0582 | 1414.7873 | -3.1461 | 1412.9451 | -3.1461 | 1404.3495 | -3.1813 | 1411.1064 | -3.1813 | 1416.1717 | -3.1989 | 1410.1882 | -3.1989 | 1401.6077
| -3.0055 | 106.661 | -3.111 | 110.4461 | -3.0407 | 109.3918 | -3.0407 | 107.1313 | -3.0934 | 110.0612 | -3.0231 | 1414.7873 | -3.111 | 1413.6355 | -3.111 | 1404.3495 | -3.1461 | 1411.1064 | -3.1461 | 1416.1717 | -3.1637 | 1411.7955 | -3.1637 | 1401.6077
| -2.9704 | 106.2721 | -3.0758 | 110.77 | -3.0055 | 109.5339 | -3.0055 | 107.07 | -3.0582 | 109.6963 | -2.9879 | 1420.6888 | -3.0758 | 1414.3264 | -3.0758 | 1405.8358 | -3.111 | 1410.3029 | -3.111 | 1411.5657 | -3.1285 | 1415.1332 | -3.1285 | 1401.9503
| -2.9352 | 105.7801 | -3.0407 | 110.9116 | -2.9704 | 109.6963 | -2.9704 | 107.07 | -3.0231 | 108.9652 | -2.9528 | 1420.805 | -3.0407 | 1414.3264 | -3.0407 | 1406.8655 | -3.0758 | 1410.3029 | -3.0758 | 1411.1064 | -3.0934 | 1413.7506 | -3.0934 | 1406.98
| -2.9001 | 105.4928 | -3.0055 | 111.2148 | -2.9352 | 109.6963 | -2.9352 | 107.3151 | -2.9879 | 108.8229 | -2.9176 | 1420.805 | -3.0055 | 1413.1751 | -3.0055 | 1405.8358 | -3.0407 | 1410.762 | -3.0407 | 1409.5001 | -3.0582 | 1413.7506 | -3.0582 | 1410.1882
| -2.8649 | 105.2875 | -2.9704 | 110.8509 | -2.9001 | 109.2497 | -2.9001 | 107.662 | -2.9528 | 108.9043 | -2.8825 | 1420.4564 | -2.9704 | 1409.9588 | -2.9704 | 1405.4927 | -3.0055 | 1406.8655 | -3.0055 | 1408.2395 | -3.0231 | 1413.7506 | -3.0231 | 1410.3029
| -2.8298 | 105.0203 | -2.9352 | 111.2754 | -2.8649 | 108.8433 | -2.8649 | 108.3346 | -2.9176 | 108.9043 | -2.8473 | 1419.6437 | -2.9352 | 1406.179 | -2.9352 | 1404.921 | -2.9704 | 1406.179 | -2.9704 | 1409.2708 | -2.9879 | 1411.9104 | -2.9879 | 1409.1562
| -2.7946 | 104.9175 | -2.9001 | 110.8104 | -2.8298 | 108.5381 | -2.8298 | 108.4567 | -2.8825 | 109.3918 | -2.8122 | 1418.6003 | -2.9001 | 1405.0354 | -2.9001 | 1404.0066 | -2.9352 | 1406.179 | -2.9352 | 1409.2708 | -2.9528 | 1411.9104 | -2.9528 | 1403.0925
| -2.7594 | 104.8969 | -2.8649 | 110.3651 | -2.7946 | 107.866 | -2.7946 | 108.4567 | -2.8473 | 109.7571 | -2.777 | 1418.2529 | -2.8649 | 1406.179 | -2.8649 | 1404.0066 | -2.9001 | 1405.8358 | -2.9001 | 1406.8655 | -2.9176 | 1410.1882 | -2.9176 | 1398.4107
| -2.7243 | 104.6705 | -2.8298 | 110.0612 | -2.7594 | 107.6212 | -2.7594 | 108.3346 | -2.8122 | 109.676 | -2.7419 | 1416.0562 | -2.8298 | 1407.0944 | -2.8298 | 1404.921 | -2.8649 | 1405.7215 | -2.8649 | 1406.8655 | -2.8825 | 1410.5324 | -2.8825 | 1395.5561
| -2.6891 | 103.7223 | -2.7946 | 109.9599 | -2.7243 | 107.2742 | -2.7243 | 106.5996 | -2.777 | 109.676 | -2.7067 | 1413.6355 | -2.7946 | 1408.9269 | -2.7946 | 1405.264 | -2.8298 | 1405.7215 | -2.8298 | 1406.8655 | -2.8473 | 1410.1882 | -2.8473 | 1394.8708
| -2.654 | 102.8746 | -2.7594 | 109.6557 | -2.6891 | 106.8247 | -2.6891 | 106.4359 | -2.7419 | 109.6557 | -2.6716 | 1413.5204 | -2.7594 | 1410.1882 | -2.7594 | 1403.6638 | -2.7946 | 1405.7215 | -2.7946 | 1406.8655 | -2.8122 | 1408.5832 | -2.8122 | 1400.9225
| -2.6188 | 102.7296 | -2.7243 | 109.2497 | -2.654 | 107.2742 | -2.654 | 106.1697 | -2.7067 | 109.3512 | -2.6364 | 1413.6355 | -2.7243 | 1408.9269 | -2.7243 | 1403.0925 | -2.7594 | 1405.7215 | -2.7594 | 1407.0944 | -2.777 | 1407.4379 | -2.777 | 1404.921
| -2.5837 | 102.7296 | -2.6891 | 108.3957 | -2.6188 | 107.2947 | -2.6188 | 106.0672 | -2.6716 | 108.9856 | -2.6013 | 1413.5204 | -2.6891 | 1407.0944 | -2.6891 | 1403.0925 | -2.7243 | 1405.264 | -2.7243 | 1406.8655 | -2.7419 | 1406.5222 | -2.7419 | 1406.6367
| -2.5485 | 102.3773 | -2.654 | 108.0494 | -2.5837 | 107.4376 | -2.5837 | 105.5955 | -2.6364 | 108.4974 | -2.5661 | 1412.4851 | -2.654 | 1405.8358 | -2.654 | 1401.9503 | -2.6891 | 1403.8923 | -2.6891 | 1406.6367 | -2.7067 | 1406.2934 | -2.7067 | 1404.1209
| -2.5134 | 102.481 | -2.6188 | 107.8048 | -2.5485 | 106.4359 | -2.5485 | 105.2875 | -2.6013 | 107.866 | -2.531 | 1414.2112 | -2.6188 | 1405.8358 | -2.6188 | 1404.0066 | -2.654 | 1403.8923 | -2.654 | 1406.6367 | -2.6716 | 1404.3495 | -2.6716 | 1403.0925
| -2.4782 | 102.8746 | -2.5837 | 107.8864 | -2.5134 | 106.7428 | -2.5134 | 105.1025 | -2.5661 | 107.5192 | -2.4958 | 1415.71 | -2.5837 | 1407.0944 | -2.5837 | 1406.0646 | -2.6188 | 1403.8923 | -2.6188 | 1405.1497 | -2.6364 | 1404.3495 | -2.6364 | 1401.8361
| -2.4431 | 103.0195 | -2.5485 | 108.0494 | -2.4782 | 105.5544 | -2.4782 | 104.8969 | -2.531 | 107.2334 | -2.4607 | 1417.4428 | -2.5485 | 1407.5523 | -2.5485 | 1406.6367 | -2.5837 | 1403.8923 | -2.5837 | 1406.2934 | -2.6013 | 1404.3495 | -2.6013 | 1399.5524
| -2.4079 | 102.7711 | -2.5134 | 107.9068 | -2.4431 | 105.3696 | -2.4431 | 104.794 | -2.4958 | 106.8042 | -2.4255 | 1417.6741 | -2.5134 | 1409.7294 | -2.5134 | 1408.4686 | -2.5485 | 1404.3495 | -2.5485 | 1407.8959 | -2.5661 | 1404.1209 | -2.5661 | 1401.0367
| -2.3728 | 102.3358 | -2.4782 | 107.6416 | -2.4079 | 105.2669 | -2.4079 | 104.5676 | -2.4607 | 106.3335 | -2.3903 | 1418.3687 | -2.4782 | 1410.3029 | -2.4782 | 1410.0735 | -2.5134 | 1407.4379 | -2.5134 | 1410.4177 | -2.531 | 1405.3784 | -2.531 | 1401.8361
| -2.3376 | 102.3358 | -2.4431 | 106.5178 | -2.3728 | 105.0614 | -2.3728 | 104.238 | -2.4255 | 106.6201 | -2.3552 | 1419.6437 | -2.4431 | 1410.3029 | -2.4431 | 1410.762 | -2.4782 | 1407.4379 | -2.4782 | 1411.4509 | -2.4958 | 1408.4686 | -2.4958 | 1401.3793
| -2.3025 | 102.3358 | -2.4079 | 106.1901 | -2.3376 | 104.4646 | -2.3376 | 104.1349 | -2.3903 | 106.3335 | -2.32 | 1419.6437 | -2.4079 | 1409.0415 | -2.4079 | 1410.762 | -2.4431 | 1405.6071 | -2.4431 | 1412.1402 | -2.4607 | 1409.0415 | -2.4607 | 1401.3793
| -2.2673 | 102.8332 | -2.3728 | 105.9852 | -2.3025 | 103.5571 | -2.3025 | 103.5571 | -2.3552 | 106.8451 | -2.2849 | 1419.1798 | -2.3728 | 1411.2212 | -2.3728 | 1410.762 | -2.4079 | 1408.8123 | -2.4079 | 1411.4509 | -2.4255 | 1408.4686 | -2.4255 | 1402.4071
| -2.2322 | 103.2677 | -2.3376 | 105.7391 | -2.2673 | 103.5571 | -2.2673 | 103.0816 | -2.32 | 106.5996 | -2.2497 | 1419.6437 | -2.3376 | 1412.2552 | -2.3376 | 1410.6472 | -2.3728 | 1409.5001 | -2.3728 | 1410.9916 | -2.3903 | 1409.6147 | -2.3903 | 1404.6924
| -2.197 | 103.3504 | -2.3025 | 105.1847 | -2.2322 | 103.619 | -2.2322 | 102.7089 | -2.2849 | 106.4973 | -2.2146 | 1420.5726 | -2.3025 | 1412.715 | -2.3025 | 1410.4177 | -2.3376 | 1407.5523 | -2.3376 | 1408.6977 | -2.3552 | 1408.9269 | -2.3552 | 1404.6924
| -2.1619 | 103.7429 | -2.2673 | 105.1847 | -2.197 | 103.7842 | -2.197 | 102.2736 | -2.2497 | 106.2926 | -2.1794 | 1421.3865 | -2.2673 | 1411.9104 | -2.2673 | 1410.4177 | -2.3025 | 1409.5001 | -2.3025 | 1408.6977 | -2.32 | 1408.6977 | -2.32 | 1404.3495
| -2.1267 | 103.9287 | -2.2322 | 105.1847 | -2.1619 | 103.7223 | -2.1619 | 101.6089 | -2.2146 | 105.8827 | -2.1443 | 1419.6437 | -2.2322 | 1411.6806 | -2.2322 | 1410.0735 | -2.2673 | 1408.9269 | -2.2673 | 1408.6977 | -2.2849 | 1408.6977 | -2.2849 | 1402.7498
| -2.0916 | 103.9287 | -2.197 | 105.6365 | -2.1267 | 103.5157 | -2.1267 | 101.4634 | -2.1794 | 105.0408 | -2.1091 | 1419.6437 | -2.197 | 1410.8768 | -2.197 | 1409.3854 | -2.2322 | 1406.6367 | -2.2322 | 1409.1562 | -2.2497 | 1409.6147 | -2.2497 | 1400.0091
| -2.0564 | 104.0937 | -2.1619 | 105.6365 | -2.0916 | 103.5157 | -2.0916 | 101.7129 | -2.1443 | 103.8255 | -2.074 | 1418.6003 | -2.1619 | 1410.762 | -2.1619 | 1409.3854 | -2.197 | 1406.0646 | -2.197 | 1409.5001 | -2.2146 | 1409.3854 | -2.2146 | 1400.0091
| -2.0213 | 104.1968 | -2.1267 | 105.6365 | -2.0564 | 103.4537 | -2.0564 | 101.7752 | -2.1091 | 103.6603 | -2.0388 | 1416.0562 | -2.1267 | 1409.7294 | -2.1267 | 1409.0415 | -2.1619 | 1406.0646 | -2.1619 | 1413.4053 | -2.1794 | 1409.6147 | -2.1794 | 1402.1787
| -1.9861 | 104.1555 | -2.0916 | 105.082 | -2.0213 | 103.5157 | -2.0213 | 101.6713 | -2.074 | 103.2057 | -2.0037 | 1414.9026 | -2.0916 | 1409.5001 | -2.0916 | 1406.5222 | -2.1267 | 1406.0646 | -2.1267 | 1409.3854 | -2.1443 | 1410.4177 | -2.1443 | 1402.864
| -1.9509 | 104.0937 | -2.0564 | 105.3285 | -1.9861 | 104.4234 | -1.9861 | 101.6921 | -2.0388 | 103.2057 | -1.9685 | 1414.3264 | -2.0564 | 1408.0104 | -2.0564 | 1406.2934 | -2.0916 | 1406.0646 | -2.0916 | 1409.0415 | -2.1091 | 1410.4177 | -2.1091 | 1403.2068
| -1.9158 | 103.8874 | -2.0213 | 105.8007 | -1.9509 | 104.7734 | -1.9509 | 101.6505 | -2.0037 | 103.619 | -1.9334 | 1412.715 | -2.0213 | 1404.2352 | -2.0213 | 1402.864 | -2.0564 | 1406.0646 | -2.0564 | 1404.4638 | -2.074 | 1410.3029 | -2.074 | 1404.3495
| -1.8806 | 103.8874 | -1.9861 | 106.3335 | -1.9158 | 104.7529 | -1.9158 | 101.5674 | -1.9685 | 103.5364 | -1.8982 | 1411.6806 | -1.9861 | 1403.5495 | -1.9861 | 1402.6356 | -2.0213 | 1401.4935 | -2.0213 | 1402.5214 | -2.0388 | 1409.6147 | -2.0388 | 1403.5495
| -1.8455 | 104.1555 | -1.9509 | 107.0904 | -1.8806 | 105.1847 | -1.8806 | 101.6921 | -1.9334 | 103.1023 | -1.8631 | 1411.6806 | -1.9509 | 1403.5495 | -1.9509 | 1401.4935 | -1.9861 | 1401.4935 | -1.9861 | 1402.5214 | -2.0037 | 1408.2395 | -2.0037 | 1400.4658
| -1.8103 | 104.3204 | -1.9158 | 107.0904 | -1.8455 | 105.2875 | -1.8455 | 102.4602 | -1.8982 | 103.1023 | -1.8279 | 1409.6147 | -1.9158 | 1404.2352 | -1.9158 | 1402.864 | -1.9509 | 1401.0367 | -1.9509 | 1402.1787 | -1.9685 | 1404.2352 | -1.9685 | 1398.8674
| -1.7752 | 105.1847 | -1.8806 | 107.0904 | -1.8103 | 105.308 | -1.8103 | 103.3297 | -1.8631 | 103.1023 | -1.7928 | 1409.6147 | -1.8806 | 1405.7215 | -1.8806 | 1405.6071 | -1.9158 | 1401.9503 | -1.9158 | 1402.7498 | -1.9334 | 1403.8923 | -1.9334 | 1400.0091
| -1.74 | 105.4723 | -1.8455 | 107.2742 | -1.7752 | 105.3491 | -1.7752 | 103.8255 | -1.8279 | 103.1023 | -1.7576 | 1411.9104 | -1.8455 | 1405.7215 | -1.8455 | 1404.2352 | -1.8806 | 1403.8923 | -1.8806 | 1404.5781 | -1.8982 | 1402.5214 | -1.8982 | 1400.0091
| -1.7049 | 105.6776 | -1.8103 | 107.1721 | -1.74 | 105.9237 | -1.74 | 104.0318 | -1.7928 | 103.1023 | -1.7225 | 1413.9809 | -1.8103 | 1405.7215 | -1.8103 | 1401.6077 | -1.8455 | 1404.8067 | -1.8455 | 1405.4927 | -1.8631 | 1402.7498 | -1.8631 | 1400.0091
| -1.6697 | 104.4646 | -1.7752 | 107.4376 | -1.7049 | 105.9647 | -1.7049 | 104.1555 | -1.7576 | 103.3297 | -1.6873 | 1411.9104 | -1.7752 | 1405.264 | -1.7752 | 1400.8083 | -1.8103 | 1403.8923 | -1.8103 | 1404.5781 | -1.8279 | 1402.7498 | -1.8279 | 1401.0367
| -1.6346 | 103.4331 | -1.74 | 107.7232 | -1.6697 | 105.9237 | -1.6697 | 104.4234 | -1.7225 | 104.1968 | -1.6522 | 1406.98 | -1.74 | 1404.5781 | -1.74 | 1398.8674 | -1.7752 | 1400.9225 | -1.7752 | 1403.8923 | -1.7928 | 1402.6356 | -1.7928 | 1401.0367
| -1.5994 | 102.9367 | -1.7049 | 107.5192 | -1.6346 | 106.0057 | -1.6346 | 104.6294 | -1.6873 | 104.1968 | -1.617 | 1404.5781 | -1.7049 | 1403.2068 | -1.7049 | 1394.071 | -1.74 | 1399.0957 | -1.74 | 1401.0367 | -1.7576 | 1400.9225 | -1.7576 | 1399.3241
| -1.5643 | 102.9367 | -1.6697 | 107.4376 | -1.5994 | 106.0057 | -1.5994 | 104.547 | -1.6522 | 104.6088 | -1.5818 | 1403.0925 | -1.6697 | 1402.5214 | -1.6697 | 1393.8425 | -1.7049 | 1398.1824 | -1.7049 | 1398.4107 | -1.7225 | 1400.9225 | -1.7225 | 1399.3241
| -1.5291 | 102.8125 | -1.6346 | 107.4172 | -1.5643 | 105.9647 | -1.5643 | 104.1968 | -1.617 | 104.6088 | -1.5467 | 1402.5214 | -1.6346 | 1401.7219 | -1.6346 | 1392.0133 | -1.6697 | 1396.6981 | -1.6697 | 1395.8987 | -1.6873 | 1400.4658 | -1.6873 | 1394.2995
| -1.494 | 102.8746 | -1.5994 | 107.4172 | -1.5291 | 106.395 | -1.5291 | 103.7223 | -1.5818 | 104.7117 | -1.5115 | 1403.5495 | -1.5994 | 1398.7532 | -1.5994 | 1391.2125 | -1.6346 | 1396.1272 | -1.6346 | 1395.8987 | -1.6522 | 1400.3516 | -1.6522 | 1391.2125
| -1.4588 | 103.0816 | -1.5643 | 107.4376 | -1.494 | 106.395 | -1.494 | 103.681 | -1.5467 | 105.616 | -1.4764 | 1405.6071 | -1.5643 | 1398.5249 | -1.5643 | 1390.7548 | -1.5994 | 1396.1272 | -1.5994 | 1395.7845 | -1.617 | 1397.9541 | -1.617 | 1384.2143
| -1.4237 | 103.4951 | -1.5291 | 108.2124 | -1.4588 | 106.0877 | -1.4588 | 103.619 | -1.5115 | 105.8827 | -1.4412 | 1404.921 | -1.5291 | 1398.5249 | -1.5291 | 1391.2125 | -1.5643 | 1394.4138 | -1.5643 | 1395.0992 | -1.5818 | 1395.2135 | -1.5818 | 1386.9727
| -1.3885 | 103.1229 | -1.494 | 108.9043 | -1.4237 | 105.1642 | -1.4237 | 103.4537 | -1.4764 | 105.8007 | -1.4061 | 1402.5214 | -1.494 | 1396.3555 | -1.494 | 1392.3564 | -1.5291 | 1393.9567 | -1.5291 | 1395.0992 | -1.5467 | 1393.0424 | -1.5467 | 1383.293
| -1.3534 | 102.7711 | -1.4588 | 108.6602 | -1.3885 | 104.794 | -1.3885 | 103.6397 | -1.4412 | 105.7391 | -1.3709 | 1400.9225 | -1.4588 | 1394.6423 | -1.4588 | 1390.9837 | -1.494 | 1393.7282 | -1.494 | 1394.6423 | -1.5115 | 1392.242 | -1.5115 | 1383.293
| -1.3182 | 102.5846 | -1.4237 | 107.6824 | -1.3534 | 104.2792 | -1.3534 | 104.0524 | -1.4061 | 105.8827 | -1.3358 | 1395.7845 | -1.4237 | 1392.3564 | -1.4237 | 1389.7244 | -1.4588 | 1392.6994 | -1.4588 | 1391.4414 | -1.4764 | 1392.3564 | -1.4764 | 1386.9727
| -1.2831 | 102.5846 | -1.3885 | 106.7838 | -1.3182 | 103.9287 | -1.3182 | 104.0524 | -1.3709 | 106.354 | -1.3006 | 1395.5561 | -1.3885 | 1392.1277 | -1.3885 | 1387.5465 | -1.4237 | 1391.4414 | -1.4237 | 1389.9534 | -1.4412 | 1392.242 | -1.4412 | 1387.7759
| -1.2479 | 102.5846 | -1.3534 | 106.7224 | -1.2831 | 103.8874 | -1.2831 | 103.8461 | -1.3358 | 106.8247 | -1.2655 | 1395.5561 | -1.3534 | 1389.8389 | -1.3534 | 1387.0875 | -1.3885 | 1389.9534 | -1.3885 | 1389.6098 | -1.4061 | 1392.3564 | -1.4061 | 1385.1346
| -1.2128 | 102.8125 | -1.3182 | 106.5792 | -1.2479 | 103.908 | -1.2479 | 103.7223 | -1.3006 | 107.07 | -1.2303 | 1397.269 | -1.3182 | 1388.3493 | -1.3182 | 1387.0875 | -1.3534 | 1388.0053 | -1.3534 | 1388.12 | -1.3709 | 1391.5558 | -1.3709 | 1384.3294
| -1.1776 | 103.2677 | -1.2831 | 106.1697 | -1.2128 | 104.4234 | -1.2128 | 103.3711 | -1.2655 | 107.07 | -1.1952 | 1397.269 | -1.2831 | 1387.0875 | -1.2831 | 1387.0875 | -1.3182 | 1387.0875 | -1.3182 | 1387.4318 | -1.3358 | 1389.7244 | -1.3358 | 1385.1346
| -1.1424 | 103.4537 | -1.2479 | 106.354 | -1.1776 | 105.2053 | -1.1776 | 103.185 | -1.2303 | 106.8247 | -1.16 | 1397.269 | -1.2479 | 1388.6932 | -1.2479 | 1387.0875 | -1.2831 | 1387.0875 | -1.2831 | 1387.4318 | -1.3006 | 1387.8906 | -1.3006 | 1384.3294
| -1.1073 | 103.7842 | -1.2128 | 106.886 | -1.1424 | 104.7117 | -1.1424 | 103.0195 | -1.1952 | 106.1082 | -1.1249 | 1396.4697 | -1.2128 | 1388.6932 | -1.2128 | 1386.3987 | -1.2479 | 1386.1689 | -1.2479 | 1387.5465 | -1.2655 | 1387.6612 | -1.2655 | 1385.1346
| -1.0721 | 103.8048 | -1.1776 | 106.9064 | -1.1073 | 104.5676 | -1.1073 | 102.6882 | -1.16 | 104.8969 | -1.0897 | 1395.8987 | -1.1776 | 1388.6932 | -1.1776 | 1385.3645 | -1.2128 | 1387.0875 | -1.2128 | 1388.0053 | -1.2303 | 1387.6612 | -1.2303 | 1386.0541
| -1.037 | 104.0318 | -1.1424 | 106.9064 | -1.0721 | 104.0318 | -1.0721 | 102.7089 | -1.1249 | 104.4852 | -1.0546 | 1394.8708 | -1.1424 | 1386.6283 | -1.1424 | 1384.3294 | -1.1776 | 1388.3493 | -1.1776 | 1388.0053 | -1.1952 | 1387.4318 | -1.1952 | 1388.5786
| -1.0018 | 104.341 | -1.1073 | 106.886 | -1.037 | 102.916 | -1.037 | 103.0195 | -1.0897 | 104.3616 | -1.0194 | 1394.2995 | -1.1073 | 1387.317 | -1.1073 | 1385.0196 | -1.1424 | 1386.8579 | -1.1424 | 1386.3987 | -1.16 | 1387.4318 | -1.16 | 1388.5786
| -0.96669 | 104.4852 | -1.0721 | 105.7596 | -1.0018 | 102.916 | -1.0018 | 103.0609 | -1.0546 | 104.3616 | -0.98426 | 1394.2995 | -1.0721 | 1384.2143 | -1.0721 | 1385.0196 | -1.1073 | 1385.2496 | -1.1073 | 1385.4795 | -1.1249 | 1387.4318 | -1.1249 | 1382.0246
| -0.93153 | 104.4852 | -1.037 | 105.7596 | -0.96669 | 102.916 | -0.96669 | 102.6675 | -1.0194 | 104.4852 | -0.94911 | 1395.2135 | -1.037 | 1384.2143 | -1.037 | 1385.9392 | -1.0721 | 1384.9046 | -1.0721 | 1384.0992 | -1.0897 | 1385.8243 | -1.0897 | 1379.5974
| -0.89638 | 105.3696 | -1.0018 | 105.7596 | -0.93153 | 103.5984 | -0.93153 | 102.626 | -0.98426 | 104.8146 | -0.91396 | 1395.5561 | -1.0018 | 1387.317 | -1.0018 | 1388.3493 | -1.037 | 1385.2496 | -1.037 | 1382.14 | -1.0546 | 1385.0196 | -1.0546 | 1374.6
| -0.86123 | 104.9997 | -0.96669 | 106.2721 | -0.89638 | 103.9287 | -0.89638 | 102.4602 | -0.94911 | 104.9997 | -0.87881 | 1396.3555 | -0.96669 | 1387.317 | -0.96669 | 1388.9224 | -1.0018 | 1385.4795 | -1.0018 | 1384.2143 | -1.0194 | 1384.7896 | -1.0194 | 1373.8993
| -0.82608 | 105.1436 | -0.93153 | 106.5792 | -0.86123 | 104.3822 | -0.86123 | 102.2113 | -0.91396 | 104.938 | -0.84365 | 1396.6981 | -0.93153 | 1386.2838 | -0.93153 | 1388.2346 | -0.96669 | 1386.7431 | -0.96669 | 1384.7896 | -0.98426 | 1384.7896 | -0.98426 | 1374.2498
| -0.79092 | 105.1642 | -0.89638 | 106.6406 | -0.82608 | 105.6981 | -0.82608 | 102.4602 | -0.87881 | 104.2998 | -0.8085 | 1396.8123 | -0.89638 | 1388.6932 | -0.89638 | 1388.2346 | -0.93153 | 1386.7431 | -0.93153 | 1384.6745 | -0.94911 | 1383.7537 | -0.94911 | 1374.0162
| -0.75577 | 105.1436 | -0.86123 | 106.8656 | -0.79092 | 106.2106 | -0.79092 | 102.6468 | -0.84365 | 104.1762 | -0.77335 | 1396.9265 | -0.86123 | 1388.2346 | -0.86123 | 1386.5135 | -0.89638 | 1386.7431 | -0.89638 | 1384.6745 | -0.91396 | 1382.2554 | -0.91396 | 1373.8993
| -0.72062 | 105.1231 | -0.82608 | 106.9678 | -0.75577 | 106.8451 | -0.75577 | 102.916 | -0.8085 | 103.5777 | -0.7382 | 1399.0957 | -0.82608 | 1388.2346 | -0.82608 | 1386.3987 | -0.86123 | 1385.9392 | -0.86123 | 1386.6283 | -0.87881 | 1383.8689 | -0.87881 | 1374.0162
| -0.68547 | 105.2258 | -0.79092 | 107.3763 | -0.72062 | 106.7428 | -0.72062 | 103.4124 | -0.77335 | 103.5984 | -0.70304 | 1400.0091 | -0.79092 | 1391.5558 | -0.79092 | 1387.5465 | -0.82608 | 1385.8243 | -0.82608 | 1387.4318 | -0.84365 | 1385.4795 | -0.84365 | 1379.1342
| -0.65032 | 105.3696 | -0.75577 | 107.5804 | -0.68547 | 106.7633 | -0.68547 | 103.4124 | -0.7382 | 104.5264 | -0.66789 | 1401.0367 | -0.75577 | 1392.4707 | -0.75577 | 1388.0053 | -0.79092 | 1385.9392 | -0.79092 | 1389.4953 | -0.8085 | 1385.5944 | -0.8085 | 1379.3658
| -0.61516 | 105.4107 | -0.72062 | 107.7844 | -0.65032 | 106.8451 | -0.65032 | 103.908 | -0.70304 | 104.5676 | -0.63274 | 1401.4935 | -0.72062 | 1392.4707 | -0.72062 | 1388.4639 | -0.75577 | 1386.3987 | -0.75577 | 1390.2969 | -0.77335 | 1385.5944 | -0.77335 | 1379.3658
| -0.58001 | 105.4312 | -0.68547 | 108.192 | -0.61516 | 106.0467 | -0.61516 | 104.8969 | -0.66789 | 104.8969 | -0.59759 | 1402.7498 | -0.68547 | 1393.1567 | -0.68547 | 1390.7548 | -0.72062 | 1386.3987 | -0.72062 | 1390.2969 | -0.7382 | 1387.2023 | -0.7382 | 1379.3658
| -0.54486 | 105.4312 | -0.65032 | 108.2531 | -0.58001 | 105.9852 | -0.58001 | 104.8969 | -0.63274 | 105.082 | -0.56244 | 1402.7498 | -0.65032 | 1393.1567 | -0.65032 | 1391.2125 | -0.68547 | 1387.0875 | -0.68547 | 1391.4414 | -0.70304 | 1387.5465 | -0.70304 | 1377.1618
| -0.50971 | 105.4312 | -0.61516 | 108.2939 | -0.54486 | 106.2311 | -0.54486 | 104.6705 | -0.59759 | 105.5339 | -0.52728 | 1403.7781 | -0.61516 | 1393.1567 | -0.61516 | 1392.0133 | -0.65032 | 1389.8389 | -0.65032 | 1392.242 | -0.66789 | 1389.2662 | -0.66789 | 1376.5805
| -0.47455 | 105.4312 | -0.58001 | 108.4364 | -0.50971 | 105.8827 | -0.50971 | 105.5544 | -0.56244 | 106.2106 | -0.49213 | 1405.7215 | -0.58001 | 1393.1567 | -0.58001 | 1392.1277 | -0.61516 | 1391.0981 | -0.61516 | 1394.528 | -0.63274 | 1389.2662 | -0.63274 | 1377.1618
| -0.4394 | 105.5544 | -0.54486 | 108.4364 | -0.47455 | 105.8622 | -0.47455 | 104.938 | -0.52728 | 106.9269 | -0.45698 | 1406.5222 | -0.54486 | 1391.8989 | -0.54486 | 1393.0424 | -0.58001 | 1391.6702 | -0.58001 | 1394.8708 | -0.59759 | 1390.9837 | -0.59759 | 1379.7132
| -0.40425 | 105.7391 | -0.50971 | 107.866 | -0.4394 | 106.2926 | -0.4394 | 104.6705 | -0.49213 | 106.8656 | -0.42183 | 1406.8655 | -0.50971 | 1393.1567 | -0.50971 | 1393.271 | -0.54486 | 1392.4707 | -0.54486 | 1394.8708 | -0.56244 | 1391.4414 | -0.56244 | 1381.5629
| -0.3691 | 107.5804 | -0.47455 | 107.9271 | -0.40425 | 106.395 | -0.40425 | 104.8969 | -0.45698 | 107.0086 | -0.38667 | 1406.5222 | -0.47455 | 1394.1853 | -0.47455 | 1395.2135 | -0.50971 | 1392.6994 | -0.50971 | 1394.8708 | -0.52728 | 1392.242 | -0.52728 | 1384.6745
| -0.33395 | 107.7436 | -0.4394 | 107.8252 | -0.3691 | 106.354 | -0.3691 | 104.7117 | -0.42183 | 106.8247 | -0.35152 | 1406.5222 | -0.4394 | 1394.985 | -0.4394 | 1396.013 | -0.47455 | 1393.3853 | -0.47455 | 1394.8708 | -0.49213 | 1392.3564 | -0.49213 | 1384.7896
| -0.29879 | 108.4974 | -0.40425 | 107.9271 | -0.33395 | 106.9269 | -0.33395 | 104.3616 | -0.38667 | 106.313 | -0.31637 | 1406.8655 | -0.40425 | 1395.3277 | -0.40425 | 1396.8123 | -0.4394 | 1394.4138 | -0.4394 | 1396.013 | -0.45698 | 1393.7282 | -0.45698 | 1384.7896
| -0.26364 | 108.4974 | -0.3691 | 108.192 | -0.29879 | 106.8656 | -0.29879 | 104.4646 | -0.35152 | 105.9442 | -0.28122 | 1407.5523 | -0.3691 | 1396.6981 | -0.3691 | 1396.9265 | -0.40425 | 1393.3853 | -0.40425 | 1394.6423 | -0.42183 | 1394.4138 | -0.42183 | 1385.0196
| -0.22849 | 107.5192 | -0.33395 | 108.6195 | -0.26364 | 106.8656 | -0.26364 | 104.8969 | -0.31637 | 106.2106 | -0.24607 | 1407.7814 | -0.33395 | 1397.0407 | -0.33395 | 1396.9265 | -0.3691 | 1394.8708 | -0.3691 | 1397.0407 | -0.38667 | 1395.7845 | -0.38667 | 1385.1346
| -0.19334 | 107.3763 | -0.29879 | 108.4974 | -0.22849 | 107.1108 | -0.22849 | 104.8352 | -0.28122 | 106.4359 | -0.21091 | 1408.2395 | -0.29879 | 1395.8987 | -0.29879 | 1396.2413 | -0.33395 | 1395.3277 | -0.33395 | 1397.269 | -0.35152 | 1395.7845 | -0.35152 | 1385.9392
| -0.15818 | 107.5192 | -0.26364 | 108.6602 | -0.19334 | 107.2947 | -0.19334 | 105.2875 | -0.24607 | 106.3335 | -0.17576 | 1409.7294 | -0.26364 | 1395.8987 | -0.26364 | 1393.1567 | -0.29879 | 1394.985 | -0.29879 | 1395.8987 | -0.31637 | 1397.3832 | -0.31637 | 1388.3493
| -0.12303 | 107.3763 | -0.22849 | 108.3142 | -0.15818 | 107.3967 | -0.15818 | 105.4107 | -0.21091 | 106.4564 | -0.14061 | 1410.4177 | -0.22849 | 1397.1548 | -0.22849 | 1392.0133 | -0.26364 | 1394.985 | -0.26364 | 1393.8425 | -0.28122 | 1398.0682 | -0.28122 | 1390.5259
| -0.087881 | 107.7232 | -0.19334 | 108.6195 | -0.12303 | 108.1513 | -0.12303 | 105.4518 | -0.17576 | 106.8656 | -0.10546 | 1410.8768 | -0.19334 | 1397.0407 | -0.19334 | 1392.9281 | -0.22849 | 1396.1272 | -0.22849 | 1394.1853 | -0.24607 | 1396.5839 | -0.24607 | 1388.3493
| -0.052728 | 108.355 | -0.15818 | 109.2294 | -0.087881 | 108.2124 | -0.087881 | 105.657 | -0.14061 | 106.8656 | -0.070304 | 1410.8768 | -0.15818 | 1393.3853 | -0.15818 | 1393.1567 | -0.19334 | 1394.985 | -0.19334 | 1396.013 | -0.21091 | 1394.4138 | -0.21091 | 1388.3493
| -0.017576 | 108.192 | -0.12303 | 109.7165 | -0.052728 | 109.5136 | -0.052728 | 105.8417 | -0.10546 | 107.0086 | -0.035152 | 1410.8768 | -0.12303 | 1394.985 | -0.12303 | 1398.5249 | -0.15818 | 1392.242 | -0.15818 | 1396.013 | -0.17576 | 1394.071 | -0.17576 | 1383.984
| 0.017576 | 108.4771 | -0.087881 | 110.3449 | -0.017576 | 109.9802 | -0.017576 | 106.2516 | -0.070304 | 107.1313 | 0 | 1411.336 | -0.087881 | 1395.6703 | -0.087881 | 1399.4382 | -0.12303 | 1390.9837 | -0.12303 | 1396.013 | -0.14061 | 1392.6994 | -0.14061 | 1383.984
| 0.052728 | 108.9043 | -0.052728 | 110.9116 | 0.017576 | 109.7165 | 0.017576 | 106.9269 | -0.035152 | 107.7028 | 0.035152 | 1410.9916 | -0.052728 | 1396.013 | -0.052728 | 1399.8949 | -0.087881 | 1390.6403 | -0.087881 | 1396.4697 | -0.10546 | 1390.6403 | -0.10546 | 1381.6783
| 0.087881 | 108.5381 | -0.017576 | 111.0329 | 0.052728 | 108.7212 | 0.052728 | 107.6824 | 0 | 107.2947 | 0.070304 | 1410.5324 | -0.017576 | 1396.2413 | -0.017576 | 1399.4382 | -0.052728 | 1390.9837 | -0.052728 | 1396.8123 | -0.070304 | 1390.1824 | -0.070304 | 1381.6783
| 0.12303 | 108.0902 | 0.017576 | 110.9318 | 0.087881 | 108.4771 | 0.087881 | 108.3142 | 0.035152 | 107.3355 | 0.10546 | 1409.7294 | 0.017576 | 1400.2374 | 0.017576 | 1399.8949 | -0.017576 | 1392.9281 | -0.017576 | 1397.269 | -0.035152 | 1390.6403 | -0.035152 | 1387.8906
| 0.15818 | 107.5192 | 0.052728 | 110.9116 | 0.12303 | 108.5788 | 0.12303 | 108.3957 | 0.070304 | 107.7232 | 0.14061 | 1410.4177 | 0.052728 | 1401.4935 | 0.052728 | 1400.6942 | 0.017576 | 1397.1548 | 0.017576 | 1398.2966 | 0 | 1394.985 | 0 | 1388.12
| 0.19334 | 106.9473 | 0.087881 | 111.0329 | 0.15818 | 108.4771 | 0.15818 | 108.3142 | 0.10546 | 108.1105 | 0.17576 | 1410.4177 | 0.087881 | 1403.5495 | 0.087881 | 1400.6942 | 0.052728 | 1399.3241 | 0.052728 | 1398.7532 | 0.035152 | 1396.3555 | 0.035152 | 1388.12
| 0.22849 | 106.7224 | 0.12303 | 111.1137 | 0.19334 | 108.7212 | 0.19334 | 108.355 | 0.14061 | 108.3346 | 0.21091 | 1410.4177 | 0.12303 | 1404.6924 | 0.12303 | 1401.4935 | 0.087881 | 1401.3793 | 0.087881 | 1399.6666 | 0.070304 | 1397.7257 | 0.070304 | 1393.1567
| 0.26364 | 106.5382 | 0.15818 | 111.3966 | 0.22849 | 108.9043 | 0.22849 | 108.3957 | 0.17576 | 109.2497 | 0.24607 | 1412.9451 | 0.15818 | 1406.0646 | 0.15818 | 1402.4071 | 0.12303 | 1401.4935 | 0.12303 | 1401.8361 | 0.10546 | 1399.6666 | 0.10546 | 1393.271
| 0.29879 | 105.9237 | 0.19334 | 111.9215 | 0.26364 | 109.7774 | 0.26364 | 108.2531 | 0.21091 | 109.4121 | 0.28122 | 1412.9451 | 0.19334 | 1406.179 | 0.19334 | 1404.3495 | 0.15818 | 1401.3793 | 0.15818 | 1402.2929 | 0.14061 | 1400.6942 | 0.14061 | 1393.1567
| 0.33395 | 105.616 | 0.22849 | 112.0224 | 0.29879 | 109.4121 | 0.29879 | 108.2124 | 0.24607 | 109.2497 | 0.31637 | 1411.6806 | 0.22849 | 1404.6924 | 0.22849 | 1403.4353 | 0.19334 | 1401.3793 | 0.19334 | 1402.2929 | 0.17576 | 1401.4935 | 0.17576 | 1393.271
| 0.3691 | 105.5544 | 0.26364 | 111.9417 | 0.33395 | 109.4121 | 0.33395 | 108.7619 | 0.28122 | 109.0465 | 0.35152 | 1409.6147 | 0.26364 | 1401.8361 | 0.26364 | 1401.4935 | 0.22849 | 1402.864 | 0.22849 | 1402.1787 | 0.21091 | 1401.6077 | 0.21091 | 1395.6703
| 0.40425 | 105.616 | 0.29879 | 111.5582 | 0.3691 | 109.27 | 0.3691 | 108.7823 | 0.31637 | 109.1888 | 0.38667 | 1408.9269 | 0.29879 | 1396.9265 | 0.29879 | 1396.013 | 0.26364 | 1402.5214 | 0.26364 | 1400.58 | 0.24607 | 1401.4935 | 0.24607 | 1395.6703
| 0.4394 | 105.5749 | 0.33395 | 111.437 | 0.40425 | 109.27 | 0.40425 | 109.0059 | 0.35152 | 109.1481 | 0.42183 | 1407.4379 | 0.33395 | 1394.6423 | 0.33395 | 1395.2135 | 0.29879 | 1399.7808 | 0.29879 | 1398.5249 | 0.28122 | 1400.58 | 0.28122 | 1392.6994
| 0.47455 | 105.616 | 0.3691 | 110.7295 | 0.4394 | 108.9043 | 0.4394 | 108.7823 | 0.38667 | 108.9856 | 0.45698 | 1406.4078 | 0.3691 | 1393.3853 | 0.3691 | 1394.4138 | 0.33395 | 1396.013 | 0.33395 | 1397.0407 | 0.31637 | 1398.8674 | 0.31637 | 1392.1277
| 0.50971 | 106.0262 | 0.40425 | 110.4866 | 0.47455 | 108.3346 | 0.47455 | 108.4974 | 0.42183 | 109.473 | 0.49213 | 1403.4353 | 0.40425 | 1393.271 | 0.40425 | 1393.7282 | 0.3691 | 1393.9567 | 0.3691 | 1395.6703 | 0.35152 | 1396.013 | 0.35152 | 1390.9837
| 0.54486 | 106.2926 | 0.4394 | 110.4866 | 0.50971 | 107.866 | 0.50971 | 107.9679 | 0.45698 | 109.6557 | 0.52728 | 1402.1787 | 0.4394 | 1392.0133 | 0.4394 | 1393.1567 | 0.40425 | 1392.4707 | 0.40425 | 1394.528 | 0.38667 | 1393.1567 | 0.38667 | 1388.5786
| 0.58001 | 106.5996 | 0.47455 | 109.818 | 0.54486 | 107.4376 | 0.54486 | 107.7028 | 0.49213 | 109.676 | 0.56244 | 1401.4935 | 0.47455 | 1391.8989 | 0.47455 | 1392.8138 | 0.4394 | 1390.4114 | 0.4394 | 1393.271 | 0.42183 | 1393.1567 | 0.42183 | 1387.2023
| 0.61516 | 107.4784 | 0.50971 | 109.7571 | 0.58001 | 106.9882 | 0.58001 | 107.5396 | 0.52728 | 109.9802 | 0.59759 | 1401.4935 | 0.50971 | 1391.8989 | 0.50971 | 1389.9534 | 0.47455 | 1390.4114 | 0.47455 | 1391.7846 | 0.45698 | 1392.4707 | 0.45698 | 1384.6745
| 0.65032 | 107.56 | 0.54486 | 109.676 | 0.61516 | 106.5792 | 0.61516 | 106.5792 | 0.56244 | 109.7368 | 0.63274 | 1401.4935 | 0.54486 | 1390.9837 | 0.54486 | 1388.5786 | 0.50971 | 1389.8389 | 0.50971 | 1391.6702 | 0.49213 | 1392.1277 | 0.49213 | 1384.0992
| 0.68547 | 108.1105 | 0.58001 | 109.3512 | 0.65032 | 106.5792 | 0.65032 | 105.9852 | 0.59759 | 109.1278 | 0.66789 | 1401.4935 | 0.58001 | 1389.2662 | 0.58001 | 1387.0875 | 0.54486 | 1389.037 | 0.54486 | 1391.6702 | 0.52728 | 1391.7846 | 0.52728 | 1383.8689
| 0.72062 | 107.56 | 0.61516 | 108.6602 | 0.68547 | 106.5792 | 0.68547 | 105.8417 | 0.63274 | 108.6602 | 0.70304 | 1399.3241 | 0.61516 | 1389.2662 | 0.61516 | 1389.1516 | 0.58001 | 1388.3493 | 0.58001 | 1391.6702 | 0.56244 | 1389.037 | 0.56244 | 1383.8689
| 0.75577 | 107.213 | 0.65032 | 108.5178 | 0.72062 | 106.5792 | 0.72062 | 105.5749 | 0.66789 | 108.3753 | 0.7382 | 1400.8083 | 0.65032 | 1391.327 | 0.65032 | 1389.9534 | 0.61516 | 1387.5465 | 0.61516 | 1391.5558 | 0.59759 | 1388.0053 | 0.59759 | 1384.0992
| 0.79092 | 106.661 | 0.68547 | 108.4771 | 0.75577 | 107.3967 | 0.75577 | 105.0408 | 0.70304 | 108.3346 | 0.77335 | 1400.8083 | 0.68547 | 1390.5259 | 0.68547 | 1390.0679 | 0.65032 | 1388.3493 | 0.65032 | 1391.8989 | 0.63274 | 1385.4795 | 0.63274 | 1382.4861
| 0.82608 | 105.5955 | 0.72062 | 108.5381 | 0.79092 | 107.3967 | 0.79092 | 105.5749 | 0.7382 | 107.7232 | 0.8085 | 1400.2374 | 0.72062 | 1391.5558 | 0.72062 | 1390.0679 | 0.68547 | 1388.5786 | 0.68547 | 1389.1516 | 0.66789 | 1385.4795 | 0.66789 | 1382.4861
| 0.86123 | 105.5955 | 0.75577 | 108.6602 | 0.82608 | 107.3967 | 0.82608 | 105.3285 | 0.77335 | 107.4376 | 0.84365 | 1400.0091 | 0.75577 | 1392.5851 | 0.75577 | 1390.0679 | 0.72062 | 1388.5786 | 0.72062 | 1388.8078 | 0.70304 | 1386.5135 | 0.70304 | 1382.4861
| 0.89638 | 105.5955 | 0.79092 | 108.7619 | 0.86123 | 107.8456 | 0.86123 | 105.2875 | 0.8085 | 107.0086 | 0.87881 | 1401.9503 | 0.79092 | 1391.5558 | 0.79092 | 1390.0679 | 0.75577 | 1391.4414 | 0.75577 | 1388.3493 | 0.7382 | 1387.5465 | 0.7382 | 1383.4082
| 0.93153 | 106.661 | 0.82608 | 109.3309 | 0.89638 | 108.5585 | 0.89638 | 105.7801 | 0.84365 | 106.5996 | 0.91396 | 1400.8083 | 0.82608 | 1391.327 | 0.82608 | 1386.8579 | 0.79092 | 1392.1277 | 0.79092 | 1388.0053 | 0.77335 | 1388.4639 | 0.77335 | 1384.6745
| 0.96669 | 107.0904 | 0.86123 | 109.9193 | 0.93153 | 108.029 | 0.93153 | 106.0467 | 0.87881 | 106.5996 | 0.94911 | 1400.2374 | 0.86123 | 1391.2125 | 0.86123 | 1387.7759 | 0.82608 | 1390.8692 | 0.82608 | 1387.5465 | 0.8085 | 1389.8389 | 0.8085 | 1384.6745
| 1.0018 | 107.2334 | 0.89638 | 110.7902 | 0.96669 | 108.5788 | 0.96669 | 105.9647 | 0.91396 | 106.9269 | 0.98426 | 1397.7257 | 0.89638 | 1387.317 | 0.89638 | 1386.6283 | 0.86123 | 1390.4114 | 0.86123 | 1385.9392 | 0.84365 | 1389.8389 | 0.84365 | 1383.4082
| 1.037 | 107.3763 | 0.93153 | 110.7902 | 1.0018 | 108.029 | 1.0018 | 106.4973 | 0.94911 | 106.5178 | 1.0194 | 1396.9265 | 0.93153 | 1385.5944 | 0.93153 | 1383.8689 | 0.89638 | 1388.2346 | 0.89638 | 1385.9392 | 0.87881 | 1388.12 | 0.87881 | 1383.0625
| 1.0721 | 107.3151 | 0.96669 | 110.2639 | 1.037 | 107.5396 | 1.037 | 107.2538 | 0.98426 | 107.0086 | 1.0546 | 1395.0992 | 0.96669 | 1385.9392 | 0.96669 | 1383.8689 | 0.93153 | 1384.0992 | 0.93153 | 1384.4444 | 0.91396 | 1386.9727 | 0.91396 | 1380.2917
| 1.1073 | 107.3151 | 1.0018 | 109.676 | 1.0721 | 107.0904 | 1.0721 | 107.6008 | 1.0194 | 107.2742 | 1.0897 | 1396.9265 | 1.0018 | 1385.3645 | 1.0018 | 1384.3294 | 0.96669 | 1383.8689 | 0.96669 | 1383.8689 | 0.94911 | 1385.2496 | 0.94911 | 1375.7658
| 1.1424 | 107.3967 | 1.037 | 109.27 | 1.1073 | 107.0495 | 1.1073 | 107.0086 | 1.0546 | 107.9271 | 1.1249 | 1398.8674 | 1.037 | 1386.5135 | 1.037 | 1384.3294 | 1.0018 | 1382.2554 | 1.0018 | 1383.4082 | 0.98426 | 1381.2164 | 0.98426 | 1375.5328
| 1.1776 | 107.3151 | 1.0721 | 108.9652 | 1.1424 | 107.0904 | 1.1424 | 106.8451 | 1.0897 | 109.1278 | 1.16 | 1399.2099 | 1.0721 | 1388.0053 | 1.0721 | 1385.2496 | 1.037 | 1382.2554 | 1.037 | 1382.7167 | 1.0194 | 1381.2164 | 1.0194 | 1372.3782
| 1.2128 | 107.1517 | 1.1073 | 108.6399 | 1.1776 | 107.5396 | 1.1776 | 106.4564 | 1.1249 | 109.1278 | 1.1952 | 1399.5524 | 1.1073 | 1388.2346 | 1.1073 | 1385.7094 | 1.0721 | 1383.8689 | 1.0721 | 1382.6014 | 1.0546 | 1382.4861 | 1.0546 | 1373.7825
| 1.2479 | 107.56 | 1.1424 | 108.6399 | 1.2128 | 108.0698 | 1.2128 | 105.1231 | 1.16 | 109.0669 | 1.2303 | 1399.5524 | 1.1424 | 1389.037 | 1.1424 | 1386.3987 | 1.1073 | 1384.9046 | 1.1073 | 1383.6386 | 1.0897 | 1383.6386 | 1.0897 | 1375.7658
| 1.2831 | 107.9271 | 1.1776 | 109.0872 | 1.2479 | 108.3346 | 1.2479 | 105.0408 | 1.1952 | 108.4974 | 1.2655 | 1399.7808 | 1.1776 | 1389.2662 | 1.1776 | 1386.0541 | 1.1424 | 1387.2023 | 1.1424 | 1384.9046 | 1.1249 | 1384.5595 | 1.1249 | 1378.0907
| 1.3182 | 109.818 | 1.2128 | 109.6963 | 1.2831 | 108.7212 | 1.2831 | 105.0408 | 1.2303 | 108.2124 | 1.3006 | 1400.9225 | 1.2128 | 1389.1516 | 1.2128 | 1385.3645 | 1.1776 | 1388.2346 | 1.1776 | 1386.3987 | 1.16 | 1384.6745 | 1.16 | 1380.4074
| 1.3534 | 110.4056 | 1.2479 | 110.0815 | 1.3182 | 108.7212 | 1.3182 | 105.7596 | 1.2655 | 107.6008 | 1.3358 | 1403.4353 | 1.2479 | 1388.0053 | 1.2479 | 1388.2346 | 1.2128 | 1388.8078 | 1.2128 | 1388.5786 | 1.1952 | 1385.2496 | 1.1952 | 1380.7542
| 1.3885 | 110.7093 | 1.2831 | 110.1828 | 1.3534 | 108.7212 | 1.3534 | 106.1492 | 1.3006 | 106.9882 | 1.3709 | 1403.6638 | 1.2831 | 1390.4114 | 1.2831 | 1390.0679 | 1.2479 | 1389.9534 | 1.2479 | 1387.4318 | 1.2303 | 1385.2496 | 1.2303 | 1380.6386
| 1.4237 | 111.0733 | 1.3182 | 110.0815 | 1.3885 | 108.7212 | 1.3885 | 107.2947 | 1.3358 | 106.7633 | 1.4061 | 1404.3495 | 1.3182 | 1391.4414 | 1.3182 | 1390.5259 | 1.2831 | 1390.4114 | 1.2831 | 1389.3807 | 1.2655 | 1385.5944 | 1.2655 | 1380.0604
| 1.4588 | 110.9722 | 1.3534 | 110.2841 | 1.4237 | 108.7212 | 1.4237 | 107.3559 | 1.3709 | 106.9882 | 1.4412 | 1401.2651 | 1.3534 | 1391.4414 | 1.3534 | 1391.4414 | 1.3182 | 1390.7548 | 1.3182 | 1390.7548 | 1.3006 | 1386.6283 | 1.3006 | 1379.8289
| 1.494 | 111.0733 | 1.3885 | 110.2841 | 1.4588 | 108.9652 | 1.4588 | 107.3559 | 1.4061 | 107.1721 | 1.4764 | 1401.2651 | 1.3885 | 1393.3853 | 1.3885 | 1391.5558 | 1.3534 | 1390.9837 | 1.3534 | 1393.8425 | 1.3358 | 1388.3493 | 1.3358 | 1379.8289
| 1.5291 | 110.9318 | 1.4237 | 110.9116 | 1.494 | 109.7165 | 1.494 | 107.3559 | 1.4412 | 108.4771 | 1.5115 | 1404.3495 | 1.4237 | 1394.7565 | 1.4237 | 1392.1277 | 1.3885 | 1392.1277 | 1.3885 | 1394.4138 | 1.3709 | 1390.2969 | 1.3709 | 1379.8289
| 1.5643 | 110.9318 | 1.4588 | 111.538 | 1.5291 | 110.77 | 1.5291 | 107.458 | 1.4764 | 108.6602 | 1.5467 | 1405.6071 | 1.4588 | 1394.7565 | 1.4588 | 1393.6139 | 1.4237 | 1391.5558 | 1.4237 | 1394.528 | 1.4061 | 1389.4953 | 1.4061 | 1382.7167
| 1.5994 | 110.7902 | 1.494 | 111.7399 | 1.5643 | 110.77 | 1.5643 | 107.213 | 1.5115 | 108.7619 | 1.5818 | 1406.0646 | 1.494 | 1394.7565 | 1.494 | 1393.9567 | 1.4588 | 1392.1277 | 1.4588 | 1392.8138 | 1.4412 | 1388.9224 | 1.4412 | 1385.2496
| 1.6346 | 110.7902 | 1.5291 | 112.1837 | 1.5994 | 111.639 | 1.5994 | 107.458 | 1.5467 | 108.7619 | 1.617 | 1406.0646 | 1.5291 | 1396.3555 | 1.5291 | 1394.7565 | 1.494 | 1392.0133 | 1.494 | 1394.4138 | 1.4764 | 1390.4114 | 1.4764 | 1385.2496
| 1.6697 | 110.4866 | 1.5643 | 113.0297 | 1.6346 | 112.224 | 1.6346 | 107.9068 | 1.5818 | 108.8839 | 1.6522 | 1406.2934 | 1.5643 | 1396.3555 | 1.5643 | 1398.2966 | 1.5291 | 1392.0133 | 1.5291 | 1396.5839 | 1.5115 | 1389.8389 | 1.5115 | 1385.1346
| 1.7049 | 110.9116 | 1.5994 | 113.5928 | 1.6697 | 112.2643 | 1.6697 | 108.3346 | 1.617 | 108.8026 | 1.6873 | 1405.264 | 1.5994 | 1396.3555 | 1.5994 | 1401.1509 | 1.5643 | 1392.4707 | 1.5643 | 1397.4974 | 1.5467 | 1388.9224 | 1.5467 | 1385.0196
| 1.74 | 111.2754 | 1.6346 | 114.7971 | 1.7049 | 111.9618 | 1.7049 | 109.6963 | 1.6522 | 108.8026 | 1.7225 | 1402.6356 | 1.6346 | 1396.2413 | 1.6346 | 1400.8083 | 1.5994 | 1394.6423 | 1.5994 | 1397.9541 | 1.5818 | 1391.5558 | 1.5818 | 1385.0196
| 1.7752 | 111.6995 | 1.6697 | 115.0176 | 1.74 | 111.9618 | 1.74 | 110.3044 | 1.6873 | 108.9856 | 1.7576 | 1402.9783 | 1.6697 | 1395.3277 | 1.6697 | 1398.7532 | 1.6346 | 1396.2413 | 1.6346 | 1397.9541 | 1.617 | 1393.0424 | 1.617 | 1385.2496
| 1.8103 | 112.224 | 1.7049 | 115.0176 | 1.7752 | 111.8004 | 1.7752 | 111.3562 | 1.7225 | 109.6354 | 1.7928 | 1406.2934 | 1.7049 | 1391.8989 | 1.7049 | 1397.6116 | 1.6697 | 1398.7532 | 1.6697 | 1397.9541 | 1.6522 | 1393.1567 | 1.6522 | 1386.8579
| 1.8455 | 112.224 | 1.74 | 114.6568 | 1.8103 | 111.7197 | 1.8103 | 111.5178 | 1.7576 | 109.9599 | 1.8279 | 1407.8959 | 1.74 | 1389.7244 | 1.74 | 1394.6423 | 1.7049 | 1398.2966 | 1.7049 | 1394.071 | 1.6873 | 1394.4138 | 1.6873 | 1387.317
| 1.8806 | 112.2442 | 1.7752 | 114.2154 | 1.8455 | 111.7197 | 1.8455 | 111.3562 | 1.7928 | 111.0127 | 1.8631 | 1410.1882 | 1.7752 | 1389.1516 | 1.7752 | 1391.7846 | 1.74 | 1398.2966 | 1.74 | 1393.6139 | 1.7225 | 1396.5839 | 1.7225 | 1387.317
| 1.9158 | 112.5465 | 1.8103 | 114.115 | 1.8806 | 111.7197 | 1.8806 | 110.9924 | 1.8279 | 111.6188 | 1.8982 | 1410.1882 | 1.8103 | 1387.2023 | 1.8103 | 1390.7548 | 1.7752 | 1398.7532 | 1.7752 | 1394.985 | 1.7576 | 1394.4138 | 1.7576 | 1388.12
| 1.9509 | 112.6472 | 1.8455 | 113.633 | 1.9158 | 111.7601 | 1.9158 | 110.3651 | 1.8631 | 112.0224 | 1.9334 | 1410.1882 | 1.8455 | 1387.6612 | 1.8455 | 1391.4414 | 1.8103 | 1400.3516 | 1.8103 | 1398.4107 | 1.7928 | 1394.4138 | 1.7928 | 1388.12
| 1.9861 | 113.1303 | 1.8806 | 113.3716 | 1.9509 | 112.0627 | 1.9509 | 110.0815 | 1.8982 | 112.9089 | 1.9685 | 1410.1882 | 1.8806 | 1393.271 | 1.8806 | 1394.985 | 1.8455 | 1401.3793 | 1.8455 | 1398.9816 | 1.8279 | 1395.7845 | 1.8279 | 1388.8078
| 2.0213 | 113.2309 | 1.9158 | 113.1303 | 1.9861 | 112.1837 | 1.9861 | 110.0004 | 1.9334 | 112.9089 | 2.0037 | 1410.1882 | 1.9158 | 1402.864 | 1.9158 | 1397.3832 | 1.8806 | 1402.0645 | 1.8806 | 1399.8949 | 1.8631 | 1395.8987 | 1.8631 | 1389.3807
| 2.0564 | 113.2107 | 1.9509 | 113.4923 | 2.0213 | 112.4659 | 2.0213 | 109.8788 | 1.9685 | 112.9089 | 2.0388 | 1409.9588 | 1.9509 | 1407.2089 | 1.9509 | 1397.3832 | 1.9158 | 1402.0645 | 1.9158 | 1399.8949 | 1.8982 | 1395.8987 | 1.8982 | 1389.7244
| 2.0916 | 113.2309 | 1.9861 | 114.0347 | 2.0564 | 113.2912 | 2.0564 | 110.3854 | 2.0037 | 112.2442 | 2.074 | 1410.1882 | 1.9861 | 1407.2089 | 1.9861 | 1403.7781 | 1.9509 | 1399.8949 | 1.9509 | 1400.6942 | 1.9334 | 1396.013 | 1.9334 | 1391.327
| 2.1267 | 113.2107 | 2.0213 | 114.2556 | 2.0916 | 113.5526 | 2.0916 | 110.5069 | 2.0388 | 112.0627 | 2.1091 | 1411.4509 | 2.0213 | 1407.2089 | 2.0213 | 1404.2352 | 1.9861 | 1399.0957 | 1.9861 | 1400.6942 | 1.9685 | 1396.9265 | 1.9685 | 1392.1277
| 2.1619 | 113.0699 | 2.0564 | 115.3982 | 2.1267 | 113.6933 | 2.1267 | 110.5069 | 2.074 | 111.5178 | 2.1443 | 1412.6 | 2.0564 | 1404.4638 | 2.0564 | 1401.7219 | 2.0213 | 1399.0957 | 2.0213 | 1399.4382 | 2.0037 | 1397.1548 | 2.0037 | 1392.1277
| 2.197 | 113.1303 | 2.0916 | 115.7585 | 2.1619 | 114.1753 | 2.1619 | 110.7093 | 2.1091 | 111.1137 | 2.1794 | 1414.2112 | 2.0916 | 1400.9225 | 2.0916 | 1395.2135 | 2.0564 | 1397.7257 | 2.0564 | 1399.0957 | 2.0388 | 1397.1548 | 2.0388 | 1392.242
| 2.2322 | 113.1303 | 2.1267 | 115.7585 | 2.197 | 114.3358 | 2.197 | 110.8711 | 2.1443 | 111.1137 | 2.2146 | 1414.4416 | 2.1267 | 1400.58 | 2.1267 | 1399.4382 | 2.0916 | 1396.1272 | 2.0916 | 1399.2099 | 2.074 | 1397.1548 | 2.074 | 1392.242
| 2.2673 | 112.5666 | 2.1619 | 115.7585 | 2.2322 | 114.3559 | 2.2322 | 110.8711 | 2.1794 | 111.982 | 2.2497 | 1415.2485 | 2.1619 | 1400.9225 | 2.1619 | 1401.7219 | 2.1267 | 1401.7219 | 2.1267 | 1399.7808 | 2.1091 | 1397.7257 | 2.1091 | 1392.242
| 2.3025 | 112.4256 | 2.197 | 116.5584 | 2.2673 | 114.5164 | 2.2673 | 110.9924 | 2.2146 | 113.0095 | 2.2849 | 1418.1371 | 2.197 | 1403.2068 | 2.197 | 1401.7219 | 2.1619 | 1404.1209 | 2.1619 | 1400.2374 | 2.1443 | 1397.7257 | 2.1443 | 1393.9567
| 2.3376 | 112.5062 | 2.2322 | 116.2986 | 2.3025 | 114.376 | 2.3025 | 111.1744 | 2.2497 | 113.1102 | 2.32 | 1418.1371 | 2.2322 | 1404.6924 | 2.2322 | 1403.4353 | 2.197 | 1401.2651 | 2.197 | 1401.8361 | 2.1794 | 1398.0682 | 2.1794 | 1394.4138
| 2.3728 | 112.2038 | 2.2673 | 116.3585 | 2.3376 | 114.376 | 2.3376 | 111.5582 | 2.2849 | 113.1303 | 2.3552 | 1417.6741 | 2.2673 | 1408.5832 | 2.2673 | 1404.8067 | 2.2322 | 1404.1209 | 2.2322 | 1402.864 | 2.2146 | 1399.5524 | 2.2146 | 1395.5561
| 2.4079 | 111.8811 | 2.3025 | 116.3585 | 2.3728 | 114.2556 | 2.3728 | 112.3853 | 2.32 | 113.1504 | 2.3903 | 1418.2529 | 2.3025 | 1411.336 | 2.3025 | 1405.264 | 2.2673 | 1405.0354 | 2.2673 | 1405.0354 | 2.2497 | 1400.58 | 2.2497 | 1396.8123
| 2.4431 | 112.6673 | 2.3376 | 116.5584 | 2.4079 | 114.1351 | 2.4079 | 112.6069 | 2.3552 | 113.5124 | 2.4255 | 1418.3687 | 2.3376 | 1412.1402 | 2.3376 | 1407.2089 | 2.3025 | 1405.0354 | 2.3025 | 1407.0944 | 2.2849 | 1401.4935 | 2.2849 | 1397.6116
| 2.4782 | 112.6673 | 2.3728 | 116.2986 | 2.4431 | 113.6129 | 2.4431 | 112.6069 | 2.3903 | 113.5124 | 2.4607 | 1418.3687 | 2.3728 | 1414.9026 | 2.3728 | 1406.7511 | 2.3376 | 1406.0646 | 2.3376 | 1408.354 | 2.32 | 1404.1209 | 2.32 | 1400.9225
| 2.5134 | 112.6673 | 2.4079 | 116.2386 | 2.4782 | 113.8741 | 2.4782 | 112.8083 | 2.4255 | 113.8339 | 2.4958 | 1416.6336 | 2.4079 | 1414.9026 | 2.4079 | 1406.7511 | 2.3728 | 1410.762 | 2.3728 | 1408.1249 | 2.3552 | 1405.0354 | 2.3552 | 1401.1509
| 2.5485 | 112.6673 | 2.4431 | 116.1386 | 2.5134 | 113.4923 | 2.5134 | 112.8083 | 2.4607 | 114.1351 | 2.531 | 1418.832 | 2.4431 | 1414.9026 | 2.4431 | 1406.5222 | 2.4079 | 1409.7294 | 2.4079 | 1410.0735 | 2.3903 | 1405.3784 | 2.3903 | 1402.6356
| 2.5837 | 112.6673 | 2.4782 | 115.7185 | 2.5485 | 112.7277 | 2.5485 | 112.6472 | 2.4958 | 114.1351 | 2.5661 | 1418.3687 | 2.4782 | 1412.1402 | 2.4782 | 1405.4927 | 2.4431 | 1410.0735 | 2.4431 | 1408.1249 | 2.4255 | 1409.8441 | 2.4255 | 1403.0925
| 2.6188 | 112.224 | 2.5134 | 115.258 | 2.5837 | 112.929 | 2.5837 | 112.1837 | 2.531 | 114.0147 | 2.6013 | 1420.9213 | 2.5134 | 1408.5832 | 2.5134 | 1404.5781 | 2.4782 | 1410.0735 | 2.4782 | 1406.179 | 2.4607 | 1409.8441 | 2.4607 | 1402.9783
| 2.654 | 112.103 | 2.5485 | 115.0377 | 2.6188 | 112.8888 | 2.6188 | 112.0627 | 2.5661 | 114.3358 | 2.6364 | 1423.3677 | 2.5485 | 1409.8441 | 2.5485 | 1406.4078 | 2.5134 | 1408.4686 | 2.5134 | 1407.4379 | 2.4958 | 1407.4379 | 2.4958 | 1403.6638
| 2.6891 | 111.5582 | 2.5837 | 115.1579 | 2.654 | 112.5666 | 2.654 | 111.7197 | 2.6013 | 114.3358 | 2.6716 | 1423.7181 | 2.5837 | 1412.0253 | 2.5837 | 1407.5523 | 2.5485 | 1406.8655 | 2.5485 | 1407.4379 | 2.531 | 1407.5523 | 2.531 | 1403.6638
| 2.7243 | 111.4774 | 2.6188 | 115.258 | 2.6891 | 112.7881 | 2.6891 | 111.7197 | 2.6364 | 114.0347 | 2.7067 | 1423.7181 | 2.6188 | 1412.0253 | 2.6188 | 1408.354 | 2.5837 | 1408.5832 | 2.5837 | 1408.9269 | 2.5661 | 1407.7814 | 2.5661 | 1403.0925
| 2.7594 | 111.3966 | 2.654 | 115.4383 | 2.7243 | 112.9089 | 2.7243 | 111.7197 | 2.6716 | 114.0347 | 2.7419 | 1424.0686 | 2.654 | 1411.7955 | 2.654 | 1408.354 | 2.6188 | 1409.8441 | 2.6188 | 1410.762 | 2.6013 | 1407.5523 | 2.6013 | 1401.6077
| 2.7946 | 111.336 | 2.6891 | 115.4583 | 2.7594 | 112.9089 | 2.7594 | 112.1434 | 2.7067 | 114.2355 | 2.777 | 1423.1343 | 2.6891 | 1412.0253 | 2.6891 | 1408.2395 | 2.654 | 1407.7814 | 2.654 | 1408.9269 | 2.6364 | 1409.2708 | 2.6364 | 1401.6077
| 2.8298 | 111.3966 | 2.7243 | 115.3382 | 2.7946 | 112.9089 | 2.7946 | 112.2442 | 2.7419 | 113.7334 | 2.8122 | 1423.0176 | 2.7243 | 1412.0253 | 2.7243 | 1407.3234 | 2.6891 | 1408.5832 | 2.6891 | 1411.7955 | 2.6716 | 1409.1562 | 2.6716 | 1399.4382
| 2.8649 | 111.5784 | 2.7594 | 115.258 | 2.8298 | 112.9089 | 2.8298 | 112.3651 | 2.777 | 113.5928 | 2.8473 | 1423.0176 | 2.7594 | 1412.6 | 2.7594 | 1407.3234 | 2.7243 | 1412.0253 | 2.7243 | 1413.0601 | 2.7067 | 1407.7814 | 2.7067 | 1399.0957
| 2.9001 | 111.134 | 2.7946 | 115.0978 | 2.8649 | 112.5062 | 2.8649 | 112.3046 | 2.8122 | 113.5928 | 2.8825 | 1423.4845 | 2.7946 | 1413.0601 | 2.7946 | 1407.6668 | 2.7594 | 1410.3029 | 2.7594 | 1410.762 | 2.7419 | 1411.2212 | 2.7419 | 1399.4382
| 2.9352 | 111.134 | 2.8298 | 114.8773 | 2.9001 | 112.3651 | 2.9001 | 112.3046 | 2.8473 | 114.0347 | 2.9176 | 1424.3025 | 2.8298 | 1415.1332 | 2.8298 | 1408.2395 | 2.7946 | 1410.3029 | 2.7946 | 1410.762 | 2.777 | 1410.6472 | 2.777 | 1400.2374
| 2.9704 | 111.982 | 2.8649 | 114.737 | 2.9352 | 112.1837 | 2.9352 | 112.2442 | 2.8825 | 113.9946 | 2.9528 | 1428.1759 | 2.8649 | 1415.1332 | 2.8649 | 1413.5204 | 2.8298 | 1412.0253 | 2.8298 | 1410.8768 | 2.8122 | 1404.3495 | 2.8122 | 1400.8083
| 3.0055 | 112.1434 | 2.9001 | 113.9544 | 2.9704 | 112.1837 | 2.9704 | 112.1232 | 2.9176 | 114.0749 | 2.9879 | 1428.1759 | 2.9001 | 1416.2871 | 2.9001 | 1414.9026 | 2.8649 | 1413.1751 | 2.8649 | 1410.762 | 2.8473 | 1410.6472 | 2.8473 | 1401.2651
| 3.0407 | 112.2442 | 2.9352 | 113.9544 | 3.0055 | 112.345 | 3.0055 | 111.4572 | 2.9528 | 114.0749 | 3.0231 | 1428.1759 | 2.9352 | 1416.7492 | 2.9352 | 1417.5585 | 2.9001 | 1411.7955 | 2.9001 | 1410.8768 | 2.8825 | 1416.5181 | 2.8825 | 1401.7219
| 3.0758 | 112.2845 | 2.9704 | 113.5928 | 3.0407 | 112.345 | 3.0407 | 110.9722 | 2.9879 | 113.854 | 3.0582 | 1426.764 | 2.9704 | 1416.8647 | 2.9704 | 1416.0562 | 2.9352 | 1411.7955 | 2.9352 | 1419.8758 | 2.9176 | 1416.5181 | 2.9176 | 1401.9503
| 3.111 | 112.8485 | 3.0055 | 113.9544 | 3.0758 | 112.4659 | 3.0758 | 110.9722 | 3.0231 | 113.6531 | 3.0934 | 1425.7076 | 3.0055 | 1417.0959 | 3.0055 | 1415.0179 | 2.9704 | 1412.6 | 2.9704 | 1425.7076 | 2.9528 | 1418.4845 | 2.9528 | 1401.9503
| 3.1461 | 112.1434 | 3.0407 | 114.396 | 3.111 | 112.5465 | 3.111 | 110.8913 | 3.0582 | 112.8284 | 3.1285 | 1423.7181 | 3.0407 | 1417.0959 | 3.0407 | 1414.7873 | 3.0055 | 1412.0253 | 3.0055 | 1418.9479 | 2.9879 | 1422.6677 | 2.9879 | 1403.7781
| 3.1813 | 111.8408 | 3.0758 | 114.4362 | 3.1461 | 113.1504 | 3.1461 | 110.8913 | 3.0934 | 112.4256 | 3.1637 | 1422.4346 | 3.0758 | 1417.0959 | 3.0758 | 1413.2902 | 3.0407 | 1412.0253 | 3.0407 | 1412.1402 | 3.0231 | 1419.5277 | 3.0231 | 1404.921
| 3.2164 | 111.336 | 3.111 | 114.4362 | 3.1813 | 113.0297 | 3.1813 | 110.952 | 3.1285 | 112.2643 | 3.1989 | 1417.9056 | 3.111 | 1416.4026 | 3.111 | 1412.83 | 3.0758 | 1411.7955 | 3.0758 | 1410.6472 | 3.0582 | 1417.9056 | 3.0582 | 1405.7215
| 3.2516 | 111.4168 | 3.1461 | 115.0577 | 3.2164 | 113.3716 | 3.2164 | 111.4572 | 3.1637 | 112.0627 | 3.234 | 1419.5277 | 3.1461 | 1414.9026 | 3.1461 | 1411.5657 | 3.111 | 1412.6 | 3.111 | 1409.2708 | 3.0934 | 1417.9056 | 3.0934 | 1405.1497
| 3.2867 | 111.336 | 3.1813 | 115.2981 | 3.2516 | 114.0749 | 3.2516 | 111.4572 | 3.1989 | 112.224 | 3.2692 | 1417.3272 | 3.1813 | 1411.7955 | 3.1813 | 1409.5001 | 3.1461 | 1411.5657 | 3.1461 | 1407.7814 | 3.1285 | 1421.7357 | 3.1285 | 1405.7215
| 3.3219 | 111.8408 | 3.2164 | 115.4383 | 3.2867 | 113.7736 | 3.2867 | 111.4572 | 3.234 | 112.2643 | 3.3043 | 1415.5946 | 3.2164 | 1411.336 | 3.2164 | 1409.2708 | 3.1813 | 1411.5657 | 3.1813 | 1409.2708 | 3.1637 | 1419.5277 | 3.1637 | 1404.5781
| 3.357 | 112.5263 | 3.2516 | 115.5984 | 3.3219 | 113.633 | 3.3219 | 111.5986 | 3.2692 | 112.5666 | 3.3395 | 1413.5204 | 3.2516 | 1409.0415 | 3.2516 | 1408.1249 | 3.2164 | 1410.9916 | 3.2164 | 1409.2708 | 3.1989 | 1414.9026 | 3.1989 | 1402.4071
| 3.3922 | 112.5263 | 3.2867 | 115.5984 | 3.357 | 113.5325 | 3.357 | 111.6995 | 3.3043 | 112.627 | 3.3746 | 1409.7294 | 3.2867 | 1408.2395 | 3.2867 | 1407.4379 | 3.2516 | 1408.5832 | 3.2516 | 1405.7215 | 3.234 | 1414.5568 | 3.234 | 1400.2374
| 3.4273 | 111.6591 | 3.3219 | 115.5184 | 3.3922 | 112.9693 | 3.3922 | 112.3853 | 3.3395 | 112.7076 | 3.4098 | 1407.3234 | 3.3219 | 1404.4638 | 3.3219 | 1405.4927 | 3.2867 | 1405.9502 | 3.2867 | 1404.4638 | 3.2692 | 1408.2395 | 3.2692 | 1398.9816
| 3.4625 | 112.2643 | 3.357 | 115.1979 | 3.4273 | 112.486 | 3.4273 | 112.4659 | 3.3746 | 113.0095 | 3.4449 | 1407.2089 | 3.357 | 1400.58 | 3.357 | 1396.6981 | 3.3219 | 1402.1787 | 3.3219 | 1403.2068 | 3.3043 | 1402.7498 | 3.3043 | 1396.8123
| 3.4976 | 111.8004 | 3.3922 | 114.6167 | 3.4625 | 112.4659 | 3.4625 | 112.3853 | 3.4098 | 112.627 | 3.4801 | 1407.2089 | 3.3922 | 1397.4974 | 3.3922 | 1396.3555 | 3.357 | 1399.7808 | 3.357 | 1399.0957 | 3.3395 | 1402.5214 | 3.3395 | 1392.9281
| 3.5328 | 111.6591 | 3.4273 | 114.4763 | 3.4976 | 112.4054 | 3.4976 | 112.4659 | 3.4449 | 113.633 | 3.5152 | 1406.7511 | 3.4273 | 1396.6981 | 3.4273 | 1395.6703 | 3.3922 | 1397.6116 | 3.3922 | 1397.0407 | 3.3746 | 1398.0682 | 3.3746 | 1391.6702
| 3.5679 | 111.8811 | 3.4625 | 114.6167 | 3.5328 | 112.2845 | 3.5328 | 112.345 | 3.4801 | 113.6933 | 3.5504 | 1406.8655 | 3.4625 | 1396.6981 | 3.4625 | 1395.6703 | 3.4273 | 1396.1272 | 3.4273 | 1394.7565 | 3.4098 | 1397.0407 | 3.4098 | 1391.6702
| 3.6031 | 111.8811 | 3.4976 | 114.6167 | 3.5679 | 112.2643 | 3.5679 | 111.3764 | 3.5152 | 113.6129 | 3.5855 | 1406.7511 | 3.4976 | 1395.8987 | 3.4976 | 1395.8987 | 3.4625 | 1395.2135 | 3.4625 | 1394.2995 | 3.4449 | 1395.8987 | 3.4449 | 1390.4114
| 3.6383 | 111.5784 | 3.5328 | 114.4963 | 3.6031 | 112.2845 | 3.6031 | 110.8104 | 3.5504 | 113.251 | 3.6207 | 1404.8067 | 3.5328 | 1396.9265 | 3.5328 | 1396.2413 | 3.4976 | 1392.6994 | 3.4976 | 1395.0992 | 3.4801 | 1391.4414 | 3.4801 | 1389.1516
| 3.6734 | 111.2956 | 3.5679 | 114.6367 | 3.6383 | 112.5867 | 3.6383 | 110.8104 | 3.5855 | 113.3314 | 3.6558 | 1404.1209 | 3.5679 | 1396.9265 | 3.5679 | 1395.4419 | 3.5328 | 1392.0133 | 3.5328 | 1396.3555 | 3.5152 | 1394.7565 | 3.5152 | 1388.4639
| 3.7086 | 111.2956 | 3.6031 | 114.6768 | 3.6734 | 112.6874 | 3.6734 | 110.8104 | 3.6207 | 112.4256 | 3.691 | 1404.1209 | 3.6031 | 1394.7565 | 3.6031 | 1395.2135 | 3.5679 | 1392.0133 | 3.5679 | 1396.3555 | 3.5504 | 1395.8987 | 3.5504 | 1385.9392
| 3.7437 | 110.9924 | 3.6383 | 114.6568 | 3.7086 | 112.7479 | 3.7086 | 110.8711 | 3.6558 | 112.345 | 3.7261 | 1403.321 | 3.6383 | 1394.1853 | 3.6383 | 1395.0992 | 3.6031 | 1389.3807 | 3.6031 | 1396.3555 | 3.5855 | 1390.9837 | 3.5855 | 1385.3645
| 3.7789 | 110.9924 | 3.6734 | 114.6768 | 3.7437 | 112.5867 | 3.7437 | 111.1137 | 3.691 | 112.345 | 3.7613 | 1403.0925 | 3.6734 | 1395.7845 | 3.6734 | 1394.6423 | 3.6383 | 1389.3807 | 3.6383 | 1395.7845 | 3.6207 | 1390.9837 | 3.6207 | 1384.3294
| 3.814 | 111.134 | 3.7086 | 114.6768 | 3.7789 | 112.4256 | 3.7789 | 111.4976 | 3.7261 | 112.345 | 3.7964 | 1403.321 | 3.7086 | 1394.7565 | 3.7086 | 1395.4419 | 3.6734 | 1389.3807 | 3.6734 | 1395.0992 | 3.6558 | 1389.4953 | 3.6558 | 1383.293
| 3.8492 | 111.538 | 3.7437 | 114.0749 | 3.814 | 112.3248 | 3.814 | 111.5582 | 3.7613 | 112.7881 | 3.8316 | 1403.321 | 3.7437 | 1395.5561 | 3.7437 | 1395.7845 | 3.7086 | 1391.327 | 3.7086 | 1393.3853 | 3.691 | 1389.4953 | 3.691 | 1384.0992
| 3.8843 | 111.6793 | 3.7789 | 113.4722 | 3.8492 | 111.9417 | 3.8492 | 111.4976 | 3.7964 | 112.8888 | 3.8667 | 1403.0925 | 3.7789 | 1396.2413 | 3.7789 | 1398.0682 | 3.7437 | 1394.4138 | 3.7437 | 1392.9281 | 3.7261 | 1390.7548 | 3.7261 | 1384.0992
| 3.9195 | 111.6793 | 3.814 | 113.5727 | 3.8843 | 111.8408 | 3.8843 | 111.4976 | 3.8316 | 113.1705 | 3.9019 | 1404.4638 | 3.814 | 1396.2413 | 3.814 | 1398.6391 | 3.7789 | 1394.985 | 3.7789 | 1392.6994 | 3.7613 | 1392.0133 | 3.7613 | 1384.9046
| 3.9546 | 111.336 | 3.8492 | 113.6129 | 3.9195 | 111.7803 | 3.9195 | 111.3158 | 3.8667 | 113.2309 | 3.937 | 1406.6367 | 3.8492 | 1397.269 | 3.8492 | 1398.6391 | 3.814 | 1394.985 | 3.814 | 1394.7565 | 3.7964 | 1392.242 | 3.7964 | 1385.9392
| 3.9898 | 111.3562 | 3.8843 | 113.7937 | 3.9546 | 111.6188 | 3.9546 | 111.3158 | 3.9019 | 113.2309 | 3.9722 | 1407.8959 | 3.8843 | 1396.8123 | 3.8843 | 1399.3241 | 3.8492 | 1394.985 | 3.8492 | 1395.8987 | 3.8316 | 1393.8425 | 3.8316 | 1386.7431
| 4.0249 | 111.4976 | 3.9195 | 114.0147 | 3.9898 | 111.2754 | 3.9898 | 110.7295 | 3.937 | 112.8284 | 4.0074 | 1409.9588 | 3.9195 | 1397.6116 | 3.9195 | 1399.7808 | 3.8843 | 1396.013 | 3.8843 | 1398.2966 | 3.8667 | 1395.3277 | 3.8667 | 1387.8906
| 4.0601 | 110.9722 | 3.9546 | 113.9142 | 4.0249 | 110.6081 | 4.0249 | 110.7295 | 3.9722 | 112.8284 | 4.0425 | 1411.5657 | 3.9546 | 1401.4935 | 3.9546 | 1400.9225 | 3.9195 | 1396.013 | 3.9195 | 1400.0091 | 3.9019 | 1395.8987 | 3.9019 | 1389.3807
| 4.0952 | 110.77 | 3.9898 | 113.7736 | 4.0601 | 111.1946 | 4.0601 | 110.3651 | 4.0074 | 112.8083 | 4.0777 | 1412.3701 | 3.9898 | 1402.1787 | 3.9898 | 1401.6077 | 3.9546 | 1398.1824 | 3.9546 | 1400.6942 | 3.937 | 1396.2413 | 3.937 | 1390.5259
| 4.1304 | 110.8509 | 4.0249 | 113.5325 | 4.0952 | 111.5986 | 4.0952 | 110.3651 | 4.0425 | 112.8284 | 4.1128 | 1413.9809 | 4.0249 | 1404.921 | 4.0249 | 1404.0066 | 3.9898 | 1399.6666 | 3.9898 | 1401.8361 | 3.9722 | 1396.2413 | 3.9722 | 1391.2125
| 4.1655 | 110.4259 | 4.0601 | 113.5928 | 4.1304 | 112.2038 | 4.1304 | 109.8788 | 4.0777 | 112.8083 | 4.148 | 1414.9026 | 4.0601 | 1406.2934 | 4.0601 | 1404.1209 | 4.0249 | 1400.3516 | 4.0249 | 1402.7498 | 4.0074 | 1396.3555 | 4.0074 | 1390.9837
| 4.2007 | 110.1828 | 4.0952 | 113.633 | 4.1655 | 112.768 | 4.1655 | 109.8788 | 4.1128 | 112.486 | 4.1831 | 1414.096 | 4.0952 | 1405.9502 | 4.0952 | 1401.6077 | 4.0601 | 1401.8361 | 4.0601 | 1402.6356 | 4.0425 | 1397.4974 | 4.0425 | 1392.0133
| 4.2358 | 110.2031 | 4.1304 | 114.737 | 4.2007 | 112.7076 | 4.2007 | 110.1423 | 4.148 | 111.639 | 4.2183 | 1414.096 | 4.1304 | 1407.4379 | 4.1304 | 1404.3495 | 4.0952 | 1401.6077 | 4.0952 | 1401.4935 | 4.0777 | 1399.5524 | 4.0777 | 1393.271
| 4.271 | 110.5271 | 4.1655 | 115.258 | 4.2358 | 112.2038 | 4.2358 | 110.6283 | 4.1831 | 111.0531 | 4.2534 | 1413.4053 | 4.1655 | 1409.1562 | 4.1655 | 1404.6924 | 4.1304 | 1400.6942 | 4.1304 | 1401.4935 | 4.1128 | 1397.8399 | 4.1128 | 1393.4996
| 4.3061 | 110.6081 | 4.2007 | 115.258 | 4.271 | 112.0425 | 4.271 | 111.2956 | 4.2183 | 110.8913 | 4.2886 | 1414.9026 | 4.2007 | 1406.5222 | 4.2007 | 1403.321 | 4.1655 | 1400.6942 | 4.1655 | 1401.8361 | 4.148 | 1399.5524 | 4.148 | 1393.7282
| 4.3413 | 111.0127 | 4.2358 | 114.6768 | 4.3061 | 111.982 | 4.3061 | 112.0829 | 4.2534 | 110.8306 | 4.3237 | 1415.8254 | 4.2358 | 1408.1249 | 4.2358 | 1404.6924 | 4.2007 | 1400.6942 | 4.2007 | 1402.6356 | 4.1831 | 1398.9816 | 4.1831 | 1394.528
| 4.3764 | 111.4572 | 4.271 | 114.5966 | 4.3413 | 111.6793 | 4.3413 | 112.1837 | 4.2886 | 111.1744 | 4.3589 | 1417.6741 | 4.271 | 1409.0415 | 4.271 | 1405.1497 | 4.2358 | 1401.6077 | 4.2358 | 1403.8923 | 4.2183 | 1400.4658 | 4.2183 | 1396.1272
| 4.4116 | 111.6793 | 4.3061 | 114.1953 | 4.3764 | 111.4168 | 4.3764 | 112.2643 | 4.3237 | 111.6995 | 4.394 | 1420.108 | 4.3061 | 1408.6977 | 4.3061 | 1404.6924 | 4.271 | 1404.0066 | 4.271 | 1407.0944 | 4.2534 | 1400.4658 | 4.2534 | 1398.6391
| 4.4468 | 112.103 | 4.3413 | 113.5727 | 4.4116 | 112.0425 | 4.4116 | 111.9013 | 4.3589 | 112.2845 | 4.4292 | 1421.5029 | 4.3413 | 1408.8123 | 4.3413 | 1405.6071 | 4.3061 | 1405.264 | 4.3061 | 1407.2089 | 4.2886 | 1402.9783 | 4.2886 | 1400.9225
| 4.4819 | 112.103 | 4.3764 | 113.251 | 4.4468 | 112.0425 | 4.4468 | 111.134 | 4.394 | 113.1303 | 4.4643 | 1422.2015 | 4.3764 | 1410.0735 | 4.3764 | 1405.8358 | 4.3413 | 1405.264 | 4.3413 | 1408.5832 | 4.3237 | 1402.6356 | 4.3237 | 1401.9503
| 4.5171 | 111.4976 | 4.4116 | 113.1705 | 4.4819 | 111.3764 | 4.4819 | 110.3854 | 4.4292 | 113.2309 | 4.4995 | 1422.2015 | 4.4116 | 1411.1064 | 4.4116 | 1406.0646 | 4.3764 | 1406.4078 | 4.3764 | 1409.5001 | 4.3589 | 1404.3495 | 4.3589 | 1402.9783
| 4.5522 | 111.4168 | 4.4468 | 113.0297 | 4.5171 | 111.3158 | 4.5171 | 110.2233 | 4.4643 | 113.4521 | 4.5346 | 1422.2015 | 4.4468 | 1410.9916 | 4.4468 | 1408.1249 | 4.4116 | 1407.3234 | 4.4116 | 1410.3029 | 4.394 | 1404.0066 | 4.394 | 1403.4353
| 4.5874 | 111.3764 | 4.4819 | 113.0297 | 4.5522 | 111.6591 | 4.5522 | 110.041 | 4.4995 | 113.1303 | 4.5698 | 1420.108 | 4.4819 | 1410.9916 | 4.4819 | 1408.1249 | 4.4468 | 1406.4078 | 4.4468 | 1410.8768 | 4.4292 | 1405.6071 | 4.4292 | 1404.8067
| 4.6225 | 111.4976 | 4.5171 | 113.1303 | 4.5874 | 111.8408 | 4.5874 | 109.9802 | 4.5346 | 112.929 | 4.6049 | 1419.8758 | 4.5171 | 1411.4509 | 4.5171 | 1408.5832 | 4.4819 | 1407.4379 | 4.4819 | 1411.2212 | 4.4643 | 1406.8655 | 4.4643 | 1403.4353
| 4.6577 | 111.639 | 4.5522 | 113.2711 | 4.6225 | 112.0627 | 4.6225 | 110.1828 | 4.5698 | 112.3248 | 4.6401 | 1419.5277 | 4.5522 | 1411.4509 | 4.5522 | 1410.3029 | 4.5171 | 1408.1249 | 4.5171 | 1411.4509 | 4.4995 | 1408.354 | 4.4995 | 1403.4353
| 4.6928 | 111.9013 | 4.5874 | 113.7535 | 4.6577 | 112.4457 | 4.6577 | 110.3854 | 4.6049 | 111.9215 | 4.6752 | 1419.5277 | 4.5874 | 1410.9916 | 4.5874 | 1410.3029 | 4.5522 | 1408.4686 | 4.5522 | 1411.9104 | 4.5346 | 1408.8123 | 4.5346 | 1403.321
| 4.728 | 112.7076 | 4.6225 | 114.2957 | 4.6928 | 113.2107 | 4.6928 | 110.6081 | 4.6401 | 111.639 | 4.7104 | 1419.6437 | 4.6225 | 1412.4851 | 4.6225 | 1408.5832 | 4.5874 | 1409.0415 | 4.5874 | 1411.9104 | 4.5698 | 1410.0735 | 4.5698 | 1403.321
| 4.7631 | 112.7881 | 4.6577 | 114.9174 | 4.728 | 113.2107 | 4.728 | 111.0935 | 4.6752 | 111.5784 | 4.7455 | 1421.1539 | 4.6577 | 1412.4851 | 4.6577 | 1408.0104 | 4.6225 | 1408.4686 | 4.6225 | 1411.2212 | 4.6049 | 1409.0415 | 4.6049 | 1403.5495
| 4.7983 | 112.7881 | 4.6928 | 115.3582 | 4.7631 | 113.1102 | 4.7631 | 111.538 | 4.7104 | 111.5986 | 4.7807 | 1421.3865 | 4.6928 | 1412.4851 | 4.6928 | 1407.4379 | 4.6577 | 1408.0104 | 4.6577 | 1411.5657 | 4.6401 | 1408.354 | 4.6401 | 1403.5495
| 4.8334 | 111.861 | 4.728 | 115.4383 | 4.7983 | 112.9693 | 4.7983 | 111.9013 | 4.7455 | 111.9215 | 4.8159 | 1420.805 | 4.728 | 1412.6 | 4.728 | 1407.2089 | 4.6928 | 1406.5222 | 4.6928 | 1411.5657 | 4.6752 | 1408.2395 | 4.6752 | 1403.6638
| 4.8686 | 111.4774 | 4.7631 | 115.1579 | 4.8334 | 112.9089 | 4.8334 | 112.3046 | 4.7807 | 112.6069 | 4.851 | 1422.6677 | 4.7631 | 1412.83 | 4.7631 | 1406.2934 | 4.728 | 1406.5222 | 4.728 | 1410.762 | 4.7104 | 1409.7294 | 4.7104 | 1403.8923
| 4.9037 | 111.4774 | 4.7983 | 114.5766 | 4.8686 | 112.7076 | 4.8686 | 112.3046 | 4.8159 | 112.929 | 4.8862 | 1424.0686 | 4.7983 | 1412.83 | 4.7983 | 1407.5523 | 4.7631 | 1406.6367 | 4.7631 | 1410.762 | 4.7455 | 1408.8123 | 4.7455 | 1404.1209
| 4.9389 | 111.7399 | 4.8334 | 114.5766 | 4.9037 | 112.5263 | 4.9037 | 112.2442 | 4.851 | 113.2107 | 4.9213 | 1424.6534 | 4.8334 | 1417.3272 | 4.8334 | 1409.1562 | 4.7983 | 1408.9269 | 4.7983 | 1412.4851 | 4.7807 | 1409.7294 | 4.7807 | 1404.1209
| 4.974 | 112.8687 | 4.8686 | 114.5766 | 4.9389 | 112.5263 | 4.9389 | 111.6793 | 4.8862 | 113.6531 | 4.9565 | 1424.1856 | 4.8686 | 1417.3272 | 4.8686 | 1412.0253 | 4.8334 | 1408.9269 | 4.8334 | 1413.5204 | 4.8159 | 1409.9588 | 4.8159 | 1403.8923
| 5.0092 | 113.2309 | 4.9037 | 114.9375 | 4.974 | 112.5465 | 4.974 | 111.336 | 4.9213 | 113.8741 | 4.9916 | 1425.2388 | 4.9037 | 1417.7899 | 4.9037 | 1414.4416 | 4.8686 | 1409.2708 | 4.8686 | 1415.0179 | 4.851 | 1410.6472 | 4.851 | 1404.8067
| 5.0443 | 113.2711 | 4.9389 | 115.0577 | 5.0092 | 111.9417 | 5.0092 | 110.4259 | 4.9565 | 113.7736 | 5.0268 | 1424.7704 | 4.9389 | 1417.7899 | 4.9389 | 1415.71 | 4.9037 | 1409.7294 | 4.9037 | 1418.0213 | 4.8862 | 1410.9916 | 4.8862 | 1405.1497
| 5.0795 | 113.2711 | 4.974 | 115.0777 | 5.0443 | 111.7197 | 5.0443 | 111.336 | 4.9916 | 113.2711 | 5.0619 | 1424.5364 | 4.974 | 1418.3687 | 4.974 | 1418.1371 | 4.9389 | 1410.3029 | 4.9389 | 1419.5277 | 4.9213 | 1411.4509 | 4.9213 | 1405.6071
| 5.1146 | 113.3515 | 5.0092 | 114.7771 | 5.0795 | 111.1137 | 5.0795 | 111.336 | 5.0268 | 112.8284 | 5.0971 | 1425.9422 | 5.0092 | 1421.7357 | 5.0092 | 1416.4026 | 4.974 | 1412.0253 | 4.974 | 1419.7598 | 4.9565 | 1412.0253 | 4.9565 | 1406.179
| 5.1498 | 114.0548 | 5.0443 | 114.5966 | 5.1146 | 111.0329 | 5.1146 | 111.3764 | 5.0619 | 112.8284 | 5.1322 | 1430.5384 | 5.0443 | 1423.9518 | 5.0443 | 1419.4117 | 5.0092 | 1414.2112 | 5.0092 | 1421.6193 | 4.9916 | 1415.4792 | 4.9916 | 1409.5001
| 5.1849 | 115.4783 | 5.0795 | 114.0548 | 5.1498 | 110.6081 | 5.1498 | 111.2552 | 5.0971 | 112.6874 | 5.1674 | 1432.6754 | 5.0795 | 1425.2388 | 5.0795 | 1419.4117 | 5.0443 | 1417.0959 | 5.0443 | 1421.3865 | 5.0268 | 1417.4428 | 5.0268 | 1410.1882
| 5.2201 | 115.5984 | 5.1146 | 113.4119 | 5.1849 | 110.5271 | 5.1849 | 111.0329 | 5.1322 | 113.09 | 5.2025 | 1436.5013 | 5.1146 | 1426.2942 | 5.1146 | 1419.8758 | 5.0795 | 1418.2529 | 5.0795 | 1421.3865 | 5.0619 | 1418.832 | 5.0619 | 1413.6355
| 5.2553 | 115.5984 | 5.1498 | 112.9089 | 5.2201 | 111.0329 | 5.2201 | 110.5271 | 5.1674 | 113.09 | 5.2377 | 1438.4282 | 5.1498 | 1427.9403 | 5.1498 | 1422.4346 | 5.1146 | 1420.6888 | 5.1146 | 1422.318 | 5.0971 | 1419.0638 | 5.0971 | 1415.3638
| 5.2904 | 115.7185 | 5.1849 | 112.9089 | 5.2553 | 112.1837 | 5.2553 | 110.4056 | 5.2025 | 113.2309 | 5.2728 | 1439.5164 | 5.1849 | 1429.4738 | 5.1849 | 1426.1768 | 5.1498 | 1422.901 | 5.1498 | 1425.5904 | 5.1322 | 1420.9213 | 5.1322 | 1418.7162
| 5.3256 | 115.3582 | 5.2201 | 112.9089 | 5.2904 | 112.5666 | 5.2904 | 110.0004 | 5.2377 | 113.0095 | 5.308 | 1439.5164 | 5.2201 | 1429.4738 | 5.2201 | 1427.8225 | 5.1849 | 1423.8349 | 5.1849 | 1427.2341 | 5.1674 | 1423.3677 | 5.1674 | 1419.2957
| 5.3607 | 115.218 | 5.2553 | 113.432 | 5.3256 | 112.7479 | 5.3256 | 110.0004 | 5.2728 | 112.8888 | 5.3431 | 1439.0324 | 5.2553 | 1429.4738 | 5.2553 | 1430.42 | 5.2201 | 1426.4116 | 5.2201 | 1429.4738 | 5.2025 | 1426.4116 | 5.2025 | 1418.1371
| 5.3959 | 115.3181 | 5.2904 | 113.854 | 5.3607 | 113.0498 | 5.3607 | 110.3246 | 5.308 | 112.2038 | 5.3783 | 1439.0324 | 5.2904 | 1431.2496 | 5.2904 | 1431.131 | 5.2553 | 1426.764 | 5.2553 | 1429.8284 | 5.2377 | 1427.7048 | 5.2377 | 1419.2957
| 5.431 | 115.8386 | 5.3256 | 113.854 | 5.3959 | 113.2309 | 5.3959 | 110.8913 | 5.3431 | 112.2038 | 5.4134 | 1440.1224 | 5.3256 | 1430.3016 | 5.3256 | 1430.1833 | 5.2904 | 1427.4694 | 5.2904 | 1429.9467 | 5.2728 | 1428.2937 | 5.2728 | 1420.5726
| 5.4662 | 116.4784 | 5.3607 | 114.1351 | 5.431 | 114.115 | 5.431 | 111.134 | 5.3783 | 111.5784 | 5.4486 | 1441.7025 | 5.3607 | 1430.8939 | 5.3607 | 1431.3682 | 5.3256 | 1427.9403 | 5.3256 | 1429.9467 | 5.308 | 1427.8225 | 5.308 | 1421.0376
| 5.5013 | 117.0377 | 5.3959 | 114.8573 | 5.4662 | 114.717 | 5.4662 | 111.6591 | 5.4134 | 111.5582 | 5.4837 | 1444.8843 | 5.3959 | 1433.1517 | 5.3959 | 1431.7243 | 5.3607 | 1428.7654 | 5.3607 | 1429.9467 | 5.3431 | 1428.7654 | 5.3431 | 1422.901
| 5.5365 | 117.0377 | 5.431 | 116.2586 | 5.5013 | 115.4583 | 5.5013 | 112.2845 | 5.4486 | 111.5582 | 5.5189 | 1445.6228 | 5.431 | 1435.4216 | 5.431 | 1431.7243 | 5.3959 | 1430.1833 | 5.3959 | 1429.9467 | 5.3783 | 1429.2376 | 5.3783 | 1424.0686
| 5.5716 | 116.4185 | 5.4662 | 116.5384 | 5.5365 | 115.6985 | 5.5365 | 112.2845 | 5.4837 | 112.103 | 5.554 | 1449.7142 | 5.4662 | 1438.3075 | 5.4662 | 1435.1821 | 5.431 | 1430.3016 | 5.431 | 1429.9467 | 5.4134 | 1430.3016 | 5.4134 | 1424.7704
| 5.6068 | 115.9786 | 5.5013 | 117.0577 | 5.5716 | 115.7385 | 5.5716 | 113.1705 | 5.5189 | 112.3046 | 5.5892 | 1452.0939 | 5.5013 | 1440.4864 | 5.5013 | 1437.2228 | 5.4662 | 1430.42 | 5.4662 | 1432.3185 | 5.4486 | 1431.2496 | 5.4486 | 1424.7704
| 5.6419 | 115.4183 | 5.5365 | 117.0976 | 5.6068 | 116.2386 | 5.6068 | 113.7937 | 5.554 | 112.9894 | 5.6244 | 1453.7327 | 5.5365 | 1442.6785 | 5.5365 | 1440.365 | 5.5013 | 1434.9427 | 5.5013 | 1436.7417 | 5.4837 | 1432.3185 | 5.4837 | 1426.4116
| 5.6771 | 115.4183 | 5.5716 | 116.838 | 5.6419 | 116.6383 | 5.6419 | 113.854 | 5.5892 | 113.3917 | 5.6595 | 1454.9994 | 5.5716 | 1443.7796 | 5.5716 | 1441.5807 | 5.5365 | 1437.3432 | 5.5365 | 1440.0011 | 5.5189 | 1434.7035 | 5.5189 | 1429.0014
| 5.7122 | 115.4583 | 5.6068 | 117.2772 | 5.6771 | 116.6183 | 5.6771 | 113.8942 | 5.6244 | 113.3917 | 5.6947 | 1456.2714 | 5.6068 | 1444.0248 | 5.6068 | 1443.4122 | 5.5716 | 1440.365 | 5.5716 | 1442.0682 | 5.554 | 1435.7812 | 5.554 | 1432.4374
| 5.7474 | 115.8786 | 5.6419 | 117.7759 | 5.7122 | 116.8979 | 5.7122 | 114.0347 | 5.6595 | 114.3559 | 5.7298 | 1456.5265 | 5.6419 | 1445.3764 | 5.6419 | 1446.61 | 5.6068 | 1442.4342 | 5.6068 | 1442.8006 | 5.5892 | 1436.5013 | 5.5892 | 1433.8672
| 5.7825 | 115.8786 | 5.6771 | 117.6762 | 5.7474 | 116.4984 | 5.7474 | 113.854 | 5.6947 | 115.0577 | 5.765 | 1457.1652 | 5.6771 | 1445.3764 | 5.6771 | 1447.6001 | 5.6419 | 1444.7614 | 5.6419 | 1443.7796 | 5.6244 | 1437.9456 | 5.6244 | 1435.1821
| 5.8177 | 116.2786 | 5.7122 | 118.1547 | 5.7825 | 116.9179 | 5.7825 | 114.0548 | 5.7298 | 115.0777 | 5.8001 | 1460.8973 | 5.7122 | 1449.4648 | 5.7122 | 1447.6001 | 5.6771 | 1443.6571 | 5.6771 | 1446.8572 | 5.6595 | 1439.8799 | 5.6595 | 1435.1821
| 5.8528 | 116.2586 | 5.7474 | 118.3738 | 5.8177 | 116.6183 | 5.8177 | 114.1552 | 5.765 | 115.9986 | 5.8353 | 1462.847 | 5.7474 | 1449.4648 | 5.7474 | 1448.2204 | 5.7122 | 1443.2898 | 5.7122 | 1446.8572 | 5.6947 | 1443.2898 | 5.6947 | 1435.4216
| 5.888 | 116.5584 | 5.7825 | 119.2696 | 5.8528 | 116.4585 | 5.8528 | 115.4383 | 5.8001 | 116.2986 | 5.8704 | 1462.847 | 5.7825 | 1449.5895 | 5.7825 | 1448.9665 | 5.7474 | 1445.9926 | 5.7474 | 1447.9722 | 5.7298 | 1443.6571 | 5.7298 | 1436.5013
| 5.9231 | 116.7981 | 5.8177 | 119.2696 | 5.888 | 116.9778 | 5.888 | 115.9186 | 5.8353 | 116.2986 | 5.9056 | 1462.847 | 5.8177 | 1451.5914 | 5.8177 | 1452.8492 | 5.7825 | 1445.0073 | 5.7825 | 1448.3447 | 5.765 | 1445.6228 | 5.765 | 1436.9822
| 5.9583 | 117.6164 | 5.8528 | 119.389 | 5.9231 | 116.9778 | 5.9231 | 116.1786 | 5.8704 | 116.3785 | 5.9407 | 1461.8055 | 5.8528 | 1453.2275 | 5.8528 | 1454.1121 | 5.8177 | 1448.842 | 5.8177 | 1449.2155 | 5.8001 | 1446.61 | 5.8001 | 1436.9822
| 5.9935 | 117.6164 | 5.888 | 119.2099 | 5.9583 | 116.4585 | 5.9583 | 116.2186 | 5.9056 | 117.1774 | 5.9759 | 1461.5457 | 5.888 | 1453.48 | 5.888 | 1452.8492 | 5.8528 | 1454.2387 | 5.8528 | 1453.2275 | 5.8353 | 1447.2284 | 5.8353 | 1437.7045
| 6.0286 | 117.8158 | 5.9231 | 118.7522 | 5.9935 | 116.5384 | 5.9935 | 116.2186 | 5.9407 | 117.5565 | 6.011 | 1456.9095 | 5.9231 | 1454.6188 | 5.9231 | 1453.48 | 5.888 | 1454.2387 | 5.888 | 1449.5895 | 5.8704 | 1448.3447 | 5.8704 | 1439.7587
| 6.0638 | 117.8158 | 5.9583 | 118.7124 | 6.0286 | 116.7981 | 6.0286 | 115.9386 | 5.9759 | 117.5565 | 6.0462 | 1455.8893 | 5.9583 | 1454.2387 | 5.9583 | 1452.3454 | 5.9231 | 1454.2387 | 5.9231 | 1451.4659 | 5.9056 | 1451.4659 | 5.9056 | 1441.8244
| 6.0989 | 116.3585 | 5.9935 | 118.792 | 6.0638 | 116.6583 | 6.0638 | 115.7986 | 6.011 | 118.314 | 6.0813 | 1454.8725 | 5.9935 | 1453.48 | 5.9935 | 1449.2155 | 5.9583 | 1453.8591 | 5.9583 | 1451.4659 | 5.9407 | 1453.3537 | 5.9407 | 1446.2394
| 6.1341 | 116.3585 | 6.0286 | 118.8318 | 6.0989 | 116.9379 | 6.0989 | 115.7986 | 6.0462 | 117.5565 | 6.1165 | 1454.1121 | 6.0286 | 1452.9752 | 6.0286 | 1449.7142 | 5.9935 | 1453.6063 | 5.9935 | 1450.2137 | 5.9759 | 1451.8425 | 5.9759 | 1446.2394
| 6.1692 | 116.3585 | 6.0638 | 118.8517 | 6.1341 | 117.2573 | 6.1341 | 115.9986 | 6.0813 | 117.2772 | 6.1516 | 1456.399 | 6.0638 | 1451.0897 | 6.0638 | 1448.9665 | 6.0286 | 1450.2137 | 6.0286 | 1448.7176 | 6.011 | 1451.4659 | 6.011 | 1446.61
| 6.2044 | 116.5983 | 6.0989 | 118.8517 | 6.1692 | 117.5765 | 6.1692 | 116.0386 | 6.1165 | 116.5384 | 6.1868 | 1457.6771 | 6.0989 | 1451.0897 | 6.0989 | 1449.2155 | 6.0638 | 1450.2137 | 6.0638 | 1448.7176 | 6.0462 | 1449.9639 | 6.0462 | 1447.6001
| 6.2395 | 117.6762 | 6.1341 | 118.8517 | 6.2044 | 118.3539 | 6.2044 | 116.0986 | 6.1516 | 116.0386 | 6.2219 | 1461.9355 | 6.1341 | 1452.8492 | 6.1341 | 1449.839 | 6.0989 | 1447.4762 | 6.0989 | 1448.0963 | 6.0813 | 1444.7614 | 6.0813 | 1445.8693
| 6.2747 | 118.573 | 6.1692 | 119.0707 | 6.2395 | 118.4933 | 6.2395 | 116.0986 | 6.1868 | 116.0386 | 6.2571 | 1463.8924 | 6.1692 | 1453.7327 | 6.1692 | 1451.9682 | 6.1341 | 1447.3523 | 6.1341 | 1448.7176 | 6.1165 | 1444.7614 | 6.1165 | 1444.7614
| 6.3098 | 120.045 | 6.2044 | 119.6674 | 6.2747 | 118.573 | 6.2747 | 115.8186 | 6.2219 | 116.1986 | 6.2922 | 1465.0733 | 6.2044 | 1454.4921 | 6.2044 | 1453.6063 | 6.1692 | 1447.3523 | 6.1692 | 1451.0897 | 6.1516 | 1444.8843 | 6.1516 | 1445.0073
| 6.345 | 120.8197 | 6.2395 | 119.9457 | 6.3098 | 118.7721 | 6.3098 | 115.8186 | 6.2571 | 116.4185 | 6.3274 | 1464.6791 | 6.2395 | 1456.7818 | 6.2395 | 1457.0373 | 6.2044 | 1450.4637 | 6.2044 | 1453.2275 | 6.1868 | 1446.4864 | 6.1868 | 1444.1475
| 6.3801 | 121.0182 | 6.2747 | 120.1444 | 6.345 | 118.7721 | 6.345 | 116.4185 | 6.2922 | 117.2174 | 6.3625 | 1464.2855 | 6.2747 | 1457.9334 | 6.2747 | 1457.8052 | 6.2395 | 1451.3404 | 6.2395 | 1454.8725 | 6.2219 | 1448.2204 | 6.2219 | 1444.1475
| 6.4153 | 121.2761 | 6.3098 | 120.0848 | 6.3801 | 118.0151 | 6.3801 | 116.9778 | 6.3274 | 117.4568 | 6.3977 | 1464.0234 | 6.3098 | 1458.0616 | 6.3098 | 1458.1899 | 6.2747 | 1455.8893 | 6.2747 | 1455.6347 | 6.2571 | 1452.0939 | 6.2571 | 1444.6385
| 6.4504 | 121.772 | 6.345 | 119.7071 | 6.4153 | 117.9553 | 6.4153 | 117.1375 | 6.3625 | 117.397 | 6.4329 | 1464.9418 | 6.345 | 1459.3471 | 6.345 | 1457.421 | 6.3098 | 1457.1652 | 6.3098 | 1456.0166 | 6.2922 | 1454.7456 | 6.2922 | 1445.9926
| 6.4856 | 121.9703 | 6.3801 | 119.6873 | 6.4504 | 117.9553 | 6.4504 | 117.1375 | 6.3977 | 117.6363 | 6.468 | 1464.6791 | 6.3801 | 1460.3796 | 6.3801 | 1457.9334 | 6.345 | 1458.5751 | 6.345 | 1456.0166 | 6.3274 | 1455.1263 | 6.3274 | 1447.7241
| 6.5207 | 121.8315 | 6.4153 | 119.7071 | 6.4856 | 118.4535 | 6.4856 | 117.2174 | 6.4329 | 117.8357 | 6.5032 | 1464.9418 | 6.4153 | 1459.3471 | 6.4153 | 1459.0895 | 6.3801 | 1458.5751 | 6.3801 | 1456.144 | 6.3625 | 1456.0166 | 6.3625 | 1450.4637
| 6.5559 | 121.772 | 6.4504 | 120.0053 | 6.5207 | 119.1701 | 6.5207 | 116.5783 | 6.468 | 118.3937 | 6.5383 | 1466.7879 | 6.4504 | 1458.7036 | 6.4504 | 1459.3471 | 6.4153 | 1460.1211 | 6.4153 | 1455.6347 | 6.3977 | 1456.0166 | 6.3977 | 1451.3404
| 6.591 | 121.7125 | 6.4856 | 120.7403 | 6.5559 | 119.8463 | 6.5559 | 116.5983 | 6.5032 | 118.3937 | 6.5735 | 1467.8485 | 6.4856 | 1461.1565 | 6.4856 | 1460.2503 | 6.4504 | 1460.7678 | 6.4504 | 1456.144 | 6.4329 | 1455.2534 | 6.4329 | 1452.4713
| 6.6262 | 120.9586 | 6.5207 | 122.0496 | 6.591 | 121.5737 | 6.591 | 116.5783 | 6.5383 | 118.4933 | 6.6086 | 1467.8485 | 6.5207 | 1461.5457 | 6.5207 | 1460.2503 | 6.4856 | 1458.3182 | 6.4856 | 1458.4467 | 6.468 | 1456.7818 | 6.468 | 1454.3654
| 6.6613 | 120.8197 | 6.5559 | 122.4064 | 6.6262 | 122.1884 | 6.6262 | 117.0577 | 6.5735 | 118.4933 | 6.6438 | 1468.6466 | 6.5559 | 1461.5457 | 6.5559 | 1463.2386 | 6.5207 | 1458.3182 | 6.5207 | 1458.7036 | 6.5032 | 1457.6771 | 6.5032 | 1452.9752
| 6.6965 | 120.9785 | 6.591 | 123.852 | 6.6613 | 123.0007 | 6.6613 | 117.7161 | 6.6086 | 118.4734 | 6.6789 | 1469.18 | 6.591 | 1464.5478 | 6.591 | 1463.2386 | 6.5559 | 1458.9608 | 6.5559 | 1459.3471 | 6.5383 | 1459.2183 | 6.5383 | 1452.9752
| 6.7316 | 121.3952 | 6.6262 | 124.2082 | 6.6965 | 122.763 | 6.6965 | 118.7124 | 6.6438 | 118.2144 | 6.7141 | 1470.3841 | 6.6262 | 1466.2592 | 6.6262 | 1463.2386 | 6.591 | 1458.3182 | 6.591 | 1458.9608 | 6.5735 | 1459.4759 | 6.5735 | 1454.4921
| 6.7668 | 121.8117 | 6.6613 | 124.4851 | 6.7316 | 123.0403 | 6.7316 | 121.5142 | 6.6789 | 118.4734 | 6.7492 | 1474.1653 | 6.6613 | 1466.7879 | 6.6613 | 1463.8924 | 6.6262 | 1457.421 | 6.6262 | 1460.5089 | 6.6086 | 1461.6755 | 6.6086 | 1455.2534
| 6.802 | 122.1686 | 6.6965 | 124.6434 | 6.7668 | 122.8819 | 6.7668 | 121.8712 | 6.7141 | 118.7721 | 6.7844 | 1479.1084 | 6.6965 | 1467.0527 | 6.6965 | 1466.7879 | 6.6613 | 1460.6383 | 6.6613 | 1461.2861 | 6.6438 | 1461.6755 | 6.6438 | 1459.3471
| 6.8371 | 123.0601 | 6.7316 | 124.4851 | 6.802 | 122.763 | 6.802 | 122.1289 | 6.7492 | 119.9854 | 6.8195 | 1482.7383 | 6.7316 | 1471.8631 | 6.7316 | 1467.4503 | 6.6965 | 1465.7315 | 6.6965 | 1464.1544 | 6.6789 | 1461.6755 | 6.6789 | 1460.3796
| 6.8723 | 124.2675 | 6.7668 | 124.3071 | 6.8371 | 122.1884 | 6.8371 | 121.8712 | 6.7844 | 121.891 | 6.8547 | 1483.0196 | 6.7668 | 1473.3505 | 6.7668 | 1469.8483 | 6.7316 | 1470.1161 | 6.7316 | 1467.1851 | 6.7141 | 1462.0655 | 6.7141 | 1461.8055
| 6.9074 | 125.9679 | 6.802 | 124.2082 | 6.8723 | 122.0496 | 6.8723 | 122.0893 | 6.8195 | 122.6045 | 6.8898 | 1484.2896 | 6.802 | 1472.944 | 6.802 | 1470.25 | 6.7668 | 1471.5936 | 6.7668 | 1470.3841 | 6.7492 | 1464.9418 | 6.7492 | 1461.8055
| 6.9426 | 125.9679 | 6.8371 | 123.9114 | 6.9074 | 121.5539 | 6.9074 | 121.7919 | 6.8547 | 123.0007 | 6.925 | 1484.7144 | 6.8371 | 1474.3014 | 6.8371 | 1471.4589 | 6.802 | 1471.8631 | 6.802 | 1470.1161 | 6.7844 | 1466.3913 | 6.7844 | 1461.8055
| 6.9777 | 125.9679 | 6.8723 | 123.5551 | 6.9426 | 121.5539 | 6.9426 | 121.6729 | 6.8898 | 122.763 | 6.9601 | 1488.7135 | 6.8723 | 1475.9395 | 6.8723 | 1476.2135 | 6.8371 | 1472.403 | 6.8371 | 1470.3841 | 6.8195 | 1467.7157 | 6.8195 | 1463.108
| 7.0129 | 124.8016 | 6.9074 | 123.3373 | 6.9777 | 122.0496 | 6.9777 | 121.1372 | 6.925 | 123.1987 | 6.9953 | 1490.1572 | 6.9074 | 1475.9395 | 6.9074 | 1479.1084 | 6.8723 | 1471.7283 | 6.8723 | 1472.6734 | 6.8547 | 1468.7799 | 6.8547 | 1466.9203
| 7.048 | 122.1289 | 6.9426 | 123.4165 | 7.0129 | 122.0694 | 7.0129 | 120.7403 | 6.9601 | 123.2977 | 7.0304 | 1491.0275 | 6.9426 | 1480.3591 | 6.9426 | 1482.7383 | 6.9074 | 1476.3507 | 6.9074 | 1472.944 | 6.8898 | 1469.5808 | 6.8898 | 1467.9813
| 7.0832 | 120.8594 | 6.9777 | 123.5155 | 7.048 | 122.8423 | 7.048 | 120.7403 | 6.9953 | 123.2977 | 7.0656 | 1491.0275 | 6.9777 | 1480.9169 | 6.9777 | 1483.7244 | 6.9426 | 1476.6251 | 6.9426 | 1474.9827 | 6.925 | 1471.0554 | 6.925 | 1469.18
| 7.1183 | 120.2636 | 7.0129 | 124.0499 | 7.0832 | 123.3571 | 7.0832 | 120.7403 | 7.0304 | 123.0997 | 7.1007 | 1490.592 | 7.0129 | 1480.7774 | 7.0129 | 1483.5833 | 6.9777 | 1477.8638 | 6.9777 | 1478.8313 | 6.9601 | 1471.7283 | 6.9601 | 1471.0554
| 7.1535 | 120.3629 | 7.048 | 124.6236 | 7.1183 | 124.4851 | 7.1183 | 121.4547 | 7.0656 | 123.1987 | 7.1359 | 1490.0125 | 7.048 | 1481.1963 | 7.048 | 1483.0196 | 7.0129 | 1478.0018 | 7.0129 | 1480.7774 | 6.9953 | 1472.1329 | 6.9953 | 1473.2149
| 7.1886 | 120.9189 | 7.0832 | 125.9284 | 7.1535 | 125.6517 | 7.1535 | 121.9703 | 7.1007 | 122.228 | 7.171 | 1489.4343 | 7.0832 | 1482.7383 | 7.0832 | 1483.0196 | 7.048 | 1478.4162 | 7.048 | 1480.7774 | 7.0304 | 1474.0293 | 7.0304 | 1473.6218
| 7.2238 | 121.4547 | 7.1183 | 127.173 | 7.1886 | 127.1532 | 7.1886 | 122.5847 | 7.1359 | 122.2082 | 7.2062 | 1487.9947 | 7.1183 | 1482.7383 | 7.1183 | 1482.5977 | 7.0832 | 1479.1084 | 7.0832 | 1480.7774 | 7.0656 | 1475.1192 | 7.0656 | 1474.8463
| 7.2589 | 121.4547 | 7.1535 | 127.9629 | 7.2238 | 127.1532 | 7.2238 | 123.4759 | 7.171 | 122.3469 | 7.2414 | 1487.9947 | 7.1535 | 1483.3013 | 7.1535 | 1480.9169 | 7.1183 | 1480.0807 | 7.1183 | 1479.6635 | 7.1007 | 1476.6251 | 7.1007 | 1475.8026
| 7.2941 | 121.6927 | 7.1886 | 128.8316 | 7.2589 | 127.1532 | 7.2589 | 124.4258 | 7.2062 | 122.5055 | 7.2765 | 1487.8512 | 7.1886 | 1485.8505 | 7.1886 | 1479.9415 | 7.1535 | 1478.9698 | 7.1535 | 1479.5247 | 7.1359 | 1475.529 | 7.1359 | 1474.8463
| 7.3292 | 120.9983 | 7.2238 | 128.8316 | 7.2941 | 126.2445 | 7.2941 | 125.8889 | 7.2414 | 123.0007 | 7.3117 | 1490.447 | 7.2238 | 1484.2896 | 7.2238 | 1478.6929 | 7.1886 | 1479.6635 | 7.1886 | 1478.278 | 7.171 | 1475.529 | 7.171 | 1475.3923
| 7.3644 | 120.502 | 7.2589 | 128.654 | 7.3292 | 123.852 | 7.3292 | 126.4224 | 7.2765 | 123.7135 | 7.3468 | 1491.1728 | 7.2589 | 1483.3013 | 7.2589 | 1478.5545 | 7.2238 | 1480.0807 | 7.2238 | 1478.1399 | 7.2062 | 1475.9395 | 7.2062 | 1475.529
| 7.3995 | 120.2835 | 7.2941 | 125.8691 | 7.3644 | 123.0205 | 7.3644 | 126.4224 | 7.3117 | 124.3071 | 7.382 | 1495.8686 | 7.2941 | 1486.1354 | 7.2941 | 1478.6929 | 7.2589 | 1478.8313 | 7.2589 | 1478.278 | 7.2414 | 1477.3125 | 7.2414 | 1475.2557
| 7.4347 | 120.3033 | 7.3292 | 124.4258 | 7.3995 | 121.6134 | 7.3995 | 126.2248 | 7.3468 | 126.2248 | 7.4171 | 1498.1006 | 7.3292 | 1488.4257 | 7.3292 | 1481.1963 | 7.2941 | 1478.8313 | 7.2941 | 1481.476 | 7.2765 | 1477.3125 | 7.2765 | 1475.2557
| 7.4698 | 120.1444 | 7.3644 | 124.2082 | 7.4347 | 121.6134 | 7.4347 | 123.7926 | 7.382 | 126.9162 | 7.4523 | 1500.5033 | 7.3644 | 1492.4846 | 7.3644 | 1488.8575 | 7.3292 | 1482.5977 | 7.3292 | 1485.2818 | 7.3117 | 1478.0018 | 7.3117 | 1476.6251
| 7.505 | 120.2835 | 7.3995 | 124.1884 | 7.4698 | 121.8315 | 7.4698 | 121.6531 | 7.4171 | 127.8247 | 7.4874 | 1499.5997 | 7.3995 | 1496.9077 | 7.3995 | 1489.4343 | 7.3644 | 1487.2779 | 7.3644 | 1491.1728 | 7.3468 | 1480.9169 | 7.3468 | 1478.4162
| 7.5401 | 120.3828 | 7.4347 | 124.4456 | 7.505 | 121.8315 | 7.505 | 119.8264 | 7.4523 | 127.8247 | 7.5226 | 1499.2992 | 7.4347 | 1498.2501 | 7.4347 | 1490.3021 | 7.3995 | 1491.4637 | 7.3995 | 1494.8339 | 7.382 | 1483.7244 | 7.382 | 1479.2471
| 7.5753 | 120.3033 | 7.4698 | 124.94 | 7.5401 | 122.01 | 7.5401 | 119.8264 | 7.4874 | 127.6075 | 7.5577 | 1497.0565 | 7.4698 | 1498.5494 | 7.4698 | 1489.29 | 7.4347 | 1493.5097 | 7.4347 | 1496.3134 | 7.4171 | 1486.8489 | 7.4171 | 1481.3361
| 7.6105 | 120.4424 | 7.505 | 125.0389 | 7.5753 | 122.6045 | 7.5753 | 119.9059 | 7.5226 | 124.3862 | 7.5929 | 1499.2992 | 7.505 | 1498.5494 | 7.505 | 1488.8575 | 7.4698 | 1494.539 | 7.4698 | 1496.1651 | 7.4523 | 1490.7371 | 7.4523 | 1482.5977
| 7.6456 | 121.3158 | 7.5401 | 124.7225 | 7.6105 | 122.3469 | 7.6105 | 121.2761 | 7.5577 | 123.2779 | 7.628 | 1496.9077 | 7.5401 | 1496.4619 | 7.5401 | 1486.4205 | 7.505 | 1494.3917 | 7.505 | 1494.8339 | 7.4874 | 1494.3917 | 7.4874 | 1482.8789
| 7.6808 | 121.3555 | 7.5753 | 124.94 | 7.6456 | 121.8712 | 7.6456 | 121.2761 | 7.5929 | 121.8514 | 7.6632 | 1496.6104 | 7.5753 | 1492.3385 | 7.5753 | 1486.2779 | 7.5401 | 1494.3917 | 7.5401 | 1493.5097 | 7.5226 | 1494.0974 | 7.5226 | 1482.5977
| 7.7159 | 121.3555 | 7.6105 | 124.3269 | 7.6808 | 121.9108 | 7.6808 | 122.3073 | 7.628 | 121.8514 | 7.6983 | 1494.2445 | 7.6105 | 1485.566 | 7.6105 | 1486.1354 | 7.5753 | 1491.1728 | 7.5753 | 1492.6308 | 7.5577 | 1492.9234 | 7.5577 | 1482.5977
| 7.7511 | 121.9108 | 7.6456 | 124.1686 | 7.7159 | 121.8315 | 7.7159 | 122.1091 | 7.6632 | 121.9307 | 7.7335 | 1494.2445 | 7.6456 | 1490.592 | 7.6456 | 1483.8656 | 7.6105 | 1489.29 | 7.6105 | 1488.282 | 7.5929 | 1490.8822 | 7.5929 | 1479.5247
| 7.7862 | 121.9108 | 7.6808 | 123.8916 | 7.7511 | 121.4547 | 7.7511 | 121.2761 | 7.6983 | 122.9413 | 7.7686 | 1491.9008 | 7.6808 | 1489.5787 | 7.6808 | 1482.5977 | 7.6456 | 1486.2779 | 7.6456 | 1486.1354 | 7.628 | 1489.29 | 7.628 | 1476.4879
| 7.8214 | 121.3158 | 7.7159 | 123.5353 | 7.7862 | 121.6531 | 7.7862 | 121.6927 | 7.7335 | 122.9413 | 7.8038 | 1493.9503 | 7.7159 | 1487.1348 | 7.7159 | 1482.5977 | 7.6808 | 1486.2779 | 7.6808 | 1483.7244 | 7.6632 | 1487.8512 | 7.6632 | 1476.0765
| 7.8565 | 121.0579 | 7.7511 | 123.6343 | 7.8214 | 122.1487 | 7.8214 | 121.772 | 7.7686 | 123.0007 | 7.8389 | 1494.0974 | 7.7511 | 1484.0069 | 7.7511 | 1480.2198 | 7.7159 | 1485.7082 | 7.7159 | 1483.0196 | 7.6983 | 1485.2818 | 7.6983 | 1476.0765
| 7.8917 | 121.0579 | 7.7862 | 124.1686 | 7.8565 | 122.5847 | 7.8565 | 120.9785 | 7.8038 | 123.0799 | 7.8741 | 1495.4246 | 7.7862 | 1485.1399 | 7.7862 | 1480.2198 | 7.7511 | 1483.3013 | 7.7511 | 1481.1963 | 7.7335 | 1483.8656 | 7.7335 | 1478.8313
| 7.9268 | 121.3555 | 7.8214 | 124.2675 | 7.8917 | 123.0997 | 7.8917 | 121.415 | 7.8389 | 122.7828 | 7.9092 | 1496.9077 | 7.8214 | 1485.1399 | 7.8214 | 1484.4311 | 7.7862 | 1482.0362 | 7.7862 | 1480.0807 | 7.7686 | 1483.4423 | 7.7686 | 1479.2471
| 7.962 | 121.6134 | 7.8565 | 124.2675 | 7.9268 | 123.3373 | 7.9268 | 121.6531 | 7.8741 | 122.5451 | 7.9444 | 1496.9077 | 7.8565 | 1485.4239 | 7.8565 | 1490.8822 | 7.8214 | 1482.8789 | 7.8214 | 1480.3591 | 7.8038 | 1481.896 | 7.8038 | 1479.2471
| 7.9971 | 122.8819 | 7.8917 | 124.7225 | 7.962 | 123.753 | 7.962 | 121.5142 | 7.9092 | 122.8423 | 7.9795 | 1498.999 | 7.8917 | 1489.1457 | 7.8917 | 1491.3182 | 7.8565 | 1482.1765 | 7.8565 | 1481.7559 | 7.8389 | 1482.1765 | 7.8389 | 1481.7559
| 8.0323 | 122.8819 | 7.9268 | 125.0784 | 7.9971 | 123.7926 | 7.9971 | 121.8117 | 7.9444 | 122.8423 | 8.0147 | 1498.999 | 7.9268 | 1490.8822 | 7.9268 | 1492.3385 | 7.8917 | 1484.7144 | 7.8917 | 1481.7559 | 7.8741 | 1483.8656 | 7.8741 | 1482.0362
| 8.0674 | 122.8819 | 7.962 | 125.0784 | 8.0323 | 123.7926 | 8.0323 | 122.7036 | 7.9795 | 122.6045 | 8.0499 | 1495.5726 | 7.962 | 1489.4343 | 7.962 | 1490.8822 | 7.9268 | 1484.8561 | 7.9268 | 1482.5977 | 7.9092 | 1483.8656 | 7.9092 | 1481.7559
| 8.1026 | 123.0799 | 7.9971 | 125.3947 | 8.0674 | 123.4957 | 8.0674 | 123.0601 | 8.0147 | 122.8423 | 8.085 | 1495.5726 | 7.9971 | 1488.8575 | 7.9971 | 1484.7144 | 7.962 | 1489.29 | 7.962 | 1482.5977 | 7.9444 | 1484.0069 | 7.9444 | 1484.4311
| 8.1377 | 123.3967 | 8.0323 | 125.6122 | 8.1026 | 123.0799 | 8.1026 | 123.1393 | 8.0499 | 122.6838 | 8.1202 | 1498.3997 | 8.0323 | 1489.4343 | 8.0323 | 1484.5727 | 7.9971 | 1489.29 | 7.9971 | 1482.3168 | 7.9795 | 1484.7144 | 7.9795 | 1485.4239
| 8.1729 | 123.7926 | 8.0674 | 125.3947 | 8.1377 | 122.664 | 8.1377 | 123.1393 | 8.085 | 122.6045 | 8.1553 | 1502.3165 | 8.0674 | 1492.9234 | 8.0674 | 1485.7082 | 8.0323 | 1488.7135 | 8.0323 | 1482.1765 | 8.0147 | 1485.1399 | 8.0147 | 1483.1604
| 8.208 | 123.852 | 8.1026 | 124.8411 | 8.1729 | 122.664 | 8.1729 | 123.0007 | 8.1202 | 122.9809 | 8.1905 | 1504.5935 | 8.1026 | 1493.6565 | 8.1026 | 1491.755 | 8.0674 | 1489.29 | 8.0674 | 1483.3013 | 8.0499 | 1483.7244 | 8.0499 | 1482.8789
| 8.2432 | 124.2082 | 8.1377 | 124.7225 | 8.208 | 123.1789 | 8.208 | 122.1686 | 8.1553 | 123.3571 | 8.2256 | 1509.3368 | 8.1377 | 1495.4246 | 8.1377 | 1495.2768 | 8.1026 | 1490.1572 | 8.1026 | 1488.8575 | 8.085 | 1483.4423 | 8.085 | 1484.8561
| 8.2783 | 124.228 | 8.1729 | 125.0191 | 8.2432 | 123.951 | 8.2432 | 121.3952 | 8.1905 | 123.6145 | 8.2608 | 1511.1864 | 8.1729 | 1498.6991 | 8.1729 | 1501.409 | 8.1377 | 1493.8034 | 8.1377 | 1491.4637 | 8.1202 | 1483.4423 | 8.1202 | 1484.998
| 8.3135 | 124.3269 | 8.208 | 125.2761 | 8.2783 | 124.129 | 8.2783 | 121.038 | 8.2256 | 123.8916 | 8.2959 | 1515.5314 | 8.208 | 1502.6194 | 8.208 | 1502.6194 | 8.1729 | 1495.8686 | 8.1729 | 1499.2992 | 8.1553 | 1485.7082 | 8.1553 | 1485.8505
| 8.3486 | 124.3862 | 8.2432 | 126.126 | 8.3135 | 124.129 | 8.3135 | 121.1968 | 8.2608 | 124.0103 | 8.3311 | 1516.6241 | 8.2432 | 1506.7294 | 8.2432 | 1504.2892 | 8.208 | 1501.8625 | 8.208 | 1501.2579 | 8.1905 | 1489.5787 | 8.1905 | 1489.8678
| 8.3838 | 124.4456 | 8.2783 | 126.126 | 8.3486 | 123.951 | 8.3486 | 121.9307 | 8.2959 | 123.7333 | 8.3662 | 1517.4061 | 8.2783 | 1508.4148 | 8.2783 | 1508.1079 | 8.2432 | 1502.3165 | 8.2432 | 1502.4679 | 8.2256 | 1497.6526 | 8.2256 | 1494.2445
| 8.419 | 124.2082 | 8.3135 | 126.2445 | 8.3838 | 123.951 | 8.3838 | 122.9809 | 8.3311 | 122.8819 | 8.4014 | 1519.2883 | 8.3135 | 1509.7985 | 8.3135 | 1508.1079 | 8.2783 | 1504.5935 | 8.2783 | 1505.5077 | 8.2608 | 1499.7501 | 8.2608 | 1495.7206
| 8.4541 | 123.6541 | 8.3486 | 126.2643 | 8.419 | 123.951 | 8.419 | 123.0205 | 8.3662 | 122.8026 | 8.4365 | 1523.2336 | 8.3486 | 1510.2607 | 8.3486 | 1508.8756 | 8.3135 | 1506.2709 | 8.3135 | 1505.8128 | 8.2959 | 1501.5601 | 8.2959 | 1497.8018
| 8.4893 | 123.8916 | 8.3838 | 126.2841 | 8.4541 | 125.197 | 8.4541 | 124.228 | 8.4014 | 123.0997 | 8.4717 | 1521.0202 | 8.3838 | 1510.4148 | 8.3838 | 1508.8756 | 8.3486 | 1508.4148 | 8.3486 | 1508.5684 | 8.3311 | 1503.6813 | 8.3311 | 1499.2992
| 8.5244 | 124.1884 | 8.419 | 127.4495 | 8.4893 | 126.4619 | 8.4893 | 124.8411 | 8.4365 | 123.3769 | 8.5068 | 1519.2883 | 8.419 | 1511.65 | 8.419 | 1508.7219 | 8.3838 | 1510.4148 | 8.3838 | 1509.3368 | 8.3662 | 1503.9852 | 8.3662 | 1500.0512
| 8.5596 | 123.3373 | 8.4541 | 127.9629 | 8.5244 | 126.6594 | 8.5244 | 124.8411 | 8.4717 | 124.0697 | 8.542 | 1518.817 | 8.4541 | 1511.8046 | 8.4541 | 1506.1181 | 8.419 | 1513.0435 | 8.419 | 1509.3368 | 8.4014 | 1503.9852 | 8.4014 | 1499.2992
| 8.5947 | 123.6541 | 8.4893 | 128.1999 | 8.5596 | 127.252 | 8.5596 | 124.8609 | 8.5068 | 124.1488 | 8.5771 | 1515.9994 | 8.4893 | 1510.2607 | 8.4893 | 1503.8332 | 8.4541 | 1515.8433 | 8.4541 | 1511.8046 | 8.4365 | 1508.8756 | 8.4365 | 1505.0504
| 8.6299 | 123.5947 | 8.5244 | 128.6342 | 8.5947 | 127.252 | 8.5947 | 125.9877 | 8.542 | 124.2477 | 8.6123 | 1511.65 | 8.5244 | 1508.2614 | 8.5244 | 1500.6541 | 8.4893 | 1513.8196 | 8.4893 | 1512.8885 | 8.4717 | 1509.7985 | 8.4717 | 1507.648
| 8.665 | 123.5947 | 8.5596 | 129.187 | 8.6299 | 127.2915 | 8.6299 | 126.8767 | 8.5771 | 124.5445 | 8.6474 | 1510.4148 | 8.5596 | 1508.2614 | 8.5596 | 1496.1651 | 8.5244 | 1506.5765 | 8.5244 | 1507.1885 | 8.5068 | 1509.7985 | 8.5068 | 1507.0354
| 8.7002 | 123.5947 | 8.5947 | 129.3449 | 8.665 | 127.3705 | 8.665 | 127.173 | 8.6123 | 124.6236 | 8.6826 | 1511.65 | 8.5947 | 1502.4679 | 8.5947 | 1495.8686 | 8.5596 | 1506.5765 | 8.5596 | 1505.9654 | 8.542 | 1504.1372 | 8.542 | 1503.8332
| 8.7353 | 124.8214 | 8.6299 | 129.187 | 8.7002 | 127.3902 | 8.7002 | 127.7062 | 8.6474 | 124.6236 | 8.7177 | 1510.8776 | 8.6299 | 1501.8625 | 8.6299 | 1496.1651 | 8.5947 | 1506.1181 | 8.5947 | 1498.6991 | 8.5771 | 1504.1372 | 8.5771 | 1495.7206
| 8.7705 | 125.3157 | 8.665 | 127.7457 | 8.7353 | 127.1137 | 8.7353 | 127.6075 | 8.6826 | 125.2959 | 8.7529 | 1511.9593 | 8.665 | 1502.4679 | 8.665 | 1497.6526 | 8.6299 | 1505.3552 | 8.6299 | 1498.2501 | 8.6123 | 1501.8625 | 8.6123 | 1495.4246
| 8.8056 | 126.0272 | 8.7002 | 127.5482 | 8.7705 | 127.2125 | 8.7705 | 127.4692 | 8.7177 | 126.4619 | 8.7881 | 1514.2858 | 8.7002 | 1506.1181 | 8.7002 | 1502.9225 | 8.665 | 1505.0504 | 8.665 | 1498.6991 | 8.6474 | 1500.805 | 8.6474 | 1494.0974
| 8.8408 | 126.7977 | 8.7353 | 127.5482 | 8.8056 | 127.1137 | 8.8056 | 126.6594 | 8.7529 | 126.6792 | 8.8232 | 1515.3755 | 8.7353 | 1508.2614 | 8.7353 | 1506.1181 | 8.7002 | 1506.7294 | 8.7002 | 1500.0512 | 8.6826 | 1500.805 | 8.6826 | 1494.0974
| 8.8759 | 126.8965 | 8.7705 | 127.5482 | 8.8408 | 126.857 | 8.8408 | 126.2841 | 8.7881 | 127.3902 | 8.8584 | 1516.7804 | 8.7705 | 1509.7985 | 8.7705 | 1508.4148 | 8.7353 | 1507.3416 | 8.7353 | 1504.898 | 8.7177 | 1502.0138 | 8.7177 | 1495.8686
| 8.9111 | 127.015 | 8.8056 | 127.7852 | 8.8759 | 126.1458 | 8.8759 | 125.9877 | 8.8232 | 127.568 | 8.8935 | 1518.1894 | 8.8056 | 1509.6446 | 8.8056 | 1508.4148 | 8.7705 | 1507.8012 | 8.7705 | 1505.3552 | 8.7529 | 1503.3776 | 8.7529 | 1498.5494
| 8.9462 | 127.2125 | 8.8408 | 127.7852 | 8.9111 | 125.5924 | 8.9111 | 126.126 | 8.8584 | 127.568 | 8.9287 | 1518.1894 | 8.8408 | 1509.6446 | 8.8408 | 1508.4148 | 8.8056 | 1507.8012 | 8.8056 | 1507.0354 | 8.7881 | 1504.5935 | 8.7881 | 1504.7458
| 8.9814 | 127.173 | 8.8759 | 127.7062 | 8.9462 | 125.5924 | 8.9462 | 126.126 | 8.8935 | 126.6792 | 8.9638 | 1520.8625 | 8.8759 | 1509.6446 | 8.8759 | 1508.8756 | 8.8408 | 1508.2614 | 8.8408 | 1508.4148 | 8.8232 | 1506.2709 | 8.8232 | 1505.6602
| 9.0165 | 127.252 | 8.9111 | 127.2323 | 8.9814 | 126.0074 | 8.9814 | 126.0469 | 8.9287 | 126.5014 | 8.999 | 1521.6516 | 8.9111 | 1512.2688 | 8.9111 | 1509.6446 | 8.8759 | 1508.2614 | 8.8759 | 1509.6446 | 8.8584 | 1506.7294 | 8.8584 | 1507.1885
| 9.0517 | 128.0222 | 8.9462 | 127.4297 | 9.0165 | 126.8372 | 9.0165 | 126.0469 | 8.9638 | 126.5409 | 9.0341 | 1526.892 | 8.9462 | 1512.8885 | 8.9462 | 1514.2858 | 8.9111 | 1508.8756 | 8.9111 | 1509.6446 | 8.8935 | 1506.8824 | 8.8935 | 1506.8824
| 9.0868 | 129.029 | 8.9814 | 128.1011 | 9.0517 | 127.4297 | 9.0517 | 126.0469 | 8.999 | 126.9162 | 9.0693 | 1529.935 | 8.9814 | 1517.2496 | 8.9814 | 1516.4678 | 8.9462 | 1510.4148 | 8.9462 | 1511.9593 | 8.9287 | 1505.3552 | 8.9287 | 1504.7458
| 9.122 | 129.7595 | 9.0165 | 129.6608 | 9.0868 | 128.2986 | 9.0868 | 126.0469 | 9.0341 | 126.9162 | 9.1044 | 1532.835 | 9.0165 | 1519.4455 | 9.0165 | 1520.3897 | 8.9814 | 1514.5969 | 8.9814 | 1515.3755 | 8.9638 | 1507.9545 | 8.9638 | 1503.3776
| 9.1571 | 130.5491 | 9.0517 | 130.1345 | 9.122 | 128.3183 | 9.122 | 126.3236 | 9.0693 | 126.5409 | 9.1396 | 1532.835 | 9.0517 | 1520.7048 | 9.0517 | 1521.3358 | 9.0165 | 1516.6241 | 9.0165 | 1518.3462 | 8.999 | 1510.4148 | 8.999 | 1503.3776
| 9.1923 | 130.2332 | 9.0868 | 130.5886 | 9.1571 | 128.496 | 9.1571 | 128.0814 | 9.1044 | 125.8691 | 9.1747 | 1533.32 | 9.0868 | 1520.7048 | 9.0868 | 1522.4419 | 9.0517 | 1521.3358 | 9.0517 | 1520.7048 | 9.0341 | 1515.5314 | 9.0341 | 1504.898
| 9.2275 | 130.5491 | 9.122 | 130.5886 | 9.1923 | 128.5947 | 9.1923 | 128.5947 | 9.1396 | 125.5529 | 9.2099 | 1535.2646 | 9.122 | 1520.7048 | 9.122 | 1523.0752 | 9.0868 | 1522.2838 | 9.0868 | 1520.7048 | 9.0693 | 1518.5031 | 9.0693 | 1505.3552
| 9.2626 | 130.7662 | 9.1571 | 130.8649 | 9.2275 | 128.9501 | 9.2275 | 128.7724 | 9.1747 | 125.5529 | 9.245 | 1536.4024 | 9.1571 | 1522.9168 | 9.1571 | 1524.8208 | 9.122 | 1525.4572 | 9.122 | 1521.8095 | 9.1044 | 1518.66 | 9.1044 | 1506.1181
| 9.2978 | 130.7662 | 9.1923 | 130.9242 | 9.2626 | 129.1672 | 9.2626 | 128.7724 | 9.2099 | 125.6715 | 9.2802 | 1543.7774 | 9.1923 | 1523.0752 | 9.1923 | 1524.8208 | 9.1571 | 1524.9799 | 9.1571 | 1523.5506 | 9.1396 | 1519.2883 | 9.1396 | 1509.3368
| 9.3329 | 130.9834 | 9.2275 | 130.5491 | 9.2978 | 129.4633 | 9.2978 | 128.7724 | 9.245 | 126.6989 | 9.3153 | 1546.259 | 9.2275 | 1529.7743 | 9.2275 | 1525.1389 | 9.1923 | 1525.4572 | 9.1923 | 1523.8679 | 9.1747 | 1519.4455 | 9.1747 | 1511.8046
| 9.3681 | 130.9834 | 9.2626 | 131.1216 | 9.3329 | 129.6015 | 9.3329 | 128.7527 | 9.2802 | 128.417 | 9.3505 | 1548.2527 | 9.2626 | 1530.7389 | 9.2626 | 1527.8509 | 9.2275 | 1528.8117 | 9.2275 | 1524.1853 | 9.2099 | 1522.2838 | 9.2099 | 1513.3538
| 9.4032 | 130.6873 | 9.2978 | 131.4177 | 9.3681 | 131.398 | 9.3681 | 128.6342 | 9.3153 | 129.0685 | 9.3856 | 1551.592 | 9.2978 | 1534.94 | 9.2978 | 1532.9967 | 9.2626 | 1532.1891 | 9.2626 | 1528.8117 | 9.245 | 1519.4455 | 9.245 | 1517.4061
| 9.4384 | 129.7595 | 9.3329 | 131.9113 | 9.4032 | 132.2273 | 9.4032 | 128.7527 | 9.3505 | 129.029 | 9.4208 | 1554.2783 | 9.3329 | 1537.869 | 9.3329 | 1539.9949 | 9.2978 | 1535.7519 | 9.2978 | 1531.8665 | 9.2802 | 1525.7757 | 9.2802 | 1521.178
| 9.4735 | 126.6792 | 9.3681 | 132.8988 | 9.4384 | 132.4643 | 9.4384 | 129.3252 | 9.3856 | 129.0685 | 9.4559 | 1554.7835 | 9.3681 | 1541.965 | 9.3681 | 1540.8148 | 9.3329 | 1536.728 | 9.3329 | 1535.7519 | 9.3153 | 1526.7324 | 9.3153 | 1525.4572
| 9.5087 | 126.4224 | 9.4032 | 133.2346 | 9.4735 | 132.5828 | 9.4735 | 130.1543 | 9.4208 | 129.029 | 9.4911 | 1555.4577 | 9.4032 | 1544.1076 | 9.4032 | 1542.9528 | 9.3681 | 1538.1955 | 9.3681 | 1540.1588 | 9.3505 | 1531.8665 | 9.3505 | 1529.6138
|  |  | 9.4384 | 133.61 | 9.5087 | 132.9581 | 9.5087 | 130.4701 | 9.4559 | 128.5355 | 9.5262 | 1556.1328 | 9.4384 | 1549.2523 | 9.4384 | 1545.0995 | 9.4032 | 1543.2825 | 9.4032 | 1540.8148 | 9.3856 | 1535.7519 | 9.3856 | 1530.578
|  |  | 9.4735 | 133.9657 | 9.5438 | 132.9581 | 9.5438 | 131.5757 | 9.4911 | 128.5355 |  |  | 9.4735 | 1551.592 | 9.4735 | 1548.9189 | 9.4384 | 1546.4249 | 9.4384 | 1543.2825 | 9.4208 | 1536.8908 | 9.4208 | 1532.3505
|  |  | 9.5087 | 133.8669 |  |  |  |  | 9.5262 | 129.2462 |  |  | 9.5087 | 1556.3017 | 9.5087 | 1553.1014 | 9.4735 | 1549.0856 | 9.4735 | 1544.7687 | 9.4559 | 1542.9528 | 9.4559 | 1534.4535
|  |  | 9.5438 | 133.9657 |  |  |  |  | 9.5614 | 130.1345 |  |  | 9.5438 | 1556.6396 | 9.5438 | 1554.2783 | 9.5087 | 1551.2572 | 9.5087 | 1546.4249 | 9.4911 | 1545.9275 | 9.4911 | 1537.0537
|  |  | 9.579 | 133.9855 |  |  |  |  | 9.5966 | 130.8255 |  |  | 9.579 | 1557.3162 | 9.579 | 1553.2694 | 9.5438 | 1551.592 | 9.5438 | 1547.2549 | 9.5262 | 1549.9198 | 9.5262 | 1538.1955
|  |  | 9.6141 | 133.4322 | [8 empty] | 9.6141 | 1557.3162 | 9.6141 | 1553.1014 | 9.579 | 1554.9519 | 9.579 | 1550.2538 | 9.5614 | 1550.0868 | 9.5614 | 1542.4587
[17 empty] | 9.6141 | 1555.2891 | 9.6141 | 1551.4246 | 9.5966 | 1552.0947 | 9.5966 | 1543.9425
[17 empty] | 9.6493 | 1555.2891 | 9.6493 | 1552.5978 | 9.6317 | 1552.4301 | 9.6317 | 1545.9275
[21 empty] | 9.6669 | 1552.0947 | 9.6669 | 1548.7523

## uniformity over various hei 2
The scan area, which is defined by the template, the inner borders of the template to be more precise, is 24cm long, 12 cm to the centre. So my plan is to measure in 5 areas across the scanner plate.
The five areas are actually planes across the scanner bed in the lateral direction; the planes are defined by the labels thusly
The idea is then to scan the step wedge film at the corresponding heights with the following particular steps that correspond to the approximate doses of 0Gy, 1Gy and 14Gy.
The heights were marked on to the scan plate template.
The steps that correspond to the three doses of 0 1 and 14 Gy are
1Gy; the fourth strip after the large black strip at the top of the film.
14Gy is the tenth strip after the large black strip at the top of the film.
it is not possible to obtain a 0Gy strip using this step wedge; a blank film can be scanned for this purpose.
height | 1Gy (4th strip) | 14Gy (tenth strip)
middle | y
4.5cm up
11.5cm up
4.5cm down
11.5cm down
| middle 1Gy |  | 4.5cm up 1Gy |  | 4.5cm down 1Gy |  | 11.5cm up 1Gy |  | 11.5cm down 1Gy |  | middle 14Gy |  | 4.5cm up 14Gy |  | 10cm up 14Gy |  | 10cm up 14Gy; scanner covered |  | 4.5cm down 14Gy |  | 4.5cm down 14Gy; covered |  | 10cm down 14Gy
| cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy
| -9.4735 | 123.951 | -9.579 | 130.1937 | -9.5087 | 127.4692 | -9.5087 | 127.41 | -9.5614 | 131.3782 | -9.4911 | 1562.2456 | -9.579 | 1552.5978 | -9.579 | 1549.2523 | -9.6141 | 1554.7835 | -9.6141 | 1562.9289 | -9.6317 | 1572.9293 | -9.6317 | 1554.7835
| -9.4384 | 123.951 | -9.5438 | 129.3252 | -9.4735 | 127.4692 | -9.4735 | 126.9557 | -9.5262 | 131.2006 | -9.4559 | 1562.2456 | -9.5438 | 1552.5978 | -9.5438 | 1548.2527 | -9.579 | 1554.7835 | -9.579 | 1560.0305 | -9.5966 | 1569.9802 | -9.5966 | 1557.147
| -9.4032 | 124.0103 | -9.5087 | 129.3054 | -9.4384 | 127.6272 | -9.4384 | 126.7385 | -9.4911 | 130.9044 | -9.4208 | 1560.8814 | -9.5087 | 1549.4191 | -9.5087 | 1547.4211 | -9.5438 | 1554.4466 | -9.5438 | 1559.0111 | -9.5614 | 1568.5976 | -9.5614 | 1552.4301
| -9.3681 | 124.0103 | -9.4735 | 129.3054 | -9.4032 | 127.8839 | -9.4032 | 126.3433 | -9.4559 | 129.8384 | -9.3856 | 1559.1808 | -9.4735 | 1546.5908 | -9.4735 | 1543.2825 | -9.5087 | 1550.7553 | -9.5087 | 1557.9935 | -9.5262 | 1564.6408 | -9.5262 | 1548.5857
| -9.3329 | 124.3467 | -9.4384 | 129.2659 | -9.3681 | 127.9629 | -9.3681 | 125.9481 | -9.4208 | 128.9698 | -9.3505 | 1555.1205 | -9.4384 | 1551.0898 | -9.4384 | 1542.6233 | -9.4735 | 1547.2549 | -9.4735 | 1556.9778 | -9.4911 | 1561.2221 | -9.4911 | 1547.7536
| -9.2978 | 124.6434 | -9.4032 | 129.3054 | -9.3329 | 128.5947 | -9.3329 | 125.6517 | -9.3856 | 128.0222 | -9.3153 | 1551.7595 | -9.4032 | 1551.0898 | -9.4032 | 1543.6124 | -9.4384 | 1546.259 | -9.4384 | 1556.9778 | -9.4559 | 1559.6905 | -9.4559 | 1545.265
| -9.2626 | 124.5445 | -9.3681 | 129.6608 | -9.2978 | 128.1011 | -9.2978 | 126.1458 | -9.3505 | 127.5285 | -9.2802 | 1546.9228 | -9.3681 | 1553.1014 | -9.3681 | 1543.2825 | -9.4032 | 1544.1076 | -9.4032 | 1555.2891 | -9.4208 | 1556.3017 | -9.4208 | 1544.9341
| -9.2275 | 124.228 | -9.3329 | 128.9501 | -9.2626 | 127.5482 | -9.2626 | 126.1458 | -9.3153 | 127.41 | -9.245 | 1545.4305 | -9.3329 | 1552.0947 | -9.3329 | 1543.2825 | -9.3681 | 1543.9425 | -9.3681 | 1552.5978 | -9.3856 | 1555.1205 | -9.3856 | 1544.9341
| -9.1923 | 124.228 | -9.2978 | 127.9827 | -9.2275 | 126.5014 | -9.2275 | 125.9086 | -9.2802 | 127.41 | -9.2099 | 1545.7618 | -9.2978 | 1551.0898 | -9.2978 | 1542.9528 | -9.3329 | 1542.9528 | -9.3329 | 1551.4246 | -9.3505 | 1553.7736 | -9.3505 | 1543.9425
| -9.1571 | 124.0697 | -9.2626 | 127.8642 | -9.1923 | 126.5014 | -9.1923 | 125.9086 | -9.245 | 127.0545 | -9.1747 | 1547.7536 | -9.2626 | 1547.2549 | -9.2626 | 1541.3074 | -9.2978 | 1542.9528 | -9.2978 | 1550.0868 | -9.3153 | 1552.0947 | -9.3153 | 1541.965
| -9.122 | 124.0103 | -9.2275 | 127.5285 | -9.1571 | 126.5014 | -9.1571 | 125.8098 | -9.2099 | 127.1335 | -9.1396 | 1548.0863 | -9.2275 | 1546.4249 | -9.2275 | 1537.869 | -9.2626 | 1542.6233 | -9.2626 | 1549.7528 | -9.2802 | 1551.592 | -9.2802 | 1540.979
| -9.0868 | 123.5155 | -9.1923 | 128.0024 | -9.122 | 127.568 | -9.122 | 125.0191 | -9.1747 | 127.6667 | -9.1044 | 1548.9189 | -9.1923 | 1545.9275 | -9.1923 | 1536.2397 | -9.2275 | 1540.979 | -9.2275 | 1550.0868 | -9.245 | 1550.5881 | -9.245 | 1538.8491
| -9.0517 | 122.9017 | -9.1571 | 128.0024 | -9.0868 | 126.9557 | -9.0868 | 124.7225 | -9.1396 | 127.7655 | -9.0693 | 1551.592 | -9.1571 | 1545.9275 | -9.1571 | 1537.2167 | -9.1923 | 1541.3074 | -9.1923 | 1549.7528 | -9.2099 | 1550.0868 | -9.2099 | 1537.0537
| -9.0165 | 122.0298 | -9.122 | 128.4565 | -9.0517 | 126.9557 | -9.0517 | 124.7422 | -9.1044 | 127.5087 | -9.0341 | 1553.7736 | -9.122 | 1545.5961 | -9.122 | 1538.8491 | -9.1571 | 1541.3074 | -9.1571 | 1549.9198 | -9.1747 | 1550.0868 | -9.1747 | 1536.5652
| -8.9814 | 122.0298 | -9.0868 | 128.3578 | -9.0165 | 125.6912 | -9.0165 | 124.7225 | -9.0693 | 127.489 | -8.999 | 1553.7736 | -9.0868 | 1544.1076 | -9.0868 | 1539.6673 | -9.122 | 1543.2825 | -9.122 | 1550.9225 | -9.1396 | 1551.9271 | -9.1396 | 1535.2646
| -8.9462 | 122.1091 | -9.0517 | 127.5087 | -8.9814 | 125.5924 | -8.9814 | 124.7225 | -9.0341 | 126.6397 | -8.9638 | 1553.7736 | -9.0517 | 1545.4305 | -9.0517 | 1541.4717 | -9.0868 | 1543.6124 | -9.0868 | 1550.4209 | -9.1044 | 1550.5881 | -9.1044 | 1536.5652
| -8.9111 | 122.8423 | -9.0165 | 127.173 | -8.9462 | 125.0982 | -8.9462 | 124.7225 | -8.999 | 126.5607 | -8.9287 | 1553.4374 | -9.0165 | 1542.4587 | -9.0165 | 1541.4717 | -9.0517 | 1546.0932 | -9.0517 | 1550.4209 | -9.0693 | 1553.6055 | -9.0693 | 1535.9145
| -8.8759 | 122.8423 | -8.9814 | 126.7582 | -8.9111 | 125.3157 | -8.9111 | 124.5247 | -8.9638 | 126.4619 | -8.8935 | 1549.4191 | -8.9814 | 1542.1295 | -8.9814 | 1540.4867 | -9.0165 | 1545.7618 | -9.0165 | 1549.7528 | -9.0341 | 1554.11 | -9.0341 | 1535.7519
| -8.8408 | 122.8026 | -8.9462 | 126.5211 | -8.8759 | 125.3157 | -8.8759 | 124.0499 | -8.9287 | 126.5607 | -8.8584 | 1545.9275 | -8.9462 | 1539.5035 | -8.9462 | 1537.0537 | -8.9814 | 1544.1076 | -8.9814 | 1546.9228 | -8.999 | 1554.11 | -8.999 | 1535.2646
| -8.8056 | 122.2875 | -8.9111 | 126.3829 | -8.8408 | 125.3157 | -8.8408 | 123.2581 | -8.8935 | 126.778 | -8.8232 | 1541.3074 | -8.9111 | 1539.5035 | -8.9111 | 1536.728 | -8.9462 | 1543.2825 | -8.9462 | 1542.788 | -8.9638 | 1550.9225 | -8.9638 | 1534.94
| -8.7705 | 122.01 | -8.8759 | 126.3829 | -8.8056 | 124.9795 | -8.8056 | 123.1195 | -8.8584 | 126.8372 | -8.7881 | 1536.5652 | -8.8759 | 1537.7059 | -8.8759 | 1532.9967 | -8.9111 | 1541.1432 | -8.9111 | 1541.8005 | -8.9287 | 1547.4211 | -8.9287 | 1535.5894
| -8.7353 | 122.01 | -8.8408 | 126.7582 | -8.7705 | 124.3467 | -8.7705 | 123.1195 | -8.8232 | 126.7187 | -8.7529 | 1535.2646 | -8.8408 | 1535.9145 | -8.8408 | 1530.578 | -8.8759 | 1537.0537 | -8.8759 | 1541.8005 | -8.8935 | 1543.6124 | -8.8935 | 1535.7519
| -8.7002 | 122.1091 | -8.8056 | 125.8493 | -8.7353 | 123.9707 | -8.7353 | 123.1393 | -8.7881 | 126.3038 | -8.7177 | 1532.9967 | -8.8056 | 1533.9674 | -8.8056 | 1528.0109 | -8.8408 | 1533.4818 | -8.8408 | 1539.5035 | -8.8584 | 1543.4474 | -8.8584 | 1535.5894
| -8.665 | 122.0893 | -8.7705 | 125.8493 | -8.7002 | 123.8916 | -8.7002 | 122.9809 | -8.7529 | 125.9877 | -8.6826 | 1531.3829 | -8.7705 | 1527.691 | -8.7705 | 1524.9799 | -8.8056 | 1530.2564 | -8.8056 | 1538.5222 | -8.8232 | 1540.3227 | -8.8232 | 1534.94
| -8.6299 | 121.9901 | -8.7353 | 126.1458 | -8.665 | 124.2082 | -8.665 | 122.8819 | -8.7177 | 125.9086 | -8.6474 | 1529.7743 | -8.7353 | 1525.6164 | -8.7353 | 1521.4936 | -8.7705 | 1529.1324 | -8.7705 | 1535.427 | -8.7881 | 1539.831 | -8.7881 | 1531.8665
| -8.5947 | 121.772 | -8.7002 | 126.0865 | -8.6299 | 124.584 | -8.6299 | 122.6442 | -8.6826 | 125.6122 | -8.6123 | 1524.8208 | -8.7002 | 1524.0266 | -8.7002 | 1519.6027 | -8.7353 | 1528.0109 | -8.7353 | 1531.8665 | -8.7529 | 1539.3398 | -8.7529 | 1524.0266
| -8.5596 | 121.415 | -8.665 | 125.5133 | -8.5947 | 124.7027 | -8.5947 | 122.5649 | -8.6474 | 125.5529 | -8.5771 | 1523.3921 | -8.665 | 1523.3921 | -8.665 | 1517.4061 | -8.7002 | 1526.892 | -8.7002 | 1529.6138 | -8.7177 | 1534.1294 | -8.7177 | 1522.4419
| -8.5244 | 121.296 | -8.6299 | 125.8493 | -8.5596 | 124.7027 | -8.5596 | 122.5649 | -8.6123 | 125.454 | -8.542 | 1521.4936 | -8.6299 | 1521.178 | -8.6299 | 1513.1986 | -8.665 | 1522.7584 | -8.665 | 1529.2928 | -8.6826 | 1532.3505 | -8.6826 | 1522.4419
| -8.4893 | 121.1372 | -8.5947 | 125.9877 | -8.5244 | 124.2082 | -8.5244 | 122.6442 | -8.5771 | 124.7027 | -8.5068 | 1516.7804 | -8.5947 | 1518.9741 | -8.5947 | 1510.4148 | -8.6299 | 1517.8759 | -8.6299 | 1527.8509 | -8.6474 | 1526.892 | -8.6474 | 1520.7048
| -8.4541 | 120.8197 | -8.5596 | 126.0074 | -8.4893 | 123.7926 | -8.4893 | 122.5649 | -8.542 | 123.6739 | -8.4717 | 1514.4413 | -8.5596 | 1516.9367 | -8.5596 | 1503.9852 | -8.5947 | 1514.7525 | -8.5947 | 1519.4455 | -8.6123 | 1524.8208 | -8.6123 | 1515.9994
| -8.419 | 120.4821 | -8.5244 | 125.9877 | -8.4541 | 123.8322 | -8.4541 | 122.5451 | -8.5068 | 123.3175 | -8.4365 | 1512.7335 | -8.5244 | 1514.5969 | -8.5244 | 1502.0138 | -8.5596 | 1510.569 | -8.5596 | 1517.2496 | -8.5771 | 1523.2336 | -8.5771 | 1514.1303
| -8.3838 | 120.1444 | -8.4893 | 125.9877 | -8.419 | 123.4561 | -8.419 | 122.2082 | -8.4717 | 123.3175 | -8.4014 | 1511.3409 | -8.4893 | 1508.7219 | -8.4893 | 1501.409 | -8.5244 | 1508.8756 | -8.5244 | 1515.8433 | -8.542 | 1518.5031 | -8.542 | 1510.1066
| -8.3486 | 119.4486 | -8.4541 | 125.5924 | -8.3838 | 123.2977 | -8.3838 | 122.1091 | -8.4365 | 124.0895 | -8.3662 | 1508.5684 | -8.4541 | 1497.8018 | -8.4541 | 1500.0512 | -8.4893 | 1506.7294 | -8.4893 | 1512.2688 | -8.5068 | 1516.3116 | -8.5068 | 1508.5684
| -8.3135 | 118.7323 | -8.419 | 124.8016 | -8.3486 | 123.2977 | -8.3486 | 121.9505 | -8.4014 | 124.1686 | -8.3311 | 1507.3416 | -8.419 | 1497.3544 | -8.419 | 1497.9512 | -8.4541 | 1503.5294 | -8.4541 | 1508.2614 | -8.4717 | 1515.8433 | -8.4717 | 1508.1079
| -8.2783 | 118.1547 | -8.3838 | 124.5049 | -8.3135 | 123.0799 | -8.3135 | 121.8117 | -8.3662 | 124.0895 | -8.2959 | 1499.4494 | -8.3838 | 1495.8686 | -8.3838 | 1494.3917 | -8.419 | 1499.4494 | -8.419 | 1507.0354 | -8.4365 | 1514.1303 | -8.4365 | 1501.409
| -8.2432 | 117.6563 | -8.3486 | 124.1488 | -8.2783 | 122.5055 | -8.2783 | 121.6134 | -8.3311 | 124.3862 | -8.2608 | 1496.4619 | -8.3486 | 1492.4846 | -8.3486 | 1491.9008 | -8.3838 | 1494.6864 | -8.3838 | 1502.6194 | -8.4014 | 1509.7985 | -8.4014 | 1499.7501
| -8.208 | 117.6563 | -8.3135 | 124.1488 | -8.2432 | 122.3073 | -8.2432 | 120.5218 | -8.2959 | 124.3862 | -8.2256 | 1493.6565 | -8.3135 | 1489.4343 | -8.3135 | 1489.29 | -8.3486 | 1493.6565 | -8.3486 | 1497.9512 | -8.3662 | 1503.2259 | -8.3662 | 1492.9234
| -8.1729 | 118.2742 | -8.2783 | 124.0103 | -8.208 | 122.5649 | -8.208 | 120.6608 | -8.2608 | 124.3862 | -8.1905 | 1490.447 | -8.2783 | 1489.0016 | -8.2783 | 1485.1399 | -8.3135 | 1492.0466 | -8.3135 | 1497.8018 | -8.3311 | 1502.1651 | -8.3311 | 1491.9008
| -8.1377 | 118.2742 | -8.2432 | 124.0103 | -8.1729 | 121.7919 | -8.1729 | 120.4225 | -8.2256 | 124.4653 | -8.1553 | 1488.282 | -8.2432 | 1489.0016 | -8.2432 | 1482.4572 | -8.2783 | 1488.1383 | -8.2783 | 1494.9814 | -8.2959 | 1499.9006 | -8.2959 | 1490.1572
| -8.1026 | 118.1945 | -8.208 | 123.8916 | -8.1377 | 121.6531 | -8.1377 | 120.1047 | -8.1905 | 124.4258 | -8.1202 | 1487.5644 | -8.208 | 1480.4985 | -8.208 | 1475.2557 | -8.2432 | 1485.7082 | -8.2432 | 1488.4257 | -8.2608 | 1496.1651 | -8.2608 | 1488.8575
| -8.0674 | 117.8955 | -8.1729 | 122.9215 | -8.1026 | 120.8594 | -8.1026 | 119.9854 | -8.1553 | 123.0601 | -8.085 | 1485.8505 | -8.1729 | 1476.8999 | -8.1729 | 1474.4375 | -8.208 | 1485.566 | -8.208 | 1485.2818 | -8.2256 | 1492.3385 | -8.2256 | 1486.706
| -8.0323 | 117.736 | -8.1377 | 121.9703 | -8.0674 | 119.7071 | -8.0674 | 119.9854 | -8.1202 | 122.2677 | -8.0499 | 1480.7774 | -8.1377 | 1475.9395 | -8.1377 | 1473.0794 | -8.1729 | 1478.8313 | -8.1729 | 1482.4572 | -8.1905 | 1489.0016 | -8.1905 | 1482.7383
| -7.9971 | 117.8955 | -8.1026 | 121.5539 | -8.0323 | 119.4685 | -8.0323 | 119.6674 | -8.085 | 122.2677 | -8.0147 | 1478.4162 | -8.1026 | 1474.8463 | -8.1026 | 1471.3244 | -8.1377 | 1476.8999 | -8.1377 | 1480.0807 | -8.1553 | 1483.1604 | -8.1553 | 1475.2557
| -7.962 | 117.5167 | -8.0674 | 121.2563 | -7.9971 | 119.389 | -7.9971 | 119.2298 | -8.0499 | 122.1686 | -7.9795 | 1478.278 | -8.0674 | 1472.944 | -8.0674 | 1470.1161 | -8.1026 | 1474.5737 | -8.1026 | 1478.278 | -8.1202 | 1481.476 | -8.1202 | 1473.3505
| -7.9268 | 117.5964 | -8.0323 | 120.7005 | -7.962 | 119.4685 | -7.962 | 119.0906 | -8.0147 | 122.0893 | -7.9444 | 1477.588 | -8.0323 | 1472.944 | -8.0323 | 1470.6524 | -8.0674 | 1474.5737 | -8.0674 | 1477.3125 | -8.085 | 1478.8313 | -8.085 | 1472.1329
| -7.8917 | 117.7959 | -7.9971 | 120.6013 | -7.9268 | 119.5481 | -7.9268 | 118.2144 | -7.9795 | 122.1686 | -7.9092 | 1477.4502 | -7.9971 | 1472.8087 | -7.9971 | 1468.5134 | -8.0323 | 1471.8631 | -8.0323 | 1476.3507 | -8.0499 | 1475.8026 | -8.0499 | 1469.5808
| -7.8565 | 117.7959 | -7.962 | 120.6013 | -7.8917 | 120.0053 | -7.8917 | 117.756 | -7.9444 | 122.0893 | -7.8741 | 1478.4162 | -7.962 | 1469.4471 | -7.962 | 1463.8924 | -7.9971 | 1471.7283 | -7.9971 | 1473.2149 | -8.0147 | 1473.2149 | -8.0147 | 1465.2048
| -7.8214 | 117.7959 | -7.9268 | 120.7005 | -7.8565 | 120.0053 | -7.8565 | 116.9978 | -7.9092 | 121.5737 | -7.8389 | 1479.5247 | -7.9268 | 1465.3364 | -7.9268 | 1461.6755 | -7.962 | 1470.6524 | -7.962 | 1470.25 | -7.9795 | 1473.7576 | -7.9795 | 1464.8104
| -7.7862 | 118.2543 | -7.8917 | 120.8395 | -7.8214 | 119.8861 | -7.8214 | 116.9978 | -7.8741 | 120.7204 | -7.8038 | 1477.588 | -7.8917 | 1466.7879 | -7.8917 | 1461.0268 | -7.9268 | 1464.6791 | -7.9268 | 1470.921 | -7.9444 | 1473.0794 | -7.9444 | 1464.5478
| -7.7511 | 118.2543 | -7.8565 | 120.8395 | -7.7862 | 119.4287 | -7.7862 | 117.1974 | -7.8389 | 119.4884 | -7.7686 | 1479.1084 | -7.8565 | 1465.0733 | -7.8565 | 1460.6383 | -7.8917 | 1462.7166 | -7.8917 | 1469.4471 | -7.9092 | 1472.8087 | -7.9092 | 1463.6307
| -7.7159 | 118.7124 | -7.8214 | 120.4622 | -7.7511 | 118.8119 | -7.7511 | 116.9778 | -7.8038 | 118.9512 | -7.7335 | 1479.1084 | -7.8214 | 1465.0733 | -7.8214 | 1462.5863 | -7.8565 | 1462.456 | -7.8565 | 1468.2473 | -7.8741 | 1472.944 | -7.8741 | 1461.4159
| -7.6808 | 119.0508 | -7.7862 | 119.9854 | -7.7159 | 118.1547 | -7.7159 | 116.9778 | -7.7686 | 118.3539 | -7.6983 | 1477.3125 | -7.7862 | 1469.18 | -7.7862 | 1464.9418 | -7.8214 | 1462.456 | -7.8214 | 1469.4471 | -7.8389 | 1472.944 | -7.8389 | 1461.1565
| -7.6456 | 118.9911 | -7.7511 | 119.9854 | -7.6808 | 117.8955 | -7.6808 | 117.0177 | -7.7335 | 118.2742 | -7.6632 | 1477.0374 | -7.7511 | 1469.18 | -7.7511 | 1465.0733 | -7.7862 | 1462.847 | -7.7862 | 1472.8087 | -7.8038 | 1471.4589 | -7.8038 | 1461.4159
| -7.6105 | 118.8517 | -7.7159 | 119.9656 | -7.6456 | 117.8955 | -7.6456 | 116.9179 | -7.6983 | 118.4336 | -7.628 | 1476.8999 | -7.7159 | 1469.18 | -7.7159 | 1465.468 | -7.7511 | 1463.108 | -7.7511 | 1470.5182 | -7.7686 | 1471.4589 | -7.7686 | 1461.4159
| -7.5753 | 118.9911 | -7.6808 | 120.045 | -7.6105 | 118.1547 | -7.6105 | 116.9778 | -7.6632 | 118.7124 | -7.5929 | 1475.2557 | -7.6808 | 1469.9821 | -7.6808 | 1465.8633 | -7.7159 | 1463.7615 | -7.7159 | 1470.5182 | -7.7335 | 1470.5182 | -7.7335 | 1463.4999
| -7.5401 | 118.6128 | -7.6456 | 120.4225 | -7.5753 | 118.334 | -7.5753 | 117.0177 | -7.628 | 118.2941 | -7.5577 | 1473.8934 | -7.6456 | 1467.0527 | -7.6456 | 1465.0733 | -7.6808 | 1464.4166 | -7.6808 | 1470.1161 | -7.6983 | 1470.25 | -7.6983 | 1464.6791
| -7.505 | 118.1148 | -7.6105 | 120.502 | -7.5401 | 119.1105 | -7.5401 | 117.2174 | -7.5929 | 118.1347 | -7.5226 | 1470.5182 | -7.6105 | 1467.0527 | -7.6105 | 1463.3692 | -7.6456 | 1463.108 | -7.6456 | 1469.7145 | -7.6632 | 1470.25 | -7.6632 | 1465.468
| -7.4698 | 117.9155 | -7.5753 | 120.641 | -7.505 | 119.1105 | -7.505 | 117.5565 | -7.5577 | 118.2941 | -7.4874 | 1470.1161 | -7.5753 | 1466.7879 | -7.5753 | 1460.2503 | -7.6105 | 1464.0234 | -7.6105 | 1468.2473 | -7.628 | 1470.25 | -7.628 | 1465.468
| -7.4347 | 118.0151 | -7.5401 | 120.7998 | -7.4698 | 119.3293 | -7.4698 | 117.2772 | -7.5226 | 118.573 | -7.4523 | 1468.6466 | -7.5401 | 1464.6791 | -7.5401 | 1459.7338 | -7.5753 | 1464.2855 | -7.5753 | 1466.5234 | -7.5929 | 1467.7157 | -7.5929 | 1463.8924
| -7.3995 | 118.1945 | -7.505 | 120.8594 | -7.4347 | 119.9457 | -7.4347 | 117.2772 | -7.4874 | 118.1746 | -7.4171 | 1468.6466 | -7.505 | 1464.5478 | -7.505 | 1458.3182 | -7.5401 | 1462.7166 | -7.5401 | 1466.5234 | -7.5577 | 1468.2473 | -7.5577 | 1464.5478
| -7.3644 | 118.2543 | -7.4698 | 121.1769 | -7.3995 | 120.4225 | -7.3995 | 117.9553 | -7.4523 | 118.5332 | -7.382 | 1470.5182 | -7.4698 | 1464.9418 | -7.4698 | 1459.7338 | -7.505 | 1460.7678 | -7.505 | 1469.0466 | -7.5226 | 1467.8485 | -7.5226 | 1462.847
| -7.3292 | 118.2742 | -7.4347 | 121.7324 | -7.3644 | 120.502 | -7.3644 | 118.1547 | -7.4171 | 118.5332 | -7.3468 | 1474.4375 | -7.4347 | 1467.9813 | -7.4347 | 1462.3258 | -7.4698 | 1460.7678 | -7.4698 | 1468.2473 | -7.4874 | 1467.7157 | -7.4874 | 1461.0268
| -7.2941 | 118.4535 | -7.3995 | 121.7324 | -7.3292 | 120.3629 | -7.3292 | 118.4336 | -7.382 | 118.4734 | -7.3117 | 1475.9395 | -7.3995 | 1468.1143 | -7.3995 | 1464.5478 | -7.4347 | 1460.7678 | -7.4347 | 1469.5808 | -7.4523 | 1467.8485 | -7.4523 | 1456.6541
| -7.2589 | 118.4535 | -7.3644 | 121.7522 | -7.2941 | 119.9854 | -7.2941 | 119.0906 | -7.3468 | 118.4734 | -7.2765 | 1476.2135 | -7.3644 | 1469.8483 | -7.3644 | 1464.8104 | -7.3995 | 1460.3796 | -7.3995 | 1469.9821 | -7.4171 | 1468.5134 | -7.4171 | 1455.8893
| -7.2238 | 118.3738 | -7.3292 | 121.4745 | -7.2589 | 119.6674 | -7.2589 | 119.2895 | -7.3117 | 118.9313 | -7.2414 | 1476.3507 | -7.3292 | 1469.7145 | -7.3292 | 1465.7315 | -7.3644 | 1461.4159 | -7.3644 | 1470.5182 | -7.382 | 1468.9132 | -7.382 | 1456.6541
| -7.1886 | 118.055 | -7.2941 | 121.3158 | -7.2238 | 119.4089 | -7.2238 | 119.4486 | -7.2765 | 119.3094 | -7.2062 | 1478.0018 | -7.2941 | 1469.7145 | -7.2941 | 1465.7315 | -7.3292 | 1462.0655 | -7.3292 | 1470.5182 | -7.3468 | 1469.3135 | -7.3468 | 1462.7166
| -7.1535 | 118.334 | -7.2589 | 121.3158 | -7.1886 | 119.4089 | -7.1886 | 119.4486 | -7.2414 | 120.2239 | -7.171 | 1479.9415 | -7.2589 | 1469.8483 | -7.2589 | 1466.5234 | -7.2941 | 1462.7166 | -7.2941 | 1471.5936 | -7.3117 | 1470.5182 | -7.3117 | 1465.2048
| -7.1183 | 118.3539 | -7.2238 | 121.296 | -7.1535 | 119.4287 | -7.1535 | 119.2895 | -7.2062 | 120.2835 | -7.1359 | 1483.5833 | -7.2238 | 1472.944 | -7.2238 | 1469.18 | -7.2589 | 1463.6307 | -7.2589 | 1471.5936 | -7.2765 | 1471.0554 | -7.2765 | 1465.2048
| -7.0832 | 118.4535 | -7.1886 | 121.296 | -7.1183 | 119.4486 | -7.1183 | 118.9313 | -7.171 | 120.3431 | -7.1007 | 1485.9929 | -7.1886 | 1474.0293 | -7.1886 | 1473.3505 | -7.2238 | 1464.6791 | -7.2238 | 1476.0765 | -7.2414 | 1472.944 | -7.2414 | 1465.0733
| -7.048 | 118.7124 | -7.1535 | 121.296 | -7.0832 | 119.6077 | -7.0832 | 118.2742 | -7.1359 | 120.2437 | -7.0656 | 1489.29 | -7.1535 | 1478.9698 | -7.1535 | 1475.3923 | -7.1886 | 1471.3244 | -7.1886 | 1478.5545 | -7.2062 | 1475.6658 | -7.2062 | 1465.5997
| -7.0129 | 118.3937 | -7.1183 | 120.7403 | -7.048 | 120.1841 | -7.048 | 117.5765 | -7.1007 | 119.7071 | -7.0304 | 1492.0466 | -7.1183 | 1481.476 | -7.1183 | 1479.5247 | -7.1535 | 1473.8934 | -7.1535 | 1481.1963 | -7.171 | 1477.4502 | -7.171 | 1466.1272
| -6.9777 | 118.7124 | -7.0832 | 120.7403 | -7.0129 | 120.3431 | -7.0129 | 117.1774 | -7.0656 | 119.2696 | -6.9953 | 1494.2445 | -7.0832 | 1484.5727 | -7.0832 | 1480.2198 | -7.1183 | 1476.4879 | -7.1183 | 1481.476 | -7.1359 | 1478.1399 | -7.1359 | 1466.7879
| -6.9426 | 118.3937 | -7.048 | 120.7403 | -6.9777 | 120.3431 | -6.9777 | 117.357 | -7.0304 | 119.1502 | -6.9601 | 1497.8018 | -7.048 | 1485.2818 | -7.048 | 1481.896 | -7.0832 | 1478.1399 | -7.0832 | 1482.1765 | -7.1007 | 1479.5247 | -7.1007 | 1470.921
| -6.9074 | 118.1148 | -7.0129 | 120.7403 | -6.9426 | 119.9258 | -6.9426 | 117.6164 | -6.9953 | 119.2298 | -6.925 | 1497.9512 | -7.0129 | 1487.4211 | -7.0129 | 1484.7144 | -7.048 | 1479.5247 | -7.048 | 1483.0196 | -7.0656 | 1480.6379 | -7.0656 | 1473.7576
| -6.8723 | 118.0949 | -6.9777 | 120.8395 | -6.9074 | 119.6276 | -6.9074 | 118.3539 | -6.9601 | 119.2696 | -6.8898 | 1500.3526 | -6.9777 | 1494.0974 | -6.9777 | 1486.5632 | -7.0129 | 1481.0566 | -7.0129 | 1485.8505 | -7.0304 | 1481.476 | -7.0304 | 1475.9395
| -6.8371 | 118.1148 | -6.9426 | 120.4821 | -6.8723 | 119.2696 | -6.8723 | 118.2543 | -6.925 | 119.4685 | -6.8547 | 1499.9006 | -6.9426 | 1494.9814 | -6.9426 | 1488.4257 | -6.9777 | 1484.5727 | -6.9777 | 1487.9947 | -6.9953 | 1482.3168 | -6.9953 | 1478.0018
| -6.802 | 118.055 | -6.9074 | 120.2835 | -6.8371 | 119.2696 | -6.8371 | 118.4535 | -6.8898 | 119.7071 | -6.8195 | 1497.9512 | -6.9074 | 1494.9814 | -6.9074 | 1487.8512 | -6.9426 | 1486.4205 | -6.9426 | 1492.4846 | -6.9601 | 1485.9929 | -6.9601 | 1479.1084
| -6.7668 | 116.9179 | -6.8723 | 120.1245 | -6.802 | 119.2099 | -6.802 | 118.2543 | -6.8547 | 120.1047 | -6.7844 | 1497.2054 | -6.8723 | 1492.6308 | -6.8723 | 1487.8512 | -6.9074 | 1486.8489 | -6.9074 | 1492.4846 | -6.925 | 1487.2779 | -6.925 | 1481.3361
| -6.7316 | 115.8786 | -6.8371 | 120.2835 | -6.7668 | 119.5878 | -6.7668 | 117.9155 | -6.8195 | 119.6674 | -6.7492 | 1497.6526 | -6.8371 | 1490.1572 | -6.8371 | 1487.8512 | -6.8723 | 1486.9918 | -6.8723 | 1492.6308 | -6.8898 | 1489.0016 | -6.8898 | 1481.896
| -6.6965 | 115.3782 | -6.802 | 120.4027 | -6.7316 | 119.5282 | -6.7316 | 117.9155 | -6.7844 | 120.045 | -6.7141 | 1494.0974 | -6.802 | 1488.8575 | -6.802 | 1484.8561 | -6.8371 | 1484.4311 | -6.8371 | 1489.0016 | -6.8547 | 1488.8575 | -6.8547 | 1483.3013
| -6.6613 | 115.0377 | -6.7668 | 120.4027 | -6.6965 | 119.5282 | -6.6965 | 118.1148 | -6.7492 | 120.0053 | -6.6789 | 1496.0168 | -6.7668 | 1487.7077 | -6.7668 | 1482.5977 | -6.802 | 1480.9169 | -6.802 | 1489.0016 | -6.8195 | 1489.4343 | -6.8195 | 1484.1482
| -6.6262 | 114.9775 | -6.7316 | 120.5615 | -6.6613 | 119.2099 | -6.6613 | 118.1148 | -6.7141 | 119.3691 | -6.6438 | 1495.2768 | -6.7316 | 1489.5787 | -6.7316 | 1488.1383 | -6.7668 | 1480.9169 | -6.7668 | 1487.9947 | -6.7844 | 1486.1354 | -6.7844 | 1484.1482
| -6.591 | 114.9575 | -6.6965 | 120.4225 | -6.6262 | 119.4685 | -6.6262 | 117.6164 | -6.6789 | 119.3691 | -6.6086 | 1496.9077 | -6.6965 | 1490.1572 | -6.6965 | 1489.7232 | -6.7316 | 1480.4985 | -6.7316 | 1488.1383 | -6.7492 | 1483.8656 | -6.7492 | 1483.0196
| -6.5559 | 114.396 | -6.6613 | 120.3629 | -6.591 | 118.4535 | -6.591 | 117.6563 | -6.6438 | 119.4685 | -6.5735 | 1493.6565 | -6.6613 | 1488.7135 | -6.6613 | 1488.7135 | -6.6965 | 1480.4985 | -6.6965 | 1488.4257 | -6.7141 | 1483.3013 | -6.7141 | 1480.0807
| -6.5207 | 114.2756 | -6.6262 | 120.204 | -6.5559 | 118.4734 | -6.5559 | 117.5565 | -6.6086 | 119.4486 | -6.5383 | 1493.6565 | -6.6262 | 1489.8678 | -6.6262 | 1489.8678 | -6.6613 | 1481.6159 | -6.6613 | 1488.4257 | -6.6789 | 1486.2779 | -6.6789 | 1482.0362
| -6.4856 | 113.4923 | -6.591 | 119.9059 | -6.5207 | 118.5132 | -6.5207 | 117.0577 | -6.5735 | 119.2497 | -6.5032 | 1493.6565 | -6.591 | 1489.8678 | -6.591 | 1490.3021 | -6.6262 | 1484.0069 | -6.6262 | 1488.4257 | -6.6438 | 1490.0125 | -6.6438 | 1480.0807
| -6.4504 | 113.2309 | -6.5559 | 119.7668 | -6.4856 | 118.5132 | -6.4856 | 117.0577 | -6.5383 | 119.1303 | -6.468 | 1493.363 | -6.5559 | 1489.0016 | -6.5559 | 1488.282 | -6.591 | 1484.8561 | -6.591 | 1489.0016 | -6.6086 | 1491.755 | -6.6086 | 1480.0807
| -6.4153 | 113.251 | -6.5207 | 119.4884 | -6.4504 | 117.736 | -6.4504 | 117.0377 | -6.5032 | 119.0508 | -6.4329 | 1492.4846 | -6.5207 | 1489.0016 | -6.5207 | 1485.4239 | -6.5559 | 1486.9918 | -6.5559 | 1490.8822 | -6.5735 | 1492.3385 | -6.5735 | 1480.6379
| -6.3801 | 113.2912 | -6.4856 | 118.9712 | -6.4153 | 117.1974 | -6.4153 | 116.7182 | -6.468 | 118.9512 | -6.3977 | 1491.9008 | -6.4856 | 1489.0016 | -6.4856 | 1487.5644 | -6.5207 | 1487.4211 | -6.5207 | 1493.0699 | -6.5383 | 1493.2164 | -6.5383 | 1481.3361
| -6.345 | 113.2711 | -6.4504 | 118.3738 | -6.3801 | 117.1176 | -6.3801 | 116.0586 | -6.4329 | 118.6526 | -6.3625 | 1491.755 | -6.4504 | 1487.2779 | -6.4504 | 1487.4211 | -6.4856 | 1486.9918 | -6.4856 | 1491.6093 | -6.5032 | 1491.3182 | -6.5032 | 1480.3591
| -6.3098 | 113.7937 | -6.4153 | 117.9553 | -6.345 | 116.858 | -6.345 | 115.6985 | -6.3977 | 118.9512 | -6.3274 | 1489.7232 | -6.4153 | 1488.7135 | -6.4153 | 1485.7082 | -6.4504 | 1487.2779 | -6.4504 | 1490.7371 | -6.468 | 1489.8678 | -6.468 | 1479.9415
| -6.2747 | 114.396 | -6.3801 | 117.5765 | -6.3098 | 116.3385 | -6.3098 | 115.4783 | -6.3625 | 118.9712 | -6.2922 | 1489.7232 | -6.3801 | 1488.7135 | -6.3801 | 1485.7082 | -6.4153 | 1486.4205 | -6.4153 | 1490.1572 | -6.4329 | 1489.1457 | -6.4329 | 1479.5247
| -6.2395 | 114.717 | -6.345 | 117.0377 | -6.2747 | 115.4583 | -6.2747 | 115.0777 | -6.3274 | 118.8318 | -6.2571 | 1489.7232 | -6.345 | 1488.282 | -6.345 | 1484.8561 | -6.3801 | 1485.9929 | -6.3801 | 1489.0016 | -6.3977 | 1488.4257 | -6.3977 | 1479.9415
| -6.2044 | 114.717 | -6.3098 | 116.4984 | -6.2395 | 114.5164 | -6.2395 | 114.9575 | -6.2922 | 118.3539 | -6.2219 | 1486.706 | -6.3098 | 1488.282 | -6.3098 | 1478.0018 | -6.345 | 1484.8561 | -6.345 | 1486.4205 | -6.3625 | 1486.9918 | -6.3625 | 1479.3858
| -6.1692 | 114.717 | -6.2747 | 116.2786 | -6.2044 | 114.3559 | -6.2044 | 114.115 | -6.2571 | 118.314 | -6.1868 | 1483.8656 | -6.2747 | 1488.282 | -6.2747 | 1474.8463 | -6.3098 | 1481.476 | -6.3098 | 1485.9929 | -6.3274 | 1485.7082 | -6.3274 | 1479.1084
| -6.1341 | 113.9544 | -6.2395 | 115.5384 | -6.1692 | 114.4562 | -6.1692 | 113.9544 | -6.2219 | 117.9753 | -6.1516 | 1483.7244 | -6.2395 | 1485.4239 | -6.2395 | 1473.0794 | -6.2747 | 1477.8638 | -6.2747 | 1483.4423 | -6.2922 | 1485.7082 | -6.2922 | 1478.278
| -6.0989 | 113.5325 | -6.2044 | 115.4183 | -6.1341 | 114.8974 | -6.1341 | 113.9142 | -6.1868 | 117.6563 | -6.1165 | 1482.4572 | -6.2044 | 1478.8313 | -6.2044 | 1472.268 | -6.2395 | 1476.3507 | -6.2395 | 1481.1963 | -6.2571 | 1483.0196 | -6.2571 | 1476.2135
| -6.0638 | 113.854 | -6.1692 | 115.6585 | -6.0989 | 115.7185 | -6.0989 | 113.0095 | -6.1516 | 117.1575 | -6.0813 | 1480.2198 | -6.1692 | 1471.5936 | -6.1692 | 1471.998 | -6.2044 | 1473.6218 | -6.2044 | 1479.5247 | -6.2219 | 1481.7559 | -6.2219 | 1475.9395
| -6.0286 | 113.9544 | -6.1341 | 115.4383 | -6.0638 | 115.7385 | -6.0638 | 112.7881 | -6.1165 | 116.6383 | -6.0462 | 1479.1084 | -6.1341 | 1470.1161 | -6.1341 | 1470.3841 | -6.1692 | 1471.998 | -6.1692 | 1479.5247 | -6.1868 | 1479.2471 | -6.1868 | 1475.529
| -5.9935 | 113.6531 | -6.0989 | 115.4383 | -6.0286 | 115.7385 | -6.0286 | 112.7277 | -6.0813 | 115.9186 | -6.011 | 1477.3125 | -6.0989 | 1468.6466 | -6.0989 | 1468.9132 | -6.1341 | 1471.7283 | -6.1341 | 1477.3125 | -6.1516 | 1476.8999 | -6.1516 | 1475.2557
| -5.9583 | 113.251 | -6.0638 | 115.8386 | -5.9935 | 115.6185 | -5.9935 | 112.3853 | -6.0462 | 115.0377 | -5.9759 | 1477.3125 | -6.0638 | 1468.5134 | -6.0638 | 1467.5829 | -6.0989 | 1471.7283 | -6.0989 | 1473.4861 | -6.1165 | 1474.5737 | -6.1165 | 1474.4375
| -5.9231 | 113.5928 | -6.0286 | 115.8386 | -5.9583 | 115.2981 | -5.9583 | 112.3853 | -6.011 | 114.4963 | -5.9407 | 1478.6929 | -6.0286 | 1468.6466 | -6.0286 | 1466.9203 | -6.0638 | 1470.6524 | -6.0638 | 1473.4861 | -6.0813 | 1474.5737 | -6.0813 | 1469.18
| -5.888 | 112.6069 | -5.9935 | 115.4383 | -5.9231 | 114.7971 | -5.9231 | 113.1906 | -5.9759 | 113.9946 | -5.9056 | 1478.278 | -5.9935 | 1469.0466 | -5.9935 | 1466.6556 | -6.0286 | 1469.9821 | -6.0286 | 1472.5382 | -6.0462 | 1473.3505 | -6.0462 | 1467.5829
| -5.8528 | 112.345 | -5.9583 | 115.8186 | -5.888 | 114.7571 | -5.888 | 112.6472 | -5.9407 | 114.5365 | -5.8704 | 1479.5247 | -5.9583 | 1469.4471 | -5.9583 | 1466.9203 | -5.9935 | 1470.1161 | -5.9935 | 1471.5936 | -6.011 | 1473.7576 | -6.011 | 1465.9952
| -5.8177 | 112.103 | -5.9231 | 115.3782 | -5.8528 | 113.8339 | -5.8528 | 112.7277 | -5.9056 | 114.4963 | -5.8353 | 1479.8025 | -5.9231 | 1470.7867 | -5.9231 | 1468.7799 | -5.9583 | 1470.5182 | -5.9583 | 1471.5936 | -5.9759 | 1473.7576 | -5.9759 | 1466.9203
| -5.7825 | 111.6793 | -5.888 | 115.3782 | -5.8177 | 113.5727 | -5.8177 | 112.768 | -5.8704 | 114.5365 | -5.8001 | 1479.5247 | -5.888 | 1471.998 | -5.888 | 1467.5829 | -5.9231 | 1466.9203 | -5.9231 | 1471.998 | -5.9407 | 1470.7867 | -5.9407 | 1466.2592
| -5.7474 | 111.4572 | -5.8528 | 115.3782 | -5.7825 | 114.8773 | -5.7825 | 112.6472 | -5.8353 | 114.6768 | -5.765 | 1473.6218 | -5.8528 | 1469.4471 | -5.8528 | 1466.6556 | -5.888 | 1468.3803 | -5.888 | 1471.998 | -5.9056 | 1472.8087 | -5.9056 | 1466.2592
| -5.7122 | 111.0531 | -5.8177 | 115.3782 | -5.7474 | 114.8773 | -5.7474 | 112.5666 | -5.8001 | 115.218 | -5.7298 | 1472.944 | -5.8177 | 1467.5829 | -5.8177 | 1466.6556 | -5.8528 | 1468.6466 | -5.8528 | 1471.998 | -5.8704 | 1472.8087 | -5.8704 | 1466.2592
| -5.6771 | 110.8306 | -5.7825 | 115.3782 | -5.7122 | 114.3358 | -5.7122 | 112.4256 | -5.765 | 115.1579 | -5.6947 | 1471.4589 | -5.7825 | 1466.1272 | -5.7825 | 1465.9952 | -5.8177 | 1468.3803 | -5.8177 | 1472.8087 | -5.8353 | 1472.8087 | -5.8353 | 1466.9203
| -5.6419 | 110.6283 | -5.7474 | 115.0577 | -5.6771 | 114.5565 | -5.6771 | 112.7076 | -5.7298 | 115.1579 | -5.6595 | 1471.3244 | -5.7474 | 1462.9775 | -5.7474 | 1463.2386 | -5.7825 | 1467.8485 | -5.7825 | 1472.1329 | -5.8001 | 1472.1329 | -5.8001 | 1467.9813
| -5.6068 | 110.5474 | -5.7122 | 114.9375 | -5.6419 | 114.5565 | -5.6419 | 112.8888 | -5.6947 | 114.2957 | -5.6244 | 1471.3244 | -5.7122 | 1463.3692 | -5.7122 | 1459.7338 | -5.7474 | 1467.8485 | -5.7474 | 1470.25 | -5.765 | 1470.3841 | -5.765 | 1466.7879
| -5.5716 | 109.9599 | -5.6771 | 114.9375 | -5.6068 | 113.7535 | -5.6068 | 112.8485 | -5.6595 | 114.8372 | -5.5892 | 1470.6524 | -5.6771 | 1463.3692 | -5.6771 | 1457.8052 | -5.7122 | 1464.5478 | -5.7122 | 1468.2473 | -5.7298 | 1469.0466 | -5.7298 | 1466.5234
| -5.5365 | 109.9599 | -5.6419 | 114.6768 | -5.5716 | 113.251 | -5.5716 | 112.7277 | -5.6244 | 114.717 | -5.554 | 1470.5182 | -5.6419 | 1462.847 | -5.6419 | 1457.421 | -5.6771 | 1463.3692 | -5.6771 | 1466.1272 | -5.6947 | 1466.3913 | -5.6947 | 1466.1272
| -5.5013 | 109.6963 | -5.6068 | 114.396 | -5.5365 | 113.0498 | -5.5365 | 112.8485 | -5.5892 | 114.6768 | -5.5189 | 1468.2473 | -5.6068 | 1462.5863 | -5.6068 | 1457.421 | -5.6419 | 1462.1956 | -5.6419 | 1464.1544 | -5.6595 | 1463.8924 | -5.6595 | 1460.2503
| -5.4662 | 109.2497 | -5.5716 | 114.3559 | -5.5013 | 112.7277 | -5.5013 | 112.2442 | -5.554 | 114.6768 | -5.4837 | 1465.3364 | -5.5716 | 1463.7615 | -5.5716 | 1456.5265 | -5.6068 | 1459.7338 | -5.6068 | 1463.2386 | -5.6244 | 1463.108 | -5.6244 | 1458.0616
| -5.431 | 109.2497 | -5.5365 | 113.9343 | -5.4662 | 112.7076 | -5.4662 | 111.6793 | -5.5189 | 114.8773 | -5.4486 | 1463.108 | -5.5365 | 1462.9775 | -5.5365 | 1455.2534 | -5.5716 | 1458.5751 | -5.5716 | 1462.5863 | -5.5892 | 1463.108 | -5.5892 | 1458.3182
| -5.3959 | 109.0465 | -5.5013 | 113.9343 | -5.431 | 112.5062 | -5.431 | 111.0531 | -5.4837 | 114.6768 | -5.4134 | 1461.6755 | -5.5013 | 1462.3258 | -5.5013 | 1454.7456 | -5.5365 | 1457.0373 | -5.5365 | 1459.3471 | -5.554 | 1462.7166 | -5.554 | 1458.0616
| -5.3607 | 109.0465 | -5.4662 | 113.5928 | -5.3959 | 112.4256 | -5.3959 | 110.77 | -5.4486 | 114.396 | -5.3783 | 1461.5457 | -5.4662 | 1460.8973 | -5.4662 | 1453.2275 | -5.5013 | 1456.2714 | -5.5013 | 1458.8322 | -5.5189 | 1462.7166 | -5.5189 | 1458.0616
| -5.3256 | 109.2091 | -5.431 | 113.7133 | -5.3607 | 112.2442 | -5.3607 | 110.77 | -5.4134 | 114.095 | -5.3431 | 1461.5457 | -5.431 | 1458.1899 | -5.431 | 1450.3387 | -5.4662 | 1454.9994 | -5.4662 | 1458.8322 | -5.4837 | 1462.0655 | -5.4837 | 1457.9334
| -5.2904 | 109.3715 | -5.3959 | 113.6129 | -5.3256 | 111.5178 | -5.3256 | 110.5271 | -5.3783 | 113.5928 | -5.308 | 1462.3258 | -5.3959 | 1456.0166 | -5.3959 | 1449.091 | -5.431 | 1454.8725 | -5.431 | 1458.1899 | -5.4486 | 1461.5457 | -5.4486 | 1456.6541
| -5.2553 | 109.8585 | -5.3607 | 113.3113 | -5.2904 | 110.8913 | -5.2904 | 110.3246 | -5.3431 | 113.251 | -5.2728 | 1462.3258 | -5.3607 | 1454.8725 | -5.3607 | 1448.4689 | -5.3959 | 1454.8725 | -5.3959 | 1458.1899 | -5.4134 | 1460.8973 | -5.4134 | 1455.762
| -5.2201 | 110.2841 | -5.3256 | 112.0829 | -5.2553 | 110.5878 | -5.2553 | 109.7165 | -5.308 | 113.1906 | -5.2377 | 1458.7036 | -5.3256 | 1452.9752 | -5.3256 | 1447.4762 | -5.3607 | 1454.3654 | -5.3607 | 1456.144 | -5.3783 | 1459.7338 | -5.3783 | 1454.7456
| -5.1849 | 110.1626 | -5.2904 | 112.1434 | -5.2201 | 110.1423 | -5.2201 | 109.0872 | -5.2728 | 113.0297 | -5.2025 | 1458.0616 | -5.2904 | 1451.9682 | -5.2904 | 1446.3629 | -5.3256 | 1453.8591 | -5.3256 | 1454.4921 | -5.3431 | 1456.9095 | -5.3431 | 1453.3537
| -5.1498 | 109.818 | -5.2553 | 111.7803 | -5.1849 | 109.5136 | -5.1849 | 108.7619 | -5.2377 | 111.982 | -5.1674 | 1455.2534 | -5.2553 | 1449.2155 | -5.2553 | 1447.4762 | -5.2904 | 1451.8425 | -5.2904 | 1454.4921 | -5.308 | 1455.6347 | -5.308 | 1448.7176
| -5.1146 | 109.5745 | -5.2201 | 111.6793 | -5.1498 | 110.1423 | -5.1498 | 108.5992 | -5.2025 | 111.336 | -5.1322 | 1453.6063 | -5.2201 | 1447.6001 | -5.2201 | 1445.8693 | -5.2553 | 1448.4689 | -5.2553 | 1454.4921 | -5.2728 | 1454.6188 | -5.2728 | 1445.9926
| -5.0795 | 109.5745 | -5.1849 | 111.639 | -5.1146 | 109.7977 | -5.1146 | 108.5992 | -5.1674 | 111.336 | -5.0971 | 1450.8392 | -5.1849 | 1445.2533 | -5.1849 | 1444.2702 | -5.2201 | 1445.8693 | -5.2201 | 1452.4713 | -5.2377 | 1453.1013 | -5.2377 | 1443.5346
| -5.0443 | 108.7823 | -5.1498 | 111.639 | -5.0795 | 109.5136 | -5.0795 | 108.4364 | -5.1322 | 111.1137 | -5.0619 | 1450.8392 | -5.1498 | 1445.2533 | -5.1498 | 1443.2898 | -5.1849 | 1444.1475 | -5.1849 | 1449.2155 | -5.2025 | 1451.5914 | -5.2025 | 1442.1902
| -5.0092 | 108.192 | -5.1146 | 111.2956 | -5.0443 | 110.2436 | -5.0443 | 108.4364 | -5.0971 | 110.952 | -5.0268 | 1450.0887 | -5.1146 | 1444.3929 | -5.1146 | 1440.6078 | -5.1498 | 1439.8799 | -5.1498 | 1445.746 | -5.1674 | 1448.2204 | -5.1674 | 1442.5563
| -4.974 | 107.9271 | -5.0795 | 111.861 | -5.0092 | 110.4056 | -5.0092 | 108.3346 | -5.0619 | 110.9924 | -4.9916 | 1449.091 | -5.0795 | 1442.0682 | -5.0795 | 1436.7417 | -5.1146 | 1439.8799 | -5.1146 | 1443.9022 | -5.1322 | 1447.3523 | -5.1322 | 1441.3373
| -4.9389 | 107.8864 | -5.0443 | 112.103 | -4.974 | 110.9924 | -4.974 | 108.3346 | -5.0268 | 110.8104 | -4.9565 | 1448.2204 | -5.0443 | 1439.8799 | -5.0443 | 1436.3812 | -5.0795 | 1437.2228 | -5.0795 | 1442.4342 | -5.0971 | 1446.4864 | -5.0971 | 1439.5164
| -4.9037 | 107.7028 | -5.0092 | 112.1232 | -4.9389 | 111.336 | -4.9389 | 108.4974 | -4.9916 | 110.7093 | -4.9213 | 1446.4864 | -5.0092 | 1439.5164 | -5.0092 | 1435.6613 | -5.0443 | 1436.9822 | -5.0443 | 1442.1902 | -5.0619 | 1444.8843 | -5.0619 | 1438.4282
| -4.8686 | 107.7028 | -4.974 | 112.345 | -4.9037 | 111.3158 | -4.9037 | 108.5992 | -4.9565 | 110.77 | -4.8862 | 1444.8843 | -4.974 | 1438.1868 | -4.974 | 1434.9427 | -5.0092 | 1434.7035 | -5.0092 | 1440.2437 | -5.0268 | 1442.0682 | -5.0268 | 1437.2228
| -4.8334 | 108.355 | -4.9389 | 112.3046 | -4.8686 | 110.9924 | -4.8686 | 108.9043 | -4.9213 | 110.7497 | -4.851 | 1443.5346 | -4.9389 | 1436.8619 | -4.9389 | 1434.2254 | -4.974 | 1434.4643 | -4.974 | 1438.9115 | -4.9916 | 1439.3954 | -4.9916 | 1433.9865
| -4.7983 | 110.122 | -4.9037 | 112.0022 | -4.8334 | 110.4056 | -4.8334 | 109.0465 | -4.8862 | 110.6688 | -4.8159 | 1441.9463 | -4.9037 | 1436.8619 | -4.9037 | 1433.9865 | -4.9389 | 1433.5093 | -4.9389 | 1438.9115 | -4.9565 | 1439.3954 | -4.9565 | 1432.1996
| -4.7631 | 111.7197 | -4.8686 | 111.8408 | -4.7983 | 109.4324 | -4.7983 | 109.1278 | -4.851 | 110.6688 | -4.7807 | 1442.0682 | -4.8686 | 1436.8619 | -4.8686 | 1432.0807 | -4.9037 | 1433.6285 | -4.9037 | 1438.9115 | -4.9213 | 1437.825 | -4.9213 | 1430.8939
| -4.728 | 112.345 | -4.8334 | 111.5582 | -4.7631 | 109.1888 | -4.7631 | 108.9449 | -4.8159 | 110.6688 | -4.7455 | 1441.3373 | -4.8334 | 1436.8619 | -4.8334 | 1432.0807 | -4.8686 | 1433.2708 | -4.8686 | 1437.825 | -4.8862 | 1437.825 | -4.8862 | 1430.42
| -4.6928 | 112.345 | -4.7983 | 110.6081 | -4.728 | 109.1888 | -4.728 | 108.9246 | -4.7807 | 110.6688 | -4.7104 | 1439.0324 | -4.7983 | 1435.1821 | -4.7983 | 1431.9619 | -4.8334 | 1433.9865 | -4.8334 | 1437.825 | -4.851 | 1437.7045 | -4.851 | 1428.6474
| -4.6577 | 112.1635 | -4.7631 | 110.6688 | -4.6928 | 108.8839 | -4.6928 | 108.8839 | -4.7455 | 110.8104 | -4.6752 | 1437.3432 | -4.7631 | 1433.1517 | -4.7631 | 1430.065 | -4.7983 | 1433.2708 | -4.7983 | 1436.9822 | -4.8159 | 1436.5013 | -4.8159 | 1428.6474
| -4.6225 | 111.8004 | -4.728 | 111.2552 | -4.6577 | 108.7619 | -4.6577 | 108.0902 | -4.7104 | 111.5784 | -4.6401 | 1436.8619 | -4.728 | 1430.1833 | -4.728 | 1429.0014 | -4.7631 | 1432.9135 | -4.7631 | 1433.6285 | -4.7807 | 1436.5013 | -4.7807 | 1428.6474
| -4.5874 | 111.3764 | -4.6928 | 110.3651 | -4.6225 | 108.9652 | -4.6225 | 107.4988 | -4.6752 | 110.5676 | -4.6049 | 1436.2612 | -4.6928 | 1429.3557 | -4.6928 | 1424.5364 | -4.728 | 1428.4116 | -4.728 | 1431.2496 | -4.7455 | 1435.5415 | -4.7455 | 1427.1166
| -4.5522 | 111.2956 | -4.6577 | 109.9599 | -4.5874 | 109.27 | -4.5874 | 107.4988 | -4.6401 | 110.2639 | -4.5698 | 1435.1821 | -4.6577 | 1427.2341 | -4.6577 | 1421.6193 | -4.6928 | 1427.4694 | -4.6928 | 1429.9467 | -4.7104 | 1432.7944 | -4.7104 | 1425.0046
| -4.5171 | 111.3966 | -4.6225 | 110.9116 | -4.5522 | 109.818 | -4.5522 | 107.4988 | -4.6049 | 109.6963 | -4.5346 | 1435.6613 | -4.6225 | 1426.2942 | -4.6225 | 1420.6888 | -4.6577 | 1425.2388 | -4.6577 | 1428.0581 | -4.6752 | 1427.9403 | -4.6752 | 1422.901
| -4.4819 | 111.2956 | -4.5874 | 110.9116 | -4.5171 | 112.103 | -4.5171 | 107.4988 | -4.5698 | 109.1684 | -4.4995 | 1435.9011 | -4.5874 | 1425.356 | -4.5874 | 1420.9213 | -4.6225 | 1422.4346 | -4.6225 | 1425.5904 | -4.6401 | 1424.7704 | -4.6401 | 1420.9213
| -4.4468 | 111.6591 | -4.5522 | 112.6673 | -4.4819 | 112.2038 | -4.4819 | 107.7844 | -4.5346 | 108.7619 | -4.4643 | 1435.9011 | -4.5522 | 1423.8349 | -4.5522 | 1421.2702 | -4.5874 | 1422.4346 | -4.5874 | 1424.6534 | -4.6049 | 1422.085 | -4.6049 | 1418.6003
| -4.4116 | 111.6591 | -4.5171 | 114.396 | -4.4468 | 112.627 | -4.4468 | 108.3346 | -4.4995 | 108.9449 | -4.4292 | 1437.3432 | -4.5171 | 1426.2942 | -4.5171 | 1423.3677 | -4.5522 | 1424.4194 | -4.5522 | 1424.6534 | -4.5698 | 1422.085 | -4.5698 | 1416.8647
| -4.3764 | 112.0829 | -4.4819 | 114.9575 | -4.4116 | 112.8888 | -4.4116 | 109.7165 | -4.4643 | 109.1684 | -4.394 | 1438.9115 | -4.4819 | 1426.999 | -4.4819 | 1424.1856 | -4.5171 | 1424.6534 | -4.5171 | 1426.999 | -4.5346 | 1423.8349 | -4.5346 | 1418.6003
| -4.3413 | 112.3248 | -4.4468 | 115.0176 | -4.3764 | 112.8687 | -4.3764 | 110.4866 | -4.4292 | 108.8636 | -4.3589 | 1440.0011 | -4.4468 | 1430.065 | -4.4468 | 1425.8249 | -4.4819 | 1426.1768 | -4.4819 | 1428.1759 | -4.4995 | 1425.4732 | -4.4995 | 1418.9479
| -4.3061 | 112.5867 | -4.4116 | 114.8573 | -4.3413 | 112.627 | -4.3413 | 111.0733 | -4.394 | 109.7977 | -4.3237 | 1440.0011 | -4.4116 | 1428.6474 | -4.4116 | 1428.4116 | -4.4468 | 1427.2341 | -4.4468 | 1429.2376 | -4.4643 | 1427.4694 | -4.4643 | 1420.2241
| -4.271 | 112.7076 | -4.3764 | 114.717 | -4.3061 | 112.5867 | -4.3061 | 111.1137 | -4.3589 | 110.041 | -4.2886 | 1442.4342 | -4.3764 | 1430.065 | -4.3764 | 1428.8834 | -4.4116 | 1428.5295 | -4.4116 | 1431.8431 | -4.4292 | 1429.8284 | -4.4292 | 1422.318
| -4.2358 | 113.1906 | -4.3413 | 114.3559 | -4.271 | 112.5867 | -4.271 | 111.134 | -4.3237 | 110.7497 | -4.2534 | 1444.3929 | -4.3413 | 1431.4869 | -4.3413 | 1429.3557 | -4.3764 | 1429.8284 | -4.3764 | 1430.1833 | -4.394 | 1430.42 | -4.394 | 1425.2388
| -4.2007 | 113.5124 | -4.3061 | 114.0147 | -4.2358 | 112.7479 | -4.2358 | 111.0733 | -4.2886 | 111.982 | -4.2183 | 1444.0248 | -4.3061 | 1430.065 | -4.3061 | 1430.8939 | -4.3413 | 1429.8284 | -4.3413 | 1431.8431 | -4.3589 | 1430.42 | -4.3589 | 1426.2942
| -4.1655 | 113.5526 | -4.271 | 114.0147 | -4.2007 | 113.2107 | -4.2007 | 110.5676 | -4.2534 | 112.2643 | -4.1831 | 1443.6571 | -4.271 | 1431.4869 | -4.271 | 1430.1833 | -4.3061 | 1430.42 | -4.3061 | 1434.3448 | -4.3237 | 1431.8431 | -4.3237 | 1424.8875
| -4.1304 | 113.2309 | -4.2358 | 113.9343 | -4.1655 | 113.4521 | -4.1655 | 110.1423 | -4.2183 | 112.4457 | -4.148 | 1443.2898 | -4.2358 | 1434.3448 | -4.2358 | 1431.9619 | -4.271 | 1430.42 | -4.271 | 1430.5384 | -4.2886 | 1431.8431 | -4.2886 | 1426.2942
| -4.0952 | 112.8284 | -4.2007 | 114.2756 | -4.1304 | 113.4521 | -4.1304 | 110.1423 | -4.1831 | 112.6069 | -4.1128 | 1441.8244 | -4.2007 | 1434.1059 | -4.2007 | 1430.8939 | -4.2358 | 1429.2376 | -4.2358 | 1431.3682 | -4.2534 | 1432.0807 | -4.2534 | 1427.7048
| -4.0601 | 112.3046 | -4.1655 | 114.4161 | -4.0952 | 113.6129 | -4.0952 | 110.3854 | -4.148 | 112.4659 | -4.0777 | 1440.6078 | -4.1655 | 1433.6285 | -4.1655 | 1429.9467 | -4.2007 | 1428.2937 | -4.2007 | 1431.8431 | -4.2183 | 1432.0807 | -4.2183 | 1424.8875
| -4.0249 | 111.336 | -4.1304 | 114.4161 | -4.0601 | 113.6732 | -4.0601 | 111.1137 | -4.1128 | 112.486 | -4.0425 | 1438.549 | -4.1304 | 1431.7243 | -4.1304 | 1429.4738 | -4.1655 | 1428.2937 | -4.1655 | 1431.8431 | -4.1831 | 1432.5564 | -4.1831 | 1425.8249
| -3.9898 | 110.4866 | -4.0952 | 114.4161 | -4.0249 | 113.6732 | -4.0249 | 111.336 | -4.0777 | 112.4659 | -4.0074 | 1436.5013 | -4.0952 | 1430.8939 | -4.0952 | 1427.1166 | -4.1304 | 1429.2376 | -4.1304 | 1430.5384 | -4.148 | 1432.5564 | -4.148 | 1426.529
| -3.9546 | 110.0004 | -4.0601 | 115.1178 | -3.9898 | 114.0347 | -3.9898 | 111.6591 | -4.0425 | 112.486 | -3.9722 | 1436.2612 | -4.0601 | 1430.3016 | -4.0601 | 1424.5364 | -4.0952 | 1428.6474 | -4.0952 | 1430.8939 | -4.1128 | 1432.7944 | -4.1128 | 1426.529
| -3.9195 | 109.7774 | -4.0249 | 115.3181 | -3.9546 | 114.0347 | -3.9546 | 111.639 | -4.0074 | 112.768 | -3.937 | 1434.4643 | -4.0249 | 1430.6569 | -4.0249 | 1425.356 | -4.0601 | 1428.1759 | -4.0601 | 1429.3557 | -4.0777 | 1431.4869 | -4.0777 | 1423.8349
| -3.8843 | 109.4527 | -3.9898 | 115.5184 | -3.9195 | 114.0347 | -3.9195 | 112.0627 | -3.9722 | 112.8687 | -3.9019 | 1432.7944 | -3.9898 | 1429.4738 | -3.9898 | 1424.8875 | -4.0249 | 1427.8225 | -4.0249 | 1429.2376 | -4.0425 | 1432.7944 | -4.0425 | 1424.1856
| -3.8492 | 109.0872 | -3.9546 | 115.9786 | -3.8843 | 114.0347 | -3.8843 | 112.2038 | -3.937 | 113.7937 | -3.8667 | 1431.131 | -3.9546 | 1429.2376 | -3.9546 | 1422.4346 | -3.9898 | 1427.3518 | -3.9898 | 1426.8815 | -4.0074 | 1430.1833 | -4.0074 | 1422.7843
| -3.814 | 108.8229 | -3.9195 | 115.9786 | -3.8492 | 113.8339 | -3.8492 | 112.2038 | -3.9019 | 114.2957 | -3.8316 | 1432.7944 | -3.9195 | 1425.0046 | -3.9195 | 1421.2702 | -3.9546 | 1425.9422 | -3.9546 | 1425.8249 | -3.9722 | 1429.3557 | -3.9722 | 1422.4346
| -3.7789 | 108.8229 | -3.8843 | 115.9786 | -3.814 | 113.8138 | -3.814 | 112.2038 | -3.8667 | 114.095 | -3.7964 | 1433.6285 | -3.8843 | 1422.4346 | -3.8843 | 1421.1539 | -3.9195 | 1423.251 | -3.9195 | 1423.0176 | -3.937 | 1428.0581 | -3.937 | 1420.3403
| -3.7437 | 109.0465 | -3.8492 | 115.7385 | -3.7789 | 113.8138 | -3.7789 | 112.1837 | -3.8316 | 113.5124 | -3.7613 | 1432.5564 | -3.8492 | 1419.1798 | -3.8492 | 1418.4845 | -3.8843 | 1421.1539 | -3.8843 | 1421.0376 | -3.9019 | 1424.6534 | -3.9019 | 1419.0638
| -3.7086 | 109.473 | -3.814 | 115.3181 | -3.7437 | 112.2643 | -3.7437 | 112.103 | -3.7964 | 113.5124 | -3.7261 | 1431.2496 | -3.814 | 1422.4346 | -3.814 | 1417.4428 | -3.8492 | 1419.4117 | -3.8492 | 1420.2241 | -3.8667 | 1421.2702 | -3.8667 | 1416.9803
| -3.6734 | 109.676 | -3.7789 | 114.0347 | -3.7086 | 111.9618 | -3.7086 | 111.7399 | -3.7613 | 113.5526 | -3.691 | 1429.9467 | -3.7789 | 1423.251 | -3.7789 | 1417.4428 | -3.814 | 1419.4117 | -3.814 | 1419.5277 | -3.8316 | 1420.108 | -3.8316 | 1416.2871
| -3.6383 | 109.7165 | -3.7437 | 113.7736 | -3.6734 | 111.4168 | -3.6734 | 111.4572 | -3.7261 | 113.5928 | -3.6558 | 1425.5904 | -3.7437 | 1423.7181 | -3.7437 | 1417.0959 | -3.7789 | 1419.4117 | -3.7789 | 1419.2957 | -3.7964 | 1419.5277 | -3.7964 | 1414.096
| -3.6031 | 109.5745 | -3.7086 | 112.9894 | -3.6383 | 111.336 | -3.6383 | 110.8509 | -3.691 | 113.5928 | -3.6207 | 1420.805 | -3.7086 | 1429.4738 | -3.7086 | 1417.0959 | -3.7437 | 1418.832 | -3.7437 | 1419.1798 | -3.7613 | 1419.5277 | -3.7613 | 1413.2902
| -3.5679 | 109.0669 | -3.6734 | 112.8284 | -3.6031 | 111.336 | -3.6031 | 110.0004 | -3.6558 | 113.5928 | -3.5855 | 1418.1371 | -3.6734 | 1429.8284 | -3.6734 | 1417.5585 | -3.7086 | 1412.715 | -3.7086 | 1421.0376 | -3.7261 | 1421.2702 | -3.7261 | 1412.1402
| -3.5328 | 108.9856 | -3.6383 | 112.345 | -3.5679 | 111.336 | -3.5679 | 109.7571 | -3.6207 | 113.3716 | -3.5504 | 1417.6741 | -3.6383 | 1421.2702 | -3.6383 | 1417.5585 | -3.6734 | 1415.9408 | -3.6734 | 1422.2015 | -3.691 | 1421.2702 | -3.691 | 1411.7955
| -3.4976 | 108.2939 | -3.6031 | 111.982 | -3.5328 | 111.2148 | -3.5328 | 109.6557 | -3.5855 | 113.3113 | -3.5152 | 1417.6741 | -3.6031 | 1418.7162 | -3.6031 | 1414.096 | -3.6383 | 1415.9408 | -3.6383 | 1422.2015 | -3.6558 | 1421.9685 | -3.6558 | 1411.7955
| -3.4625 | 107.866 | -3.5679 | 111.538 | -3.4976 | 110.77 | -3.4976 | 109.473 | -3.5504 | 113.2309 | -3.4801 | 1419.0638 | -3.5679 | 1410.5324 | -3.5679 | 1416.0562 | -3.6031 | 1415.9408 | -3.6031 | 1417.9056 | -3.6207 | 1421.9685 | -3.6207 | 1411.6806
| -3.4273 | 107.866 | -3.5328 | 111.5986 | -3.4625 | 110.77 | -3.4625 | 108.9856 | -3.5152 | 113.0297 | -3.4449 | 1425.8249 | -3.5328 | 1410.5324 | -3.5328 | 1415.8254 | -3.5679 | 1415.9408 | -3.5679 | 1413.2902 | -3.5855 | 1419.1798 | -3.5855 | 1410.5324
| -3.3922 | 107.866 | -3.4976 | 111.9417 | -3.4273 | 110.6486 | -3.4273 | 108.7416 | -3.4801 | 112.3853 | -3.4098 | 1428.1759 | -3.4976 | 1406.98 | -3.4976 | 1414.096 | -3.5328 | 1416.8647 | -3.5328 | 1411.2212 | -3.5504 | 1415.3638 | -3.5504 | 1409.8441
| -3.357 | 107.4376 | -3.4625 | 112.4457 | -3.3922 | 110.3044 | -3.3922 | 108.1513 | -3.4449 | 111.6995 | -3.3746 | 1431.7243 | -3.4625 | 1412.1402 | -3.4625 | 1414.096 | -3.4976 | 1419.5277 | -3.4976 | 1411.2212 | -3.5152 | 1413.9809 | -3.5152 | 1409.8441
| -3.3219 | 108.2328 | -3.4273 | 112.9492 | -3.357 | 110.0815 | -3.357 | 108.1513 | -3.4098 | 111.6995 | -3.3395 | 1430.42 | -3.4273 | 1416.2871 | -3.4273 | 1415.8254 | -3.4625 | 1413.8657 | -3.4625 | 1418.3687 | -3.4801 | 1413.9809 | -3.4801 | 1407.7814
| -3.2867 | 108.4364 | -3.3922 | 112.6874 | -3.3219 | 110.1423 | -3.3219 | 108.1513 | -3.3746 | 111.0935 | -3.3043 | 1430.42 | -3.3922 | 1416.2871 | -3.3922 | 1411.4509 | -3.4273 | 1412.3701 | -3.4273 | 1423.9518 | -3.4449 | 1418.1371 | -3.4449 | 1408.1249
| -3.2516 | 108.5992 | -3.357 | 112.7277 | -3.2867 | 110.0612 | -3.2867 | 108.6602 | -3.3395 | 110.7902 | -3.2692 | 1426.8815 | -3.357 | 1416.2871 | -3.357 | 1412.1402 | -3.3922 | 1412.3701 | -3.3922 | 1425.356 | -3.4098 | 1421.3865 | -3.4098 | 1408.1249
| -3.2164 | 108.5992 | -3.3219 | 111.7803 | -3.2516 | 109.6557 | -3.2516 | 108.7416 | -3.3043 | 110.2233 | -3.234 | 1426.6465 | -3.3219 | 1416.2871 | -3.3219 | 1415.9408 | -3.357 | 1414.7873 | -3.357 | 1423.9518 | -3.3746 | 1428.5295 | -3.3746 | 1409.9588
| -3.1813 | 108.5992 | -3.2867 | 111.4572 | -3.2164 | 109.4527 | -3.2164 | 108.9043 | -3.2692 | 109.8788 | -3.1989 | 1423.251 | -3.2867 | 1415.8254 | -3.2867 | 1417.0959 | -3.3219 | 1414.9026 | -3.3219 | 1413.5204 | -3.3395 | 1428.5295 | -3.3395 | 1409.9588
| -3.1461 | 108.2939 | -3.2516 | 111.3562 | -3.1813 | 109.2497 | -3.1813 | 108.9043 | -3.234 | 109.9802 | -3.1637 | 1422.318 | -3.2516 | 1415.71 | -3.2516 | 1413.6355 | -3.2867 | 1414.9026 | -3.2867 | 1412.2552 | -3.3043 | 1421.3865 | -3.3043 | 1410.5324
| -3.111 | 108.0698 | -3.2164 | 111.0329 | -3.1461 | 108.8229 | -3.1461 | 108.7416 | -3.1989 | 110.2233 | -3.1285 | 1416.8647 | -3.2164 | 1413.7506 | -3.2164 | 1412.1402 | -3.2516 | 1412.9451 | -3.2516 | 1409.6147 | -3.2692 | 1415.1332 | -3.2692 | 1408.1249
| -3.0758 | 107.5396 | -3.1813 | 110.7497 | -3.111 | 108.8229 | -3.111 | 107.7436 | -3.1637 | 110.1626 | -3.0934 | 1414.096 | -3.1813 | 1413.6355 | -3.1813 | 1406.6367 | -3.2164 | 1411.336 | -3.2164 | 1413.4053 | -3.234 | 1410.1882 | -3.234 | 1406.6367
| -3.0407 | 106.9269 | -3.1461 | 110.689 | -3.0758 | 109.2497 | -3.0758 | 107.3151 | -3.1285 | 110.0612 | -3.0582 | 1414.7873 | -3.1461 | 1412.9451 | -3.1461 | 1404.3495 | -3.1813 | 1411.1064 | -3.1813 | 1416.1717 | -3.1989 | 1410.1882 | -3.1989 | 1401.6077
| -3.0055 | 106.661 | -3.111 | 110.4461 | -3.0407 | 109.3918 | -3.0407 | 107.1313 | -3.0934 | 110.0612 | -3.0231 | 1414.7873 | -3.111 | 1413.6355 | -3.111 | 1404.3495 | -3.1461 | 1411.1064 | -3.1461 | 1416.1717 | -3.1637 | 1411.7955 | -3.1637 | 1401.6077
| -2.9704 | 106.2721 | -3.0758 | 110.77 | -3.0055 | 109.5339 | -3.0055 | 107.07 | -3.0582 | 109.6963 | -2.9879 | 1420.6888 | -3.0758 | 1414.3264 | -3.0758 | 1405.8358 | -3.111 | 1410.3029 | -3.111 | 1411.5657 | -3.1285 | 1415.1332 | -3.1285 | 1401.9503
| -2.9352 | 105.7801 | -3.0407 | 110.9116 | -2.9704 | 109.6963 | -2.9704 | 107.07 | -3.0231 | 108.9652 | -2.9528 | 1420.805 | -3.0407 | 1414.3264 | -3.0407 | 1406.8655 | -3.0758 | 1410.3029 | -3.0758 | 1411.1064 | -3.0934 | 1413.7506 | -3.0934 | 1406.98
| -2.9001 | 105.4928 | -3.0055 | 111.2148 | -2.9352 | 109.6963 | -2.9352 | 107.3151 | -2.9879 | 108.8229 | -2.9176 | 1420.805 | -3.0055 | 1413.1751 | -3.0055 | 1405.8358 | -3.0407 | 1410.762 | -3.0407 | 1409.5001 | -3.0582 | 1413.7506 | -3.0582 | 1410.1882
| -2.8649 | 105.2875 | -2.9704 | 110.8509 | -2.9001 | 109.2497 | -2.9001 | 107.662 | -2.9528 | 108.9043 | -2.8825 | 1420.4564 | -2.9704 | 1409.9588 | -2.9704 | 1405.4927 | -3.0055 | 1406.8655 | -3.0055 | 1408.2395 | -3.0231 | 1413.7506 | -3.0231 | 1410.3029
| -2.8298 | 105.0203 | -2.9352 | 111.2754 | -2.8649 | 108.8433 | -2.8649 | 108.3346 | -2.9176 | 108.9043 | -2.8473 | 1419.6437 | -2.9352 | 1406.179 | -2.9352 | 1404.921 | -2.9704 | 1406.179 | -2.9704 | 1409.2708 | -2.9879 | 1411.9104 | -2.9879 | 1409.1562
| -2.7946 | 104.9175 | -2.9001 | 110.8104 | -2.8298 | 108.5381 | -2.8298 | 108.4567 | -2.8825 | 109.3918 | -2.8122 | 1418.6003 | -2.9001 | 1405.0354 | -2.9001 | 1404.0066 | -2.9352 | 1406.179 | -2.9352 | 1409.2708 | -2.9528 | 1411.9104 | -2.9528 | 1403.0925
| -2.7594 | 104.8969 | -2.8649 | 110.3651 | -2.7946 | 107.866 | -2.7946 | 108.4567 | -2.8473 | 109.7571 | -2.777 | 1418.2529 | -2.8649 | 1406.179 | -2.8649 | 1404.0066 | -2.9001 | 1405.8358 | -2.9001 | 1406.8655 | -2.9176 | 1410.1882 | -2.9176 | 1398.4107
| -2.7243 | 104.6705 | -2.8298 | 110.0612 | -2.7594 | 107.6212 | -2.7594 | 108.3346 | -2.8122 | 109.676 | -2.7419 | 1416.0562 | -2.8298 | 1407.0944 | -2.8298 | 1404.921 | -2.8649 | 1405.7215 | -2.8649 | 1406.8655 | -2.8825 | 1410.5324 | -2.8825 | 1395.5561
| -2.6891 | 103.7223 | -2.7946 | 109.9599 | -2.7243 | 107.2742 | -2.7243 | 106.5996 | -2.777 | 109.676 | -2.7067 | 1413.6355 | -2.7946 | 1408.9269 | -2.7946 | 1405.264 | -2.8298 | 1405.7215 | -2.8298 | 1406.8655 | -2.8473 | 1410.1882 | -2.8473 | 1394.8708
| -2.654 | 102.8746 | -2.7594 | 109.6557 | -2.6891 | 106.8247 | -2.6891 | 106.4359 | -2.7419 | 109.6557 | -2.6716 | 1413.5204 | -2.7594 | 1410.1882 | -2.7594 | 1403.6638 | -2.7946 | 1405.7215 | -2.7946 | 1406.8655 | -2.8122 | 1408.5832 | -2.8122 | 1400.9225
| -2.6188 | 102.7296 | -2.7243 | 109.2497 | -2.654 | 107.2742 | -2.654 | 106.1697 | -2.7067 | 109.3512 | -2.6364 | 1413.6355 | -2.7243 | 1408.9269 | -2.7243 | 1403.0925 | -2.7594 | 1405.7215 | -2.7594 | 1407.0944 | -2.777 | 1407.4379 | -2.777 | 1404.921
| -2.5837 | 102.7296 | -2.6891 | 108.3957 | -2.6188 | 107.2947 | -2.6188 | 106.0672 | -2.6716 | 108.9856 | -2.6013 | 1413.5204 | -2.6891 | 1407.0944 | -2.6891 | 1403.0925 | -2.7243 | 1405.264 | -2.7243 | 1406.8655 | -2.7419 | 1406.5222 | -2.7419 | 1406.6367
| -2.5485 | 102.3773 | -2.654 | 108.0494 | -2.5837 | 107.4376 | -2.5837 | 105.5955 | -2.6364 | 108.4974 | -2.5661 | 1412.4851 | -2.654 | 1405.8358 | -2.654 | 1401.9503 | -2.6891 | 1403.8923 | -2.6891 | 1406.6367 | -2.7067 | 1406.2934 | -2.7067 | 1404.1209
| -2.5134 | 102.481 | -2.6188 | 107.8048 | -2.5485 | 106.4359 | -2.5485 | 105.2875 | -2.6013 | 107.866 | -2.531 | 1414.2112 | -2.6188 | 1405.8358 | -2.6188 | 1404.0066 | -2.654 | 1403.8923 | -2.654 | 1406.6367 | -2.6716 | 1404.3495 | -2.6716 | 1403.0925
| -2.4782 | 102.8746 | -2.5837 | 107.8864 | -2.5134 | 106.7428 | -2.5134 | 105.1025 | -2.5661 | 107.5192 | -2.4958 | 1415.71 | -2.5837 | 1407.0944 | -2.5837 | 1406.0646 | -2.6188 | 1403.8923 | -2.6188 | 1405.1497 | -2.6364 | 1404.3495 | -2.6364 | 1401.8361
| -2.4431 | 103.0195 | -2.5485 | 108.0494 | -2.4782 | 105.5544 | -2.4782 | 104.8969 | -2.531 | 107.2334 | -2.4607 | 1417.4428 | -2.5485 | 1407.5523 | -2.5485 | 1406.6367 | -2.5837 | 1403.8923 | -2.5837 | 1406.2934 | -2.6013 | 1404.3495 | -2.6013 | 1399.5524
| -2.4079 | 102.7711 | -2.5134 | 107.9068 | -2.4431 | 105.3696 | -2.4431 | 104.794 | -2.4958 | 106.8042 | -2.4255 | 1417.6741 | -2.5134 | 1409.7294 | -2.5134 | 1408.4686 | -2.5485 | 1404.3495 | -2.5485 | 1407.8959 | -2.5661 | 1404.1209 | -2.5661 | 1401.0367
| -2.3728 | 102.3358 | -2.4782 | 107.6416 | -2.4079 | 105.2669 | -2.4079 | 104.5676 | -2.4607 | 106.3335 | -2.3903 | 1418.3687 | -2.4782 | 1410.3029 | -2.4782 | 1410.0735 | -2.5134 | 1407.4379 | -2.5134 | 1410.4177 | -2.531 | 1405.3784 | -2.531 | 1401.8361
| -2.3376 | 102.3358 | -2.4431 | 106.5178 | -2.3728 | 105.0614 | -2.3728 | 104.238 | -2.4255 | 106.6201 | -2.3552 | 1419.6437 | -2.4431 | 1410.3029 | -2.4431 | 1410.762 | -2.4782 | 1407.4379 | -2.4782 | 1411.4509 | -2.4958 | 1408.4686 | -2.4958 | 1401.3793
| -2.3025 | 102.3358 | -2.4079 | 106.1901 | -2.3376 | 104.4646 | -2.3376 | 104.1349 | -2.3903 | 106.3335 | -2.32 | 1419.6437 | -2.4079 | 1409.0415 | -2.4079 | 1410.762 | -2.4431 | 1405.6071 | -2.4431 | 1412.1402 | -2.4607 | 1409.0415 | -2.4607 | 1401.3793
| -2.2673 | 102.8332 | -2.3728 | 105.9852 | -2.3025 | 103.5571 | -2.3025 | 103.5571 | -2.3552 | 106.8451 | -2.2849 | 1419.1798 | -2.3728 | 1411.2212 | -2.3728 | 1410.762 | -2.4079 | 1408.8123 | -2.4079 | 1411.4509 | -2.4255 | 1408.4686 | -2.4255 | 1402.4071
| -2.2322 | 103.2677 | -2.3376 | 105.7391 | -2.2673 | 103.5571 | -2.2673 | 103.0816 | -2.32 | 106.5996 | -2.2497 | 1419.6437 | -2.3376 | 1412.2552 | -2.3376 | 1410.6472 | -2.3728 | 1409.5001 | -2.3728 | 1410.9916 | -2.3903 | 1409.6147 | -2.3903 | 1404.6924
| -2.197 | 103.3504 | -2.3025 | 105.1847 | -2.2322 | 103.619 | -2.2322 | 102.7089 | -2.2849 | 106.4973 | -2.2146 | 1420.5726 | -2.3025 | 1412.715 | -2.3025 | 1410.4177 | -2.3376 | 1407.5523 | -2.3376 | 1408.6977 | -2.3552 | 1408.9269 | -2.3552 | 1404.6924
| -2.1619 | 103.7429 | -2.2673 | 105.1847 | -2.197 | 103.7842 | -2.197 | 102.2736 | -2.2497 | 106.2926 | -2.1794 | 1421.3865 | -2.2673 | 1411.9104 | -2.2673 | 1410.4177 | -2.3025 | 1409.5001 | -2.3025 | 1408.6977 | -2.32 | 1408.6977 | -2.32 | 1404.3495
| -2.1267 | 103.9287 | -2.2322 | 105.1847 | -2.1619 | 103.7223 | -2.1619 | 101.6089 | -2.2146 | 105.8827 | -2.1443 | 1419.6437 | -2.2322 | 1411.6806 | -2.2322 | 1410.0735 | -2.2673 | 1408.9269 | -2.2673 | 1408.6977 | -2.2849 | 1408.6977 | -2.2849 | 1402.7498
| -2.0916 | 103.9287 | -2.197 | 105.6365 | -2.1267 | 103.5157 | -2.1267 | 101.4634 | -2.1794 | 105.0408 | -2.1091 | 1419.6437 | -2.197 | 1410.8768 | -2.197 | 1409.3854 | -2.2322 | 1406.6367 | -2.2322 | 1409.1562 | -2.2497 | 1409.6147 | -2.2497 | 1400.0091
| -2.0564 | 104.0937 | -2.1619 | 105.6365 | -2.0916 | 103.5157 | -2.0916 | 101.7129 | -2.1443 | 103.8255 | -2.074 | 1418.6003 | -2.1619 | 1410.762 | -2.1619 | 1409.3854 | -2.197 | 1406.0646 | -2.197 | 1409.5001 | -2.2146 | 1409.3854 | -2.2146 | 1400.0091
| -2.0213 | 104.1968 | -2.1267 | 105.6365 | -2.0564 | 103.4537 | -2.0564 | 101.7752 | -2.1091 | 103.6603 | -2.0388 | 1416.0562 | -2.1267 | 1409.7294 | -2.1267 | 1409.0415 | -2.1619 | 1406.0646 | -2.1619 | 1413.4053 | -2.1794 | 1409.6147 | -2.1794 | 1402.1787
| -1.9861 | 104.1555 | -2.0916 | 105.082 | -2.0213 | 103.5157 | -2.0213 | 101.6713 | -2.074 | 103.2057 | -2.0037 | 1414.9026 | -2.0916 | 1409.5001 | -2.0916 | 1406.5222 | -2.1267 | 1406.0646 | -2.1267 | 1409.3854 | -2.1443 | 1410.4177 | -2.1443 | 1402.864
| -1.9509 | 104.0937 | -2.0564 | 105.3285 | -1.9861 | 104.4234 | -1.9861 | 101.6921 | -2.0388 | 103.2057 | -1.9685 | 1414.3264 | -2.0564 | 1408.0104 | -2.0564 | 1406.2934 | -2.0916 | 1406.0646 | -2.0916 | 1409.0415 | -2.1091 | 1410.4177 | -2.1091 | 1403.2068
| -1.9158 | 103.8874 | -2.0213 | 105.8007 | -1.9509 | 104.7734 | -1.9509 | 101.6505 | -2.0037 | 103.619 | -1.9334 | 1412.715 | -2.0213 | 1404.2352 | -2.0213 | 1402.864 | -2.0564 | 1406.0646 | -2.0564 | 1404.4638 | -2.074 | 1410.3029 | -2.074 | 1404.3495
| -1.8806 | 103.8874 | -1.9861 | 106.3335 | -1.9158 | 104.7529 | -1.9158 | 101.5674 | -1.9685 | 103.5364 | -1.8982 | 1411.6806 | -1.9861 | 1403.5495 | -1.9861 | 1402.6356 | -2.0213 | 1401.4935 | -2.0213 | 1402.5214 | -2.0388 | 1409.6147 | -2.0388 | 1403.5495
| -1.8455 | 104.1555 | -1.9509 | 107.0904 | -1.8806 | 105.1847 | -1.8806 | 101.6921 | -1.9334 | 103.1023 | -1.8631 | 1411.6806 | -1.9509 | 1403.5495 | -1.9509 | 1401.4935 | -1.9861 | 1401.4935 | -1.9861 | 1402.5214 | -2.0037 | 1408.2395 | -2.0037 | 1400.4658
| -1.8103 | 104.3204 | -1.9158 | 107.0904 | -1.8455 | 105.2875 | -1.8455 | 102.4602 | -1.8982 | 103.1023 | -1.8279 | 1409.6147 | -1.9158 | 1404.2352 | -1.9158 | 1402.864 | -1.9509 | 1401.0367 | -1.9509 | 1402.1787 | -1.9685 | 1404.2352 | -1.9685 | 1398.8674
| -1.7752 | 105.1847 | -1.8806 | 107.0904 | -1.8103 | 105.308 | -1.8103 | 103.3297 | -1.8631 | 103.1023 | -1.7928 | 1409.6147 | -1.8806 | 1405.7215 | -1.8806 | 1405.6071 | -1.9158 | 1401.9503 | -1.9158 | 1402.7498 | -1.9334 | 1403.8923 | -1.9334 | 1400.0091
| -1.74 | 105.4723 | -1.8455 | 107.2742 | -1.7752 | 105.3491 | -1.7752 | 103.8255 | -1.8279 | 103.1023 | -1.7576 | 1411.9104 | -1.8455 | 1405.7215 | -1.8455 | 1404.2352 | -1.8806 | 1403.8923 | -1.8806 | 1404.5781 | -1.8982 | 1402.5214 | -1.8982 | 1400.0091
| -1.7049 | 105.6776 | -1.8103 | 107.1721 | -1.74 | 105.9237 | -1.74 | 104.0318 | -1.7928 | 103.1023 | -1.7225 | 1413.9809 | -1.8103 | 1405.7215 | -1.8103 | 1401.6077 | -1.8455 | 1404.8067 | -1.8455 | 1405.4927 | -1.8631 | 1402.7498 | -1.8631 | 1400.0091
| -1.6697 | 104.4646 | -1.7752 | 107.4376 | -1.7049 | 105.9647 | -1.7049 | 104.1555 | -1.7576 | 103.3297 | -1.6873 | 1411.9104 | -1.7752 | 1405.264 | -1.7752 | 1400.8083 | -1.8103 | 1403.8923 | -1.8103 | 1404.5781 | -1.8279 | 1402.7498 | -1.8279 | 1401.0367
| -1.6346 | 103.4331 | -1.74 | 107.7232 | -1.6697 | 105.9237 | -1.6697 | 104.4234 | -1.7225 | 104.1968 | -1.6522 | 1406.98 | -1.74 | 1404.5781 | -1.74 | 1398.8674 | -1.7752 | 1400.9225 | -1.7752 | 1403.8923 | -1.7928 | 1402.6356 | -1.7928 | 1401.0367
| -1.5994 | 102.9367 | -1.7049 | 107.5192 | -1.6346 | 106.0057 | -1.6346 | 104.6294 | -1.6873 | 104.1968 | -1.617 | 1404.5781 | -1.7049 | 1403.2068 | -1.7049 | 1394.071 | -1.74 | 1399.0957 | -1.74 | 1401.0367 | -1.7576 | 1400.9225 | -1.7576 | 1399.3241
| -1.5643 | 102.9367 | -1.6697 | 107.4376 | -1.5994 | 106.0057 | -1.5994 | 104.547 | -1.6522 | 104.6088 | -1.5818 | 1403.0925 | -1.6697 | 1402.5214 | -1.6697 | 1393.8425 | -1.7049 | 1398.1824 | -1.7049 | 1398.4107 | -1.7225 | 1400.9225 | -1.7225 | 1399.3241
| -1.5291 | 102.8125 | -1.6346 | 107.4172 | -1.5643 | 105.9647 | -1.5643 | 104.1968 | -1.617 | 104.6088 | -1.5467 | 1402.5214 | -1.6346 | 1401.7219 | -1.6346 | 1392.0133 | -1.6697 | 1396.6981 | -1.6697 | 1395.8987 | -1.6873 | 1400.4658 | -1.6873 | 1394.2995
| -1.494 | 102.8746 | -1.5994 | 107.4172 | -1.5291 | 106.395 | -1.5291 | 103.7223 | -1.5818 | 104.7117 | -1.5115 | 1403.5495 | -1.5994 | 1398.7532 | -1.5994 | 1391.2125 | -1.6346 | 1396.1272 | -1.6346 | 1395.8987 | -1.6522 | 1400.3516 | -1.6522 | 1391.2125
| -1.4588 | 103.0816 | -1.5643 | 107.4376 | -1.494 | 106.395 | -1.494 | 103.681 | -1.5467 | 105.616 | -1.4764 | 1405.6071 | -1.5643 | 1398.5249 | -1.5643 | 1390.7548 | -1.5994 | 1396.1272 | -1.5994 | 1395.7845 | -1.617 | 1397.9541 | -1.617 | 1384.2143
| -1.4237 | 103.4951 | -1.5291 | 108.2124 | -1.4588 | 106.0877 | -1.4588 | 103.619 | -1.5115 | 105.8827 | -1.4412 | 1404.921 | -1.5291 | 1398.5249 | -1.5291 | 1391.2125 | -1.5643 | 1394.4138 | -1.5643 | 1395.0992 | -1.5818 | 1395.2135 | -1.5818 | 1386.9727
| -1.3885 | 103.1229 | -1.494 | 108.9043 | -1.4237 | 105.1642 | -1.4237 | 103.4537 | -1.4764 | 105.8007 | -1.4061 | 1402.5214 | -1.494 | 1396.3555 | -1.494 | 1392.3564 | -1.5291 | 1393.9567 | -1.5291 | 1395.0992 | -1.5467 | 1393.0424 | -1.5467 | 1383.293
| -1.3534 | 102.7711 | -1.4588 | 108.6602 | -1.3885 | 104.794 | -1.3885 | 103.6397 | -1.4412 | 105.7391 | -1.3709 | 1400.9225 | -1.4588 | 1394.6423 | -1.4588 | 1390.9837 | -1.494 | 1393.7282 | -1.494 | 1394.6423 | -1.5115 | 1392.242 | -1.5115 | 1383.293
| -1.3182 | 102.5846 | -1.4237 | 107.6824 | -1.3534 | 104.2792 | -1.3534 | 104.0524 | -1.4061 | 105.8827 | -1.3358 | 1395.7845 | -1.4237 | 1392.3564 | -1.4237 | 1389.7244 | -1.4588 | 1392.6994 | -1.4588 | 1391.4414 | -1.4764 | 1392.3564 | -1.4764 | 1386.9727
| -1.2831 | 102.5846 | -1.3885 | 106.7838 | -1.3182 | 103.9287 | -1.3182 | 104.0524 | -1.3709 | 106.354 | -1.3006 | 1395.5561 | -1.3885 | 1392.1277 | -1.3885 | 1387.5465 | -1.4237 | 1391.4414 | -1.4237 | 1389.9534 | -1.4412 | 1392.242 | -1.4412 | 1387.7759
| -1.2479 | 102.5846 | -1.3534 | 106.7224 | -1.2831 | 103.8874 | -1.2831 | 103.8461 | -1.3358 | 106.8247 | -1.2655 | 1395.5561 | -1.3534 | 1389.8389 | -1.3534 | 1387.0875 | -1.3885 | 1389.9534 | -1.3885 | 1389.6098 | -1.4061 | 1392.3564 | -1.4061 | 1385.1346
| -1.2128 | 102.8125 | -1.3182 | 106.5792 | -1.2479 | 103.908 | -1.2479 | 103.7223 | -1.3006 | 107.07 | -1.2303 | 1397.269 | -1.3182 | 1388.3493 | -1.3182 | 1387.0875 | -1.3534 | 1388.0053 | -1.3534 | 1388.12 | -1.3709 | 1391.5558 | -1.3709 | 1384.3294
| -1.1776 | 103.2677 | -1.2831 | 106.1697 | -1.2128 | 104.4234 | -1.2128 | 103.3711 | -1.2655 | 107.07 | -1.1952 | 1397.269 | -1.2831 | 1387.0875 | -1.2831 | 1387.0875 | -1.3182 | 1387.0875 | -1.3182 | 1387.4318 | -1.3358 | 1389.7244 | -1.3358 | 1385.1346
| -1.1424 | 103.4537 | -1.2479 | 106.354 | -1.1776 | 105.2053 | -1.1776 | 103.185 | -1.2303 | 106.8247 | -1.16 | 1397.269 | -1.2479 | 1388.6932 | -1.2479 | 1387.0875 | -1.2831 | 1387.0875 | -1.2831 | 1387.4318 | -1.3006 | 1387.8906 | -1.3006 | 1384.3294
| -1.1073 | 103.7842 | -1.2128 | 106.886 | -1.1424 | 104.7117 | -1.1424 | 103.0195 | -1.1952 | 106.1082 | -1.1249 | 1396.4697 | -1.2128 | 1388.6932 | -1.2128 | 1386.3987 | -1.2479 | 1386.1689 | -1.2479 | 1387.5465 | -1.2655 | 1387.6612 | -1.2655 | 1385.1346
| -1.0721 | 103.8048 | -1.1776 | 106.9064 | -1.1073 | 104.5676 | -1.1073 | 102.6882 | -1.16 | 104.8969 | -1.0897 | 1395.8987 | -1.1776 | 1388.6932 | -1.1776 | 1385.3645 | -1.2128 | 1387.0875 | -1.2128 | 1388.0053 | -1.2303 | 1387.6612 | -1.2303 | 1386.0541
| -1.037 | 104.0318 | -1.1424 | 106.9064 | -1.0721 | 104.0318 | -1.0721 | 102.7089 | -1.1249 | 104.4852 | -1.0546 | 1394.8708 | -1.1424 | 1386.6283 | -1.1424 | 1384.3294 | -1.1776 | 1388.3493 | -1.1776 | 1388.0053 | -1.1952 | 1387.4318 | -1.1952 | 1388.5786
| -1.0018 | 104.341 | -1.1073 | 106.886 | -1.037 | 102.916 | -1.037 | 103.0195 | -1.0897 | 104.3616 | -1.0194 | 1394.2995 | -1.1073 | 1387.317 | -1.1073 | 1385.0196 | -1.1424 | 1386.8579 | -1.1424 | 1386.3987 | -1.16 | 1387.4318 | -1.16 | 1388.5786
| -0.96669 | 104.4852 | -1.0721 | 105.7596 | -1.0018 | 102.916 | -1.0018 | 103.0609 | -1.0546 | 104.3616 | -0.98426 | 1394.2995 | -1.0721 | 1384.2143 | -1.0721 | 1385.0196 | -1.1073 | 1385.2496 | -1.1073 | 1385.4795 | -1.1249 | 1387.4318 | -1.1249 | 1382.0246
| -0.93153 | 104.4852 | -1.037 | 105.7596 | -0.96669 | 102.916 | -0.96669 | 102.6675 | -1.0194 | 104.4852 | -0.94911 | 1395.2135 | -1.037 | 1384.2143 | -1.037 | 1385.9392 | -1.0721 | 1384.9046 | -1.0721 | 1384.0992 | -1.0897 | 1385.8243 | -1.0897 | 1379.5974
| -0.89638 | 105.3696 | -1.0018 | 105.7596 | -0.93153 | 103.5984 | -0.93153 | 102.626 | -0.98426 | 104.8146 | -0.91396 | 1395.5561 | -1.0018 | 1387.317 | -1.0018 | 1388.3493 | -1.037 | 1385.2496 | -1.037 | 1382.14 | -1.0546 | 1385.0196 | -1.0546 | 1374.6
| -0.86123 | 104.9997 | -0.96669 | 106.2721 | -0.89638 | 103.9287 | -0.89638 | 102.4602 | -0.94911 | 104.9997 | -0.87881 | 1396.3555 | -0.96669 | 1387.317 | -0.96669 | 1388.9224 | -1.0018 | 1385.4795 | -1.0018 | 1384.2143 | -1.0194 | 1384.7896 | -1.0194 | 1373.8993
| -0.82608 | 105.1436 | -0.93153 | 106.5792 | -0.86123 | 104.3822 | -0.86123 | 102.2113 | -0.91396 | 104.938 | -0.84365 | 1396.6981 | -0.93153 | 1386.2838 | -0.93153 | 1388.2346 | -0.96669 | 1386.7431 | -0.96669 | 1384.7896 | -0.98426 | 1384.7896 | -0.98426 | 1374.2498
| -0.79092 | 105.1642 | -0.89638 | 106.6406 | -0.82608 | 105.6981 | -0.82608 | 102.4602 | -0.87881 | 104.2998 | -0.8085 | 1396.8123 | -0.89638 | 1388.6932 | -0.89638 | 1388.2346 | -0.93153 | 1386.7431 | -0.93153 | 1384.6745 | -0.94911 | 1383.7537 | -0.94911 | 1374.0162
| -0.75577 | 105.1436 | -0.86123 | 106.8656 | -0.79092 | 106.2106 | -0.79092 | 102.6468 | -0.84365 | 104.1762 | -0.77335 | 1396.9265 | -0.86123 | 1388.2346 | -0.86123 | 1386.5135 | -0.89638 | 1386.7431 | -0.89638 | 1384.6745 | -0.91396 | 1382.2554 | -0.91396 | 1373.8993
| -0.72062 | 105.1231 | -0.82608 | 106.9678 | -0.75577 | 106.8451 | -0.75577 | 102.916 | -0.8085 | 103.5777 | -0.7382 | 1399.0957 | -0.82608 | 1388.2346 | -0.82608 | 1386.3987 | -0.86123 | 1385.9392 | -0.86123 | 1386.6283 | -0.87881 | 1383.8689 | -0.87881 | 1374.0162
| -0.68547 | 105.2258 | -0.79092 | 107.3763 | -0.72062 | 106.7428 | -0.72062 | 103.4124 | -0.77335 | 103.5984 | -0.70304 | 1400.0091 | -0.79092 | 1391.5558 | -0.79092 | 1387.5465 | -0.82608 | 1385.8243 | -0.82608 | 1387.4318 | -0.84365 | 1385.4795 | -0.84365 | 1379.1342
| -0.65032 | 105.3696 | -0.75577 | 107.5804 | -0.68547 | 106.7633 | -0.68547 | 103.4124 | -0.7382 | 104.5264 | -0.66789 | 1401.0367 | -0.75577 | 1392.4707 | -0.75577 | 1388.0053 | -0.79092 | 1385.9392 | -0.79092 | 1389.4953 | -0.8085 | 1385.5944 | -0.8085 | 1379.3658
| -0.61516 | 105.4107 | -0.72062 | 107.7844 | -0.65032 | 106.8451 | -0.65032 | 103.908 | -0.70304 | 104.5676 | -0.63274 | 1401.4935 | -0.72062 | 1392.4707 | -0.72062 | 1388.4639 | -0.75577 | 1386.3987 | -0.75577 | 1390.2969 | -0.77335 | 1385.5944 | -0.77335 | 1379.3658
| -0.58001 | 105.4312 | -0.68547 | 108.192 | -0.61516 | 106.0467 | -0.61516 | 104.8969 | -0.66789 | 104.8969 | -0.59759 | 1402.7498 | -0.68547 | 1393.1567 | -0.68547 | 1390.7548 | -0.72062 | 1386.3987 | -0.72062 | 1390.2969 | -0.7382 | 1387.2023 | -0.7382 | 1379.3658
| -0.54486 | 105.4312 | -0.65032 | 108.2531 | -0.58001 | 105.9852 | -0.58001 | 104.8969 | -0.63274 | 105.082 | -0.56244 | 1402.7498 | -0.65032 | 1393.1567 | -0.65032 | 1391.2125 | -0.68547 | 1387.0875 | -0.68547 | 1391.4414 | -0.70304 | 1387.5465 | -0.70304 | 1377.1618
| -0.50971 | 105.4312 | -0.61516 | 108.2939 | -0.54486 | 106.2311 | -0.54486 | 104.6705 | -0.59759 | 105.5339 | -0.52728 | 1403.7781 | -0.61516 | 1393.1567 | -0.61516 | 1392.0133 | -0.65032 | 1389.8389 | -0.65032 | 1392.242 | -0.66789 | 1389.2662 | -0.66789 | 1376.5805
| -0.47455 | 105.4312 | -0.58001 | 108.4364 | -0.50971 | 105.8827 | -0.50971 | 105.5544 | -0.56244 | 106.2106 | -0.49213 | 1405.7215 | -0.58001 | 1393.1567 | -0.58001 | 1392.1277 | -0.61516 | 1391.0981 | -0.61516 | 1394.528 | -0.63274 | 1389.2662 | -0.63274 | 1377.1618
| -0.4394 | 105.5544 | -0.54486 | 108.4364 | -0.47455 | 105.8622 | -0.47455 | 104.938 | -0.52728 | 106.9269 | -0.45698 | 1406.5222 | -0.54486 | 1391.8989 | -0.54486 | 1393.0424 | -0.58001 | 1391.6702 | -0.58001 | 1394.8708 | -0.59759 | 1390.9837 | -0.59759 | 1379.7132
| -0.40425 | 105.7391 | -0.50971 | 107.866 | -0.4394 | 106.2926 | -0.4394 | 104.6705 | -0.49213 | 106.8656 | -0.42183 | 1406.8655 | -0.50971 | 1393.1567 | -0.50971 | 1393.271 | -0.54486 | 1392.4707 | -0.54486 | 1394.8708 | -0.56244 | 1391.4414 | -0.56244 | 1381.5629
| -0.3691 | 107.5804 | -0.47455 | 107.9271 | -0.40425 | 106.395 | -0.40425 | 104.8969 | -0.45698 | 107.0086 | -0.38667 | 1406.5222 | -0.47455 | 1394.1853 | -0.47455 | 1395.2135 | -0.50971 | 1392.6994 | -0.50971 | 1394.8708 | -0.52728 | 1392.242 | -0.52728 | 1384.6745
| -0.33395 | 107.7436 | -0.4394 | 107.8252 | -0.3691 | 106.354 | -0.3691 | 104.7117 | -0.42183 | 106.8247 | -0.35152 | 1406.5222 | -0.4394 | 1394.985 | -0.4394 | 1396.013 | -0.47455 | 1393.3853 | -0.47455 | 1394.8708 | -0.49213 | 1392.3564 | -0.49213 | 1384.7896
| -0.29879 | 108.4974 | -0.40425 | 107.9271 | -0.33395 | 106.9269 | -0.33395 | 104.3616 | -0.38667 | 106.313 | -0.31637 | 1406.8655 | -0.40425 | 1395.3277 | -0.40425 | 1396.8123 | -0.4394 | 1394.4138 | -0.4394 | 1396.013 | -0.45698 | 1393.7282 | -0.45698 | 1384.7896
| -0.26364 | 108.4974 | -0.3691 | 108.192 | -0.29879 | 106.8656 | -0.29879 | 104.4646 | -0.35152 | 105.9442 | -0.28122 | 1407.5523 | -0.3691 | 1396.6981 | -0.3691 | 1396.9265 | -0.40425 | 1393.3853 | -0.40425 | 1394.6423 | -0.42183 | 1394.4138 | -0.42183 | 1385.0196
| -0.22849 | 107.5192 | -0.33395 | 108.6195 | -0.26364 | 106.8656 | -0.26364 | 104.8969 | -0.31637 | 106.2106 | -0.24607 | 1407.7814 | -0.33395 | 1397.0407 | -0.33395 | 1396.9265 | -0.3691 | 1394.8708 | -0.3691 | 1397.0407 | -0.38667 | 1395.7845 | -0.38667 | 1385.1346
| -0.19334 | 107.3763 | -0.29879 | 108.4974 | -0.22849 | 107.1108 | -0.22849 | 104.8352 | -0.28122 | 106.4359 | -0.21091 | 1408.2395 | -0.29879 | 1395.8987 | -0.29879 | 1396.2413 | -0.33395 | 1395.3277 | -0.33395 | 1397.269 | -0.35152 | 1395.7845 | -0.35152 | 1385.9392
| -0.15818 | 107.5192 | -0.26364 | 108.6602 | -0.19334 | 107.2947 | -0.19334 | 105.2875 | -0.24607 | 106.3335 | -0.17576 | 1409.7294 | -0.26364 | 1395.8987 | -0.26364 | 1393.1567 | -0.29879 | 1394.985 | -0.29879 | 1395.8987 | -0.31637 | 1397.3832 | -0.31637 | 1388.3493
| -0.12303 | 107.3763 | -0.22849 | 108.3142 | -0.15818 | 107.3967 | -0.15818 | 105.4107 | -0.21091 | 106.4564 | -0.14061 | 1410.4177 | -0.22849 | 1397.1548 | -0.22849 | 1392.0133 | -0.26364 | 1394.985 | -0.26364 | 1393.8425 | -0.28122 | 1398.0682 | -0.28122 | 1390.5259
| -0.087881 | 107.7232 | -0.19334 | 108.6195 | -0.12303 | 108.1513 | -0.12303 | 105.4518 | -0.17576 | 106.8656 | -0.10546 | 1410.8768 | -0.19334 | 1397.0407 | -0.19334 | 1392.9281 | -0.22849 | 1396.1272 | -0.22849 | 1394.1853 | -0.24607 | 1396.5839 | -0.24607 | 1388.3493
| -0.052728 | 108.355 | -0.15818 | 109.2294 | -0.087881 | 108.2124 | -0.087881 | 105.657 | -0.14061 | 106.8656 | -0.070304 | 1410.8768 | -0.15818 | 1393.3853 | -0.15818 | 1393.1567 | -0.19334 | 1394.985 | -0.19334 | 1396.013 | -0.21091 | 1394.4138 | -0.21091 | 1388.3493
| -0.017576 | 108.192 | -0.12303 | 109.7165 | -0.052728 | 109.5136 | -0.052728 | 105.8417 | -0.10546 | 107.0086 | -0.035152 | 1410.8768 | -0.12303 | 1394.985 | -0.12303 | 1398.5249 | -0.15818 | 1392.242 | -0.15818 | 1396.013 | -0.17576 | 1394.071 | -0.17576 | 1383.984
| 0.017576 | 108.4771 | -0.087881 | 110.3449 | -0.017576 | 109.9802 | -0.017576 | 106.2516 | -0.070304 | 107.1313 | 0 | 1411.336 | -0.087881 | 1395.6703 | -0.087881 | 1399.4382 | -0.12303 | 1390.9837 | -0.12303 | 1396.013 | -0.14061 | 1392.6994 | -0.14061 | 1383.984
| 0.052728 | 108.9043 | -0.052728 | 110.9116 | 0.017576 | 109.7165 | 0.017576 | 106.9269 | -0.035152 | 107.7028 | 0.035152 | 1410.9916 | -0.052728 | 1396.013 | -0.052728 | 1399.8949 | -0.087881 | 1390.6403 | -0.087881 | 1396.4697 | -0.10546 | 1390.6403 | -0.10546 | 1381.6783
| 0.087881 | 108.5381 | -0.017576 | 111.0329 | 0.052728 | 108.7212 | 0.052728 | 107.6824 | 0 | 107.2947 | 0.070304 | 1410.5324 | -0.017576 | 1396.2413 | -0.017576 | 1399.4382 | -0.052728 | 1390.9837 | -0.052728 | 1396.8123 | -0.070304 | 1390.1824 | -0.070304 | 1381.6783
| 0.12303 | 108.0902 | 0.017576 | 110.9318 | 0.087881 | 108.4771 | 0.087881 | 108.3142 | 0.035152 | 107.3355 | 0.10546 | 1409.7294 | 0.017576 | 1400.2374 | 0.017576 | 1399.8949 | -0.017576 | 1392.9281 | -0.017576 | 1397.269 | -0.035152 | 1390.6403 | -0.035152 | 1387.8906
| 0.15818 | 107.5192 | 0.052728 | 110.9116 | 0.12303 | 108.5788 | 0.12303 | 108.3957 | 0.070304 | 107.7232 | 0.14061 | 1410.4177 | 0.052728 | 1401.4935 | 0.052728 | 1400.6942 | 0.017576 | 1397.1548 | 0.017576 | 1398.2966 | 0 | 1394.985 | 0 | 1388.12
| 0.19334 | 106.9473 | 0.087881 | 111.0329 | 0.15818 | 108.4771 | 0.15818 | 108.3142 | 0.10546 | 108.1105 | 0.17576 | 1410.4177 | 0.087881 | 1403.5495 | 0.087881 | 1400.6942 | 0.052728 | 1399.3241 | 0.052728 | 1398.7532 | 0.035152 | 1396.3555 | 0.035152 | 1388.12
| 0.22849 | 106.7224 | 0.12303 | 111.1137 | 0.19334 | 108.7212 | 0.19334 | 108.355 | 0.14061 | 108.3346 | 0.21091 | 1410.4177 | 0.12303 | 1404.6924 | 0.12303 | 1401.4935 | 0.087881 | 1401.3793 | 0.087881 | 1399.6666 | 0.070304 | 1397.7257 | 0.070304 | 1393.1567
| 0.26364 | 106.5382 | 0.15818 | 111.3966 | 0.22849 | 108.9043 | 0.22849 | 108.3957 | 0.17576 | 109.2497 | 0.24607 | 1412.9451 | 0.15818 | 1406.0646 | 0.15818 | 1402.4071 | 0.12303 | 1401.4935 | 0.12303 | 1401.8361 | 0.10546 | 1399.6666 | 0.10546 | 1393.271
| 0.29879 | 105.9237 | 0.19334 | 111.9215 | 0.26364 | 109.7774 | 0.26364 | 108.2531 | 0.21091 | 109.4121 | 0.28122 | 1412.9451 | 0.19334 | 1406.179 | 0.19334 | 1404.3495 | 0.15818 | 1401.3793 | 0.15818 | 1402.2929 | 0.14061 | 1400.6942 | 0.14061 | 1393.1567
| 0.33395 | 105.616 | 0.22849 | 112.0224 | 0.29879 | 109.4121 | 0.29879 | 108.2124 | 0.24607 | 109.2497 | 0.31637 | 1411.6806 | 0.22849 | 1404.6924 | 0.22849 | 1403.4353 | 0.19334 | 1401.3793 | 0.19334 | 1402.2929 | 0.17576 | 1401.4935 | 0.17576 | 1393.271
| 0.3691 | 105.5544 | 0.26364 | 111.9417 | 0.33395 | 109.4121 | 0.33395 | 108.7619 | 0.28122 | 109.0465 | 0.35152 | 1409.6147 | 0.26364 | 1401.8361 | 0.26364 | 1401.4935 | 0.22849 | 1402.864 | 0.22849 | 1402.1787 | 0.21091 | 1401.6077 | 0.21091 | 1395.6703
| 0.40425 | 105.616 | 0.29879 | 111.5582 | 0.3691 | 109.27 | 0.3691 | 108.7823 | 0.31637 | 109.1888 | 0.38667 | 1408.9269 | 0.29879 | 1396.9265 | 0.29879 | 1396.013 | 0.26364 | 1402.5214 | 0.26364 | 1400.58 | 0.24607 | 1401.4935 | 0.24607 | 1395.6703
| 0.4394 | 105.5749 | 0.33395 | 111.437 | 0.40425 | 109.27 | 0.40425 | 109.0059 | 0.35152 | 109.1481 | 0.42183 | 1407.4379 | 0.33395 | 1394.6423 | 0.33395 | 1395.2135 | 0.29879 | 1399.7808 | 0.29879 | 1398.5249 | 0.28122 | 1400.58 | 0.28122 | 1392.6994
| 0.47455 | 105.616 | 0.3691 | 110.7295 | 0.4394 | 108.9043 | 0.4394 | 108.7823 | 0.38667 | 108.9856 | 0.45698 | 1406.4078 | 0.3691 | 1393.3853 | 0.3691 | 1394.4138 | 0.33395 | 1396.013 | 0.33395 | 1397.0407 | 0.31637 | 1398.8674 | 0.31637 | 1392.1277
| 0.50971 | 106.0262 | 0.40425 | 110.4866 | 0.47455 | 108.3346 | 0.47455 | 108.4974 | 0.42183 | 109.473 | 0.49213 | 1403.4353 | 0.40425 | 1393.271 | 0.40425 | 1393.7282 | 0.3691 | 1393.9567 | 0.3691 | 1395.6703 | 0.35152 | 1396.013 | 0.35152 | 1390.9837
| 0.54486 | 106.2926 | 0.4394 | 110.4866 | 0.50971 | 107.866 | 0.50971 | 107.9679 | 0.45698 | 109.6557 | 0.52728 | 1402.1787 | 0.4394 | 1392.0133 | 0.4394 | 1393.1567 | 0.40425 | 1392.4707 | 0.40425 | 1394.528 | 0.38667 | 1393.1567 | 0.38667 | 1388.5786
| 0.58001 | 106.5996 | 0.47455 | 109.818 | 0.54486 | 107.4376 | 0.54486 | 107.7028 | 0.49213 | 109.676 | 0.56244 | 1401.4935 | 0.47455 | 1391.8989 | 0.47455 | 1392.8138 | 0.4394 | 1390.4114 | 0.4394 | 1393.271 | 0.42183 | 1393.1567 | 0.42183 | 1387.2023
| 0.61516 | 107.4784 | 0.50971 | 109.7571 | 0.58001 | 106.9882 | 0.58001 | 107.5396 | 0.52728 | 109.9802 | 0.59759 | 1401.4935 | 0.50971 | 1391.8989 | 0.50971 | 1389.9534 | 0.47455 | 1390.4114 | 0.47455 | 1391.7846 | 0.45698 | 1392.4707 | 0.45698 | 1384.6745
| 0.65032 | 107.56 | 0.54486 | 109.676 | 0.61516 | 106.5792 | 0.61516 | 106.5792 | 0.56244 | 109.7368 | 0.63274 | 1401.4935 | 0.54486 | 1390.9837 | 0.54486 | 1388.5786 | 0.50971 | 1389.8389 | 0.50971 | 1391.6702 | 0.49213 | 1392.1277 | 0.49213 | 1384.0992
| 0.68547 | 108.1105 | 0.58001 | 109.3512 | 0.65032 | 106.5792 | 0.65032 | 105.9852 | 0.59759 | 109.1278 | 0.66789 | 1401.4935 | 0.58001 | 1389.2662 | 0.58001 | 1387.0875 | 0.54486 | 1389.037 | 0.54486 | 1391.6702 | 0.52728 | 1391.7846 | 0.52728 | 1383.8689
| 0.72062 | 107.56 | 0.61516 | 108.6602 | 0.68547 | 106.5792 | 0.68547 | 105.8417 | 0.63274 | 108.6602 | 0.70304 | 1399.3241 | 0.61516 | 1389.2662 | 0.61516 | 1389.1516 | 0.58001 | 1388.3493 | 0.58001 | 1391.6702 | 0.56244 | 1389.037 | 0.56244 | 1383.8689
| 0.75577 | 107.213 | 0.65032 | 108.5178 | 0.72062 | 106.5792 | 0.72062 | 105.5749 | 0.66789 | 108.3753 | 0.7382 | 1400.8083 | 0.65032 | 1391.327 | 0.65032 | 1389.9534 | 0.61516 | 1387.5465 | 0.61516 | 1391.5558 | 0.59759 | 1388.0053 | 0.59759 | 1384.0992
| 0.79092 | 106.661 | 0.68547 | 108.4771 | 0.75577 | 107.3967 | 0.75577 | 105.0408 | 0.70304 | 108.3346 | 0.77335 | 1400.8083 | 0.68547 | 1390.5259 | 0.68547 | 1390.0679 | 0.65032 | 1388.3493 | 0.65032 | 1391.8989 | 0.63274 | 1385.4795 | 0.63274 | 1382.4861
| 0.82608 | 105.5955 | 0.72062 | 108.5381 | 0.79092 | 107.3967 | 0.79092 | 105.5749 | 0.7382 | 107.7232 | 0.8085 | 1400.2374 | 0.72062 | 1391.5558 | 0.72062 | 1390.0679 | 0.68547 | 1388.5786 | 0.68547 | 1389.1516 | 0.66789 | 1385.4795 | 0.66789 | 1382.4861
| 0.86123 | 105.5955 | 0.75577 | 108.6602 | 0.82608 | 107.3967 | 0.82608 | 105.3285 | 0.77335 | 107.4376 | 0.84365 | 1400.0091 | 0.75577 | 1392.5851 | 0.75577 | 1390.0679 | 0.72062 | 1388.5786 | 0.72062 | 1388.8078 | 0.70304 | 1386.5135 | 0.70304 | 1382.4861
| 0.89638 | 105.5955 | 0.79092 | 108.7619 | 0.86123 | 107.8456 | 0.86123 | 105.2875 | 0.8085 | 107.0086 | 0.87881 | 1401.9503 | 0.79092 | 1391.5558 | 0.79092 | 1390.0679 | 0.75577 | 1391.4414 | 0.75577 | 1388.3493 | 0.7382 | 1387.5465 | 0.7382 | 1383.4082
| 0.93153 | 106.661 | 0.82608 | 109.3309 | 0.89638 | 108.5585 | 0.89638 | 105.7801 | 0.84365 | 106.5996 | 0.91396 | 1400.8083 | 0.82608 | 1391.327 | 0.82608 | 1386.8579 | 0.79092 | 1392.1277 | 0.79092 | 1388.0053 | 0.77335 | 1388.4639 | 0.77335 | 1384.6745
| 0.96669 | 107.0904 | 0.86123 | 109.9193 | 0.93153 | 108.029 | 0.93153 | 106.0467 | 0.87881 | 106.5996 | 0.94911 | 1400.2374 | 0.86123 | 1391.2125 | 0.86123 | 1387.7759 | 0.82608 | 1390.8692 | 0.82608 | 1387.5465 | 0.8085 | 1389.8389 | 0.8085 | 1384.6745
| 1.0018 | 107.2334 | 0.89638 | 110.7902 | 0.96669 | 108.5788 | 0.96669 | 105.9647 | 0.91396 | 106.9269 | 0.98426 | 1397.7257 | 0.89638 | 1387.317 | 0.89638 | 1386.6283 | 0.86123 | 1390.4114 | 0.86123 | 1385.9392 | 0.84365 | 1389.8389 | 0.84365 | 1383.4082
| 1.037 | 107.3763 | 0.93153 | 110.7902 | 1.0018 | 108.029 | 1.0018 | 106.4973 | 0.94911 | 106.5178 | 1.0194 | 1396.9265 | 0.93153 | 1385.5944 | 0.93153 | 1383.8689 | 0.89638 | 1388.2346 | 0.89638 | 1385.9392 | 0.87881 | 1388.12 | 0.87881 | 1383.0625
| 1.0721 | 107.3151 | 0.96669 | 110.2639 | 1.037 | 107.5396 | 1.037 | 107.2538 | 0.98426 | 107.0086 | 1.0546 | 1395.0992 | 0.96669 | 1385.9392 | 0.96669 | 1383.8689 | 0.93153 | 1384.0992 | 0.93153 | 1384.4444 | 0.91396 | 1386.9727 | 0.91396 | 1380.2917
| 1.1073 | 107.3151 | 1.0018 | 109.676 | 1.0721 | 107.0904 | 1.0721 | 107.6008 | 1.0194 | 107.2742 | 1.0897 | 1396.9265 | 1.0018 | 1385.3645 | 1.0018 | 1384.3294 | 0.96669 | 1383.8689 | 0.96669 | 1383.8689 | 0.94911 | 1385.2496 | 0.94911 | 1375.7658
| 1.1424 | 107.3967 | 1.037 | 109.27 | 1.1073 | 107.0495 | 1.1073 | 107.0086 | 1.0546 | 107.9271 | 1.1249 | 1398.8674 | 1.037 | 1386.5135 | 1.037 | 1384.3294 | 1.0018 | 1382.2554 | 1.0018 | 1383.4082 | 0.98426 | 1381.2164 | 0.98426 | 1375.5328
| 1.1776 | 107.3151 | 1.0721 | 108.9652 | 1.1424 | 107.0904 | 1.1424 | 106.8451 | 1.0897 | 109.1278 | 1.16 | 1399.2099 | 1.0721 | 1388.0053 | 1.0721 | 1385.2496 | 1.037 | 1382.2554 | 1.037 | 1382.7167 | 1.0194 | 1381.2164 | 1.0194 | 1372.3782
| 1.2128 | 107.1517 | 1.1073 | 108.6399 | 1.1776 | 107.5396 | 1.1776 | 106.4564 | 1.1249 | 109.1278 | 1.1952 | 1399.5524 | 1.1073 | 1388.2346 | 1.1073 | 1385.7094 | 1.0721 | 1383.8689 | 1.0721 | 1382.6014 | 1.0546 | 1382.4861 | 1.0546 | 1373.7825
| 1.2479 | 107.56 | 1.1424 | 108.6399 | 1.2128 | 108.0698 | 1.2128 | 105.1231 | 1.16 | 109.0669 | 1.2303 | 1399.5524 | 1.1424 | 1389.037 | 1.1424 | 1386.3987 | 1.1073 | 1384.9046 | 1.1073 | 1383.6386 | 1.0897 | 1383.6386 | 1.0897 | 1375.7658
| 1.2831 | 107.9271 | 1.1776 | 109.0872 | 1.2479 | 108.3346 | 1.2479 | 105.0408 | 1.1952 | 108.4974 | 1.2655 | 1399.7808 | 1.1776 | 1389.2662 | 1.1776 | 1386.0541 | 1.1424 | 1387.2023 | 1.1424 | 1384.9046 | 1.1249 | 1384.5595 | 1.1249 | 1378.0907
| 1.3182 | 109.818 | 1.2128 | 109.6963 | 1.2831 | 108.7212 | 1.2831 | 105.0408 | 1.2303 | 108.2124 | 1.3006 | 1400.9225 | 1.2128 | 1389.1516 | 1.2128 | 1385.3645 | 1.1776 | 1388.2346 | 1.1776 | 1386.3987 | 1.16 | 1384.6745 | 1.16 | 1380.4074
| 1.3534 | 110.4056 | 1.2479 | 110.0815 | 1.3182 | 108.7212 | 1.3182 | 105.7596 | 1.2655 | 107.6008 | 1.3358 | 1403.4353 | 1.2479 | 1388.0053 | 1.2479 | 1388.2346 | 1.2128 | 1388.8078 | 1.2128 | 1388.5786 | 1.1952 | 1385.2496 | 1.1952 | 1380.7542
| 1.3885 | 110.7093 | 1.2831 | 110.1828 | 1.3534 | 108.7212 | 1.3534 | 106.1492 | 1.3006 | 106.9882 | 1.3709 | 1403.6638 | 1.2831 | 1390.4114 | 1.2831 | 1390.0679 | 1.2479 | 1389.9534 | 1.2479 | 1387.4318 | 1.2303 | 1385.2496 | 1.2303 | 1380.6386
| 1.4237 | 111.0733 | 1.3182 | 110.0815 | 1.3885 | 108.7212 | 1.3885 | 107.2947 | 1.3358 | 106.7633 | 1.4061 | 1404.3495 | 1.3182 | 1391.4414 | 1.3182 | 1390.5259 | 1.2831 | 1390.4114 | 1.2831 | 1389.3807 | 1.2655 | 1385.5944 | 1.2655 | 1380.0604
| 1.4588 | 110.9722 | 1.3534 | 110.2841 | 1.4237 | 108.7212 | 1.4237 | 107.3559 | 1.3709 | 106.9882 | 1.4412 | 1401.2651 | 1.3534 | 1391.4414 | 1.3534 | 1391.4414 | 1.3182 | 1390.7548 | 1.3182 | 1390.7548 | 1.3006 | 1386.6283 | 1.3006 | 1379.8289
| 1.494 | 111.0733 | 1.3885 | 110.2841 | 1.4588 | 108.9652 | 1.4588 | 107.3559 | 1.4061 | 107.1721 | 1.4764 | 1401.2651 | 1.3885 | 1393.3853 | 1.3885 | 1391.5558 | 1.3534 | 1390.9837 | 1.3534 | 1393.8425 | 1.3358 | 1388.3493 | 1.3358 | 1379.8289
| 1.5291 | 110.9318 | 1.4237 | 110.9116 | 1.494 | 109.7165 | 1.494 | 107.3559 | 1.4412 | 108.4771 | 1.5115 | 1404.3495 | 1.4237 | 1394.7565 | 1.4237 | 1392.1277 | 1.3885 | 1392.1277 | 1.3885 | 1394.4138 | 1.3709 | 1390.2969 | 1.3709 | 1379.8289
| 1.5643 | 110.9318 | 1.4588 | 111.538 | 1.5291 | 110.77 | 1.5291 | 107.458 | 1.4764 | 108.6602 | 1.5467 | 1405.6071 | 1.4588 | 1394.7565 | 1.4588 | 1393.6139 | 1.4237 | 1391.5558 | 1.4237 | 1394.528 | 1.4061 | 1389.4953 | 1.4061 | 1382.7167
| 1.5994 | 110.7902 | 1.494 | 111.7399 | 1.5643 | 110.77 | 1.5643 | 107.213 | 1.5115 | 108.7619 | 1.5818 | 1406.0646 | 1.494 | 1394.7565 | 1.494 | 1393.9567 | 1.4588 | 1392.1277 | 1.4588 | 1392.8138 | 1.4412 | 1388.9224 | 1.4412 | 1385.2496
| 1.6346 | 110.7902 | 1.5291 | 112.1837 | 1.5994 | 111.639 | 1.5994 | 107.458 | 1.5467 | 108.7619 | 1.617 | 1406.0646 | 1.5291 | 1396.3555 | 1.5291 | 1394.7565 | 1.494 | 1392.0133 | 1.494 | 1394.4138 | 1.4764 | 1390.4114 | 1.4764 | 1385.2496
| 1.6697 | 110.4866 | 1.5643 | 113.0297 | 1.6346 | 112.224 | 1.6346 | 107.9068 | 1.5818 | 108.8839 | 1.6522 | 1406.2934 | 1.5643 | 1396.3555 | 1.5643 | 1398.2966 | 1.5291 | 1392.0133 | 1.5291 | 1396.5839 | 1.5115 | 1389.8389 | 1.5115 | 1385.1346
| 1.7049 | 110.9116 | 1.5994 | 113.5928 | 1.6697 | 112.2643 | 1.6697 | 108.3346 | 1.617 | 108.8026 | 1.6873 | 1405.264 | 1.5994 | 1396.3555 | 1.5994 | 1401.1509 | 1.5643 | 1392.4707 | 1.5643 | 1397.4974 | 1.5467 | 1388.9224 | 1.5467 | 1385.0196
| 1.74 | 111.2754 | 1.6346 | 114.7971 | 1.7049 | 111.9618 | 1.7049 | 109.6963 | 1.6522 | 108.8026 | 1.7225 | 1402.6356 | 1.6346 | 1396.2413 | 1.6346 | 1400.8083 | 1.5994 | 1394.6423 | 1.5994 | 1397.9541 | 1.5818 | 1391.5558 | 1.5818 | 1385.0196
| 1.7752 | 111.6995 | 1.6697 | 115.0176 | 1.74 | 111.9618 | 1.74 | 110.3044 | 1.6873 | 108.9856 | 1.7576 | 1402.9783 | 1.6697 | 1395.3277 | 1.6697 | 1398.7532 | 1.6346 | 1396.2413 | 1.6346 | 1397.9541 | 1.617 | 1393.0424 | 1.617 | 1385.2496
| 1.8103 | 112.224 | 1.7049 | 115.0176 | 1.7752 | 111.8004 | 1.7752 | 111.3562 | 1.7225 | 109.6354 | 1.7928 | 1406.2934 | 1.7049 | 1391.8989 | 1.7049 | 1397.6116 | 1.6697 | 1398.7532 | 1.6697 | 1397.9541 | 1.6522 | 1393.1567 | 1.6522 | 1386.8579
| 1.8455 | 112.224 | 1.74 | 114.6568 | 1.8103 | 111.7197 | 1.8103 | 111.5178 | 1.7576 | 109.9599 | 1.8279 | 1407.8959 | 1.74 | 1389.7244 | 1.74 | 1394.6423 | 1.7049 | 1398.2966 | 1.7049 | 1394.071 | 1.6873 | 1394.4138 | 1.6873 | 1387.317
| 1.8806 | 112.2442 | 1.7752 | 114.2154 | 1.8455 | 111.7197 | 1.8455 | 111.3562 | 1.7928 | 111.0127 | 1.8631 | 1410.1882 | 1.7752 | 1389.1516 | 1.7752 | 1391.7846 | 1.74 | 1398.2966 | 1.74 | 1393.6139 | 1.7225 | 1396.5839 | 1.7225 | 1387.317
| 1.9158 | 112.5465 | 1.8103 | 114.115 | 1.8806 | 111.7197 | 1.8806 | 110.9924 | 1.8279 | 111.6188 | 1.8982 | 1410.1882 | 1.8103 | 1387.2023 | 1.8103 | 1390.7548 | 1.7752 | 1398.7532 | 1.7752 | 1394.985 | 1.7576 | 1394.4138 | 1.7576 | 1388.12
| 1.9509 | 112.6472 | 1.8455 | 113.633 | 1.9158 | 111.7601 | 1.9158 | 110.3651 | 1.8631 | 112.0224 | 1.9334 | 1410.1882 | 1.8455 | 1387.6612 | 1.8455 | 1391.4414 | 1.8103 | 1400.3516 | 1.8103 | 1398.4107 | 1.7928 | 1394.4138 | 1.7928 | 1388.12
| 1.9861 | 113.1303 | 1.8806 | 113.3716 | 1.9509 | 112.0627 | 1.9509 | 110.0815 | 1.8982 | 112.9089 | 1.9685 | 1410.1882 | 1.8806 | 1393.271 | 1.8806 | 1394.985 | 1.8455 | 1401.3793 | 1.8455 | 1398.9816 | 1.8279 | 1395.7845 | 1.8279 | 1388.8078
| 2.0213 | 113.2309 | 1.9158 | 113.1303 | 1.9861 | 112.1837 | 1.9861 | 110.0004 | 1.9334 | 112.9089 | 2.0037 | 1410.1882 | 1.9158 | 1402.864 | 1.9158 | 1397.3832 | 1.8806 | 1402.0645 | 1.8806 | 1399.8949 | 1.8631 | 1395.8987 | 1.8631 | 1389.3807
| 2.0564 | 113.2107 | 1.9509 | 113.4923 | 2.0213 | 112.4659 | 2.0213 | 109.8788 | 1.9685 | 112.9089 | 2.0388 | 1409.9588 | 1.9509 | 1407.2089 | 1.9509 | 1397.3832 | 1.9158 | 1402.0645 | 1.9158 | 1399.8949 | 1.8982 | 1395.8987 | 1.8982 | 1389.7244
| 2.0916 | 113.2309 | 1.9861 | 114.0347 | 2.0564 | 113.2912 | 2.0564 | 110.3854 | 2.0037 | 112.2442 | 2.074 | 1410.1882 | 1.9861 | 1407.2089 | 1.9861 | 1403.7781 | 1.9509 | 1399.8949 | 1.9509 | 1400.6942 | 1.9334 | 1396.013 | 1.9334 | 1391.327
| 2.1267 | 113.2107 | 2.0213 | 114.2556 | 2.0916 | 113.5526 | 2.0916 | 110.5069 | 2.0388 | 112.0627 | 2.1091 | 1411.4509 | 2.0213 | 1407.2089 | 2.0213 | 1404.2352 | 1.9861 | 1399.0957 | 1.9861 | 1400.6942 | 1.9685 | 1396.9265 | 1.9685 | 1392.1277
| 2.1619 | 113.0699 | 2.0564 | 115.3982 | 2.1267 | 113.6933 | 2.1267 | 110.5069 | 2.074 | 111.5178 | 2.1443 | 1412.6 | 2.0564 | 1404.4638 | 2.0564 | 1401.7219 | 2.0213 | 1399.0957 | 2.0213 | 1399.4382 | 2.0037 | 1397.1548 | 2.0037 | 1392.1277
| 2.197 | 113.1303 | 2.0916 | 115.7585 | 2.1619 | 114.1753 | 2.1619 | 110.7093 | 2.1091 | 111.1137 | 2.1794 | 1414.2112 | 2.0916 | 1400.9225 | 2.0916 | 1395.2135 | 2.0564 | 1397.7257 | 2.0564 | 1399.0957 | 2.0388 | 1397.1548 | 2.0388 | 1392.242
| 2.2322 | 113.1303 | 2.1267 | 115.7585 | 2.197 | 114.3358 | 2.197 | 110.8711 | 2.1443 | 111.1137 | 2.2146 | 1414.4416 | 2.1267 | 1400.58 | 2.1267 | 1399.4382 | 2.0916 | 1396.1272 | 2.0916 | 1399.2099 | 2.074 | 1397.1548 | 2.074 | 1392.242
| 2.2673 | 112.5666 | 2.1619 | 115.7585 | 2.2322 | 114.3559 | 2.2322 | 110.8711 | 2.1794 | 111.982 | 2.2497 | 1415.2485 | 2.1619 | 1400.9225 | 2.1619 | 1401.7219 | 2.1267 | 1401.7219 | 2.1267 | 1399.7808 | 2.1091 | 1397.7257 | 2.1091 | 1392.242
| 2.3025 | 112.4256 | 2.197 | 116.5584 | 2.2673 | 114.5164 | 2.2673 | 110.9924 | 2.2146 | 113.0095 | 2.2849 | 1418.1371 | 2.197 | 1403.2068 | 2.197 | 1401.7219 | 2.1619 | 1404.1209 | 2.1619 | 1400.2374 | 2.1443 | 1397.7257 | 2.1443 | 1393.9567
| 2.3376 | 112.5062 | 2.2322 | 116.2986 | 2.3025 | 114.376 | 2.3025 | 111.1744 | 2.2497 | 113.1102 | 2.32 | 1418.1371 | 2.2322 | 1404.6924 | 2.2322 | 1403.4353 | 2.197 | 1401.2651 | 2.197 | 1401.8361 | 2.1794 | 1398.0682 | 2.1794 | 1394.4138
| 2.3728 | 112.2038 | 2.2673 | 116.3585 | 2.3376 | 114.376 | 2.3376 | 111.5582 | 2.2849 | 113.1303 | 2.3552 | 1417.6741 | 2.2673 | 1408.5832 | 2.2673 | 1404.8067 | 2.2322 | 1404.1209 | 2.2322 | 1402.864 | 2.2146 | 1399.5524 | 2.2146 | 1395.5561
| 2.4079 | 111.8811 | 2.3025 | 116.3585 | 2.3728 | 114.2556 | 2.3728 | 112.3853 | 2.32 | 113.1504 | 2.3903 | 1418.2529 | 2.3025 | 1411.336 | 2.3025 | 1405.264 | 2.2673 | 1405.0354 | 2.2673 | 1405.0354 | 2.2497 | 1400.58 | 2.2497 | 1396.8123
| 2.4431 | 112.6673 | 2.3376 | 116.5584 | 2.4079 | 114.1351 | 2.4079 | 112.6069 | 2.3552 | 113.5124 | 2.4255 | 1418.3687 | 2.3376 | 1412.1402 | 2.3376 | 1407.2089 | 2.3025 | 1405.0354 | 2.3025 | 1407.0944 | 2.2849 | 1401.4935 | 2.2849 | 1397.6116
| 2.4782 | 112.6673 | 2.3728 | 116.2986 | 2.4431 | 113.6129 | 2.4431 | 112.6069 | 2.3903 | 113.5124 | 2.4607 | 1418.3687 | 2.3728 | 1414.9026 | 2.3728 | 1406.7511 | 2.3376 | 1406.0646 | 2.3376 | 1408.354 | 2.32 | 1404.1209 | 2.32 | 1400.9225
| 2.5134 | 112.6673 | 2.4079 | 116.2386 | 2.4782 | 113.8741 | 2.4782 | 112.8083 | 2.4255 | 113.8339 | 2.4958 | 1416.6336 | 2.4079 | 1414.9026 | 2.4079 | 1406.7511 | 2.3728 | 1410.762 | 2.3728 | 1408.1249 | 2.3552 | 1405.0354 | 2.3552 | 1401.1509
| 2.5485 | 112.6673 | 2.4431 | 116.1386 | 2.5134 | 113.4923 | 2.5134 | 112.8083 | 2.4607 | 114.1351 | 2.531 | 1418.832 | 2.4431 | 1414.9026 | 2.4431 | 1406.5222 | 2.4079 | 1409.7294 | 2.4079 | 1410.0735 | 2.3903 | 1405.3784 | 2.3903 | 1402.6356
| 2.5837 | 112.6673 | 2.4782 | 115.7185 | 2.5485 | 112.7277 | 2.5485 | 112.6472 | 2.4958 | 114.1351 | 2.5661 | 1418.3687 | 2.4782 | 1412.1402 | 2.4782 | 1405.4927 | 2.4431 | 1410.0735 | 2.4431 | 1408.1249 | 2.4255 | 1409.8441 | 2.4255 | 1403.0925
| 2.6188 | 112.224 | 2.5134 | 115.258 | 2.5837 | 112.929 | 2.5837 | 112.1837 | 2.531 | 114.0147 | 2.6013 | 1420.9213 | 2.5134 | 1408.5832 | 2.5134 | 1404.5781 | 2.4782 | 1410.0735 | 2.4782 | 1406.179 | 2.4607 | 1409.8441 | 2.4607 | 1402.9783
| 2.654 | 112.103 | 2.5485 | 115.0377 | 2.6188 | 112.8888 | 2.6188 | 112.0627 | 2.5661 | 114.3358 | 2.6364 | 1423.3677 | 2.5485 | 1409.8441 | 2.5485 | 1406.4078 | 2.5134 | 1408.4686 | 2.5134 | 1407.4379 | 2.4958 | 1407.4379 | 2.4958 | 1403.6638
| 2.6891 | 111.5582 | 2.5837 | 115.1579 | 2.654 | 112.5666 | 2.654 | 111.7197 | 2.6013 | 114.3358 | 2.6716 | 1423.7181 | 2.5837 | 1412.0253 | 2.5837 | 1407.5523 | 2.5485 | 1406.8655 | 2.5485 | 1407.4379 | 2.531 | 1407.5523 | 2.531 | 1403.6638
| 2.7243 | 111.4774 | 2.6188 | 115.258 | 2.6891 | 112.7881 | 2.6891 | 111.7197 | 2.6364 | 114.0347 | 2.7067 | 1423.7181 | 2.6188 | 1412.0253 | 2.6188 | 1408.354 | 2.5837 | 1408.5832 | 2.5837 | 1408.9269 | 2.5661 | 1407.7814 | 2.5661 | 1403.0925
| 2.7594 | 111.3966 | 2.654 | 115.4383 | 2.7243 | 112.9089 | 2.7243 | 111.7197 | 2.6716 | 114.0347 | 2.7419 | 1424.0686 | 2.654 | 1411.7955 | 2.654 | 1408.354 | 2.6188 | 1409.8441 | 2.6188 | 1410.762 | 2.6013 | 1407.5523 | 2.6013 | 1401.6077
| 2.7946 | 111.336 | 2.6891 | 115.4583 | 2.7594 | 112.9089 | 2.7594 | 112.1434 | 2.7067 | 114.2355 | 2.777 | 1423.1343 | 2.6891 | 1412.0253 | 2.6891 | 1408.2395 | 2.654 | 1407.7814 | 2.654 | 1408.9269 | 2.6364 | 1409.2708 | 2.6364 | 1401.6077
| 2.8298 | 111.3966 | 2.7243 | 115.3382 | 2.7946 | 112.9089 | 2.7946 | 112.2442 | 2.7419 | 113.7334 | 2.8122 | 1423.0176 | 2.7243 | 1412.0253 | 2.7243 | 1407.3234 | 2.6891 | 1408.5832 | 2.6891 | 1411.7955 | 2.6716 | 1409.1562 | 2.6716 | 1399.4382
| 2.8649 | 111.5784 | 2.7594 | 115.258 | 2.8298 | 112.9089 | 2.8298 | 112.3651 | 2.777 | 113.5928 | 2.8473 | 1423.0176 | 2.7594 | 1412.6 | 2.7594 | 1407.3234 | 2.7243 | 1412.0253 | 2.7243 | 1413.0601 | 2.7067 | 1407.7814 | 2.7067 | 1399.0957
| 2.9001 | 111.134 | 2.7946 | 115.0978 | 2.8649 | 112.5062 | 2.8649 | 112.3046 | 2.8122 | 113.5928 | 2.8825 | 1423.4845 | 2.7946 | 1413.0601 | 2.7946 | 1407.6668 | 2.7594 | 1410.3029 | 2.7594 | 1410.762 | 2.7419 | 1411.2212 | 2.7419 | 1399.4382
| 2.9352 | 111.134 | 2.8298 | 114.8773 | 2.9001 | 112.3651 | 2.9001 | 112.3046 | 2.8473 | 114.0347 | 2.9176 | 1424.3025 | 2.8298 | 1415.1332 | 2.8298 | 1408.2395 | 2.7946 | 1410.3029 | 2.7946 | 1410.762 | 2.777 | 1410.6472 | 2.777 | 1400.2374
| 2.9704 | 111.982 | 2.8649 | 114.737 | 2.9352 | 112.1837 | 2.9352 | 112.2442 | 2.8825 | 113.9946 | 2.9528 | 1428.1759 | 2.8649 | 1415.1332 | 2.8649 | 1413.5204 | 2.8298 | 1412.0253 | 2.8298 | 1410.8768 | 2.8122 | 1404.3495 | 2.8122 | 1400.8083
| 3.0055 | 112.1434 | 2.9001 | 113.9544 | 2.9704 | 112.1837 | 2.9704 | 112.1232 | 2.9176 | 114.0749 | 2.9879 | 1428.1759 | 2.9001 | 1416.2871 | 2.9001 | 1414.9026 | 2.8649 | 1413.1751 | 2.8649 | 1410.762 | 2.8473 | 1410.6472 | 2.8473 | 1401.2651
| 3.0407 | 112.2442 | 2.9352 | 113.9544 | 3.0055 | 112.345 | 3.0055 | 111.4572 | 2.9528 | 114.0749 | 3.0231 | 1428.1759 | 2.9352 | 1416.7492 | 2.9352 | 1417.5585 | 2.9001 | 1411.7955 | 2.9001 | 1410.8768 | 2.8825 | 1416.5181 | 2.8825 | 1401.7219
| 3.0758 | 112.2845 | 2.9704 | 113.5928 | 3.0407 | 112.345 | 3.0407 | 110.9722 | 2.9879 | 113.854 | 3.0582 | 1426.764 | 2.9704 | 1416.8647 | 2.9704 | 1416.0562 | 2.9352 | 1411.7955 | 2.9352 | 1419.8758 | 2.9176 | 1416.5181 | 2.9176 | 1401.9503
| 3.111 | 112.8485 | 3.0055 | 113.9544 | 3.0758 | 112.4659 | 3.0758 | 110.9722 | 3.0231 | 113.6531 | 3.0934 | 1425.7076 | 3.0055 | 1417.0959 | 3.0055 | 1415.0179 | 2.9704 | 1412.6 | 2.9704 | 1425.7076 | 2.9528 | 1418.4845 | 2.9528 | 1401.9503
| 3.1461 | 112.1434 | 3.0407 | 114.396 | 3.111 | 112.5465 | 3.111 | 110.8913 | 3.0582 | 112.8284 | 3.1285 | 1423.7181 | 3.0407 | 1417.0959 | 3.0407 | 1414.7873 | 3.0055 | 1412.0253 | 3.0055 | 1418.9479 | 2.9879 | 1422.6677 | 2.9879 | 1403.7781
| 3.1813 | 111.8408 | 3.0758 | 114.4362 | 3.1461 | 113.1504 | 3.1461 | 110.8913 | 3.0934 | 112.4256 | 3.1637 | 1422.4346 | 3.0758 | 1417.0959 | 3.0758 | 1413.2902 | 3.0407 | 1412.0253 | 3.0407 | 1412.1402 | 3.0231 | 1419.5277 | 3.0231 | 1404.921
| 3.2164 | 111.336 | 3.111 | 114.4362 | 3.1813 | 113.0297 | 3.1813 | 110.952 | 3.1285 | 112.2643 | 3.1989 | 1417.9056 | 3.111 | 1416.4026 | 3.111 | 1412.83 | 3.0758 | 1411.7955 | 3.0758 | 1410.6472 | 3.0582 | 1417.9056 | 3.0582 | 1405.7215
| 3.2516 | 111.4168 | 3.1461 | 115.0577 | 3.2164 | 113.3716 | 3.2164 | 111.4572 | 3.1637 | 112.0627 | 3.234 | 1419.5277 | 3.1461 | 1414.9026 | 3.1461 | 1411.5657 | 3.111 | 1412.6 | 3.111 | 1409.2708 | 3.0934 | 1417.9056 | 3.0934 | 1405.1497
| 3.2867 | 111.336 | 3.1813 | 115.2981 | 3.2516 | 114.0749 | 3.2516 | 111.4572 | 3.1989 | 112.224 | 3.2692 | 1417.3272 | 3.1813 | 1411.7955 | 3.1813 | 1409.5001 | 3.1461 | 1411.5657 | 3.1461 | 1407.7814 | 3.1285 | 1421.7357 | 3.1285 | 1405.7215
| 3.3219 | 111.8408 | 3.2164 | 115.4383 | 3.2867 | 113.7736 | 3.2867 | 111.4572 | 3.234 | 112.2643 | 3.3043 | 1415.5946 | 3.2164 | 1411.336 | 3.2164 | 1409.2708 | 3.1813 | 1411.5657 | 3.1813 | 1409.2708 | 3.1637 | 1419.5277 | 3.1637 | 1404.5781
| 3.357 | 112.5263 | 3.2516 | 115.5984 | 3.3219 | 113.633 | 3.3219 | 111.5986 | 3.2692 | 112.5666 | 3.3395 | 1413.5204 | 3.2516 | 1409.0415 | 3.2516 | 1408.1249 | 3.2164 | 1410.9916 | 3.2164 | 1409.2708 | 3.1989 | 1414.9026 | 3.1989 | 1402.4071
| 3.3922 | 112.5263 | 3.2867 | 115.5984 | 3.357 | 113.5325 | 3.357 | 111.6995 | 3.3043 | 112.627 | 3.3746 | 1409.7294 | 3.2867 | 1408.2395 | 3.2867 | 1407.4379 | 3.2516 | 1408.5832 | 3.2516 | 1405.7215 | 3.234 | 1414.5568 | 3.234 | 1400.2374
| 3.4273 | 111.6591 | 3.3219 | 115.5184 | 3.3922 | 112.9693 | 3.3922 | 112.3853 | 3.3395 | 112.7076 | 3.4098 | 1407.3234 | 3.3219 | 1404.4638 | 3.3219 | 1405.4927 | 3.2867 | 1405.9502 | 3.2867 | 1404.4638 | 3.2692 | 1408.2395 | 3.2692 | 1398.9816
| 3.4625 | 112.2643 | 3.357 | 115.1979 | 3.4273 | 112.486 | 3.4273 | 112.4659 | 3.3746 | 113.0095 | 3.4449 | 1407.2089 | 3.357 | 1400.58 | 3.357 | 1396.6981 | 3.3219 | 1402.1787 | 3.3219 | 1403.2068 | 3.3043 | 1402.7498 | 3.3043 | 1396.8123
| 3.4976 | 111.8004 | 3.3922 | 114.6167 | 3.4625 | 112.4659 | 3.4625 | 112.3853 | 3.4098 | 112.627 | 3.4801 | 1407.2089 | 3.3922 | 1397.4974 | 3.3922 | 1396.3555 | 3.357 | 1399.7808 | 3.357 | 1399.0957 | 3.3395 | 1402.5214 | 3.3395 | 1392.9281
| 3.5328 | 111.6591 | 3.4273 | 114.4763 | 3.4976 | 112.4054 | 3.4976 | 112.4659 | 3.4449 | 113.633 | 3.5152 | 1406.7511 | 3.4273 | 1396.6981 | 3.4273 | 1395.6703 | 3.3922 | 1397.6116 | 3.3922 | 1397.0407 | 3.3746 | 1398.0682 | 3.3746 | 1391.6702
| 3.5679 | 111.8811 | 3.4625 | 114.6167 | 3.5328 | 112.2845 | 3.5328 | 112.345 | 3.4801 | 113.6933 | 3.5504 | 1406.8655 | 3.4625 | 1396.6981 | 3.4625 | 1395.6703 | 3.4273 | 1396.1272 | 3.4273 | 1394.7565 | 3.4098 | 1397.0407 | 3.4098 | 1391.6702
| 3.6031 | 111.8811 | 3.4976 | 114.6167 | 3.5679 | 112.2643 | 3.5679 | 111.3764 | 3.5152 | 113.6129 | 3.5855 | 1406.7511 | 3.4976 | 1395.8987 | 3.4976 | 1395.8987 | 3.4625 | 1395.2135 | 3.4625 | 1394.2995 | 3.4449 | 1395.8987 | 3.4449 | 1390.4114
| 3.6383 | 111.5784 | 3.5328 | 114.4963 | 3.6031 | 112.2845 | 3.6031 | 110.8104 | 3.5504 | 113.251 | 3.6207 | 1404.8067 | 3.5328 | 1396.9265 | 3.5328 | 1396.2413 | 3.4976 | 1392.6994 | 3.4976 | 1395.0992 | 3.4801 | 1391.4414 | 3.4801 | 1389.1516
| 3.6734 | 111.2956 | 3.5679 | 114.6367 | 3.6383 | 112.5867 | 3.6383 | 110.8104 | 3.5855 | 113.3314 | 3.6558 | 1404.1209 | 3.5679 | 1396.9265 | 3.5679 | 1395.4419 | 3.5328 | 1392.0133 | 3.5328 | 1396.3555 | 3.5152 | 1394.7565 | 3.5152 | 1388.4639
| 3.7086 | 111.2956 | 3.6031 | 114.6768 | 3.6734 | 112.6874 | 3.6734 | 110.8104 | 3.6207 | 112.4256 | 3.691 | 1404.1209 | 3.6031 | 1394.7565 | 3.6031 | 1395.2135 | 3.5679 | 1392.0133 | 3.5679 | 1396.3555 | 3.5504 | 1395.8987 | 3.5504 | 1385.9392
| 3.7437 | 110.9924 | 3.6383 | 114.6568 | 3.7086 | 112.7479 | 3.7086 | 110.8711 | 3.6558 | 112.345 | 3.7261 | 1403.321 | 3.6383 | 1394.1853 | 3.6383 | 1395.0992 | 3.6031 | 1389.3807 | 3.6031 | 1396.3555 | 3.5855 | 1390.9837 | 3.5855 | 1385.3645
| 3.7789 | 110.9924 | 3.6734 | 114.6768 | 3.7437 | 112.5867 | 3.7437 | 111.1137 | 3.691 | 112.345 | 3.7613 | 1403.0925 | 3.6734 | 1395.7845 | 3.6734 | 1394.6423 | 3.6383 | 1389.3807 | 3.6383 | 1395.7845 | 3.6207 | 1390.9837 | 3.6207 | 1384.3294
| 3.814 | 111.134 | 3.7086 | 114.6768 | 3.7789 | 112.4256 | 3.7789 | 111.4976 | 3.7261 | 112.345 | 3.7964 | 1403.321 | 3.7086 | 1394.7565 | 3.7086 | 1395.4419 | 3.6734 | 1389.3807 | 3.6734 | 1395.0992 | 3.6558 | 1389.4953 | 3.6558 | 1383.293
| 3.8492 | 111.538 | 3.7437 | 114.0749 | 3.814 | 112.3248 | 3.814 | 111.5582 | 3.7613 | 112.7881 | 3.8316 | 1403.321 | 3.7437 | 1395.5561 | 3.7437 | 1395.7845 | 3.7086 | 1391.327 | 3.7086 | 1393.3853 | 3.691 | 1389.4953 | 3.691 | 1384.0992
| 3.8843 | 111.6793 | 3.7789 | 113.4722 | 3.8492 | 111.9417 | 3.8492 | 111.4976 | 3.7964 | 112.8888 | 3.8667 | 1403.0925 | 3.7789 | 1396.2413 | 3.7789 | 1398.0682 | 3.7437 | 1394.4138 | 3.7437 | 1392.9281 | 3.7261 | 1390.7548 | 3.7261 | 1384.0992
| 3.9195 | 111.6793 | 3.814 | 113.5727 | 3.8843 | 111.8408 | 3.8843 | 111.4976 | 3.8316 | 113.1705 | 3.9019 | 1404.4638 | 3.814 | 1396.2413 | 3.814 | 1398.6391 | 3.7789 | 1394.985 | 3.7789 | 1392.6994 | 3.7613 | 1392.0133 | 3.7613 | 1384.9046
| 3.9546 | 111.336 | 3.8492 | 113.6129 | 3.9195 | 111.7803 | 3.9195 | 111.3158 | 3.8667 | 113.2309 | 3.937 | 1406.6367 | 3.8492 | 1397.269 | 3.8492 | 1398.6391 | 3.814 | 1394.985 | 3.814 | 1394.7565 | 3.7964 | 1392.242 | 3.7964 | 1385.9392
| 3.9898 | 111.3562 | 3.8843 | 113.7937 | 3.9546 | 111.6188 | 3.9546 | 111.3158 | 3.9019 | 113.2309 | 3.9722 | 1407.8959 | 3.8843 | 1396.8123 | 3.8843 | 1399.3241 | 3.8492 | 1394.985 | 3.8492 | 1395.8987 | 3.8316 | 1393.8425 | 3.8316 | 1386.7431
| 4.0249 | 111.4976 | 3.9195 | 114.0147 | 3.9898 | 111.2754 | 3.9898 | 110.7295 | 3.937 | 112.8284 | 4.0074 | 1409.9588 | 3.9195 | 1397.6116 | 3.9195 | 1399.7808 | 3.8843 | 1396.013 | 3.8843 | 1398.2966 | 3.8667 | 1395.3277 | 3.8667 | 1387.8906
| 4.0601 | 110.9722 | 3.9546 | 113.9142 | 4.0249 | 110.6081 | 4.0249 | 110.7295 | 3.9722 | 112.8284 | 4.0425 | 1411.5657 | 3.9546 | 1401.4935 | 3.9546 | 1400.9225 | 3.9195 | 1396.013 | 3.9195 | 1400.0091 | 3.9019 | 1395.8987 | 3.9019 | 1389.3807
| 4.0952 | 110.77 | 3.9898 | 113.7736 | 4.0601 | 111.1946 | 4.0601 | 110.3651 | 4.0074 | 112.8083 | 4.0777 | 1412.3701 | 3.9898 | 1402.1787 | 3.9898 | 1401.6077 | 3.9546 | 1398.1824 | 3.9546 | 1400.6942 | 3.937 | 1396.2413 | 3.937 | 1390.5259
| 4.1304 | 110.8509 | 4.0249 | 113.5325 | 4.0952 | 111.5986 | 4.0952 | 110.3651 | 4.0425 | 112.8284 | 4.1128 | 1413.9809 | 4.0249 | 1404.921 | 4.0249 | 1404.0066 | 3.9898 | 1399.6666 | 3.9898 | 1401.8361 | 3.9722 | 1396.2413 | 3.9722 | 1391.2125
| 4.1655 | 110.4259 | 4.0601 | 113.5928 | 4.1304 | 112.2038 | 4.1304 | 109.8788 | 4.0777 | 112.8083 | 4.148 | 1414.9026 | 4.0601 | 1406.2934 | 4.0601 | 1404.1209 | 4.0249 | 1400.3516 | 4.0249 | 1402.7498 | 4.0074 | 1396.3555 | 4.0074 | 1390.9837
| 4.2007 | 110.1828 | 4.0952 | 113.633 | 4.1655 | 112.768 | 4.1655 | 109.8788 | 4.1128 | 112.486 | 4.1831 | 1414.096 | 4.0952 | 1405.9502 | 4.0952 | 1401.6077 | 4.0601 | 1401.8361 | 4.0601 | 1402.6356 | 4.0425 | 1397.4974 | 4.0425 | 1392.0133
| 4.2358 | 110.2031 | 4.1304 | 114.737 | 4.2007 | 112.7076 | 4.2007 | 110.1423 | 4.148 | 111.639 | 4.2183 | 1414.096 | 4.1304 | 1407.4379 | 4.1304 | 1404.3495 | 4.0952 | 1401.6077 | 4.0952 | 1401.4935 | 4.0777 | 1399.5524 | 4.0777 | 1393.271
| 4.271 | 110.5271 | 4.1655 | 115.258 | 4.2358 | 112.2038 | 4.2358 | 110.6283 | 4.1831 | 111.0531 | 4.2534 | 1413.4053 | 4.1655 | 1409.1562 | 4.1655 | 1404.6924 | 4.1304 | 1400.6942 | 4.1304 | 1401.4935 | 4.1128 | 1397.8399 | 4.1128 | 1393.4996
| 4.3061 | 110.6081 | 4.2007 | 115.258 | 4.271 | 112.0425 | 4.271 | 111.2956 | 4.2183 | 110.8913 | 4.2886 | 1414.9026 | 4.2007 | 1406.5222 | 4.2007 | 1403.321 | 4.1655 | 1400.6942 | 4.1655 | 1401.8361 | 4.148 | 1399.5524 | 4.148 | 1393.7282
| 4.3413 | 111.0127 | 4.2358 | 114.6768 | 4.3061 | 111.982 | 4.3061 | 112.0829 | 4.2534 | 110.8306 | 4.3237 | 1415.8254 | 4.2358 | 1408.1249 | 4.2358 | 1404.6924 | 4.2007 | 1400.6942 | 4.2007 | 1402.6356 | 4.1831 | 1398.9816 | 4.1831 | 1394.528
| 4.3764 | 111.4572 | 4.271 | 114.5966 | 4.3413 | 111.6793 | 4.3413 | 112.1837 | 4.2886 | 111.1744 | 4.3589 | 1417.6741 | 4.271 | 1409.0415 | 4.271 | 1405.1497 | 4.2358 | 1401.6077 | 4.2358 | 1403.8923 | 4.2183 | 1400.4658 | 4.2183 | 1396.1272
| 4.4116 | 111.6793 | 4.3061 | 114.1953 | 4.3764 | 111.4168 | 4.3764 | 112.2643 | 4.3237 | 111.6995 | 4.394 | 1420.108 | 4.3061 | 1408.6977 | 4.3061 | 1404.6924 | 4.271 | 1404.0066 | 4.271 | 1407.0944 | 4.2534 | 1400.4658 | 4.2534 | 1398.6391
| 4.4468 | 112.103 | 4.3413 | 113.5727 | 4.4116 | 112.0425 | 4.4116 | 111.9013 | 4.3589 | 112.2845 | 4.4292 | 1421.5029 | 4.3413 | 1408.8123 | 4.3413 | 1405.6071 | 4.3061 | 1405.264 | 4.3061 | 1407.2089 | 4.2886 | 1402.9783 | 4.2886 | 1400.9225
| 4.4819 | 112.103 | 4.3764 | 113.251 | 4.4468 | 112.0425 | 4.4468 | 111.134 | 4.394 | 113.1303 | 4.4643 | 1422.2015 | 4.3764 | 1410.0735 | 4.3764 | 1405.8358 | 4.3413 | 1405.264 | 4.3413 | 1408.5832 | 4.3237 | 1402.6356 | 4.3237 | 1401.9503
| 4.5171 | 111.4976 | 4.4116 | 113.1705 | 4.4819 | 111.3764 | 4.4819 | 110.3854 | 4.4292 | 113.2309 | 4.4995 | 1422.2015 | 4.4116 | 1411.1064 | 4.4116 | 1406.0646 | 4.3764 | 1406.4078 | 4.3764 | 1409.5001 | 4.3589 | 1404.3495 | 4.3589 | 1402.9783
| 4.5522 | 111.4168 | 4.4468 | 113.0297 | 4.5171 | 111.3158 | 4.5171 | 110.2233 | 4.4643 | 113.4521 | 4.5346 | 1422.2015 | 4.4468 | 1410.9916 | 4.4468 | 1408.1249 | 4.4116 | 1407.3234 | 4.4116 | 1410.3029 | 4.394 | 1404.0066 | 4.394 | 1403.4353
| 4.5874 | 111.3764 | 4.4819 | 113.0297 | 4.5522 | 111.6591 | 4.5522 | 110.041 | 4.4995 | 113.1303 | 4.5698 | 1420.108 | 4.4819 | 1410.9916 | 4.4819 | 1408.1249 | 4.4468 | 1406.4078 | 4.4468 | 1410.8768 | 4.4292 | 1405.6071 | 4.4292 | 1404.8067
| 4.6225 | 111.4976 | 4.5171 | 113.1303 | 4.5874 | 111.8408 | 4.5874 | 109.9802 | 4.5346 | 112.929 | 4.6049 | 1419.8758 | 4.5171 | 1411.4509 | 4.5171 | 1408.5832 | 4.4819 | 1407.4379 | 4.4819 | 1411.2212 | 4.4643 | 1406.8655 | 4.4643 | 1403.4353
| 4.6577 | 111.639 | 4.5522 | 113.2711 | 4.6225 | 112.0627 | 4.6225 | 110.1828 | 4.5698 | 112.3248 | 4.6401 | 1419.5277 | 4.5522 | 1411.4509 | 4.5522 | 1410.3029 | 4.5171 | 1408.1249 | 4.5171 | 1411.4509 | 4.4995 | 1408.354 | 4.4995 | 1403.4353
| 4.6928 | 111.9013 | 4.5874 | 113.7535 | 4.6577 | 112.4457 | 4.6577 | 110.3854 | 4.6049 | 111.9215 | 4.6752 | 1419.5277 | 4.5874 | 1410.9916 | 4.5874 | 1410.3029 | 4.5522 | 1408.4686 | 4.5522 | 1411.9104 | 4.5346 | 1408.8123 | 4.5346 | 1403.321
| 4.728 | 112.7076 | 4.6225 | 114.2957 | 4.6928 | 113.2107 | 4.6928 | 110.6081 | 4.6401 | 111.639 | 4.7104 | 1419.6437 | 4.6225 | 1412.4851 | 4.6225 | 1408.5832 | 4.5874 | 1409.0415 | 4.5874 | 1411.9104 | 4.5698 | 1410.0735 | 4.5698 | 1403.321
| 4.7631 | 112.7881 | 4.6577 | 114.9174 | 4.728 | 113.2107 | 4.728 | 111.0935 | 4.6752 | 111.5784 | 4.7455 | 1421.1539 | 4.6577 | 1412.4851 | 4.6577 | 1408.0104 | 4.6225 | 1408.4686 | 4.6225 | 1411.2212 | 4.6049 | 1409.0415 | 4.6049 | 1403.5495
| 4.7983 | 112.7881 | 4.6928 | 115.3582 | 4.7631 | 113.1102 | 4.7631 | 111.538 | 4.7104 | 111.5986 | 4.7807 | 1421.3865 | 4.6928 | 1412.4851 | 4.6928 | 1407.4379 | 4.6577 | 1408.0104 | 4.6577 | 1411.5657 | 4.6401 | 1408.354 | 4.6401 | 1403.5495
| 4.8334 | 111.861 | 4.728 | 115.4383 | 4.7983 | 112.9693 | 4.7983 | 111.9013 | 4.7455 | 111.9215 | 4.8159 | 1420.805 | 4.728 | 1412.6 | 4.728 | 1407.2089 | 4.6928 | 1406.5222 | 4.6928 | 1411.5657 | 4.6752 | 1408.2395 | 4.6752 | 1403.6638
| 4.8686 | 111.4774 | 4.7631 | 115.1579 | 4.8334 | 112.9089 | 4.8334 | 112.3046 | 4.7807 | 112.6069 | 4.851 | 1422.6677 | 4.7631 | 1412.83 | 4.7631 | 1406.2934 | 4.728 | 1406.5222 | 4.728 | 1410.762 | 4.7104 | 1409.7294 | 4.7104 | 1403.8923
| 4.9037 | 111.4774 | 4.7983 | 114.5766 | 4.8686 | 112.7076 | 4.8686 | 112.3046 | 4.8159 | 112.929 | 4.8862 | 1424.0686 | 4.7983 | 1412.83 | 4.7983 | 1407.5523 | 4.7631 | 1406.6367 | 4.7631 | 1410.762 | 4.7455 | 1408.8123 | 4.7455 | 1404.1209
| 4.9389 | 111.7399 | 4.8334 | 114.5766 | 4.9037 | 112.5263 | 4.9037 | 112.2442 | 4.851 | 113.2107 | 4.9213 | 1424.6534 | 4.8334 | 1417.3272 | 4.8334 | 1409.1562 | 4.7983 | 1408.9269 | 4.7983 | 1412.4851 | 4.7807 | 1409.7294 | 4.7807 | 1404.1209
| 4.974 | 112.8687 | 4.8686 | 114.5766 | 4.9389 | 112.5263 | 4.9389 | 111.6793 | 4.8862 | 113.6531 | 4.9565 | 1424.1856 | 4.8686 | 1417.3272 | 4.8686 | 1412.0253 | 4.8334 | 1408.9269 | 4.8334 | 1413.5204 | 4.8159 | 1409.9588 | 4.8159 | 1403.8923
| 5.0092 | 113.2309 | 4.9037 | 114.9375 | 4.974 | 112.5465 | 4.974 | 111.336 | 4.9213 | 113.8741 | 4.9916 | 1425.2388 | 4.9037 | 1417.7899 | 4.9037 | 1414.4416 | 4.8686 | 1409.2708 | 4.8686 | 1415.0179 | 4.851 | 1410.6472 | 4.851 | 1404.8067
| 5.0443 | 113.2711 | 4.9389 | 115.0577 | 5.0092 | 111.9417 | 5.0092 | 110.4259 | 4.9565 | 113.7736 | 5.0268 | 1424.7704 | 4.9389 | 1417.7899 | 4.9389 | 1415.71 | 4.9037 | 1409.7294 | 4.9037 | 1418.0213 | 4.8862 | 1410.9916 | 4.8862 | 1405.1497
| 5.0795 | 113.2711 | 4.974 | 115.0777 | 5.0443 | 111.7197 | 5.0443 | 111.336 | 4.9916 | 113.2711 | 5.0619 | 1424.5364 | 4.974 | 1418.3687 | 4.974 | 1418.1371 | 4.9389 | 1410.3029 | 4.9389 | 1419.5277 | 4.9213 | 1411.4509 | 4.9213 | 1405.6071
| 5.1146 | 113.3515 | 5.0092 | 114.7771 | 5.0795 | 111.1137 | 5.0795 | 111.336 | 5.0268 | 112.8284 | 5.0971 | 1425.9422 | 5.0092 | 1421.7357 | 5.0092 | 1416.4026 | 4.974 | 1412.0253 | 4.974 | 1419.7598 | 4.9565 | 1412.0253 | 4.9565 | 1406.179
| 5.1498 | 114.0548 | 5.0443 | 114.5966 | 5.1146 | 111.0329 | 5.1146 | 111.3764 | 5.0619 | 112.8284 | 5.1322 | 1430.5384 | 5.0443 | 1423.9518 | 5.0443 | 1419.4117 | 5.0092 | 1414.2112 | 5.0092 | 1421.6193 | 4.9916 | 1415.4792 | 4.9916 | 1409.5001
| 5.1849 | 115.4783 | 5.0795 | 114.0548 | 5.1498 | 110.6081 | 5.1498 | 111.2552 | 5.0971 | 112.6874 | 5.1674 | 1432.6754 | 5.0795 | 1425.2388 | 5.0795 | 1419.4117 | 5.0443 | 1417.0959 | 5.0443 | 1421.3865 | 5.0268 | 1417.4428 | 5.0268 | 1410.1882
| 5.2201 | 115.5984 | 5.1146 | 113.4119 | 5.1849 | 110.5271 | 5.1849 | 111.0329 | 5.1322 | 113.09 | 5.2025 | 1436.5013 | 5.1146 | 1426.2942 | 5.1146 | 1419.8758 | 5.0795 | 1418.2529 | 5.0795 | 1421.3865 | 5.0619 | 1418.832 | 5.0619 | 1413.6355
| 5.2553 | 115.5984 | 5.1498 | 112.9089 | 5.2201 | 111.0329 | 5.2201 | 110.5271 | 5.1674 | 113.09 | 5.2377 | 1438.4282 | 5.1498 | 1427.9403 | 5.1498 | 1422.4346 | 5.1146 | 1420.6888 | 5.1146 | 1422.318 | 5.0971 | 1419.0638 | 5.0971 | 1415.3638
| 5.2904 | 115.7185 | 5.1849 | 112.9089 | 5.2553 | 112.1837 | 5.2553 | 110.4056 | 5.2025 | 113.2309 | 5.2728 | 1439.5164 | 5.1849 | 1429.4738 | 5.1849 | 1426.1768 | 5.1498 | 1422.901 | 5.1498 | 1425.5904 | 5.1322 | 1420.9213 | 5.1322 | 1418.7162
| 5.3256 | 115.3582 | 5.2201 | 112.9089 | 5.2904 | 112.5666 | 5.2904 | 110.0004 | 5.2377 | 113.0095 | 5.308 | 1439.5164 | 5.2201 | 1429.4738 | 5.2201 | 1427.8225 | 5.1849 | 1423.8349 | 5.1849 | 1427.2341 | 5.1674 | 1423.3677 | 5.1674 | 1419.2957
| 5.3607 | 115.218 | 5.2553 | 113.432 | 5.3256 | 112.7479 | 5.3256 | 110.0004 | 5.2728 | 112.8888 | 5.3431 | 1439.0324 | 5.2553 | 1429.4738 | 5.2553 | 1430.42 | 5.2201 | 1426.4116 | 5.2201 | 1429.4738 | 5.2025 | 1426.4116 | 5.2025 | 1418.1371
| 5.3959 | 115.3181 | 5.2904 | 113.854 | 5.3607 | 113.0498 | 5.3607 | 110.3246 | 5.308 | 112.2038 | 5.3783 | 1439.0324 | 5.2904 | 1431.2496 | 5.2904 | 1431.131 | 5.2553 | 1426.764 | 5.2553 | 1429.8284 | 5.2377 | 1427.7048 | 5.2377 | 1419.2957
| 5.431 | 115.8386 | 5.3256 | 113.854 | 5.3959 | 113.2309 | 5.3959 | 110.8913 | 5.3431 | 112.2038 | 5.4134 | 1440.1224 | 5.3256 | 1430.3016 | 5.3256 | 1430.1833 | 5.2904 | 1427.4694 | 5.2904 | 1429.9467 | 5.2728 | 1428.2937 | 5.2728 | 1420.5726
| 5.4662 | 116.4784 | 5.3607 | 114.1351 | 5.431 | 114.115 | 5.431 | 111.134 | 5.3783 | 111.5784 | 5.4486 | 1441.7025 | 5.3607 | 1430.8939 | 5.3607 | 1431.3682 | 5.3256 | 1427.9403 | 5.3256 | 1429.9467 | 5.308 | 1427.8225 | 5.308 | 1421.0376
| 5.5013 | 117.0377 | 5.3959 | 114.8573 | 5.4662 | 114.717 | 5.4662 | 111.6591 | 5.4134 | 111.5582 | 5.4837 | 1444.8843 | 5.3959 | 1433.1517 | 5.3959 | 1431.7243 | 5.3607 | 1428.7654 | 5.3607 | 1429.9467 | 5.3431 | 1428.7654 | 5.3431 | 1422.901
| 5.5365 | 117.0377 | 5.431 | 116.2586 | 5.5013 | 115.4583 | 5.5013 | 112.2845 | 5.4486 | 111.5582 | 5.5189 | 1445.6228 | 5.431 | 1435.4216 | 5.431 | 1431.7243 | 5.3959 | 1430.1833 | 5.3959 | 1429.9467 | 5.3783 | 1429.2376 | 5.3783 | 1424.0686
| 5.5716 | 116.4185 | 5.4662 | 116.5384 | 5.5365 | 115.6985 | 5.5365 | 112.2845 | 5.4837 | 112.103 | 5.554 | 1449.7142 | 5.4662 | 1438.3075 | 5.4662 | 1435.1821 | 5.431 | 1430.3016 | 5.431 | 1429.9467 | 5.4134 | 1430.3016 | 5.4134 | 1424.7704
| 5.6068 | 115.9786 | 5.5013 | 117.0577 | 5.5716 | 115.7385 | 5.5716 | 113.1705 | 5.5189 | 112.3046 | 5.5892 | 1452.0939 | 5.5013 | 1440.4864 | 5.5013 | 1437.2228 | 5.4662 | 1430.42 | 5.4662 | 1432.3185 | 5.4486 | 1431.2496 | 5.4486 | 1424.7704
| 5.6419 | 115.4183 | 5.5365 | 117.0976 | 5.6068 | 116.2386 | 5.6068 | 113.7937 | 5.554 | 112.9894 | 5.6244 | 1453.7327 | 5.5365 | 1442.6785 | 5.5365 | 1440.365 | 5.5013 | 1434.9427 | 5.5013 | 1436.7417 | 5.4837 | 1432.3185 | 5.4837 | 1426.4116
| 5.6771 | 115.4183 | 5.5716 | 116.838 | 5.6419 | 116.6383 | 5.6419 | 113.854 | 5.5892 | 113.3917 | 5.6595 | 1454.9994 | 5.5716 | 1443.7796 | 5.5716 | 1441.5807 | 5.5365 | 1437.3432 | 5.5365 | 1440.0011 | 5.5189 | 1434.7035 | 5.5189 | 1429.0014
| 5.7122 | 115.4583 | 5.6068 | 117.2772 | 5.6771 | 116.6183 | 5.6771 | 113.8942 | 5.6244 | 113.3917 | 5.6947 | 1456.2714 | 5.6068 | 1444.0248 | 5.6068 | 1443.4122 | 5.5716 | 1440.365 | 5.5716 | 1442.0682 | 5.554 | 1435.7812 | 5.554 | 1432.4374
| 5.7474 | 115.8786 | 5.6419 | 117.7759 | 5.7122 | 116.8979 | 5.7122 | 114.0347 | 5.6595 | 114.3559 | 5.7298 | 1456.5265 | 5.6419 | 1445.3764 | 5.6419 | 1446.61 | 5.6068 | 1442.4342 | 5.6068 | 1442.8006 | 5.5892 | 1436.5013 | 5.5892 | 1433.8672
| 5.7825 | 115.8786 | 5.6771 | 117.6762 | 5.7474 | 116.4984 | 5.7474 | 113.854 | 5.6947 | 115.0577 | 5.765 | 1457.1652 | 5.6771 | 1445.3764 | 5.6771 | 1447.6001 | 5.6419 | 1444.7614 | 5.6419 | 1443.7796 | 5.6244 | 1437.9456 | 5.6244 | 1435.1821
| 5.8177 | 116.2786 | 5.7122 | 118.1547 | 5.7825 | 116.9179 | 5.7825 | 114.0548 | 5.7298 | 115.0777 | 5.8001 | 1460.8973 | 5.7122 | 1449.4648 | 5.7122 | 1447.6001 | 5.6771 | 1443.6571 | 5.6771 | 1446.8572 | 5.6595 | 1439.8799 | 5.6595 | 1435.1821
| 5.8528 | 116.2586 | 5.7474 | 118.3738 | 5.8177 | 116.6183 | 5.8177 | 114.1552 | 5.765 | 115.9986 | 5.8353 | 1462.847 | 5.7474 | 1449.4648 | 5.7474 | 1448.2204 | 5.7122 | 1443.2898 | 5.7122 | 1446.8572 | 5.6947 | 1443.2898 | 5.6947 | 1435.4216
| 5.888 | 116.5584 | 5.7825 | 119.2696 | 5.8528 | 116.4585 | 5.8528 | 115.4383 | 5.8001 | 116.2986 | 5.8704 | 1462.847 | 5.7825 | 1449.5895 | 5.7825 | 1448.9665 | 5.7474 | 1445.9926 | 5.7474 | 1447.9722 | 5.7298 | 1443.6571 | 5.7298 | 1436.5013
| 5.9231 | 116.7981 | 5.8177 | 119.2696 | 5.888 | 116.9778 | 5.888 | 115.9186 | 5.8353 | 116.2986 | 5.9056 | 1462.847 | 5.8177 | 1451.5914 | 5.8177 | 1452.8492 | 5.7825 | 1445.0073 | 5.7825 | 1448.3447 | 5.765 | 1445.6228 | 5.765 | 1436.9822
| 5.9583 | 117.6164 | 5.8528 | 119.389 | 5.9231 | 116.9778 | 5.9231 | 116.1786 | 5.8704 | 116.3785 | 5.9407 | 1461.8055 | 5.8528 | 1453.2275 | 5.8528 | 1454.1121 | 5.8177 | 1448.842 | 5.8177 | 1449.2155 | 5.8001 | 1446.61 | 5.8001 | 1436.9822
| 5.9935 | 117.6164 | 5.888 | 119.2099 | 5.9583 | 116.4585 | 5.9583 | 116.2186 | 5.9056 | 117.1774 | 5.9759 | 1461.5457 | 5.888 | 1453.48 | 5.888 | 1452.8492 | 5.8528 | 1454.2387 | 5.8528 | 1453.2275 | 5.8353 | 1447.2284 | 5.8353 | 1437.7045
| 6.0286 | 117.8158 | 5.9231 | 118.7522 | 5.9935 | 116.5384 | 5.9935 | 116.2186 | 5.9407 | 117.5565 | 6.011 | 1456.9095 | 5.9231 | 1454.6188 | 5.9231 | 1453.48 | 5.888 | 1454.2387 | 5.888 | 1449.5895 | 5.8704 | 1448.3447 | 5.8704 | 1439.7587
| 6.0638 | 117.8158 | 5.9583 | 118.7124 | 6.0286 | 116.7981 | 6.0286 | 115.9386 | 5.9759 | 117.5565 | 6.0462 | 1455.8893 | 5.9583 | 1454.2387 | 5.9583 | 1452.3454 | 5.9231 | 1454.2387 | 5.9231 | 1451.4659 | 5.9056 | 1451.4659 | 5.9056 | 1441.8244
| 6.0989 | 116.3585 | 5.9935 | 118.792 | 6.0638 | 116.6583 | 6.0638 | 115.7986 | 6.011 | 118.314 | 6.0813 | 1454.8725 | 5.9935 | 1453.48 | 5.9935 | 1449.2155 | 5.9583 | 1453.8591 | 5.9583 | 1451.4659 | 5.9407 | 1453.3537 | 5.9407 | 1446.2394
| 6.1341 | 116.3585 | 6.0286 | 118.8318 | 6.0989 | 116.9379 | 6.0989 | 115.7986 | 6.0462 | 117.5565 | 6.1165 | 1454.1121 | 6.0286 | 1452.9752 | 6.0286 | 1449.7142 | 5.9935 | 1453.6063 | 5.9935 | 1450.2137 | 5.9759 | 1451.8425 | 5.9759 | 1446.2394
| 6.1692 | 116.3585 | 6.0638 | 118.8517 | 6.1341 | 117.2573 | 6.1341 | 115.9986 | 6.0813 | 117.2772 | 6.1516 | 1456.399 | 6.0638 | 1451.0897 | 6.0638 | 1448.9665 | 6.0286 | 1450.2137 | 6.0286 | 1448.7176 | 6.011 | 1451.4659 | 6.011 | 1446.61
| 6.2044 | 116.5983 | 6.0989 | 118.8517 | 6.1692 | 117.5765 | 6.1692 | 116.0386 | 6.1165 | 116.5384 | 6.1868 | 1457.6771 | 6.0989 | 1451.0897 | 6.0989 | 1449.2155 | 6.0638 | 1450.2137 | 6.0638 | 1448.7176 | 6.0462 | 1449.9639 | 6.0462 | 1447.6001
| 6.2395 | 117.6762 | 6.1341 | 118.8517 | 6.2044 | 118.3539 | 6.2044 | 116.0986 | 6.1516 | 116.0386 | 6.2219 | 1461.9355 | 6.1341 | 1452.8492 | 6.1341 | 1449.839 | 6.0989 | 1447.4762 | 6.0989 | 1448.0963 | 6.0813 | 1444.7614 | 6.0813 | 1445.8693
| 6.2747 | 118.573 | 6.1692 | 119.0707 | 6.2395 | 118.4933 | 6.2395 | 116.0986 | 6.1868 | 116.0386 | 6.2571 | 1463.8924 | 6.1692 | 1453.7327 | 6.1692 | 1451.9682 | 6.1341 | 1447.3523 | 6.1341 | 1448.7176 | 6.1165 | 1444.7614 | 6.1165 | 1444.7614
| 6.3098 | 120.045 | 6.2044 | 119.6674 | 6.2747 | 118.573 | 6.2747 | 115.8186 | 6.2219 | 116.1986 | 6.2922 | 1465.0733 | 6.2044 | 1454.4921 | 6.2044 | 1453.6063 | 6.1692 | 1447.3523 | 6.1692 | 1451.0897 | 6.1516 | 1444.8843 | 6.1516 | 1445.0073
| 6.345 | 120.8197 | 6.2395 | 119.9457 | 6.3098 | 118.7721 | 6.3098 | 115.8186 | 6.2571 | 116.4185 | 6.3274 | 1464.6791 | 6.2395 | 1456.7818 | 6.2395 | 1457.0373 | 6.2044 | 1450.4637 | 6.2044 | 1453.2275 | 6.1868 | 1446.4864 | 6.1868 | 1444.1475
| 6.3801 | 121.0182 | 6.2747 | 120.1444 | 6.345 | 118.7721 | 6.345 | 116.4185 | 6.2922 | 117.2174 | 6.3625 | 1464.2855 | 6.2747 | 1457.9334 | 6.2747 | 1457.8052 | 6.2395 | 1451.3404 | 6.2395 | 1454.8725 | 6.2219 | 1448.2204 | 6.2219 | 1444.1475
| 6.4153 | 121.2761 | 6.3098 | 120.0848 | 6.3801 | 118.0151 | 6.3801 | 116.9778 | 6.3274 | 117.4568 | 6.3977 | 1464.0234 | 6.3098 | 1458.0616 | 6.3098 | 1458.1899 | 6.2747 | 1455.8893 | 6.2747 | 1455.6347 | 6.2571 | 1452.0939 | 6.2571 | 1444.6385
| 6.4504 | 121.772 | 6.345 | 119.7071 | 6.4153 | 117.9553 | 6.4153 | 117.1375 | 6.3625 | 117.397 | 6.4329 | 1464.9418 | 6.345 | 1459.3471 | 6.345 | 1457.421 | 6.3098 | 1457.1652 | 6.3098 | 1456.0166 | 6.2922 | 1454.7456 | 6.2922 | 1445.9926
| 6.4856 | 121.9703 | 6.3801 | 119.6873 | 6.4504 | 117.9553 | 6.4504 | 117.1375 | 6.3977 | 117.6363 | 6.468 | 1464.6791 | 6.3801 | 1460.3796 | 6.3801 | 1457.9334 | 6.345 | 1458.5751 | 6.345 | 1456.0166 | 6.3274 | 1455.1263 | 6.3274 | 1447.7241
| 6.5207 | 121.8315 | 6.4153 | 119.7071 | 6.4856 | 118.4535 | 6.4856 | 117.2174 | 6.4329 | 117.8357 | 6.5032 | 1464.9418 | 6.4153 | 1459.3471 | 6.4153 | 1459.0895 | 6.3801 | 1458.5751 | 6.3801 | 1456.144 | 6.3625 | 1456.0166 | 6.3625 | 1450.4637
| 6.5559 | 121.772 | 6.4504 | 120.0053 | 6.5207 | 119.1701 | 6.5207 | 116.5783 | 6.468 | 118.3937 | 6.5383 | 1466.7879 | 6.4504 | 1458.7036 | 6.4504 | 1459.3471 | 6.4153 | 1460.1211 | 6.4153 | 1455.6347 | 6.3977 | 1456.0166 | 6.3977 | 1451.3404
| 6.591 | 121.7125 | 6.4856 | 120.7403 | 6.5559 | 119.8463 | 6.5559 | 116.5983 | 6.5032 | 118.3937 | 6.5735 | 1467.8485 | 6.4856 | 1461.1565 | 6.4856 | 1460.2503 | 6.4504 | 1460.7678 | 6.4504 | 1456.144 | 6.4329 | 1455.2534 | 6.4329 | 1452.4713
| 6.6262 | 120.9586 | 6.5207 | 122.0496 | 6.591 | 121.5737 | 6.591 | 116.5783 | 6.5383 | 118.4933 | 6.6086 | 1467.8485 | 6.5207 | 1461.5457 | 6.5207 | 1460.2503 | 6.4856 | 1458.3182 | 6.4856 | 1458.4467 | 6.468 | 1456.7818 | 6.468 | 1454.3654
| 6.6613 | 120.8197 | 6.5559 | 122.4064 | 6.6262 | 122.1884 | 6.6262 | 117.0577 | 6.5735 | 118.4933 | 6.6438 | 1468.6466 | 6.5559 | 1461.5457 | 6.5559 | 1463.2386 | 6.5207 | 1458.3182 | 6.5207 | 1458.7036 | 6.5032 | 1457.6771 | 6.5032 | 1452.9752
| 6.6965 | 120.9785 | 6.591 | 123.852 | 6.6613 | 123.0007 | 6.6613 | 117.7161 | 6.6086 | 118.4734 | 6.6789 | 1469.18 | 6.591 | 1464.5478 | 6.591 | 1463.2386 | 6.5559 | 1458.9608 | 6.5559 | 1459.3471 | 6.5383 | 1459.2183 | 6.5383 | 1452.9752
| 6.7316 | 121.3952 | 6.6262 | 124.2082 | 6.6965 | 122.763 | 6.6965 | 118.7124 | 6.6438 | 118.2144 | 6.7141 | 1470.3841 | 6.6262 | 1466.2592 | 6.6262 | 1463.2386 | 6.591 | 1458.3182 | 6.591 | 1458.9608 | 6.5735 | 1459.4759 | 6.5735 | 1454.4921
| 6.7668 | 121.8117 | 6.6613 | 124.4851 | 6.7316 | 123.0403 | 6.7316 | 121.5142 | 6.6789 | 118.4734 | 6.7492 | 1474.1653 | 6.6613 | 1466.7879 | 6.6613 | 1463.8924 | 6.6262 | 1457.421 | 6.6262 | 1460.5089 | 6.6086 | 1461.6755 | 6.6086 | 1455.2534
| 6.802 | 122.1686 | 6.6965 | 124.6434 | 6.7668 | 122.8819 | 6.7668 | 121.8712 | 6.7141 | 118.7721 | 6.7844 | 1479.1084 | 6.6965 | 1467.0527 | 6.6965 | 1466.7879 | 6.6613 | 1460.6383 | 6.6613 | 1461.2861 | 6.6438 | 1461.6755 | 6.6438 | 1459.3471
| 6.8371 | 123.0601 | 6.7316 | 124.4851 | 6.802 | 122.763 | 6.802 | 122.1289 | 6.7492 | 119.9854 | 6.8195 | 1482.7383 | 6.7316 | 1471.8631 | 6.7316 | 1467.4503 | 6.6965 | 1465.7315 | 6.6965 | 1464.1544 | 6.6789 | 1461.6755 | 6.6789 | 1460.3796
| 6.8723 | 124.2675 | 6.7668 | 124.3071 | 6.8371 | 122.1884 | 6.8371 | 121.8712 | 6.7844 | 121.891 | 6.8547 | 1483.0196 | 6.7668 | 1473.3505 | 6.7668 | 1469.8483 | 6.7316 | 1470.1161 | 6.7316 | 1467.1851 | 6.7141 | 1462.0655 | 6.7141 | 1461.8055
| 6.9074 | 125.9679 | 6.802 | 124.2082 | 6.8723 | 122.0496 | 6.8723 | 122.0893 | 6.8195 | 122.6045 | 6.8898 | 1484.2896 | 6.802 | 1472.944 | 6.802 | 1470.25 | 6.7668 | 1471.5936 | 6.7668 | 1470.3841 | 6.7492 | 1464.9418 | 6.7492 | 1461.8055
| 6.9426 | 125.9679 | 6.8371 | 123.9114 | 6.9074 | 121.5539 | 6.9074 | 121.7919 | 6.8547 | 123.0007 | 6.925 | 1484.7144 | 6.8371 | 1474.3014 | 6.8371 | 1471.4589 | 6.802 | 1471.8631 | 6.802 | 1470.1161 | 6.7844 | 1466.3913 | 6.7844 | 1461.8055
| 6.9777 | 125.9679 | 6.8723 | 123.5551 | 6.9426 | 121.5539 | 6.9426 | 121.6729 | 6.8898 | 122.763 | 6.9601 | 1488.7135 | 6.8723 | 1475.9395 | 6.8723 | 1476.2135 | 6.8371 | 1472.403 | 6.8371 | 1470.3841 | 6.8195 | 1467.7157 | 6.8195 | 1463.108
| 7.0129 | 124.8016 | 6.9074 | 123.3373 | 6.9777 | 122.0496 | 6.9777 | 121.1372 | 6.925 | 123.1987 | 6.9953 | 1490.1572 | 6.9074 | 1475.9395 | 6.9074 | 1479.1084 | 6.8723 | 1471.7283 | 6.8723 | 1472.6734 | 6.8547 | 1468.7799 | 6.8547 | 1466.9203
| 7.048 | 122.1289 | 6.9426 | 123.4165 | 7.0129 | 122.0694 | 7.0129 | 120.7403 | 6.9601 | 123.2977 | 7.0304 | 1491.0275 | 6.9426 | 1480.3591 | 6.9426 | 1482.7383 | 6.9074 | 1476.3507 | 6.9074 | 1472.944 | 6.8898 | 1469.5808 | 6.8898 | 1467.9813
| 7.0832 | 120.8594 | 6.9777 | 123.5155 | 7.048 | 122.8423 | 7.048 | 120.7403 | 6.9953 | 123.2977 | 7.0656 | 1491.0275 | 6.9777 | 1480.9169 | 6.9777 | 1483.7244 | 6.9426 | 1476.6251 | 6.9426 | 1474.9827 | 6.925 | 1471.0554 | 6.925 | 1469.18
| 7.1183 | 120.2636 | 7.0129 | 124.0499 | 7.0832 | 123.3571 | 7.0832 | 120.7403 | 7.0304 | 123.0997 | 7.1007 | 1490.592 | 7.0129 | 1480.7774 | 7.0129 | 1483.5833 | 6.9777 | 1477.8638 | 6.9777 | 1478.8313 | 6.9601 | 1471.7283 | 6.9601 | 1471.0554
| 7.1535 | 120.3629 | 7.048 | 124.6236 | 7.1183 | 124.4851 | 7.1183 | 121.4547 | 7.0656 | 123.1987 | 7.1359 | 1490.0125 | 7.048 | 1481.1963 | 7.048 | 1483.0196 | 7.0129 | 1478.0018 | 7.0129 | 1480.7774 | 6.9953 | 1472.1329 | 6.9953 | 1473.2149
| 7.1886 | 120.9189 | 7.0832 | 125.9284 | 7.1535 | 125.6517 | 7.1535 | 121.9703 | 7.1007 | 122.228 | 7.171 | 1489.4343 | 7.0832 | 1482.7383 | 7.0832 | 1483.0196 | 7.048 | 1478.4162 | 7.048 | 1480.7774 | 7.0304 | 1474.0293 | 7.0304 | 1473.6218
| 7.2238 | 121.4547 | 7.1183 | 127.173 | 7.1886 | 127.1532 | 7.1886 | 122.5847 | 7.1359 | 122.2082 | 7.2062 | 1487.9947 | 7.1183 | 1482.7383 | 7.1183 | 1482.5977 | 7.0832 | 1479.1084 | 7.0832 | 1480.7774 | 7.0656 | 1475.1192 | 7.0656 | 1474.8463
| 7.2589 | 121.4547 | 7.1535 | 127.9629 | 7.2238 | 127.1532 | 7.2238 | 123.4759 | 7.171 | 122.3469 | 7.2414 | 1487.9947 | 7.1535 | 1483.3013 | 7.1535 | 1480.9169 | 7.1183 | 1480.0807 | 7.1183 | 1479.6635 | 7.1007 | 1476.6251 | 7.1007 | 1475.8026
| 7.2941 | 121.6927 | 7.1886 | 128.8316 | 7.2589 | 127.1532 | 7.2589 | 124.4258 | 7.2062 | 122.5055 | 7.2765 | 1487.8512 | 7.1886 | 1485.8505 | 7.1886 | 1479.9415 | 7.1535 | 1478.9698 | 7.1535 | 1479.5247 | 7.1359 | 1475.529 | 7.1359 | 1474.8463
| 7.3292 | 120.9983 | 7.2238 | 128.8316 | 7.2941 | 126.2445 | 7.2941 | 125.8889 | 7.2414 | 123.0007 | 7.3117 | 1490.447 | 7.2238 | 1484.2896 | 7.2238 | 1478.6929 | 7.1886 | 1479.6635 | 7.1886 | 1478.278 | 7.171 | 1475.529 | 7.171 | 1475.3923
| 7.3644 | 120.502 | 7.2589 | 128.654 | 7.3292 | 123.852 | 7.3292 | 126.4224 | 7.2765 | 123.7135 | 7.3468 | 1491.1728 | 7.2589 | 1483.3013 | 7.2589 | 1478.5545 | 7.2238 | 1480.0807 | 7.2238 | 1478.1399 | 7.2062 | 1475.9395 | 7.2062 | 1475.529
| 7.3995 | 120.2835 | 7.2941 | 125.8691 | 7.3644 | 123.0205 | 7.3644 | 126.4224 | 7.3117 | 124.3071 | 7.382 | 1495.8686 | 7.2941 | 1486.1354 | 7.2941 | 1478.6929 | 7.2589 | 1478.8313 | 7.2589 | 1478.278 | 7.2414 | 1477.3125 | 7.2414 | 1475.2557
| 7.4347 | 120.3033 | 7.3292 | 124.4258 | 7.3995 | 121.6134 | 7.3995 | 126.2248 | 7.3468 | 126.2248 | 7.4171 | 1498.1006 | 7.3292 | 1488.4257 | 7.3292 | 1481.1963 | 7.2941 | 1478.8313 | 7.2941 | 1481.476 | 7.2765 | 1477.3125 | 7.2765 | 1475.2557
| 7.4698 | 120.1444 | 7.3644 | 124.2082 | 7.4347 | 121.6134 | 7.4347 | 123.7926 | 7.382 | 126.9162 | 7.4523 | 1500.5033 | 7.3644 | 1492.4846 | 7.3644 | 1488.8575 | 7.3292 | 1482.5977 | 7.3292 | 1485.2818 | 7.3117 | 1478.0018 | 7.3117 | 1476.6251
| 7.505 | 120.2835 | 7.3995 | 124.1884 | 7.4698 | 121.8315 | 7.4698 | 121.6531 | 7.4171 | 127.8247 | 7.4874 | 1499.5997 | 7.3995 | 1496.9077 | 7.3995 | 1489.4343 | 7.3644 | 1487.2779 | 7.3644 | 1491.1728 | 7.3468 | 1480.9169 | 7.3468 | 1478.4162
| 7.5401 | 120.3828 | 7.4347 | 124.4456 | 7.505 | 121.8315 | 7.505 | 119.8264 | 7.4523 | 127.8247 | 7.5226 | 1499.2992 | 7.4347 | 1498.2501 | 7.4347 | 1490.3021 | 7.3995 | 1491.4637 | 7.3995 | 1494.8339 | 7.382 | 1483.7244 | 7.382 | 1479.2471
| 7.5753 | 120.3033 | 7.4698 | 124.94 | 7.5401 | 122.01 | 7.5401 | 119.8264 | 7.4874 | 127.6075 | 7.5577 | 1497.0565 | 7.4698 | 1498.5494 | 7.4698 | 1489.29 | 7.4347 | 1493.5097 | 7.4347 | 1496.3134 | 7.4171 | 1486.8489 | 7.4171 | 1481.3361
| 7.6105 | 120.4424 | 7.505 | 125.0389 | 7.5753 | 122.6045 | 7.5753 | 119.9059 | 7.5226 | 124.3862 | 7.5929 | 1499.2992 | 7.505 | 1498.5494 | 7.505 | 1488.8575 | 7.4698 | 1494.539 | 7.4698 | 1496.1651 | 7.4523 | 1490.7371 | 7.4523 | 1482.5977
| 7.6456 | 121.3158 | 7.5401 | 124.7225 | 7.6105 | 122.3469 | 7.6105 | 121.2761 | 7.5577 | 123.2779 | 7.628 | 1496.9077 | 7.5401 | 1496.4619 | 7.5401 | 1486.4205 | 7.505 | 1494.3917 | 7.505 | 1494.8339 | 7.4874 | 1494.3917 | 7.4874 | 1482.8789
| 7.6808 | 121.3555 | 7.5753 | 124.94 | 7.6456 | 121.8712 | 7.6456 | 121.2761 | 7.5929 | 121.8514 | 7.6632 | 1496.6104 | 7.5753 | 1492.3385 | 7.5753 | 1486.2779 | 7.5401 | 1494.3917 | 7.5401 | 1493.5097 | 7.5226 | 1494.0974 | 7.5226 | 1482.5977
| 7.7159 | 121.3555 | 7.6105 | 124.3269 | 7.6808 | 121.9108 | 7.6808 | 122.3073 | 7.628 | 121.8514 | 7.6983 | 1494.2445 | 7.6105 | 1485.566 | 7.6105 | 1486.1354 | 7.5753 | 1491.1728 | 7.5753 | 1492.6308 | 7.5577 | 1492.9234 | 7.5577 | 1482.5977
| 7.7511 | 121.9108 | 7.6456 | 124.1686 | 7.7159 | 121.8315 | 7.7159 | 122.1091 | 7.6632 | 121.9307 | 7.7335 | 1494.2445 | 7.6456 | 1490.592 | 7.6456 | 1483.8656 | 7.6105 | 1489.29 | 7.6105 | 1488.282 | 7.5929 | 1490.8822 | 7.5929 | 1479.5247
| 7.7862 | 121.9108 | 7.6808 | 123.8916 | 7.7511 | 121.4547 | 7.7511 | 121.2761 | 7.6983 | 122.9413 | 7.7686 | 1491.9008 | 7.6808 | 1489.5787 | 7.6808 | 1482.5977 | 7.6456 | 1486.2779 | 7.6456 | 1486.1354 | 7.628 | 1489.29 | 7.628 | 1476.4879
| 7.8214 | 121.3158 | 7.7159 | 123.5353 | 7.7862 | 121.6531 | 7.7862 | 121.6927 | 7.7335 | 122.9413 | 7.8038 | 1493.9503 | 7.7159 | 1487.1348 | 7.7159 | 1482.5977 | 7.6808 | 1486.2779 | 7.6808 | 1483.7244 | 7.6632 | 1487.8512 | 7.6632 | 1476.0765
| 7.8565 | 121.0579 | 7.7511 | 123.6343 | 7.8214 | 122.1487 | 7.8214 | 121.772 | 7.7686 | 123.0007 | 7.8389 | 1494.0974 | 7.7511 | 1484.0069 | 7.7511 | 1480.2198 | 7.7159 | 1485.7082 | 7.7159 | 1483.0196 | 7.6983 | 1485.2818 | 7.6983 | 1476.0765
| 7.8917 | 121.0579 | 7.7862 | 124.1686 | 7.8565 | 122.5847 | 7.8565 | 120.9785 | 7.8038 | 123.0799 | 7.8741 | 1495.4246 | 7.7862 | 1485.1399 | 7.7862 | 1480.2198 | 7.7511 | 1483.3013 | 7.7511 | 1481.1963 | 7.7335 | 1483.8656 | 7.7335 | 1478.8313
| 7.9268 | 121.3555 | 7.8214 | 124.2675 | 7.8917 | 123.0997 | 7.8917 | 121.415 | 7.8389 | 122.7828 | 7.9092 | 1496.9077 | 7.8214 | 1485.1399 | 7.8214 | 1484.4311 | 7.7862 | 1482.0362 | 7.7862 | 1480.0807 | 7.7686 | 1483.4423 | 7.7686 | 1479.2471
| 7.962 | 121.6134 | 7.8565 | 124.2675 | 7.9268 | 123.3373 | 7.9268 | 121.6531 | 7.8741 | 122.5451 | 7.9444 | 1496.9077 | 7.8565 | 1485.4239 | 7.8565 | 1490.8822 | 7.8214 | 1482.8789 | 7.8214 | 1480.3591 | 7.8038 | 1481.896 | 7.8038 | 1479.2471
| 7.9971 | 122.8819 | 7.8917 | 124.7225 | 7.962 | 123.753 | 7.962 | 121.5142 | 7.9092 | 122.8423 | 7.9795 | 1498.999 | 7.8917 | 1489.1457 | 7.8917 | 1491.3182 | 7.8565 | 1482.1765 | 7.8565 | 1481.7559 | 7.8389 | 1482.1765 | 7.8389 | 1481.7559
| 8.0323 | 122.8819 | 7.9268 | 125.0784 | 7.9971 | 123.7926 | 7.9971 | 121.8117 | 7.9444 | 122.8423 | 8.0147 | 1498.999 | 7.9268 | 1490.8822 | 7.9268 | 1492.3385 | 7.8917 | 1484.7144 | 7.8917 | 1481.7559 | 7.8741 | 1483.8656 | 7.8741 | 1482.0362
| 8.0674 | 122.8819 | 7.962 | 125.0784 | 8.0323 | 123.7926 | 8.0323 | 122.7036 | 7.9795 | 122.6045 | 8.0499 | 1495.5726 | 7.962 | 1489.4343 | 7.962 | 1490.8822 | 7.9268 | 1484.8561 | 7.9268 | 1482.5977 | 7.9092 | 1483.8656 | 7.9092 | 1481.7559
| 8.1026 | 123.0799 | 7.9971 | 125.3947 | 8.0674 | 123.4957 | 8.0674 | 123.0601 | 8.0147 | 122.8423 | 8.085 | 1495.5726 | 7.9971 | 1488.8575 | 7.9971 | 1484.7144 | 7.962 | 1489.29 | 7.962 | 1482.5977 | 7.9444 | 1484.0069 | 7.9444 | 1484.4311
| 8.1377 | 123.3967 | 8.0323 | 125.6122 | 8.1026 | 123.0799 | 8.1026 | 123.1393 | 8.0499 | 122.6838 | 8.1202 | 1498.3997 | 8.0323 | 1489.4343 | 8.0323 | 1484.5727 | 7.9971 | 1489.29 | 7.9971 | 1482.3168 | 7.9795 | 1484.7144 | 7.9795 | 1485.4239
| 8.1729 | 123.7926 | 8.0674 | 125.3947 | 8.1377 | 122.664 | 8.1377 | 123.1393 | 8.085 | 122.6045 | 8.1553 | 1502.3165 | 8.0674 | 1492.9234 | 8.0674 | 1485.7082 | 8.0323 | 1488.7135 | 8.0323 | 1482.1765 | 8.0147 | 1485.1399 | 8.0147 | 1483.1604
| 8.208 | 123.852 | 8.1026 | 124.8411 | 8.1729 | 122.664 | 8.1729 | 123.0007 | 8.1202 | 122.9809 | 8.1905 | 1504.5935 | 8.1026 | 1493.6565 | 8.1026 | 1491.755 | 8.0674 | 1489.29 | 8.0674 | 1483.3013 | 8.0499 | 1483.7244 | 8.0499 | 1482.8789
| 8.2432 | 124.2082 | 8.1377 | 124.7225 | 8.208 | 123.1789 | 8.208 | 122.1686 | 8.1553 | 123.3571 | 8.2256 | 1509.3368 | 8.1377 | 1495.4246 | 8.1377 | 1495.2768 | 8.1026 | 1490.1572 | 8.1026 | 1488.8575 | 8.085 | 1483.4423 | 8.085 | 1484.8561
| 8.2783 | 124.228 | 8.1729 | 125.0191 | 8.2432 | 123.951 | 8.2432 | 121.3952 | 8.1905 | 123.6145 | 8.2608 | 1511.1864 | 8.1729 | 1498.6991 | 8.1729 | 1501.409 | 8.1377 | 1493.8034 | 8.1377 | 1491.4637 | 8.1202 | 1483.4423 | 8.1202 | 1484.998
| 8.3135 | 124.3269 | 8.208 | 125.2761 | 8.2783 | 124.129 | 8.2783 | 121.038 | 8.2256 | 123.8916 | 8.2959 | 1515.5314 | 8.208 | 1502.6194 | 8.208 | 1502.6194 | 8.1729 | 1495.8686 | 8.1729 | 1499.2992 | 8.1553 | 1485.7082 | 8.1553 | 1485.8505
| 8.3486 | 124.3862 | 8.2432 | 126.126 | 8.3135 | 124.129 | 8.3135 | 121.1968 | 8.2608 | 124.0103 | 8.3311 | 1516.6241 | 8.2432 | 1506.7294 | 8.2432 | 1504.2892 | 8.208 | 1501.8625 | 8.208 | 1501.2579 | 8.1905 | 1489.5787 | 8.1905 | 1489.8678
| 8.3838 | 124.4456 | 8.2783 | 126.126 | 8.3486 | 123.951 | 8.3486 | 121.9307 | 8.2959 | 123.7333 | 8.3662 | 1517.4061 | 8.2783 | 1508.4148 | 8.2783 | 1508.1079 | 8.2432 | 1502.3165 | 8.2432 | 1502.4679 | 8.2256 | 1497.6526 | 8.2256 | 1494.2445
| 8.419 | 124.2082 | 8.3135 | 126.2445 | 8.3838 | 123.951 | 8.3838 | 122.9809 | 8.3311 | 122.8819 | 8.4014 | 1519.2883 | 8.3135 | 1509.7985 | 8.3135 | 1508.1079 | 8.2783 | 1504.5935 | 8.2783 | 1505.5077 | 8.2608 | 1499.7501 | 8.2608 | 1495.7206
| 8.4541 | 123.6541 | 8.3486 | 126.2643 | 8.419 | 123.951 | 8.419 | 123.0205 | 8.3662 | 122.8026 | 8.4365 | 1523.2336 | 8.3486 | 1510.2607 | 8.3486 | 1508.8756 | 8.3135 | 1506.2709 | 8.3135 | 1505.8128 | 8.2959 | 1501.5601 | 8.2959 | 1497.8018
| 8.4893 | 123.8916 | 8.3838 | 126.2841 | 8.4541 | 125.197 | 8.4541 | 124.228 | 8.4014 | 123.0997 | 8.4717 | 1521.0202 | 8.3838 | 1510.4148 | 8.3838 | 1508.8756 | 8.3486 | 1508.4148 | 8.3486 | 1508.5684 | 8.3311 | 1503.6813 | 8.3311 | 1499.2992
| 8.5244 | 124.1884 | 8.419 | 127.4495 | 8.4893 | 126.4619 | 8.4893 | 124.8411 | 8.4365 | 123.3769 | 8.5068 | 1519.2883 | 8.419 | 1511.65 | 8.419 | 1508.7219 | 8.3838 | 1510.4148 | 8.3838 | 1509.3368 | 8.3662 | 1503.9852 | 8.3662 | 1500.0512
| 8.5596 | 123.3373 | 8.4541 | 127.9629 | 8.5244 | 126.6594 | 8.5244 | 124.8411 | 8.4717 | 124.0697 | 8.542 | 1518.817 | 8.4541 | 1511.8046 | 8.4541 | 1506.1181 | 8.419 | 1513.0435 | 8.419 | 1509.3368 | 8.4014 | 1503.9852 | 8.4014 | 1499.2992
| 8.5947 | 123.6541 | 8.4893 | 128.1999 | 8.5596 | 127.252 | 8.5596 | 124.8609 | 8.5068 | 124.1488 | 8.5771 | 1515.9994 | 8.4893 | 1510.2607 | 8.4893 | 1503.8332 | 8.4541 | 1515.8433 | 8.4541 | 1511.8046 | 8.4365 | 1508.8756 | 8.4365 | 1505.0504
| 8.6299 | 123.5947 | 8.5244 | 128.6342 | 8.5947 | 127.252 | 8.5947 | 125.9877 | 8.542 | 124.2477 | 8.6123 | 1511.65 | 8.5244 | 1508.2614 | 8.5244 | 1500.6541 | 8.4893 | 1513.8196 | 8.4893 | 1512.8885 | 8.4717 | 1509.7985 | 8.4717 | 1507.648
| 8.665 | 123.5947 | 8.5596 | 129.187 | 8.6299 | 127.2915 | 8.6299 | 126.8767 | 8.5771 | 124.5445 | 8.6474 | 1510.4148 | 8.5596 | 1508.2614 | 8.5596 | 1496.1651 | 8.5244 | 1506.5765 | 8.5244 | 1507.1885 | 8.5068 | 1509.7985 | 8.5068 | 1507.0354
| 8.7002 | 123.5947 | 8.5947 | 129.3449 | 8.665 | 127.3705 | 8.665 | 127.173 | 8.6123 | 124.6236 | 8.6826 | 1511.65 | 8.5947 | 1502.4679 | 8.5947 | 1495.8686 | 8.5596 | 1506.5765 | 8.5596 | 1505.9654 | 8.542 | 1504.1372 | 8.542 | 1503.8332
| 8.7353 | 124.8214 | 8.6299 | 129.187 | 8.7002 | 127.3902 | 8.7002 | 127.7062 | 8.6474 | 124.6236 | 8.7177 | 1510.8776 | 8.6299 | 1501.8625 | 8.6299 | 1496.1651 | 8.5947 | 1506.1181 | 8.5947 | 1498.6991 | 8.5771 | 1504.1372 | 8.5771 | 1495.7206
| 8.7705 | 125.3157 | 8.665 | 127.7457 | 8.7353 | 127.1137 | 8.7353 | 127.6075 | 8.6826 | 125.2959 | 8.7529 | 1511.9593 | 8.665 | 1502.4679 | 8.665 | 1497.6526 | 8.6299 | 1505.3552 | 8.6299 | 1498.2501 | 8.6123 | 1501.8625 | 8.6123 | 1495.4246
| 8.8056 | 126.0272 | 8.7002 | 127.5482 | 8.7705 | 127.2125 | 8.7705 | 127.4692 | 8.7177 | 126.4619 | 8.7881 | 1514.2858 | 8.7002 | 1506.1181 | 8.7002 | 1502.9225 | 8.665 | 1505.0504 | 8.665 | 1498.6991 | 8.6474 | 1500.805 | 8.6474 | 1494.0974
| 8.8408 | 126.7977 | 8.7353 | 127.5482 | 8.8056 | 127.1137 | 8.8056 | 126.6594 | 8.7529 | 126.6792 | 8.8232 | 1515.3755 | 8.7353 | 1508.2614 | 8.7353 | 1506.1181 | 8.7002 | 1506.7294 | 8.7002 | 1500.0512 | 8.6826 | 1500.805 | 8.6826 | 1494.0974
| 8.8759 | 126.8965 | 8.7705 | 127.5482 | 8.8408 | 126.857 | 8.8408 | 126.2841 | 8.7881 | 127.3902 | 8.8584 | 1516.7804 | 8.7705 | 1509.7985 | 8.7705 | 1508.4148 | 8.7353 | 1507.3416 | 8.7353 | 1504.898 | 8.7177 | 1502.0138 | 8.7177 | 1495.8686
| 8.9111 | 127.015 | 8.8056 | 127.7852 | 8.8759 | 126.1458 | 8.8759 | 125.9877 | 8.8232 | 127.568 | 8.8935 | 1518.1894 | 8.8056 | 1509.6446 | 8.8056 | 1508.4148 | 8.7705 | 1507.8012 | 8.7705 | 1505.3552 | 8.7529 | 1503.3776 | 8.7529 | 1498.5494
| 8.9462 | 127.2125 | 8.8408 | 127.7852 | 8.9111 | 125.5924 | 8.9111 | 126.126 | 8.8584 | 127.568 | 8.9287 | 1518.1894 | 8.8408 | 1509.6446 | 8.8408 | 1508.4148 | 8.8056 | 1507.8012 | 8.8056 | 1507.0354 | 8.7881 | 1504.5935 | 8.7881 | 1504.7458
| 8.9814 | 127.173 | 8.8759 | 127.7062 | 8.9462 | 125.5924 | 8.9462 | 126.126 | 8.8935 | 126.6792 | 8.9638 | 1520.8625 | 8.8759 | 1509.6446 | 8.8759 | 1508.8756 | 8.8408 | 1508.2614 | 8.8408 | 1508.4148 | 8.8232 | 1506.2709 | 8.8232 | 1505.6602
| 9.0165 | 127.252 | 8.9111 | 127.2323 | 8.9814 | 126.0074 | 8.9814 | 126.0469 | 8.9287 | 126.5014 | 8.999 | 1521.6516 | 8.9111 | 1512.2688 | 8.9111 | 1509.6446 | 8.8759 | 1508.2614 | 8.8759 | 1509.6446 | 8.8584 | 1506.7294 | 8.8584 | 1507.1885
| 9.0517 | 128.0222 | 8.9462 | 127.4297 | 9.0165 | 126.8372 | 9.0165 | 126.0469 | 8.9638 | 126.5409 | 9.0341 | 1526.892 | 8.9462 | 1512.8885 | 8.9462 | 1514.2858 | 8.9111 | 1508.8756 | 8.9111 | 1509.6446 | 8.8935 | 1506.8824 | 8.8935 | 1506.8824
| 9.0868 | 129.029 | 8.9814 | 128.1011 | 9.0517 | 127.4297 | 9.0517 | 126.0469 | 8.999 | 126.9162 | 9.0693 | 1529.935 | 8.9814 | 1517.2496 | 8.9814 | 1516.4678 | 8.9462 | 1510.4148 | 8.9462 | 1511.9593 | 8.9287 | 1505.3552 | 8.9287 | 1504.7458
| 9.122 | 129.7595 | 9.0165 | 129.6608 | 9.0868 | 128.2986 | 9.0868 | 126.0469 | 9.0341 | 126.9162 | 9.1044 | 1532.835 | 9.0165 | 1519.4455 | 9.0165 | 1520.3897 | 8.9814 | 1514.5969 | 8.9814 | 1515.3755 | 8.9638 | 1507.9545 | 8.9638 | 1503.3776
| 9.1571 | 130.5491 | 9.0517 | 130.1345 | 9.122 | 128.3183 | 9.122 | 126.3236 | 9.0693 | 126.5409 | 9.1396 | 1532.835 | 9.0517 | 1520.7048 | 9.0517 | 1521.3358 | 9.0165 | 1516.6241 | 9.0165 | 1518.3462 | 8.999 | 1510.4148 | 8.999 | 1503.3776
| 9.1923 | 130.2332 | 9.0868 | 130.5886 | 9.1571 | 128.496 | 9.1571 | 128.0814 | 9.1044 | 125.8691 | 9.1747 | 1533.32 | 9.0868 | 1520.7048 | 9.0868 | 1522.4419 | 9.0517 | 1521.3358 | 9.0517 | 1520.7048 | 9.0341 | 1515.5314 | 9.0341 | 1504.898
| 9.2275 | 130.5491 | 9.122 | 130.5886 | 9.1923 | 128.5947 | 9.1923 | 128.5947 | 9.1396 | 125.5529 | 9.2099 | 1535.2646 | 9.122 | 1520.7048 | 9.122 | 1523.0752 | 9.0868 | 1522.2838 | 9.0868 | 1520.7048 | 9.0693 | 1518.5031 | 9.0693 | 1505.3552
| 9.2626 | 130.7662 | 9.1571 | 130.8649 | 9.2275 | 128.9501 | 9.2275 | 128.7724 | 9.1747 | 125.5529 | 9.245 | 1536.4024 | 9.1571 | 1522.9168 | 9.1571 | 1524.8208 | 9.122 | 1525.4572 | 9.122 | 1521.8095 | 9.1044 | 1518.66 | 9.1044 | 1506.1181
| 9.2978 | 130.7662 | 9.1923 | 130.9242 | 9.2626 | 129.1672 | 9.2626 | 128.7724 | 9.2099 | 125.6715 | 9.2802 | 1543.7774 | 9.1923 | 1523.0752 | 9.1923 | 1524.8208 | 9.1571 | 1524.9799 | 9.1571 | 1523.5506 | 9.1396 | 1519.2883 | 9.1396 | 1509.3368
| 9.3329 | 130.9834 | 9.2275 | 130.5491 | 9.2978 | 129.4633 | 9.2978 | 128.7724 | 9.245 | 126.6989 | 9.3153 | 1546.259 | 9.2275 | 1529.7743 | 9.2275 | 1525.1389 | 9.1923 | 1525.4572 | 9.1923 | 1523.8679 | 9.1747 | 1519.4455 | 9.1747 | 1511.8046
| 9.3681 | 130.9834 | 9.2626 | 131.1216 | 9.3329 | 129.6015 | 9.3329 | 128.7527 | 9.2802 | 128.417 | 9.3505 | 1548.2527 | 9.2626 | 1530.7389 | 9.2626 | 1527.8509 | 9.2275 | 1528.8117 | 9.2275 | 1524.1853 | 9.2099 | 1522.2838 | 9.2099 | 1513.3538
| 9.4032 | 130.6873 | 9.2978 | 131.4177 | 9.3681 | 131.398 | 9.3681 | 128.6342 | 9.3153 | 129.0685 | 9.3856 | 1551.592 | 9.2978 | 1534.94 | 9.2978 | 1532.9967 | 9.2626 | 1532.1891 | 9.2626 | 1528.8117 | 9.245 | 1519.4455 | 9.245 | 1517.4061
| 9.4384 | 129.7595 | 9.3329 | 131.9113 | 9.4032 | 132.2273 | 9.4032 | 128.7527 | 9.3505 | 129.029 | 9.4208 | 1554.2783 | 9.3329 | 1537.869 | 9.3329 | 1539.9949 | 9.2978 | 1535.7519 | 9.2978 | 1531.8665 | 9.2802 | 1525.7757 | 9.2802 | 1521.178
| 9.4735 | 126.6792 | 9.3681 | 132.8988 | 9.4384 | 132.4643 | 9.4384 | 129.3252 | 9.3856 | 129.0685 | 9.4559 | 1554.7835 | 9.3681 | 1541.965 | 9.3681 | 1540.8148 | 9.3329 | 1536.728 | 9.3329 | 1535.7519 | 9.3153 | 1526.7324 | 9.3153 | 1525.4572
| 9.5087 | 126.4224 | 9.4032 | 133.2346 | 9.4735 | 132.5828 | 9.4735 | 130.1543 | 9.4208 | 129.029 | 9.4911 | 1555.4577 | 9.4032 | 1544.1076 | 9.4032 | 1542.9528 | 9.3681 | 1538.1955 | 9.3681 | 1540.1588 | 9.3505 | 1531.8665 | 9.3505 | 1529.6138
|  |  | 9.4384 | 133.61 | 9.5087 | 132.9581 | 9.5087 | 130.4701 | 9.4559 | 128.5355 | 9.5262 | 1556.1328 | 9.4384 | 1549.2523 | 9.4384 | 1545.0995 | 9.4032 | 1543.2825 | 9.4032 | 1540.8148 | 9.3856 | 1535.7519 | 9.3856 | 1530.578
|  |  | 9.4735 | 133.9657 | 9.5438 | 132.9581 | 9.5438 | 131.5757 | 9.4911 | 128.5355 |  |  | 9.4735 | 1551.592 | 9.4735 | 1548.9189 | 9.4384 | 1546.4249 | 9.4384 | 1543.2825 | 9.4208 | 1536.8908 | 9.4208 | 1532.3505
|  |  | 9.5087 | 133.8669 |  |  |  |  | 9.5262 | 129.2462 |  |  | 9.5087 | 1556.3017 | 9.5087 | 1553.1014 | 9.4735 | 1549.0856 | 9.4735 | 1544.7687 | 9.4559 | 1542.9528 | 9.4559 | 1534.4535
|  |  | 9.5438 | 133.9657 |  |  |  |  | 9.5614 | 130.1345 |  |  | 9.5438 | 1556.6396 | 9.5438 | 1554.2783 | 9.5087 | 1551.2572 | 9.5087 | 1546.4249 | 9.4911 | 1545.9275 | 9.4911 | 1537.0537
|  |  | 9.579 | 133.9855 |  |  |  |  | 9.5966 | 130.8255 |  |  | 9.579 | 1557.3162 | 9.579 | 1553.2694 | 9.5438 | 1551.592 | 9.5438 | 1547.2549 | 9.5262 | 1549.9198 | 9.5262 | 1538.1955
|  |  | 9.6141 | 133.4322 | [8 empty] | 9.6141 | 1557.3162 | 9.6141 | 1553.1014 | 9.579 | 1554.9519 | 9.579 | 1550.2538 | 9.5614 | 1550.0868 | 9.5614 | 1542.4587
[17 empty] | 9.6141 | 1555.2891 | 9.6141 | 1551.4246 | 9.5966 | 1552.0947 | 9.5966 | 1543.9425
[17 empty] | 9.6493 | 1555.2891 | 9.6493 | 1552.5978 | 9.6317 | 1552.4301 | 9.6317 | 1545.9275
[21 empty] | 9.6669 | 1552.0947 | 9.6669 | 1548.7523

## uniformity position corrected
The scan area, which is defined by the template, the inner borders of the template to be more precise, is 24cm long, 12 cm to the centre. So my plan is to measure in 5 areas across the scanner plate.
The five areas are actually planes across the scanner bed in the lateral direction; the planes are defined by the labels thusly
The idea is then to scan the step wedge film at the corresponding heights with the following particular steps that correspond to the approximate doses of 0Gy, 1Gy and 14Gy.
The heights were marked on to the scan plate template.
The steps that correspond to the three doses of 0 1 and 14 Gy are
1Gy; the fourth strip after the large black strip at the top of the film.
14Gy is the tenth strip after the large black strip at the top of the film.
it is not possible to obtain a 0Gy strip using this step wedge; a blank film can be scanned for this purpose.
height | 1Gy (4th strip) | 14Gy (tenth strip)
middle | y
4.5cm up
11.5cm up
4.5cm down
11.5cm down
| middle 1Gy |  | 4.5cm up 1Gy |  | 4.5cm down 1Gy |  | 11.5cm up 1Gy |  | 11.5cm down 1Gy |  | middle 14Gy |  | 4.5cm up 14Gy |  | 10cm up 14Gy |  | 10cm up 14Gy; scanner covered |  | 4.5cm down 14Gy |  | 4.5cm down 14Gy; covered |  | 10cm down 14Gy
| cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy
| -9.4735 | 123.951 | -9.579 | 130.1937 | -9.5087 | 127.4692 | -9.5087 | 127.41 | -9.5614 | 131.3782 | -9.4911 | 1562.2456 | -9.579 | 1552.5978 | -9.579 | 1549.2523 | -9.6141 | 1554.7835 | -9.6141 | 1562.9289 | -9.6317 | 1572.9293 | -9.6317 | 1554.7835
| -9.4384 | 123.951 | -9.5438 | 129.3252 | -9.4735 | 127.4692 | -9.4735 | 126.9557 | -9.5262 | 131.2006 | -9.4559 | 1562.2456 | -9.5438 | 1552.5978 | -9.5438 | 1548.2527 | -9.579 | 1554.7835 | -9.579 | 1560.0305 | -9.5966 | 1569.9802 | -9.5966 | 1557.147
| -9.4032 | 124.0103 | -9.5087 | 129.3054 | -9.4384 | 127.6272 | -9.4384 | 126.7385 | -9.4911 | 130.9044 | -9.4208 | 1560.8814 | -9.5087 | 1549.4191 | -9.5087 | 1547.4211 | -9.5438 | 1554.4466 | -9.5438 | 1559.0111 | -9.5614 | 1568.5976 | -9.5614 | 1552.4301
| -9.3681 | 124.0103 | -9.4735 | 129.3054 | -9.4032 | 127.8839 | -9.4032 | 126.3433 | -9.4559 | 129.8384 | -9.3856 | 1559.1808 | -9.4735 | 1546.5908 | -9.4735 | 1543.2825 | -9.5087 | 1550.7553 | -9.5087 | 1557.9935 | -9.5262 | 1564.6408 | -9.5262 | 1548.5857
| -9.3329 | 124.3467 | -9.4384 | 129.2659 | -9.3681 | 127.9629 | -9.3681 | 125.9481 | -9.4208 | 128.9698 | -9.3505 | 1555.1205 | -9.4384 | 1551.0898 | -9.4384 | 1542.6233 | -9.4735 | 1547.2549 | -9.4735 | 1556.9778 | -9.4911 | 1561.2221 | -9.4911 | 1547.7536
| -9.2978 | 124.6434 | -9.4032 | 129.3054 | -9.3329 | 128.5947 | -9.3329 | 125.6517 | -9.3856 | 128.0222 | -9.3153 | 1551.7595 | -9.4032 | 1551.0898 | -9.4032 | 1543.6124 | -9.4384 | 1546.259 | -9.4384 | 1556.9778 | -9.4559 | 1559.6905 | -9.4559 | 1545.265
| -9.2626 | 124.5445 | -9.3681 | 129.6608 | -9.2978 | 128.1011 | -9.2978 | 126.1458 | -9.3505 | 127.5285 | -9.2802 | 1546.9228 | -9.3681 | 1553.1014 | -9.3681 | 1543.2825 | -9.4032 | 1544.1076 | -9.4032 | 1555.2891 | -9.4208 | 1556.3017 | -9.4208 | 1544.9341
| -9.2275 | 124.228 | -9.3329 | 128.9501 | -9.2626 | 127.5482 | -9.2626 | 126.1458 | -9.3153 | 127.41 | -9.245 | 1545.4305 | -9.3329 | 1552.0947 | -9.3329 | 1543.2825 | -9.3681 | 1543.9425 | -9.3681 | 1552.5978 | -9.3856 | 1555.1205 | -9.3856 | 1544.9341
| -9.1923 | 124.228 | -9.2978 | 127.9827 | -9.2275 | 126.5014 | -9.2275 | 125.9086 | -9.2802 | 127.41 | -9.2099 | 1545.7618 | -9.2978 | 1551.0898 | -9.2978 | 1542.9528 | -9.3329 | 1542.9528 | -9.3329 | 1551.4246 | -9.3505 | 1553.7736 | -9.3505 | 1543.9425
| -9.1571 | 124.0697 | -9.2626 | 127.8642 | -9.1923 | 126.5014 | -9.1923 | 125.9086 | -9.245 | 127.0545 | -9.1747 | 1547.7536 | -9.2626 | 1547.2549 | -9.2626 | 1541.3074 | -9.2978 | 1542.9528 | -9.2978 | 1550.0868 | -9.3153 | 1552.0947 | -9.3153 | 1541.965
| -9.122 | 124.0103 | -9.2275 | 127.5285 | -9.1571 | 126.5014 | -9.1571 | 125.8098 | -9.2099 | 127.1335 | -9.1396 | 1548.0863 | -9.2275 | 1546.4249 | -9.2275 | 1537.869 | -9.2626 | 1542.6233 | -9.2626 | 1549.7528 | -9.2802 | 1551.592 | -9.2802 | 1540.979
| -9.0868 | 123.5155 | -9.1923 | 128.0024 | -9.122 | 127.568 | -9.122 | 125.0191 | -9.1747 | 127.6667 | -9.1044 | 1548.9189 | -9.1923 | 1545.9275 | -9.1923 | 1536.2397 | -9.2275 | 1540.979 | -9.2275 | 1550.0868 | -9.245 | 1550.5881 | -9.245 | 1538.8491
| -9.0517 | 122.9017 | -9.1571 | 128.0024 | -9.0868 | 126.9557 | -9.0868 | 124.7225 | -9.1396 | 127.7655 | -9.0693 | 1551.592 | -9.1571 | 1545.9275 | -9.1571 | 1537.2167 | -9.1923 | 1541.3074 | -9.1923 | 1549.7528 | -9.2099 | 1550.0868 | -9.2099 | 1537.0537
| -9.0165 | 122.0298 | -9.122 | 128.4565 | -9.0517 | 126.9557 | -9.0517 | 124.7422 | -9.1044 | 127.5087 | -9.0341 | 1553.7736 | -9.122 | 1545.5961 | -9.122 | 1538.8491 | -9.1571 | 1541.3074 | -9.1571 | 1549.9198 | -9.1747 | 1550.0868 | -9.1747 | 1536.5652
| -8.9814 | 122.0298 | -9.0868 | 128.3578 | -9.0165 | 125.6912 | -9.0165 | 124.7225 | -9.0693 | 127.489 | -8.999 | 1553.7736 | -9.0868 | 1544.1076 | -9.0868 | 1539.6673 | -9.122 | 1543.2825 | -9.122 | 1550.9225 | -9.1396 | 1551.9271 | -9.1396 | 1535.2646
| -8.9462 | 122.1091 | -9.0517 | 127.5087 | -8.9814 | 125.5924 | -8.9814 | 124.7225 | -9.0341 | 126.6397 | -8.9638 | 1553.7736 | -9.0517 | 1545.4305 | -9.0517 | 1541.4717 | -9.0868 | 1543.6124 | -9.0868 | 1550.4209 | -9.1044 | 1550.5881 | -9.1044 | 1536.5652
| -8.9111 | 122.8423 | -9.0165 | 127.173 | -8.9462 | 125.0982 | -8.9462 | 124.7225 | -8.999 | 126.5607 | -8.9287 | 1553.4374 | -9.0165 | 1542.4587 | -9.0165 | 1541.4717 | -9.0517 | 1546.0932 | -9.0517 | 1550.4209 | -9.0693 | 1553.6055 | -9.0693 | 1535.9145
| -8.8759 | 122.8423 | -8.9814 | 126.7582 | -8.9111 | 125.3157 | -8.9111 | 124.5247 | -8.9638 | 126.4619 | -8.8935 | 1549.4191 | -8.9814 | 1542.1295 | -8.9814 | 1540.4867 | -9.0165 | 1545.7618 | -9.0165 | 1549.7528 | -9.0341 | 1554.11 | -9.0341 | 1535.7519
| -8.8408 | 122.8026 | -8.9462 | 126.5211 | -8.8759 | 125.3157 | -8.8759 | 124.0499 | -8.9287 | 126.5607 | -8.8584 | 1545.9275 | -8.9462 | 1539.5035 | -8.9462 | 1537.0537 | -8.9814 | 1544.1076 | -8.9814 | 1546.9228 | -8.999 | 1554.11 | -8.999 | 1535.2646
| -8.8056 | 122.2875 | -8.9111 | 126.3829 | -8.8408 | 125.3157 | -8.8408 | 123.2581 | -8.8935 | 126.778 | -8.8232 | 1541.3074 | -8.9111 | 1539.5035 | -8.9111 | 1536.728 | -8.9462 | 1543.2825 | -8.9462 | 1542.788 | -8.9638 | 1550.9225 | -8.9638 | 1534.94
| -8.7705 | 122.01 | -8.8759 | 126.3829 | -8.8056 | 124.9795 | -8.8056 | 123.1195 | -8.8584 | 126.8372 | -8.7881 | 1536.5652 | -8.8759 | 1537.7059 | -8.8759 | 1532.9967 | -8.9111 | 1541.1432 | -8.9111 | 1541.8005 | -8.9287 | 1547.4211 | -8.9287 | 1535.5894
| -8.7353 | 122.01 | -8.8408 | 126.7582 | -8.7705 | 124.3467 | -8.7705 | 123.1195 | -8.8232 | 126.7187 | -8.7529 | 1535.2646 | -8.8408 | 1535.9145 | -8.8408 | 1530.578 | -8.8759 | 1537.0537 | -8.8759 | 1541.8005 | -8.8935 | 1543.6124 | -8.8935 | 1535.7519
| -8.7002 | 122.1091 | -8.8056 | 125.8493 | -8.7353 | 123.9707 | -8.7353 | 123.1393 | -8.7881 | 126.3038 | -8.7177 | 1532.9967 | -8.8056 | 1533.9674 | -8.8056 | 1528.0109 | -8.8408 | 1533.4818 | -8.8408 | 1539.5035 | -8.8584 | 1543.4474 | -8.8584 | 1535.5894
| -8.665 | 122.0893 | -8.7705 | 125.8493 | -8.7002 | 123.8916 | -8.7002 | 122.9809 | -8.7529 | 125.9877 | -8.6826 | 1531.3829 | -8.7705 | 1527.691 | -8.7705 | 1524.9799 | -8.8056 | 1530.2564 | -8.8056 | 1538.5222 | -8.8232 | 1540.3227 | -8.8232 | 1534.94
| -8.6299 | 121.9901 | -8.7353 | 126.1458 | -8.665 | 124.2082 | -8.665 | 122.8819 | -8.7177 | 125.9086 | -8.6474 | 1529.7743 | -8.7353 | 1525.6164 | -8.7353 | 1521.4936 | -8.7705 | 1529.1324 | -8.7705 | 1535.427 | -8.7881 | 1539.831 | -8.7881 | 1531.8665
| -8.5947 | 121.772 | -8.7002 | 126.0865 | -8.6299 | 124.584 | -8.6299 | 122.6442 | -8.6826 | 125.6122 | -8.6123 | 1524.8208 | -8.7002 | 1524.0266 | -8.7002 | 1519.6027 | -8.7353 | 1528.0109 | -8.7353 | 1531.8665 | -8.7529 | 1539.3398 | -8.7529 | 1524.0266
| -8.5596 | 121.415 | -8.665 | 125.5133 | -8.5947 | 124.7027 | -8.5947 | 122.5649 | -8.6474 | 125.5529 | -8.5771 | 1523.3921 | -8.665 | 1523.3921 | -8.665 | 1517.4061 | -8.7002 | 1526.892 | -8.7002 | 1529.6138 | -8.7177 | 1534.1294 | -8.7177 | 1522.4419
| -8.5244 | 121.296 | -8.6299 | 125.8493 | -8.5596 | 124.7027 | -8.5596 | 122.5649 | -8.6123 | 125.454 | -8.542 | 1521.4936 | -8.6299 | 1521.178 | -8.6299 | 1513.1986 | -8.665 | 1522.7584 | -8.665 | 1529.2928 | -8.6826 | 1532.3505 | -8.6826 | 1522.4419
| -8.4893 | 121.1372 | -8.5947 | 125.9877 | -8.5244 | 124.2082 | -8.5244 | 122.6442 | -8.5771 | 124.7027 | -8.5068 | 1516.7804 | -8.5947 | 1518.9741 | -8.5947 | 1510.4148 | -8.6299 | 1517.8759 | -8.6299 | 1527.8509 | -8.6474 | 1526.892 | -8.6474 | 1520.7048
| -8.4541 | 120.8197 | -8.5596 | 126.0074 | -8.4893 | 123.7926 | -8.4893 | 122.5649 | -8.542 | 123.6739 | -8.4717 | 1514.4413 | -8.5596 | 1516.9367 | -8.5596 | 1503.9852 | -8.5947 | 1514.7525 | -8.5947 | 1519.4455 | -8.6123 | 1524.8208 | -8.6123 | 1515.9994
| -8.419 | 120.4821 | -8.5244 | 125.9877 | -8.4541 | 123.8322 | -8.4541 | 122.5451 | -8.5068 | 123.3175 | -8.4365 | 1512.7335 | -8.5244 | 1514.5969 | -8.5244 | 1502.0138 | -8.5596 | 1510.569 | -8.5596 | 1517.2496 | -8.5771 | 1523.2336 | -8.5771 | 1514.1303
| -8.3838 | 120.1444 | -8.4893 | 125.9877 | -8.419 | 123.4561 | -8.419 | 122.2082 | -8.4717 | 123.3175 | -8.4014 | 1511.3409 | -8.4893 | 1508.7219 | -8.4893 | 1501.409 | -8.5244 | 1508.8756 | -8.5244 | 1515.8433 | -8.542 | 1518.5031 | -8.542 | 1510.1066
| -8.3486 | 119.4486 | -8.4541 | 125.5924 | -8.3838 | 123.2977 | -8.3838 | 122.1091 | -8.4365 | 124.0895 | -8.3662 | 1508.5684 | -8.4541 | 1497.8018 | -8.4541 | 1500.0512 | -8.4893 | 1506.7294 | -8.4893 | 1512.2688 | -8.5068 | 1516.3116 | -8.5068 | 1508.5684
| -8.3135 | 118.7323 | -8.419 | 124.8016 | -8.3486 | 123.2977 | -8.3486 | 121.9505 | -8.4014 | 124.1686 | -8.3311 | 1507.3416 | -8.419 | 1497.3544 | -8.419 | 1497.9512 | -8.4541 | 1503.5294 | -8.4541 | 1508.2614 | -8.4717 | 1515.8433 | -8.4717 | 1508.1079
| -8.2783 | 118.1547 | -8.3838 | 124.5049 | -8.3135 | 123.0799 | -8.3135 | 121.8117 | -8.3662 | 124.0895 | -8.2959 | 1499.4494 | -8.3838 | 1495.8686 | -8.3838 | 1494.3917 | -8.419 | 1499.4494 | -8.419 | 1507.0354 | -8.4365 | 1514.1303 | -8.4365 | 1501.409
| -8.2432 | 117.6563 | -8.3486 | 124.1488 | -8.2783 | 122.5055 | -8.2783 | 121.6134 | -8.3311 | 124.3862 | -8.2608 | 1496.4619 | -8.3486 | 1492.4846 | -8.3486 | 1491.9008 | -8.3838 | 1494.6864 | -8.3838 | 1502.6194 | -8.4014 | 1509.7985 | -8.4014 | 1499.7501
| -8.208 | 117.6563 | -8.3135 | 124.1488 | -8.2432 | 122.3073 | -8.2432 | 120.5218 | -8.2959 | 124.3862 | -8.2256 | 1493.6565 | -8.3135 | 1489.4343 | -8.3135 | 1489.29 | -8.3486 | 1493.6565 | -8.3486 | 1497.9512 | -8.3662 | 1503.2259 | -8.3662 | 1492.9234
| -8.1729 | 118.2742 | -8.2783 | 124.0103 | -8.208 | 122.5649 | -8.208 | 120.6608 | -8.2608 | 124.3862 | -8.1905 | 1490.447 | -8.2783 | 1489.0016 | -8.2783 | 1485.1399 | -8.3135 | 1492.0466 | -8.3135 | 1497.8018 | -8.3311 | 1502.1651 | -8.3311 | 1491.9008
| -8.1377 | 118.2742 | -8.2432 | 124.0103 | -8.1729 | 121.7919 | -8.1729 | 120.4225 | -8.2256 | 124.4653 | -8.1553 | 1488.282 | -8.2432 | 1489.0016 | -8.2432 | 1482.4572 | -8.2783 | 1488.1383 | -8.2783 | 1494.9814 | -8.2959 | 1499.9006 | -8.2959 | 1490.1572
| -8.1026 | 118.1945 | -8.208 | 123.8916 | -8.1377 | 121.6531 | -8.1377 | 120.1047 | -8.1905 | 124.4258 | -8.1202 | 1487.5644 | -8.208 | 1480.4985 | -8.208 | 1475.2557 | -8.2432 | 1485.7082 | -8.2432 | 1488.4257 | -8.2608 | 1496.1651 | -8.2608 | 1488.8575
| -8.0674 | 117.8955 | -8.1729 | 122.9215 | -8.1026 | 120.8594 | -8.1026 | 119.9854 | -8.1553 | 123.0601 | -8.085 | 1485.8505 | -8.1729 | 1476.8999 | -8.1729 | 1474.4375 | -8.208 | 1485.566 | -8.208 | 1485.2818 | -8.2256 | 1492.3385 | -8.2256 | 1486.706
| -8.0323 | 117.736 | -8.1377 | 121.9703 | -8.0674 | 119.7071 | -8.0674 | 119.9854 | -8.1202 | 122.2677 | -8.0499 | 1480.7774 | -8.1377 | 1475.9395 | -8.1377 | 1473.0794 | -8.1729 | 1478.8313 | -8.1729 | 1482.4572 | -8.1905 | 1489.0016 | -8.1905 | 1482.7383
| -7.9971 | 117.8955 | -8.1026 | 121.5539 | -8.0323 | 119.4685 | -8.0323 | 119.6674 | -8.085 | 122.2677 | -8.0147 | 1478.4162 | -8.1026 | 1474.8463 | -8.1026 | 1471.3244 | -8.1377 | 1476.8999 | -8.1377 | 1480.0807 | -8.1553 | 1483.1604 | -8.1553 | 1475.2557
| -7.962 | 117.5167 | -8.0674 | 121.2563 | -7.9971 | 119.389 | -7.9971 | 119.2298 | -8.0499 | 122.1686 | -7.9795 | 1478.278 | -8.0674 | 1472.944 | -8.0674 | 1470.1161 | -8.1026 | 1474.5737 | -8.1026 | 1478.278 | -8.1202 | 1481.476 | -8.1202 | 1473.3505
| -7.9268 | 117.5964 | -8.0323 | 120.7005 | -7.962 | 119.4685 | -7.962 | 119.0906 | -8.0147 | 122.0893 | -7.9444 | 1477.588 | -8.0323 | 1472.944 | -8.0323 | 1470.6524 | -8.0674 | 1474.5737 | -8.0674 | 1477.3125 | -8.085 | 1478.8313 | -8.085 | 1472.1329
| -7.8917 | 117.7959 | -7.9971 | 120.6013 | -7.9268 | 119.5481 | -7.9268 | 118.2144 | -7.9795 | 122.1686 | -7.9092 | 1477.4502 | -7.9971 | 1472.8087 | -7.9971 | 1468.5134 | -8.0323 | 1471.8631 | -8.0323 | 1476.3507 | -8.0499 | 1475.8026 | -8.0499 | 1469.5808
| -7.8565 | 117.7959 | -7.962 | 120.6013 | -7.8917 | 120.0053 | -7.8917 | 117.756 | -7.9444 | 122.0893 | -7.8741 | 1478.4162 | -7.962 | 1469.4471 | -7.962 | 1463.8924 | -7.9971 | 1471.7283 | -7.9971 | 1473.2149 | -8.0147 | 1473.2149 | -8.0147 | 1465.2048
| -7.8214 | 117.7959 | -7.9268 | 120.7005 | -7.8565 | 120.0053 | -7.8565 | 116.9978 | -7.9092 | 121.5737 | -7.8389 | 1479.5247 | -7.9268 | 1465.3364 | -7.9268 | 1461.6755 | -7.962 | 1470.6524 | -7.962 | 1470.25 | -7.9795 | 1473.7576 | -7.9795 | 1464.8104
| -7.7862 | 118.2543 | -7.8917 | 120.8395 | -7.8214 | 119.8861 | -7.8214 | 116.9978 | -7.8741 | 120.7204 | -7.8038 | 1477.588 | -7.8917 | 1466.7879 | -7.8917 | 1461.0268 | -7.9268 | 1464.6791 | -7.9268 | 1470.921 | -7.9444 | 1473.0794 | -7.9444 | 1464.5478
| -7.7511 | 118.2543 | -7.8565 | 120.8395 | -7.7862 | 119.4287 | -7.7862 | 117.1974 | -7.8389 | 119.4884 | -7.7686 | 1479.1084 | -7.8565 | 1465.0733 | -7.8565 | 1460.6383 | -7.8917 | 1462.7166 | -7.8917 | 1469.4471 | -7.9092 | 1472.8087 | -7.9092 | 1463.6307
| -7.7159 | 118.7124 | -7.8214 | 120.4622 | -7.7511 | 118.8119 | -7.7511 | 116.9778 | -7.8038 | 118.9512 | -7.7335 | 1479.1084 | -7.8214 | 1465.0733 | -7.8214 | 1462.5863 | -7.8565 | 1462.456 | -7.8565 | 1468.2473 | -7.8741 | 1472.944 | -7.8741 | 1461.4159
| -7.6808 | 119.0508 | -7.7862 | 119.9854 | -7.7159 | 118.1547 | -7.7159 | 116.9778 | -7.7686 | 118.3539 | -7.6983 | 1477.3125 | -7.7862 | 1469.18 | -7.7862 | 1464.9418 | -7.8214 | 1462.456 | -7.8214 | 1469.4471 | -7.8389 | 1472.944 | -7.8389 | 1461.1565
| -7.6456 | 118.9911 | -7.7511 | 119.9854 | -7.6808 | 117.8955 | -7.6808 | 117.0177 | -7.7335 | 118.2742 | -7.6632 | 1477.0374 | -7.7511 | 1469.18 | -7.7511 | 1465.0733 | -7.7862 | 1462.847 | -7.7862 | 1472.8087 | -7.8038 | 1471.4589 | -7.8038 | 1461.4159
| -7.6105 | 118.8517 | -7.7159 | 119.9656 | -7.6456 | 117.8955 | -7.6456 | 116.9179 | -7.6983 | 118.4336 | -7.628 | 1476.8999 | -7.7159 | 1469.18 | -7.7159 | 1465.468 | -7.7511 | 1463.108 | -7.7511 | 1470.5182 | -7.7686 | 1471.4589 | -7.7686 | 1461.4159
| -7.5753 | 118.9911 | -7.6808 | 120.045 | -7.6105 | 118.1547 | -7.6105 | 116.9778 | -7.6632 | 118.7124 | -7.5929 | 1475.2557 | -7.6808 | 1469.9821 | -7.6808 | 1465.8633 | -7.7159 | 1463.7615 | -7.7159 | 1470.5182 | -7.7335 | 1470.5182 | -7.7335 | 1463.4999
| -7.5401 | 118.6128 | -7.6456 | 120.4225 | -7.5753 | 118.334 | -7.5753 | 117.0177 | -7.628 | 118.2941 | -7.5577 | 1473.8934 | -7.6456 | 1467.0527 | -7.6456 | 1465.0733 | -7.6808 | 1464.4166 | -7.6808 | 1470.1161 | -7.6983 | 1470.25 | -7.6983 | 1464.6791
| -7.505 | 118.1148 | -7.6105 | 120.502 | -7.5401 | 119.1105 | -7.5401 | 117.2174 | -7.5929 | 118.1347 | -7.5226 | 1470.5182 | -7.6105 | 1467.0527 | -7.6105 | 1463.3692 | -7.6456 | 1463.108 | -7.6456 | 1469.7145 | -7.6632 | 1470.25 | -7.6632 | 1465.468
| -7.4698 | 117.9155 | -7.5753 | 120.641 | -7.505 | 119.1105 | -7.505 | 117.5565 | -7.5577 | 118.2941 | -7.4874 | 1470.1161 | -7.5753 | 1466.7879 | -7.5753 | 1460.2503 | -7.6105 | 1464.0234 | -7.6105 | 1468.2473 | -7.628 | 1470.25 | -7.628 | 1465.468
| -7.4347 | 118.0151 | -7.5401 | 120.7998 | -7.4698 | 119.3293 | -7.4698 | 117.2772 | -7.5226 | 118.573 | -7.4523 | 1468.6466 | -7.5401 | 1464.6791 | -7.5401 | 1459.7338 | -7.5753 | 1464.2855 | -7.5753 | 1466.5234 | -7.5929 | 1467.7157 | -7.5929 | 1463.8924
| -7.3995 | 118.1945 | -7.505 | 120.8594 | -7.4347 | 119.9457 | -7.4347 | 117.2772 | -7.4874 | 118.1746 | -7.4171 | 1468.6466 | -7.505 | 1464.5478 | -7.505 | 1458.3182 | -7.5401 | 1462.7166 | -7.5401 | 1466.5234 | -7.5577 | 1468.2473 | -7.5577 | 1464.5478
| -7.3644 | 118.2543 | -7.4698 | 121.1769 | -7.3995 | 120.4225 | -7.3995 | 117.9553 | -7.4523 | 118.5332 | -7.382 | 1470.5182 | -7.4698 | 1464.9418 | -7.4698 | 1459.7338 | -7.505 | 1460.7678 | -7.505 | 1469.0466 | -7.5226 | 1467.8485 | -7.5226 | 1462.847
| -7.3292 | 118.2742 | -7.4347 | 121.7324 | -7.3644 | 120.502 | -7.3644 | 118.1547 | -7.4171 | 118.5332 | -7.3468 | 1474.4375 | -7.4347 | 1467.9813 | -7.4347 | 1462.3258 | -7.4698 | 1460.7678 | -7.4698 | 1468.2473 | -7.4874 | 1467.7157 | -7.4874 | 1461.0268
| -7.2941 | 118.4535 | -7.3995 | 121.7324 | -7.3292 | 120.3629 | -7.3292 | 118.4336 | -7.382 | 118.4734 | -7.3117 | 1475.9395 | -7.3995 | 1468.1143 | -7.3995 | 1464.5478 | -7.4347 | 1460.7678 | -7.4347 | 1469.5808 | -7.4523 | 1467.8485 | -7.4523 | 1456.6541
| -7.2589 | 118.4535 | -7.3644 | 121.7522 | -7.2941 | 119.9854 | -7.2941 | 119.0906 | -7.3468 | 118.4734 | -7.2765 | 1476.2135 | -7.3644 | 1469.8483 | -7.3644 | 1464.8104 | -7.3995 | 1460.3796 | -7.3995 | 1469.9821 | -7.4171 | 1468.5134 | -7.4171 | 1455.8893
| -7.2238 | 118.3738 | -7.3292 | 121.4745 | -7.2589 | 119.6674 | -7.2589 | 119.2895 | -7.3117 | 118.9313 | -7.2414 | 1476.3507 | -7.3292 | 1469.7145 | -7.3292 | 1465.7315 | -7.3644 | 1461.4159 | -7.3644 | 1470.5182 | -7.382 | 1468.9132 | -7.382 | 1456.6541
| -7.1886 | 118.055 | -7.2941 | 121.3158 | -7.2238 | 119.4089 | -7.2238 | 119.4486 | -7.2765 | 119.3094 | -7.2062 | 1478.0018 | -7.2941 | 1469.7145 | -7.2941 | 1465.7315 | -7.3292 | 1462.0655 | -7.3292 | 1470.5182 | -7.3468 | 1469.3135 | -7.3468 | 1462.7166
| -7.1535 | 118.334 | -7.2589 | 121.3158 | -7.1886 | 119.4089 | -7.1886 | 119.4486 | -7.2414 | 120.2239 | -7.171 | 1479.9415 | -7.2589 | 1469.8483 | -7.2589 | 1466.5234 | -7.2941 | 1462.7166 | -7.2941 | 1471.5936 | -7.3117 | 1470.5182 | -7.3117 | 1465.2048
| -7.1183 | 118.3539 | -7.2238 | 121.296 | -7.1535 | 119.4287 | -7.1535 | 119.2895 | -7.2062 | 120.2835 | -7.1359 | 1483.5833 | -7.2238 | 1472.944 | -7.2238 | 1469.18 | -7.2589 | 1463.6307 | -7.2589 | 1471.5936 | -7.2765 | 1471.0554 | -7.2765 | 1465.2048
| -7.0832 | 118.4535 | -7.1886 | 121.296 | -7.1183 | 119.4486 | -7.1183 | 118.9313 | -7.171 | 120.3431 | -7.1007 | 1485.9929 | -7.1886 | 1474.0293 | -7.1886 | 1473.3505 | -7.2238 | 1464.6791 | -7.2238 | 1476.0765 | -7.2414 | 1472.944 | -7.2414 | 1465.0733
| -7.048 | 118.7124 | -7.1535 | 121.296 | -7.0832 | 119.6077 | -7.0832 | 118.2742 | -7.1359 | 120.2437 | -7.0656 | 1489.29 | -7.1535 | 1478.9698 | -7.1535 | 1475.3923 | -7.1886 | 1471.3244 | -7.1886 | 1478.5545 | -7.2062 | 1475.6658 | -7.2062 | 1465.5997
| -7.0129 | 118.3937 | -7.1183 | 120.7403 | -7.048 | 120.1841 | -7.048 | 117.5765 | -7.1007 | 119.7071 | -7.0304 | 1492.0466 | -7.1183 | 1481.476 | -7.1183 | 1479.5247 | -7.1535 | 1473.8934 | -7.1535 | 1481.1963 | -7.171 | 1477.4502 | -7.171 | 1466.1272
| -6.9777 | 118.7124 | -7.0832 | 120.7403 | -7.0129 | 120.3431 | -7.0129 | 117.1774 | -7.0656 | 119.2696 | -6.9953 | 1494.2445 | -7.0832 | 1484.5727 | -7.0832 | 1480.2198 | -7.1183 | 1476.4879 | -7.1183 | 1481.476 | -7.1359 | 1478.1399 | -7.1359 | 1466.7879
| -6.9426 | 118.3937 | -7.048 | 120.7403 | -6.9777 | 120.3431 | -6.9777 | 117.357 | -7.0304 | 119.1502 | -6.9601 | 1497.8018 | -7.048 | 1485.2818 | -7.048 | 1481.896 | -7.0832 | 1478.1399 | -7.0832 | 1482.1765 | -7.1007 | 1479.5247 | -7.1007 | 1470.921
| -6.9074 | 118.1148 | -7.0129 | 120.7403 | -6.9426 | 119.9258 | -6.9426 | 117.6164 | -6.9953 | 119.2298 | -6.925 | 1497.9512 | -7.0129 | 1487.4211 | -7.0129 | 1484.7144 | -7.048 | 1479.5247 | -7.048 | 1483.0196 | -7.0656 | 1480.6379 | -7.0656 | 1473.7576
| -6.8723 | 118.0949 | -6.9777 | 120.8395 | -6.9074 | 119.6276 | -6.9074 | 118.3539 | -6.9601 | 119.2696 | -6.8898 | 1500.3526 | -6.9777 | 1494.0974 | -6.9777 | 1486.5632 | -7.0129 | 1481.0566 | -7.0129 | 1485.8505 | -7.0304 | 1481.476 | -7.0304 | 1475.9395
| -6.8371 | 118.1148 | -6.9426 | 120.4821 | -6.8723 | 119.2696 | -6.8723 | 118.2543 | -6.925 | 119.4685 | -6.8547 | 1499.9006 | -6.9426 | 1494.9814 | -6.9426 | 1488.4257 | -6.9777 | 1484.5727 | -6.9777 | 1487.9947 | -6.9953 | 1482.3168 | -6.9953 | 1478.0018
| -6.802 | 118.055 | -6.9074 | 120.2835 | -6.8371 | 119.2696 | -6.8371 | 118.4535 | -6.8898 | 119.7071 | -6.8195 | 1497.9512 | -6.9074 | 1494.9814 | -6.9074 | 1487.8512 | -6.9426 | 1486.4205 | -6.9426 | 1492.4846 | -6.9601 | 1485.9929 | -6.9601 | 1479.1084
| -6.7668 | 116.9179 | -6.8723 | 120.1245 | -6.802 | 119.2099 | -6.802 | 118.2543 | -6.8547 | 120.1047 | -6.7844 | 1497.2054 | -6.8723 | 1492.6308 | -6.8723 | 1487.8512 | -6.9074 | 1486.8489 | -6.9074 | 1492.4846 | -6.925 | 1487.2779 | -6.925 | 1481.3361
| -6.7316 | 115.8786 | -6.8371 | 120.2835 | -6.7668 | 119.5878 | -6.7668 | 117.9155 | -6.8195 | 119.6674 | -6.7492 | 1497.6526 | -6.8371 | 1490.1572 | -6.8371 | 1487.8512 | -6.8723 | 1486.9918 | -6.8723 | 1492.6308 | -6.8898 | 1489.0016 | -6.8898 | 1481.896
| -6.6965 | 115.3782 | -6.802 | 120.4027 | -6.7316 | 119.5282 | -6.7316 | 117.9155 | -6.7844 | 120.045 | -6.7141 | 1494.0974 | -6.802 | 1488.8575 | -6.802 | 1484.8561 | -6.8371 | 1484.4311 | -6.8371 | 1489.0016 | -6.8547 | 1488.8575 | -6.8547 | 1483.3013
| -6.6613 | 115.0377 | -6.7668 | 120.4027 | -6.6965 | 119.5282 | -6.6965 | 118.1148 | -6.7492 | 120.0053 | -6.6789 | 1496.0168 | -6.7668 | 1487.7077 | -6.7668 | 1482.5977 | -6.802 | 1480.9169 | -6.802 | 1489.0016 | -6.8195 | 1489.4343 | -6.8195 | 1484.1482
| -6.6262 | 114.9775 | -6.7316 | 120.5615 | -6.6613 | 119.2099 | -6.6613 | 118.1148 | -6.7141 | 119.3691 | -6.6438 | 1495.2768 | -6.7316 | 1489.5787 | -6.7316 | 1488.1383 | -6.7668 | 1480.9169 | -6.7668 | 1487.9947 | -6.7844 | 1486.1354 | -6.7844 | 1484.1482
| -6.591 | 114.9575 | -6.6965 | 120.4225 | -6.6262 | 119.4685 | -6.6262 | 117.6164 | -6.6789 | 119.3691 | -6.6086 | 1496.9077 | -6.6965 | 1490.1572 | -6.6965 | 1489.7232 | -6.7316 | 1480.4985 | -6.7316 | 1488.1383 | -6.7492 | 1483.8656 | -6.7492 | 1483.0196
| -6.5559 | 114.396 | -6.6613 | 120.3629 | -6.591 | 118.4535 | -6.591 | 117.6563 | -6.6438 | 119.4685 | -6.5735 | 1493.6565 | -6.6613 | 1488.7135 | -6.6613 | 1488.7135 | -6.6965 | 1480.4985 | -6.6965 | 1488.4257 | -6.7141 | 1483.3013 | -6.7141 | 1480.0807
| -6.5207 | 114.2756 | -6.6262 | 120.204 | -6.5559 | 118.4734 | -6.5559 | 117.5565 | -6.6086 | 119.4486 | -6.5383 | 1493.6565 | -6.6262 | 1489.8678 | -6.6262 | 1489.8678 | -6.6613 | 1481.6159 | -6.6613 | 1488.4257 | -6.6789 | 1486.2779 | -6.6789 | 1482.0362
| -6.4856 | 113.4923 | -6.591 | 119.9059 | -6.5207 | 118.5132 | -6.5207 | 117.0577 | -6.5735 | 119.2497 | -6.5032 | 1493.6565 | -6.591 | 1489.8678 | -6.591 | 1490.3021 | -6.6262 | 1484.0069 | -6.6262 | 1488.4257 | -6.6438 | 1490.0125 | -6.6438 | 1480.0807
| -6.4504 | 113.2309 | -6.5559 | 119.7668 | -6.4856 | 118.5132 | -6.4856 | 117.0577 | -6.5383 | 119.1303 | -6.468 | 1493.363 | -6.5559 | 1489.0016 | -6.5559 | 1488.282 | -6.591 | 1484.8561 | -6.591 | 1489.0016 | -6.6086 | 1491.755 | -6.6086 | 1480.0807
| -6.4153 | 113.251 | -6.5207 | 119.4884 | -6.4504 | 117.736 | -6.4504 | 117.0377 | -6.5032 | 119.0508 | -6.4329 | 1492.4846 | -6.5207 | 1489.0016 | -6.5207 | 1485.4239 | -6.5559 | 1486.9918 | -6.5559 | 1490.8822 | -6.5735 | 1492.3385 | -6.5735 | 1480.6379
| -6.3801 | 113.2912 | -6.4856 | 118.9712 | -6.4153 | 117.1974 | -6.4153 | 116.7182 | -6.468 | 118.9512 | -6.3977 | 1491.9008 | -6.4856 | 1489.0016 | -6.4856 | 1487.5644 | -6.5207 | 1487.4211 | -6.5207 | 1493.0699 | -6.5383 | 1493.2164 | -6.5383 | 1481.3361
| -6.345 | 113.2711 | -6.4504 | 118.3738 | -6.3801 | 117.1176 | -6.3801 | 116.0586 | -6.4329 | 118.6526 | -6.3625 | 1491.755 | -6.4504 | 1487.2779 | -6.4504 | 1487.4211 | -6.4856 | 1486.9918 | -6.4856 | 1491.6093 | -6.5032 | 1491.3182 | -6.5032 | 1480.3591
| -6.3098 | 113.7937 | -6.4153 | 117.9553 | -6.345 | 116.858 | -6.345 | 115.6985 | -6.3977 | 118.9512 | -6.3274 | 1489.7232 | -6.4153 | 1488.7135 | -6.4153 | 1485.7082 | -6.4504 | 1487.2779 | -6.4504 | 1490.7371 | -6.468 | 1489.8678 | -6.468 | 1479.9415
| -6.2747 | 114.396 | -6.3801 | 117.5765 | -6.3098 | 116.3385 | -6.3098 | 115.4783 | -6.3625 | 118.9712 | -6.2922 | 1489.7232 | -6.3801 | 1488.7135 | -6.3801 | 1485.7082 | -6.4153 | 1486.4205 | -6.4153 | 1490.1572 | -6.4329 | 1489.1457 | -6.4329 | 1479.5247
| -6.2395 | 114.717 | -6.345 | 117.0377 | -6.2747 | 115.4583 | -6.2747 | 115.0777 | -6.3274 | 118.8318 | -6.2571 | 1489.7232 | -6.345 | 1488.282 | -6.345 | 1484.8561 | -6.3801 | 1485.9929 | -6.3801 | 1489.0016 | -6.3977 | 1488.4257 | -6.3977 | 1479.9415
| -6.2044 | 114.717 | -6.3098 | 116.4984 | -6.2395 | 114.5164 | -6.2395 | 114.9575 | -6.2922 | 118.3539 | -6.2219 | 1486.706 | -6.3098 | 1488.282 | -6.3098 | 1478.0018 | -6.345 | 1484.8561 | -6.345 | 1486.4205 | -6.3625 | 1486.9918 | -6.3625 | 1479.3858
| -6.1692 | 114.717 | -6.2747 | 116.2786 | -6.2044 | 114.3559 | -6.2044 | 114.115 | -6.2571 | 118.314 | -6.1868 | 1483.8656 | -6.2747 | 1488.282 | -6.2747 | 1474.8463 | -6.3098 | 1481.476 | -6.3098 | 1485.9929 | -6.3274 | 1485.7082 | -6.3274 | 1479.1084
| -6.1341 | 113.9544 | -6.2395 | 115.5384 | -6.1692 | 114.4562 | -6.1692 | 113.9544 | -6.2219 | 117.9753 | -6.1516 | 1483.7244 | -6.2395 | 1485.4239 | -6.2395 | 1473.0794 | -6.2747 | 1477.8638 | -6.2747 | 1483.4423 | -6.2922 | 1485.7082 | -6.2922 | 1478.278
| -6.0989 | 113.5325 | -6.2044 | 115.4183 | -6.1341 | 114.8974 | -6.1341 | 113.9142 | -6.1868 | 117.6563 | -6.1165 | 1482.4572 | -6.2044 | 1478.8313 | -6.2044 | 1472.268 | -6.2395 | 1476.3507 | -6.2395 | 1481.1963 | -6.2571 | 1483.0196 | -6.2571 | 1476.2135
| -6.0638 | 113.854 | -6.1692 | 115.6585 | -6.0989 | 115.7185 | -6.0989 | 113.0095 | -6.1516 | 117.1575 | -6.0813 | 1480.2198 | -6.1692 | 1471.5936 | -6.1692 | 1471.998 | -6.2044 | 1473.6218 | -6.2044 | 1479.5247 | -6.2219 | 1481.7559 | -6.2219 | 1475.9395
| -6.0286 | 113.9544 | -6.1341 | 115.4383 | -6.0638 | 115.7385 | -6.0638 | 112.7881 | -6.1165 | 116.6383 | -6.0462 | 1479.1084 | -6.1341 | 1470.1161 | -6.1341 | 1470.3841 | -6.1692 | 1471.998 | -6.1692 | 1479.5247 | -6.1868 | 1479.2471 | -6.1868 | 1475.529
| -5.9935 | 113.6531 | -6.0989 | 115.4383 | -6.0286 | 115.7385 | -6.0286 | 112.7277 | -6.0813 | 115.9186 | -6.011 | 1477.3125 | -6.0989 | 1468.6466 | -6.0989 | 1468.9132 | -6.1341 | 1471.7283 | -6.1341 | 1477.3125 | -6.1516 | 1476.8999 | -6.1516 | 1475.2557
| -5.9583 | 113.251 | -6.0638 | 115.8386 | -5.9935 | 115.6185 | -5.9935 | 112.3853 | -6.0462 | 115.0377 | -5.9759 | 1477.3125 | -6.0638 | 1468.5134 | -6.0638 | 1467.5829 | -6.0989 | 1471.7283 | -6.0989 | 1473.4861 | -6.1165 | 1474.5737 | -6.1165 | 1474.4375
| -5.9231 | 113.5928 | -6.0286 | 115.8386 | -5.9583 | 115.2981 | -5.9583 | 112.3853 | -6.011 | 114.4963 | -5.9407 | 1478.6929 | -6.0286 | 1468.6466 | -6.0286 | 1466.9203 | -6.0638 | 1470.6524 | -6.0638 | 1473.4861 | -6.0813 | 1474.5737 | -6.0813 | 1469.18
| -5.888 | 112.6069 | -5.9935 | 115.4383 | -5.9231 | 114.7971 | -5.9231 | 113.1906 | -5.9759 | 113.9946 | -5.9056 | 1478.278 | -5.9935 | 1469.0466 | -5.9935 | 1466.6556 | -6.0286 | 1469.9821 | -6.0286 | 1472.5382 | -6.0462 | 1473.3505 | -6.0462 | 1467.5829
| -5.8528 | 112.345 | -5.9583 | 115.8186 | -5.888 | 114.7571 | -5.888 | 112.6472 | -5.9407 | 114.5365 | -5.8704 | 1479.5247 | -5.9583 | 1469.4471 | -5.9583 | 1466.9203 | -5.9935 | 1470.1161 | -5.9935 | 1471.5936 | -6.011 | 1473.7576 | -6.011 | 1465.9952
| -5.8177 | 112.103 | -5.9231 | 115.3782 | -5.8528 | 113.8339 | -5.8528 | 112.7277 | -5.9056 | 114.4963 | -5.8353 | 1479.8025 | -5.9231 | 1470.7867 | -5.9231 | 1468.7799 | -5.9583 | 1470.5182 | -5.9583 | 1471.5936 | -5.9759 | 1473.7576 | -5.9759 | 1466.9203
| -5.7825 | 111.6793 | -5.888 | 115.3782 | -5.8177 | 113.5727 | -5.8177 | 112.768 | -5.8704 | 114.5365 | -5.8001 | 1479.5247 | -5.888 | 1471.998 | -5.888 | 1467.5829 | -5.9231 | 1466.9203 | -5.9231 | 1471.998 | -5.9407 | 1470.7867 | -5.9407 | 1466.2592
| -5.7474 | 111.4572 | -5.8528 | 115.3782 | -5.7825 | 114.8773 | -5.7825 | 112.6472 | -5.8353 | 114.6768 | -5.765 | 1473.6218 | -5.8528 | 1469.4471 | -5.8528 | 1466.6556 | -5.888 | 1468.3803 | -5.888 | 1471.998 | -5.9056 | 1472.8087 | -5.9056 | 1466.2592
| -5.7122 | 111.0531 | -5.8177 | 115.3782 | -5.7474 | 114.8773 | -5.7474 | 112.5666 | -5.8001 | 115.218 | -5.7298 | 1472.944 | -5.8177 | 1467.5829 | -5.8177 | 1466.6556 | -5.8528 | 1468.6466 | -5.8528 | 1471.998 | -5.8704 | 1472.8087 | -5.8704 | 1466.2592
| -5.6771 | 110.8306 | -5.7825 | 115.3782 | -5.7122 | 114.3358 | -5.7122 | 112.4256 | -5.765 | 115.1579 | -5.6947 | 1471.4589 | -5.7825 | 1466.1272 | -5.7825 | 1465.9952 | -5.8177 | 1468.3803 | -5.8177 | 1472.8087 | -5.8353 | 1472.8087 | -5.8353 | 1466.9203
| -5.6419 | 110.6283 | -5.7474 | 115.0577 | -5.6771 | 114.5565 | -5.6771 | 112.7076 | -5.7298 | 115.1579 | -5.6595 | 1471.3244 | -5.7474 | 1462.9775 | -5.7474 | 1463.2386 | -5.7825 | 1467.8485 | -5.7825 | 1472.1329 | -5.8001 | 1472.1329 | -5.8001 | 1467.9813
| -5.6068 | 110.5474 | -5.7122 | 114.9375 | -5.6419 | 114.5565 | -5.6419 | 112.8888 | -5.6947 | 114.2957 | -5.6244 | 1471.3244 | -5.7122 | 1463.3692 | -5.7122 | 1459.7338 | -5.7474 | 1467.8485 | -5.7474 | 1470.25 | -5.765 | 1470.3841 | -5.765 | 1466.7879
| -5.5716 | 109.9599 | -5.6771 | 114.9375 | -5.6068 | 113.7535 | -5.6068 | 112.8485 | -5.6595 | 114.8372 | -5.5892 | 1470.6524 | -5.6771 | 1463.3692 | -5.6771 | 1457.8052 | -5.7122 | 1464.5478 | -5.7122 | 1468.2473 | -5.7298 | 1469.0466 | -5.7298 | 1466.5234
| -5.5365 | 109.9599 | -5.6419 | 114.6768 | -5.5716 | 113.251 | -5.5716 | 112.7277 | -5.6244 | 114.717 | -5.554 | 1470.5182 | -5.6419 | 1462.847 | -5.6419 | 1457.421 | -5.6771 | 1463.3692 | -5.6771 | 1466.1272 | -5.6947 | 1466.3913 | -5.6947 | 1466.1272
| -5.5013 | 109.6963 | -5.6068 | 114.396 | -5.5365 | 113.0498 | -5.5365 | 112.8485 | -5.5892 | 114.6768 | -5.5189 | 1468.2473 | -5.6068 | 1462.5863 | -5.6068 | 1457.421 | -5.6419 | 1462.1956 | -5.6419 | 1464.1544 | -5.6595 | 1463.8924 | -5.6595 | 1460.2503
| -5.4662 | 109.2497 | -5.5716 | 114.3559 | -5.5013 | 112.7277 | -5.5013 | 112.2442 | -5.554 | 114.6768 | -5.4837 | 1465.3364 | -5.5716 | 1463.7615 | -5.5716 | 1456.5265 | -5.6068 | 1459.7338 | -5.6068 | 1463.2386 | -5.6244 | 1463.108 | -5.6244 | 1458.0616
| -5.431 | 109.2497 | -5.5365 | 113.9343 | -5.4662 | 112.7076 | -5.4662 | 111.6793 | -5.5189 | 114.8773 | -5.4486 | 1463.108 | -5.5365 | 1462.9775 | -5.5365 | 1455.2534 | -5.5716 | 1458.5751 | -5.5716 | 1462.5863 | -5.5892 | 1463.108 | -5.5892 | 1458.3182
| -5.3959 | 109.0465 | -5.5013 | 113.9343 | -5.431 | 112.5062 | -5.431 | 111.0531 | -5.4837 | 114.6768 | -5.4134 | 1461.6755 | -5.5013 | 1462.3258 | -5.5013 | 1454.7456 | -5.5365 | 1457.0373 | -5.5365 | 1459.3471 | -5.554 | 1462.7166 | -5.554 | 1458.0616
| -5.3607 | 109.0465 | -5.4662 | 113.5928 | -5.3959 | 112.4256 | -5.3959 | 110.77 | -5.4486 | 114.396 | -5.3783 | 1461.5457 | -5.4662 | 1460.8973 | -5.4662 | 1453.2275 | -5.5013 | 1456.2714 | -5.5013 | 1458.8322 | -5.5189 | 1462.7166 | -5.5189 | 1458.0616
| -5.3256 | 109.2091 | -5.431 | 113.7133 | -5.3607 | 112.2442 | -5.3607 | 110.77 | -5.4134 | 114.095 | -5.3431 | 1461.5457 | -5.431 | 1458.1899 | -5.431 | 1450.3387 | -5.4662 | 1454.9994 | -5.4662 | 1458.8322 | -5.4837 | 1462.0655 | -5.4837 | 1457.9334
| -5.2904 | 109.3715 | -5.3959 | 113.6129 | -5.3256 | 111.5178 | -5.3256 | 110.5271 | -5.3783 | 113.5928 | -5.308 | 1462.3258 | -5.3959 | 1456.0166 | -5.3959 | 1449.091 | -5.431 | 1454.8725 | -5.431 | 1458.1899 | -5.4486 | 1461.5457 | -5.4486 | 1456.6541
| -5.2553 | 109.8585 | -5.3607 | 113.3113 | -5.2904 | 110.8913 | -5.2904 | 110.3246 | -5.3431 | 113.251 | -5.2728 | 1462.3258 | -5.3607 | 1454.8725 | -5.3607 | 1448.4689 | -5.3959 | 1454.8725 | -5.3959 | 1458.1899 | -5.4134 | 1460.8973 | -5.4134 | 1455.762
| -5.2201 | 110.2841 | -5.3256 | 112.0829 | -5.2553 | 110.5878 | -5.2553 | 109.7165 | -5.308 | 113.1906 | -5.2377 | 1458.7036 | -5.3256 | 1452.9752 | -5.3256 | 1447.4762 | -5.3607 | 1454.3654 | -5.3607 | 1456.144 | -5.3783 | 1459.7338 | -5.3783 | 1454.7456
| -5.1849 | 110.1626 | -5.2904 | 112.1434 | -5.2201 | 110.1423 | -5.2201 | 109.0872 | -5.2728 | 113.0297 | -5.2025 | 1458.0616 | -5.2904 | 1451.9682 | -5.2904 | 1446.3629 | -5.3256 | 1453.8591 | -5.3256 | 1454.4921 | -5.3431 | 1456.9095 | -5.3431 | 1453.3537
| -5.1498 | 109.818 | -5.2553 | 111.7803 | -5.1849 | 109.5136 | -5.1849 | 108.7619 | -5.2377 | 111.982 | -5.1674 | 1455.2534 | -5.2553 | 1449.2155 | -5.2553 | 1447.4762 | -5.2904 | 1451.8425 | -5.2904 | 1454.4921 | -5.308 | 1455.6347 | -5.308 | 1448.7176
| -5.1146 | 109.5745 | -5.2201 | 111.6793 | -5.1498 | 110.1423 | -5.1498 | 108.5992 | -5.2025 | 111.336 | -5.1322 | 1453.6063 | -5.2201 | 1447.6001 | -5.2201 | 1445.8693 | -5.2553 | 1448.4689 | -5.2553 | 1454.4921 | -5.2728 | 1454.6188 | -5.2728 | 1445.9926
| -5.0795 | 109.5745 | -5.1849 | 111.639 | -5.1146 | 109.7977 | -5.1146 | 108.5992 | -5.1674 | 111.336 | -5.0971 | 1450.8392 | -5.1849 | 1445.2533 | -5.1849 | 1444.2702 | -5.2201 | 1445.8693 | -5.2201 | 1452.4713 | -5.2377 | 1453.1013 | -5.2377 | 1443.5346
| -5.0443 | 108.7823 | -5.1498 | 111.639 | -5.0795 | 109.5136 | -5.0795 | 108.4364 | -5.1322 | 111.1137 | -5.0619 | 1450.8392 | -5.1498 | 1445.2533 | -5.1498 | 1443.2898 | -5.1849 | 1444.1475 | -5.1849 | 1449.2155 | -5.2025 | 1451.5914 | -5.2025 | 1442.1902
| -5.0092 | 108.192 | -5.1146 | 111.2956 | -5.0443 | 110.2436 | -5.0443 | 108.4364 | -5.0971 | 110.952 | -5.0268 | 1450.0887 | -5.1146 | 1444.3929 | -5.1146 | 1440.6078 | -5.1498 | 1439.8799 | -5.1498 | 1445.746 | -5.1674 | 1448.2204 | -5.1674 | 1442.5563
| -4.974 | 107.9271 | -5.0795 | 111.861 | -5.0092 | 110.4056 | -5.0092 | 108.3346 | -5.0619 | 110.9924 | -4.9916 | 1449.091 | -5.0795 | 1442.0682 | -5.0795 | 1436.7417 | -5.1146 | 1439.8799 | -5.1146 | 1443.9022 | -5.1322 | 1447.3523 | -5.1322 | 1441.3373
| -4.9389 | 107.8864 | -5.0443 | 112.103 | -4.974 | 110.9924 | -4.974 | 108.3346 | -5.0268 | 110.8104 | -4.9565 | 1448.2204 | -5.0443 | 1439.8799 | -5.0443 | 1436.3812 | -5.0795 | 1437.2228 | -5.0795 | 1442.4342 | -5.0971 | 1446.4864 | -5.0971 | 1439.5164
| -4.9037 | 107.7028 | -5.0092 | 112.1232 | -4.9389 | 111.336 | -4.9389 | 108.4974 | -4.9916 | 110.7093 | -4.9213 | 1446.4864 | -5.0092 | 1439.5164 | -5.0092 | 1435.6613 | -5.0443 | 1436.9822 | -5.0443 | 1442.1902 | -5.0619 | 1444.8843 | -5.0619 | 1438.4282
| -4.8686 | 107.7028 | -4.974 | 112.345 | -4.9037 | 111.3158 | -4.9037 | 108.5992 | -4.9565 | 110.77 | -4.8862 | 1444.8843 | -4.974 | 1438.1868 | -4.974 | 1434.9427 | -5.0092 | 1434.7035 | -5.0092 | 1440.2437 | -5.0268 | 1442.0682 | -5.0268 | 1437.2228
| -4.8334 | 108.355 | -4.9389 | 112.3046 | -4.8686 | 110.9924 | -4.8686 | 108.9043 | -4.9213 | 110.7497 | -4.851 | 1443.5346 | -4.9389 | 1436.8619 | -4.9389 | 1434.2254 | -4.974 | 1434.4643 | -4.974 | 1438.9115 | -4.9916 | 1439.3954 | -4.9916 | 1433.9865
| -4.7983 | 110.122 | -4.9037 | 112.0022 | -4.8334 | 110.4056 | -4.8334 | 109.0465 | -4.8862 | 110.6688 | -4.8159 | 1441.9463 | -4.9037 | 1436.8619 | -4.9037 | 1433.9865 | -4.9389 | 1433.5093 | -4.9389 | 1438.9115 | -4.9565 | 1439.3954 | -4.9565 | 1432.1996
| -4.7631 | 111.7197 | -4.8686 | 111.8408 | -4.7983 | 109.4324 | -4.7983 | 109.1278 | -4.851 | 110.6688 | -4.7807 | 1442.0682 | -4.8686 | 1436.8619 | -4.8686 | 1432.0807 | -4.9037 | 1433.6285 | -4.9037 | 1438.9115 | -4.9213 | 1437.825 | -4.9213 | 1430.8939
| -4.728 | 112.345 | -4.8334 | 111.5582 | -4.7631 | 109.1888 | -4.7631 | 108.9449 | -4.8159 | 110.6688 | -4.7455 | 1441.3373 | -4.8334 | 1436.8619 | -4.8334 | 1432.0807 | -4.8686 | 1433.2708 | -4.8686 | 1437.825 | -4.8862 | 1437.825 | -4.8862 | 1430.42
| -4.6928 | 112.345 | -4.7983 | 110.6081 | -4.728 | 109.1888 | -4.728 | 108.9246 | -4.7807 | 110.6688 | -4.7104 | 1439.0324 | -4.7983 | 1435.1821 | -4.7983 | 1431.9619 | -4.8334 | 1433.9865 | -4.8334 | 1437.825 | -4.851 | 1437.7045 | -4.851 | 1428.6474
| -4.6577 | 112.1635 | -4.7631 | 110.6688 | -4.6928 | 108.8839 | -4.6928 | 108.8839 | -4.7455 | 110.8104 | -4.6752 | 1437.3432 | -4.7631 | 1433.1517 | -4.7631 | 1430.065 | -4.7983 | 1433.2708 | -4.7983 | 1436.9822 | -4.8159 | 1436.5013 | -4.8159 | 1428.6474
| -4.6225 | 111.8004 | -4.728 | 111.2552 | -4.6577 | 108.7619 | -4.6577 | 108.0902 | -4.7104 | 111.5784 | -4.6401 | 1436.8619 | -4.728 | 1430.1833 | -4.728 | 1429.0014 | -4.7631 | 1432.9135 | -4.7631 | 1433.6285 | -4.7807 | 1436.5013 | -4.7807 | 1428.6474
| -4.5874 | 111.3764 | -4.6928 | 110.3651 | -4.6225 | 108.9652 | -4.6225 | 107.4988 | -4.6752 | 110.5676 | -4.6049 | 1436.2612 | -4.6928 | 1429.3557 | -4.6928 | 1424.5364 | -4.728 | 1428.4116 | -4.728 | 1431.2496 | -4.7455 | 1435.5415 | -4.7455 | 1427.1166
| -4.5522 | 111.2956 | -4.6577 | 109.9599 | -4.5874 | 109.27 | -4.5874 | 107.4988 | -4.6401 | 110.2639 | -4.5698 | 1435.1821 | -4.6577 | 1427.2341 | -4.6577 | 1421.6193 | -4.6928 | 1427.4694 | -4.6928 | 1429.9467 | -4.7104 | 1432.7944 | -4.7104 | 1425.0046
| -4.5171 | 111.3966 | -4.6225 | 110.9116 | -4.5522 | 109.818 | -4.5522 | 107.4988 | -4.6049 | 109.6963 | -4.5346 | 1435.6613 | -4.6225 | 1426.2942 | -4.6225 | 1420.6888 | -4.6577 | 1425.2388 | -4.6577 | 1428.0581 | -4.6752 | 1427.9403 | -4.6752 | 1422.901
| -4.4819 | 111.2956 | -4.5874 | 110.9116 | -4.5171 | 112.103 | -4.5171 | 107.4988 | -4.5698 | 109.1684 | -4.4995 | 1435.9011 | -4.5874 | 1425.356 | -4.5874 | 1420.9213 | -4.6225 | 1422.4346 | -4.6225 | 1425.5904 | -4.6401 | 1424.7704 | -4.6401 | 1420.9213
| -4.4468 | 111.6591 | -4.5522 | 112.6673 | -4.4819 | 112.2038 | -4.4819 | 107.7844 | -4.5346 | 108.7619 | -4.4643 | 1435.9011 | -4.5522 | 1423.8349 | -4.5522 | 1421.2702 | -4.5874 | 1422.4346 | -4.5874 | 1424.6534 | -4.6049 | 1422.085 | -4.6049 | 1418.6003
| -4.4116 | 111.6591 | -4.5171 | 114.396 | -4.4468 | 112.627 | -4.4468 | 108.3346 | -4.4995 | 108.9449 | -4.4292 | 1437.3432 | -4.5171 | 1426.2942 | -4.5171 | 1423.3677 | -4.5522 | 1424.4194 | -4.5522 | 1424.6534 | -4.5698 | 1422.085 | -4.5698 | 1416.8647
| -4.3764 | 112.0829 | -4.4819 | 114.9575 | -4.4116 | 112.8888 | -4.4116 | 109.7165 | -4.4643 | 109.1684 | -4.394 | 1438.9115 | -4.4819 | 1426.999 | -4.4819 | 1424.1856 | -4.5171 | 1424.6534 | -4.5171 | 1426.999 | -4.5346 | 1423.8349 | -4.5346 | 1418.6003
| -4.3413 | 112.3248 | -4.4468 | 115.0176 | -4.3764 | 112.8687 | -4.3764 | 110.4866 | -4.4292 | 108.8636 | -4.3589 | 1440.0011 | -4.4468 | 1430.065 | -4.4468 | 1425.8249 | -4.4819 | 1426.1768 | -4.4819 | 1428.1759 | -4.4995 | 1425.4732 | -4.4995 | 1418.9479
| -4.3061 | 112.5867 | -4.4116 | 114.8573 | -4.3413 | 112.627 | -4.3413 | 111.0733 | -4.394 | 109.7977 | -4.3237 | 1440.0011 | -4.4116 | 1428.6474 | -4.4116 | 1428.4116 | -4.4468 | 1427.2341 | -4.4468 | 1429.2376 | -4.4643 | 1427.4694 | -4.4643 | 1420.2241
| -4.271 | 112.7076 | -4.3764 | 114.717 | -4.3061 | 112.5867 | -4.3061 | 111.1137 | -4.3589 | 110.041 | -4.2886 | 1442.4342 | -4.3764 | 1430.065 | -4.3764 | 1428.8834 | -4.4116 | 1428.5295 | -4.4116 | 1431.8431 | -4.4292 | 1429.8284 | -4.4292 | 1422.318
| -4.2358 | 113.1906 | -4.3413 | 114.3559 | -4.271 | 112.5867 | -4.271 | 111.134 | -4.3237 | 110.7497 | -4.2534 | 1444.3929 | -4.3413 | 1431.4869 | -4.3413 | 1429.3557 | -4.3764 | 1429.8284 | -4.3764 | 1430.1833 | -4.394 | 1430.42 | -4.394 | 1425.2388
| -4.2007 | 113.5124 | -4.3061 | 114.0147 | -4.2358 | 112.7479 | -4.2358 | 111.0733 | -4.2886 | 111.982 | -4.2183 | 1444.0248 | -4.3061 | 1430.065 | -4.3061 | 1430.8939 | -4.3413 | 1429.8284 | -4.3413 | 1431.8431 | -4.3589 | 1430.42 | -4.3589 | 1426.2942
| -4.1655 | 113.5526 | -4.271 | 114.0147 | -4.2007 | 113.2107 | -4.2007 | 110.5676 | -4.2534 | 112.2643 | -4.1831 | 1443.6571 | -4.271 | 1431.4869 | -4.271 | 1430.1833 | -4.3061 | 1430.42 | -4.3061 | 1434.3448 | -4.3237 | 1431.8431 | -4.3237 | 1424.8875
| -4.1304 | 113.2309 | -4.2358 | 113.9343 | -4.1655 | 113.4521 | -4.1655 | 110.1423 | -4.2183 | 112.4457 | -4.148 | 1443.2898 | -4.2358 | 1434.3448 | -4.2358 | 1431.9619 | -4.271 | 1430.42 | -4.271 | 1430.5384 | -4.2886 | 1431.8431 | -4.2886 | 1426.2942
| -4.0952 | 112.8284 | -4.2007 | 114.2756 | -4.1304 | 113.4521 | -4.1304 | 110.1423 | -4.1831 | 112.6069 | -4.1128 | 1441.8244 | -4.2007 | 1434.1059 | -4.2007 | 1430.8939 | -4.2358 | 1429.2376 | -4.2358 | 1431.3682 | -4.2534 | 1432.0807 | -4.2534 | 1427.7048
| -4.0601 | 112.3046 | -4.1655 | 114.4161 | -4.0952 | 113.6129 | -4.0952 | 110.3854 | -4.148 | 112.4659 | -4.0777 | 1440.6078 | -4.1655 | 1433.6285 | -4.1655 | 1429.9467 | -4.2007 | 1428.2937 | -4.2007 | 1431.8431 | -4.2183 | 1432.0807 | -4.2183 | 1424.8875
| -4.0249 | 111.336 | -4.1304 | 114.4161 | -4.0601 | 113.6732 | -4.0601 | 111.1137 | -4.1128 | 112.486 | -4.0425 | 1438.549 | -4.1304 | 1431.7243 | -4.1304 | 1429.4738 | -4.1655 | 1428.2937 | -4.1655 | 1431.8431 | -4.1831 | 1432.5564 | -4.1831 | 1425.8249
| -3.9898 | 110.4866 | -4.0952 | 114.4161 | -4.0249 | 113.6732 | -4.0249 | 111.336 | -4.0777 | 112.4659 | -4.0074 | 1436.5013 | -4.0952 | 1430.8939 | -4.0952 | 1427.1166 | -4.1304 | 1429.2376 | -4.1304 | 1430.5384 | -4.148 | 1432.5564 | -4.148 | 1426.529
| -3.9546 | 110.0004 | -4.0601 | 115.1178 | -3.9898 | 114.0347 | -3.9898 | 111.6591 | -4.0425 | 112.486 | -3.9722 | 1436.2612 | -4.0601 | 1430.3016 | -4.0601 | 1424.5364 | -4.0952 | 1428.6474 | -4.0952 | 1430.8939 | -4.1128 | 1432.7944 | -4.1128 | 1426.529
| -3.9195 | 109.7774 | -4.0249 | 115.3181 | -3.9546 | 114.0347 | -3.9546 | 111.639 | -4.0074 | 112.768 | -3.937 | 1434.4643 | -4.0249 | 1430.6569 | -4.0249 | 1425.356 | -4.0601 | 1428.1759 | -4.0601 | 1429.3557 | -4.0777 | 1431.4869 | -4.0777 | 1423.8349
| -3.8843 | 109.4527 | -3.9898 | 115.5184 | -3.9195 | 114.0347 | -3.9195 | 112.0627 | -3.9722 | 112.8687 | -3.9019 | 1432.7944 | -3.9898 | 1429.4738 | -3.9898 | 1424.8875 | -4.0249 | 1427.8225 | -4.0249 | 1429.2376 | -4.0425 | 1432.7944 | -4.0425 | 1424.1856
| -3.8492 | 109.0872 | -3.9546 | 115.9786 | -3.8843 | 114.0347 | -3.8843 | 112.2038 | -3.937 | 113.7937 | -3.8667 | 1431.131 | -3.9546 | 1429.2376 | -3.9546 | 1422.4346 | -3.9898 | 1427.3518 | -3.9898 | 1426.8815 | -4.0074 | 1430.1833 | -4.0074 | 1422.7843
| -3.814 | 108.8229 | -3.9195 | 115.9786 | -3.8492 | 113.8339 | -3.8492 | 112.2038 | -3.9019 | 114.2957 | -3.8316 | 1432.7944 | -3.9195 | 1425.0046 | -3.9195 | 1421.2702 | -3.9546 | 1425.9422 | -3.9546 | 1425.8249 | -3.9722 | 1429.3557 | -3.9722 | 1422.4346
| -3.7789 | 108.8229 | -3.8843 | 115.9786 | -3.814 | 113.8138 | -3.814 | 112.2038 | -3.8667 | 114.095 | -3.7964 | 1433.6285 | -3.8843 | 1422.4346 | -3.8843 | 1421.1539 | -3.9195 | 1423.251 | -3.9195 | 1423.0176 | -3.937 | 1428.0581 | -3.937 | 1420.3403
| -3.7437 | 109.0465 | -3.8492 | 115.7385 | -3.7789 | 113.8138 | -3.7789 | 112.1837 | -3.8316 | 113.5124 | -3.7613 | 1432.5564 | -3.8492 | 1419.1798 | -3.8492 | 1418.4845 | -3.8843 | 1421.1539 | -3.8843 | 1421.0376 | -3.9019 | 1424.6534 | -3.9019 | 1419.0638
| -3.7086 | 109.473 | -3.814 | 115.3181 | -3.7437 | 112.2643 | -3.7437 | 112.103 | -3.7964 | 113.5124 | -3.7261 | 1431.2496 | -3.814 | 1422.4346 | -3.814 | 1417.4428 | -3.8492 | 1419.4117 | -3.8492 | 1420.2241 | -3.8667 | 1421.2702 | -3.8667 | 1416.9803
| -3.6734 | 109.676 | -3.7789 | 114.0347 | -3.7086 | 111.9618 | -3.7086 | 111.7399 | -3.7613 | 113.5526 | -3.691 | 1429.9467 | -3.7789 | 1423.251 | -3.7789 | 1417.4428 | -3.814 | 1419.4117 | -3.814 | 1419.5277 | -3.8316 | 1420.108 | -3.8316 | 1416.2871
| -3.6383 | 109.7165 | -3.7437 | 113.7736 | -3.6734 | 111.4168 | -3.6734 | 111.4572 | -3.7261 | 113.5928 | -3.6558 | 1425.5904 | -3.7437 | 1423.7181 | -3.7437 | 1417.0959 | -3.7789 | 1419.4117 | -3.7789 | 1419.2957 | -3.7964 | 1419.5277 | -3.7964 | 1414.096
| -3.6031 | 109.5745 | -3.7086 | 112.9894 | -3.6383 | 111.336 | -3.6383 | 110.8509 | -3.691 | 113.5928 | -3.6207 | 1420.805 | -3.7086 | 1429.4738 | -3.7086 | 1417.0959 | -3.7437 | 1418.832 | -3.7437 | 1419.1798 | -3.7613 | 1419.5277 | -3.7613 | 1413.2902
| -3.5679 | 109.0669 | -3.6734 | 112.8284 | -3.6031 | 111.336 | -3.6031 | 110.0004 | -3.6558 | 113.5928 | -3.5855 | 1418.1371 | -3.6734 | 1429.8284 | -3.6734 | 1417.5585 | -3.7086 | 1412.715 | -3.7086 | 1421.0376 | -3.7261 | 1421.2702 | -3.7261 | 1412.1402
| -3.5328 | 108.9856 | -3.6383 | 112.345 | -3.5679 | 111.336 | -3.5679 | 109.7571 | -3.6207 | 113.3716 | -3.5504 | 1417.6741 | -3.6383 | 1421.2702 | -3.6383 | 1417.5585 | -3.6734 | 1415.9408 | -3.6734 | 1422.2015 | -3.691 | 1421.2702 | -3.691 | 1411.7955
| -3.4976 | 108.2939 | -3.6031 | 111.982 | -3.5328 | 111.2148 | -3.5328 | 109.6557 | -3.5855 | 113.3113 | -3.5152 | 1417.6741 | -3.6031 | 1418.7162 | -3.6031 | 1414.096 | -3.6383 | 1415.9408 | -3.6383 | 1422.2015 | -3.6558 | 1421.9685 | -3.6558 | 1411.7955
| -3.4625 | 107.866 | -3.5679 | 111.538 | -3.4976 | 110.77 | -3.4976 | 109.473 | -3.5504 | 113.2309 | -3.4801 | 1419.0638 | -3.5679 | 1410.5324 | -3.5679 | 1416.0562 | -3.6031 | 1415.9408 | -3.6031 | 1417.9056 | -3.6207 | 1421.9685 | -3.6207 | 1411.6806
| -3.4273 | 107.866 | -3.5328 | 111.5986 | -3.4625 | 110.77 | -3.4625 | 108.9856 | -3.5152 | 113.0297 | -3.4449 | 1425.8249 | -3.5328 | 1410.5324 | -3.5328 | 1415.8254 | -3.5679 | 1415.9408 | -3.5679 | 1413.2902 | -3.5855 | 1419.1798 | -3.5855 | 1410.5324
| -3.3922 | 107.866 | -3.4976 | 111.9417 | -3.4273 | 110.6486 | -3.4273 | 108.7416 | -3.4801 | 112.3853 | -3.4098 | 1428.1759 | -3.4976 | 1406.98 | -3.4976 | 1414.096 | -3.5328 | 1416.8647 | -3.5328 | 1411.2212 | -3.5504 | 1415.3638 | -3.5504 | 1409.8441
| -3.357 | 107.4376 | -3.4625 | 112.4457 | -3.3922 | 110.3044 | -3.3922 | 108.1513 | -3.4449 | 111.6995 | -3.3746 | 1431.7243 | -3.4625 | 1412.1402 | -3.4625 | 1414.096 | -3.4976 | 1419.5277 | -3.4976 | 1411.2212 | -3.5152 | 1413.9809 | -3.5152 | 1409.8441
| -3.3219 | 108.2328 | -3.4273 | 112.9492 | -3.357 | 110.0815 | -3.357 | 108.1513 | -3.4098 | 111.6995 | -3.3395 | 1430.42 | -3.4273 | 1416.2871 | -3.4273 | 1415.8254 | -3.4625 | 1413.8657 | -3.4625 | 1418.3687 | -3.4801 | 1413.9809 | -3.4801 | 1407.7814
| -3.2867 | 108.4364 | -3.3922 | 112.6874 | -3.3219 | 110.1423 | -3.3219 | 108.1513 | -3.3746 | 111.0935 | -3.3043 | 1430.42 | -3.3922 | 1416.2871 | -3.3922 | 1411.4509 | -3.4273 | 1412.3701 | -3.4273 | 1423.9518 | -3.4449 | 1418.1371 | -3.4449 | 1408.1249
| -3.2516 | 108.5992 | -3.357 | 112.7277 | -3.2867 | 110.0612 | -3.2867 | 108.6602 | -3.3395 | 110.7902 | -3.2692 | 1426.8815 | -3.357 | 1416.2871 | -3.357 | 1412.1402 | -3.3922 | 1412.3701 | -3.3922 | 1425.356 | -3.4098 | 1421.3865 | -3.4098 | 1408.1249
| -3.2164 | 108.5992 | -3.3219 | 111.7803 | -3.2516 | 109.6557 | -3.2516 | 108.7416 | -3.3043 | 110.2233 | -3.234 | 1426.6465 | -3.3219 | 1416.2871 | -3.3219 | 1415.9408 | -3.357 | 1414.7873 | -3.357 | 1423.9518 | -3.3746 | 1428.5295 | -3.3746 | 1409.9588
| -3.1813 | 108.5992 | -3.2867 | 111.4572 | -3.2164 | 109.4527 | -3.2164 | 108.9043 | -3.2692 | 109.8788 | -3.1989 | 1423.251 | -3.2867 | 1415.8254 | -3.2867 | 1417.0959 | -3.3219 | 1414.9026 | -3.3219 | 1413.5204 | -3.3395 | 1428.5295 | -3.3395 | 1409.9588
| -3.1461 | 108.2939 | -3.2516 | 111.3562 | -3.1813 | 109.2497 | -3.1813 | 108.9043 | -3.234 | 109.9802 | -3.1637 | 1422.318 | -3.2516 | 1415.71 | -3.2516 | 1413.6355 | -3.2867 | 1414.9026 | -3.2867 | 1412.2552 | -3.3043 | 1421.3865 | -3.3043 | 1410.5324
| -3.111 | 108.0698 | -3.2164 | 111.0329 | -3.1461 | 108.8229 | -3.1461 | 108.7416 | -3.1989 | 110.2233 | -3.1285 | 1416.8647 | -3.2164 | 1413.7506 | -3.2164 | 1412.1402 | -3.2516 | 1412.9451 | -3.2516 | 1409.6147 | -3.2692 | 1415.1332 | -3.2692 | 1408.1249
| -3.0758 | 107.5396 | -3.1813 | 110.7497 | -3.111 | 108.8229 | -3.111 | 107.7436 | -3.1637 | 110.1626 | -3.0934 | 1414.096 | -3.1813 | 1413.6355 | -3.1813 | 1406.6367 | -3.2164 | 1411.336 | -3.2164 | 1413.4053 | -3.234 | 1410.1882 | -3.234 | 1406.6367
| -3.0407 | 106.9269 | -3.1461 | 110.689 | -3.0758 | 109.2497 | -3.0758 | 107.3151 | -3.1285 | 110.0612 | -3.0582 | 1414.7873 | -3.1461 | 1412.9451 | -3.1461 | 1404.3495 | -3.1813 | 1411.1064 | -3.1813 | 1416.1717 | -3.1989 | 1410.1882 | -3.1989 | 1401.6077
| -3.0055 | 106.661 | -3.111 | 110.4461 | -3.0407 | 109.3918 | -3.0407 | 107.1313 | -3.0934 | 110.0612 | -3.0231 | 1414.7873 | -3.111 | 1413.6355 | -3.111 | 1404.3495 | -3.1461 | 1411.1064 | -3.1461 | 1416.1717 | -3.1637 | 1411.7955 | -3.1637 | 1401.6077
| -2.9704 | 106.2721 | -3.0758 | 110.77 | -3.0055 | 109.5339 | -3.0055 | 107.07 | -3.0582 | 109.6963 | -2.9879 | 1420.6888 | -3.0758 | 1414.3264 | -3.0758 | 1405.8358 | -3.111 | 1410.3029 | -3.111 | 1411.5657 | -3.1285 | 1415.1332 | -3.1285 | 1401.9503
| -2.9352 | 105.7801 | -3.0407 | 110.9116 | -2.9704 | 109.6963 | -2.9704 | 107.07 | -3.0231 | 108.9652 | -2.9528 | 1420.805 | -3.0407 | 1414.3264 | -3.0407 | 1406.8655 | -3.0758 | 1410.3029 | -3.0758 | 1411.1064 | -3.0934 | 1413.7506 | -3.0934 | 1406.98
| -2.9001 | 105.4928 | -3.0055 | 111.2148 | -2.9352 | 109.6963 | -2.9352 | 107.3151 | -2.9879 | 108.8229 | -2.9176 | 1420.805 | -3.0055 | 1413.1751 | -3.0055 | 1405.8358 | -3.0407 | 1410.762 | -3.0407 | 1409.5001 | -3.0582 | 1413.7506 | -3.0582 | 1410.1882
| -2.8649 | 105.2875 | -2.9704 | 110.8509 | -2.9001 | 109.2497 | -2.9001 | 107.662 | -2.9528 | 108.9043 | -2.8825 | 1420.4564 | -2.9704 | 1409.9588 | -2.9704 | 1405.4927 | -3.0055 | 1406.8655 | -3.0055 | 1408.2395 | -3.0231 | 1413.7506 | -3.0231 | 1410.3029
| -2.8298 | 105.0203 | -2.9352 | 111.2754 | -2.8649 | 108.8433 | -2.8649 | 108.3346 | -2.9176 | 108.9043 | -2.8473 | 1419.6437 | -2.9352 | 1406.179 | -2.9352 | 1404.921 | -2.9704 | 1406.179 | -2.9704 | 1409.2708 | -2.9879 | 1411.9104 | -2.9879 | 1409.1562
| -2.7946 | 104.9175 | -2.9001 | 110.8104 | -2.8298 | 108.5381 | -2.8298 | 108.4567 | -2.8825 | 109.3918 | -2.8122 | 1418.6003 | -2.9001 | 1405.0354 | -2.9001 | 1404.0066 | -2.9352 | 1406.179 | -2.9352 | 1409.2708 | -2.9528 | 1411.9104 | -2.9528 | 1403.0925
| -2.7594 | 104.8969 | -2.8649 | 110.3651 | -2.7946 | 107.866 | -2.7946 | 108.4567 | -2.8473 | 109.7571 | -2.777 | 1418.2529 | -2.8649 | 1406.179 | -2.8649 | 1404.0066 | -2.9001 | 1405.8358 | -2.9001 | 1406.8655 | -2.9176 | 1410.1882 | -2.9176 | 1398.4107
| -2.7243 | 104.6705 | -2.8298 | 110.0612 | -2.7594 | 107.6212 | -2.7594 | 108.3346 | -2.8122 | 109.676 | -2.7419 | 1416.0562 | -2.8298 | 1407.0944 | -2.8298 | 1404.921 | -2.8649 | 1405.7215 | -2.8649 | 1406.8655 | -2.8825 | 1410.5324 | -2.8825 | 1395.5561
| -2.6891 | 103.7223 | -2.7946 | 109.9599 | -2.7243 | 107.2742 | -2.7243 | 106.5996 | -2.777 | 109.676 | -2.7067 | 1413.6355 | -2.7946 | 1408.9269 | -2.7946 | 1405.264 | -2.8298 | 1405.7215 | -2.8298 | 1406.8655 | -2.8473 | 1410.1882 | -2.8473 | 1394.8708
| -2.654 | 102.8746 | -2.7594 | 109.6557 | -2.6891 | 106.8247 | -2.6891 | 106.4359 | -2.7419 | 109.6557 | -2.6716 | 1413.5204 | -2.7594 | 1410.1882 | -2.7594 | 1403.6638 | -2.7946 | 1405.7215 | -2.7946 | 1406.8655 | -2.8122 | 1408.5832 | -2.8122 | 1400.9225
| -2.6188 | 102.7296 | -2.7243 | 109.2497 | -2.654 | 107.2742 | -2.654 | 106.1697 | -2.7067 | 109.3512 | -2.6364 | 1413.6355 | -2.7243 | 1408.9269 | -2.7243 | 1403.0925 | -2.7594 | 1405.7215 | -2.7594 | 1407.0944 | -2.777 | 1407.4379 | -2.777 | 1404.921
| -2.5837 | 102.7296 | -2.6891 | 108.3957 | -2.6188 | 107.2947 | -2.6188 | 106.0672 | -2.6716 | 108.9856 | -2.6013 | 1413.5204 | -2.6891 | 1407.0944 | -2.6891 | 1403.0925 | -2.7243 | 1405.264 | -2.7243 | 1406.8655 | -2.7419 | 1406.5222 | -2.7419 | 1406.6367
| -2.5485 | 102.3773 | -2.654 | 108.0494 | -2.5837 | 107.4376 | -2.5837 | 105.5955 | -2.6364 | 108.4974 | -2.5661 | 1412.4851 | -2.654 | 1405.8358 | -2.654 | 1401.9503 | -2.6891 | 1403.8923 | -2.6891 | 1406.6367 | -2.7067 | 1406.2934 | -2.7067 | 1404.1209
| -2.5134 | 102.481 | -2.6188 | 107.8048 | -2.5485 | 106.4359 | -2.5485 | 105.2875 | -2.6013 | 107.866 | -2.531 | 1414.2112 | -2.6188 | 1405.8358 | -2.6188 | 1404.0066 | -2.654 | 1403.8923 | -2.654 | 1406.6367 | -2.6716 | 1404.3495 | -2.6716 | 1403.0925
| -2.4782 | 102.8746 | -2.5837 | 107.8864 | -2.5134 | 106.7428 | -2.5134 | 105.1025 | -2.5661 | 107.5192 | -2.4958 | 1415.71 | -2.5837 | 1407.0944 | -2.5837 | 1406.0646 | -2.6188 | 1403.8923 | -2.6188 | 1405.1497 | -2.6364 | 1404.3495 | -2.6364 | 1401.8361
| -2.4431 | 103.0195 | -2.5485 | 108.0494 | -2.4782 | 105.5544 | -2.4782 | 104.8969 | -2.531 | 107.2334 | -2.4607 | 1417.4428 | -2.5485 | 1407.5523 | -2.5485 | 1406.6367 | -2.5837 | 1403.8923 | -2.5837 | 1406.2934 | -2.6013 | 1404.3495 | -2.6013 | 1399.5524
| -2.4079 | 102.7711 | -2.5134 | 107.9068 | -2.4431 | 105.3696 | -2.4431 | 104.794 | -2.4958 | 106.8042 | -2.4255 | 1417.6741 | -2.5134 | 1409.7294 | -2.5134 | 1408.4686 | -2.5485 | 1404.3495 | -2.5485 | 1407.8959 | -2.5661 | 1404.1209 | -2.5661 | 1401.0367
| -2.3728 | 102.3358 | -2.4782 | 107.6416 | -2.4079 | 105.2669 | -2.4079 | 104.5676 | -2.4607 | 106.3335 | -2.3903 | 1418.3687 | -2.4782 | 1410.3029 | -2.4782 | 1410.0735 | -2.5134 | 1407.4379 | -2.5134 | 1410.4177 | -2.531 | 1405.3784 | -2.531 | 1401.8361
| -2.3376 | 102.3358 | -2.4431 | 106.5178 | -2.3728 | 105.0614 | -2.3728 | 104.238 | -2.4255 | 106.6201 | -2.3552 | 1419.6437 | -2.4431 | 1410.3029 | -2.4431 | 1410.762 | -2.4782 | 1407.4379 | -2.4782 | 1411.4509 | -2.4958 | 1408.4686 | -2.4958 | 1401.3793
| -2.3025 | 102.3358 | -2.4079 | 106.1901 | -2.3376 | 104.4646 | -2.3376 | 104.1349 | -2.3903 | 106.3335 | -2.32 | 1419.6437 | -2.4079 | 1409.0415 | -2.4079 | 1410.762 | -2.4431 | 1405.6071 | -2.4431 | 1412.1402 | -2.4607 | 1409.0415 | -2.4607 | 1401.3793
| -2.2673 | 102.8332 | -2.3728 | 105.9852 | -2.3025 | 103.5571 | -2.3025 | 103.5571 | -2.3552 | 106.8451 | -2.2849 | 1419.1798 | -2.3728 | 1411.2212 | -2.3728 | 1410.762 | -2.4079 | 1408.8123 | -2.4079 | 1411.4509 | -2.4255 | 1408.4686 | -2.4255 | 1402.4071
| -2.2322 | 103.2677 | -2.3376 | 105.7391 | -2.2673 | 103.5571 | -2.2673 | 103.0816 | -2.32 | 106.5996 | -2.2497 | 1419.6437 | -2.3376 | 1412.2552 | -2.3376 | 1410.6472 | -2.3728 | 1409.5001 | -2.3728 | 1410.9916 | -2.3903 | 1409.6147 | -2.3903 | 1404.6924
| -2.197 | 103.3504 | -2.3025 | 105.1847 | -2.2322 | 103.619 | -2.2322 | 102.7089 | -2.2849 | 106.4973 | -2.2146 | 1420.5726 | -2.3025 | 1412.715 | -2.3025 | 1410.4177 | -2.3376 | 1407.5523 | -2.3376 | 1408.6977 | -2.3552 | 1408.9269 | -2.3552 | 1404.6924
| -2.1619 | 103.7429 | -2.2673 | 105.1847 | -2.197 | 103.7842 | -2.197 | 102.2736 | -2.2497 | 106.2926 | -2.1794 | 1421.3865 | -2.2673 | 1411.9104 | -2.2673 | 1410.4177 | -2.3025 | 1409.5001 | -2.3025 | 1408.6977 | -2.32 | 1408.6977 | -2.32 | 1404.3495
| -2.1267 | 103.9287 | -2.2322 | 105.1847 | -2.1619 | 103.7223 | -2.1619 | 101.6089 | -2.2146 | 105.8827 | -2.1443 | 1419.6437 | -2.2322 | 1411.6806 | -2.2322 | 1410.0735 | -2.2673 | 1408.9269 | -2.2673 | 1408.6977 | -2.2849 | 1408.6977 | -2.2849 | 1402.7498
| -2.0916 | 103.9287 | -2.197 | 105.6365 | -2.1267 | 103.5157 | -2.1267 | 101.4634 | -2.1794 | 105.0408 | -2.1091 | 1419.6437 | -2.197 | 1410.8768 | -2.197 | 1409.3854 | -2.2322 | 1406.6367 | -2.2322 | 1409.1562 | -2.2497 | 1409.6147 | -2.2497 | 1400.0091
| -2.0564 | 104.0937 | -2.1619 | 105.6365 | -2.0916 | 103.5157 | -2.0916 | 101.7129 | -2.1443 | 103.8255 | -2.074 | 1418.6003 | -2.1619 | 1410.762 | -2.1619 | 1409.3854 | -2.197 | 1406.0646 | -2.197 | 1409.5001 | -2.2146 | 1409.3854 | -2.2146 | 1400.0091
| -2.0213 | 104.1968 | -2.1267 | 105.6365 | -2.0564 | 103.4537 | -2.0564 | 101.7752 | -2.1091 | 103.6603 | -2.0388 | 1416.0562 | -2.1267 | 1409.7294 | -2.1267 | 1409.0415 | -2.1619 | 1406.0646 | -2.1619 | 1413.4053 | -2.1794 | 1409.6147 | -2.1794 | 1402.1787
| -1.9861 | 104.1555 | -2.0916 | 105.082 | -2.0213 | 103.5157 | -2.0213 | 101.6713 | -2.074 | 103.2057 | -2.0037 | 1414.9026 | -2.0916 | 1409.5001 | -2.0916 | 1406.5222 | -2.1267 | 1406.0646 | -2.1267 | 1409.3854 | -2.1443 | 1410.4177 | -2.1443 | 1402.864
| -1.9509 | 104.0937 | -2.0564 | 105.3285 | -1.9861 | 104.4234 | -1.9861 | 101.6921 | -2.0388 | 103.2057 | -1.9685 | 1414.3264 | -2.0564 | 1408.0104 | -2.0564 | 1406.2934 | -2.0916 | 1406.0646 | -2.0916 | 1409.0415 | -2.1091 | 1410.4177 | -2.1091 | 1403.2068
| -1.9158 | 103.8874 | -2.0213 | 105.8007 | -1.9509 | 104.7734 | -1.9509 | 101.6505 | -2.0037 | 103.619 | -1.9334 | 1412.715 | -2.0213 | 1404.2352 | -2.0213 | 1402.864 | -2.0564 | 1406.0646 | -2.0564 | 1404.4638 | -2.074 | 1410.3029 | -2.074 | 1404.3495
| -1.8806 | 103.8874 | -1.9861 | 106.3335 | -1.9158 | 104.7529 | -1.9158 | 101.5674 | -1.9685 | 103.5364 | -1.8982 | 1411.6806 | -1.9861 | 1403.5495 | -1.9861 | 1402.6356 | -2.0213 | 1401.4935 | -2.0213 | 1402.5214 | -2.0388 | 1409.6147 | -2.0388 | 1403.5495
| -1.8455 | 104.1555 | -1.9509 | 107.0904 | -1.8806 | 105.1847 | -1.8806 | 101.6921 | -1.9334 | 103.1023 | -1.8631 | 1411.6806 | -1.9509 | 1403.5495 | -1.9509 | 1401.4935 | -1.9861 | 1401.4935 | -1.9861 | 1402.5214 | -2.0037 | 1408.2395 | -2.0037 | 1400.4658
| -1.8103 | 104.3204 | -1.9158 | 107.0904 | -1.8455 | 105.2875 | -1.8455 | 102.4602 | -1.8982 | 103.1023 | -1.8279 | 1409.6147 | -1.9158 | 1404.2352 | -1.9158 | 1402.864 | -1.9509 | 1401.0367 | -1.9509 | 1402.1787 | -1.9685 | 1404.2352 | -1.9685 | 1398.8674
| -1.7752 | 105.1847 | -1.8806 | 107.0904 | -1.8103 | 105.308 | -1.8103 | 103.3297 | -1.8631 | 103.1023 | -1.7928 | 1409.6147 | -1.8806 | 1405.7215 | -1.8806 | 1405.6071 | -1.9158 | 1401.9503 | -1.9158 | 1402.7498 | -1.9334 | 1403.8923 | -1.9334 | 1400.0091
| -1.74 | 105.4723 | -1.8455 | 107.2742 | -1.7752 | 105.3491 | -1.7752 | 103.8255 | -1.8279 | 103.1023 | -1.7576 | 1411.9104 | -1.8455 | 1405.7215 | -1.8455 | 1404.2352 | -1.8806 | 1403.8923 | -1.8806 | 1404.5781 | -1.8982 | 1402.5214 | -1.8982 | 1400.0091
| -1.7049 | 105.6776 | -1.8103 | 107.1721 | -1.74 | 105.9237 | -1.74 | 104.0318 | -1.7928 | 103.1023 | -1.7225 | 1413.9809 | -1.8103 | 1405.7215 | -1.8103 | 1401.6077 | -1.8455 | 1404.8067 | -1.8455 | 1405.4927 | -1.8631 | 1402.7498 | -1.8631 | 1400.0091
| -1.6697 | 104.4646 | -1.7752 | 107.4376 | -1.7049 | 105.9647 | -1.7049 | 104.1555 | -1.7576 | 103.3297 | -1.6873 | 1411.9104 | -1.7752 | 1405.264 | -1.7752 | 1400.8083 | -1.8103 | 1403.8923 | -1.8103 | 1404.5781 | -1.8279 | 1402.7498 | -1.8279 | 1401.0367
| -1.6346 | 103.4331 | -1.74 | 107.7232 | -1.6697 | 105.9237 | -1.6697 | 104.4234 | -1.7225 | 104.1968 | -1.6522 | 1406.98 | -1.74 | 1404.5781 | -1.74 | 1398.8674 | -1.7752 | 1400.9225 | -1.7752 | 1403.8923 | -1.7928 | 1402.6356 | -1.7928 | 1401.0367
| -1.5994 | 102.9367 | -1.7049 | 107.5192 | -1.6346 | 106.0057 | -1.6346 | 104.6294 | -1.6873 | 104.1968 | -1.617 | 1404.5781 | -1.7049 | 1403.2068 | -1.7049 | 1394.071 | -1.74 | 1399.0957 | -1.74 | 1401.0367 | -1.7576 | 1400.9225 | -1.7576 | 1399.3241
| -1.5643 | 102.9367 | -1.6697 | 107.4376 | -1.5994 | 106.0057 | -1.5994 | 104.547 | -1.6522 | 104.6088 | -1.5818 | 1403.0925 | -1.6697 | 1402.5214 | -1.6697 | 1393.8425 | -1.7049 | 1398.1824 | -1.7049 | 1398.4107 | -1.7225 | 1400.9225 | -1.7225 | 1399.3241
| -1.5291 | 102.8125 | -1.6346 | 107.4172 | -1.5643 | 105.9647 | -1.5643 | 104.1968 | -1.617 | 104.6088 | -1.5467 | 1402.5214 | -1.6346 | 1401.7219 | -1.6346 | 1392.0133 | -1.6697 | 1396.6981 | -1.6697 | 1395.8987 | -1.6873 | 1400.4658 | -1.6873 | 1394.2995
| -1.494 | 102.8746 | -1.5994 | 107.4172 | -1.5291 | 106.395 | -1.5291 | 103.7223 | -1.5818 | 104.7117 | -1.5115 | 1403.5495 | -1.5994 | 1398.7532 | -1.5994 | 1391.2125 | -1.6346 | 1396.1272 | -1.6346 | 1395.8987 | -1.6522 | 1400.3516 | -1.6522 | 1391.2125
| -1.4588 | 103.0816 | -1.5643 | 107.4376 | -1.494 | 106.395 | -1.494 | 103.681 | -1.5467 | 105.616 | -1.4764 | 1405.6071 | -1.5643 | 1398.5249 | -1.5643 | 1390.7548 | -1.5994 | 1396.1272 | -1.5994 | 1395.7845 | -1.617 | 1397.9541 | -1.617 | 1384.2143
| -1.4237 | 103.4951 | -1.5291 | 108.2124 | -1.4588 | 106.0877 | -1.4588 | 103.619 | -1.5115 | 105.8827 | -1.4412 | 1404.921 | -1.5291 | 1398.5249 | -1.5291 | 1391.2125 | -1.5643 | 1394.4138 | -1.5643 | 1395.0992 | -1.5818 | 1395.2135 | -1.5818 | 1386.9727
| -1.3885 | 103.1229 | -1.494 | 108.9043 | -1.4237 | 105.1642 | -1.4237 | 103.4537 | -1.4764 | 105.8007 | -1.4061 | 1402.5214 | -1.494 | 1396.3555 | -1.494 | 1392.3564 | -1.5291 | 1393.9567 | -1.5291 | 1395.0992 | -1.5467 | 1393.0424 | -1.5467 | 1383.293
| -1.3534 | 102.7711 | -1.4588 | 108.6602 | -1.3885 | 104.794 | -1.3885 | 103.6397 | -1.4412 | 105.7391 | -1.3709 | 1400.9225 | -1.4588 | 1394.6423 | -1.4588 | 1390.9837 | -1.494 | 1393.7282 | -1.494 | 1394.6423 | -1.5115 | 1392.242 | -1.5115 | 1383.293
| -1.3182 | 102.5846 | -1.4237 | 107.6824 | -1.3534 | 104.2792 | -1.3534 | 104.0524 | -1.4061 | 105.8827 | -1.3358 | 1395.7845 | -1.4237 | 1392.3564 | -1.4237 | 1389.7244 | -1.4588 | 1392.6994 | -1.4588 | 1391.4414 | -1.4764 | 1392.3564 | -1.4764 | 1386.9727
| -1.2831 | 102.5846 | -1.3885 | 106.7838 | -1.3182 | 103.9287 | -1.3182 | 104.0524 | -1.3709 | 106.354 | -1.3006 | 1395.5561 | -1.3885 | 1392.1277 | -1.3885 | 1387.5465 | -1.4237 | 1391.4414 | -1.4237 | 1389.9534 | -1.4412 | 1392.242 | -1.4412 | 1387.7759
| -1.2479 | 102.5846 | -1.3534 | 106.7224 | -1.2831 | 103.8874 | -1.2831 | 103.8461 | -1.3358 | 106.8247 | -1.2655 | 1395.5561 | -1.3534 | 1389.8389 | -1.3534 | 1387.0875 | -1.3885 | 1389.9534 | -1.3885 | 1389.6098 | -1.4061 | 1392.3564 | -1.4061 | 1385.1346
| -1.2128 | 102.8125 | -1.3182 | 106.5792 | -1.2479 | 103.908 | -1.2479 | 103.7223 | -1.3006 | 107.07 | -1.2303 | 1397.269 | -1.3182 | 1388.3493 | -1.3182 | 1387.0875 | -1.3534 | 1388.0053 | -1.3534 | 1388.12 | -1.3709 | 1391.5558 | -1.3709 | 1384.3294
| -1.1776 | 103.2677 | -1.2831 | 106.1697 | -1.2128 | 104.4234 | -1.2128 | 103.3711 | -1.2655 | 107.07 | -1.1952 | 1397.269 | -1.2831 | 1387.0875 | -1.2831 | 1387.0875 | -1.3182 | 1387.0875 | -1.3182 | 1387.4318 | -1.3358 | 1389.7244 | -1.3358 | 1385.1346
| -1.1424 | 103.4537 | -1.2479 | 106.354 | -1.1776 | 105.2053 | -1.1776 | 103.185 | -1.2303 | 106.8247 | -1.16 | 1397.269 | -1.2479 | 1388.6932 | -1.2479 | 1387.0875 | -1.2831 | 1387.0875 | -1.2831 | 1387.4318 | -1.3006 | 1387.8906 | -1.3006 | 1384.3294
| -1.1073 | 103.7842 | -1.2128 | 106.886 | -1.1424 | 104.7117 | -1.1424 | 103.0195 | -1.1952 | 106.1082 | -1.1249 | 1396.4697 | -1.2128 | 1388.6932 | -1.2128 | 1386.3987 | -1.2479 | 1386.1689 | -1.2479 | 1387.5465 | -1.2655 | 1387.6612 | -1.2655 | 1385.1346
| -1.0721 | 103.8048 | -1.1776 | 106.9064 | -1.1073 | 104.5676 | -1.1073 | 102.6882 | -1.16 | 104.8969 | -1.0897 | 1395.8987 | -1.1776 | 1388.6932 | -1.1776 | 1385.3645 | -1.2128 | 1387.0875 | -1.2128 | 1388.0053 | -1.2303 | 1387.6612 | -1.2303 | 1386.0541
| -1.037 | 104.0318 | -1.1424 | 106.9064 | -1.0721 | 104.0318 | -1.0721 | 102.7089 | -1.1249 | 104.4852 | -1.0546 | 1394.8708 | -1.1424 | 1386.6283 | -1.1424 | 1384.3294 | -1.1776 | 1388.3493 | -1.1776 | 1388.0053 | -1.1952 | 1387.4318 | -1.1952 | 1388.5786
| -1.0018 | 104.341 | -1.1073 | 106.886 | -1.037 | 102.916 | -1.037 | 103.0195 | -1.0897 | 104.3616 | -1.0194 | 1394.2995 | -1.1073 | 1387.317 | -1.1073 | 1385.0196 | -1.1424 | 1386.8579 | -1.1424 | 1386.3987 | -1.16 | 1387.4318 | -1.16 | 1388.5786
| -0.96669 | 104.4852 | -1.0721 | 105.7596 | -1.0018 | 102.916 | -1.0018 | 103.0609 | -1.0546 | 104.3616 | -0.98426 | 1394.2995 | -1.0721 | 1384.2143 | -1.0721 | 1385.0196 | -1.1073 | 1385.2496 | -1.1073 | 1385.4795 | -1.1249 | 1387.4318 | -1.1249 | 1382.0246
| -0.93153 | 104.4852 | -1.037 | 105.7596 | -0.96669 | 102.916 | -0.96669 | 102.6675 | -1.0194 | 104.4852 | -0.94911 | 1395.2135 | -1.037 | 1384.2143 | -1.037 | 1385.9392 | -1.0721 | 1384.9046 | -1.0721 | 1384.0992 | -1.0897 | 1385.8243 | -1.0897 | 1379.5974
| -0.89638 | 105.3696 | -1.0018 | 105.7596 | -0.93153 | 103.5984 | -0.93153 | 102.626 | -0.98426 | 104.8146 | -0.91396 | 1395.5561 | -1.0018 | 1387.317 | -1.0018 | 1388.3493 | -1.037 | 1385.2496 | -1.037 | 1382.14 | -1.0546 | 1385.0196 | -1.0546 | 1374.6
| -0.86123 | 104.9997 | -0.96669 | 106.2721 | -0.89638 | 103.9287 | -0.89638 | 102.4602 | -0.94911 | 104.9997 | -0.87881 | 1396.3555 | -0.96669 | 1387.317 | -0.96669 | 1388.9224 | -1.0018 | 1385.4795 | -1.0018 | 1384.2143 | -1.0194 | 1384.7896 | -1.0194 | 1373.8993
| -0.82608 | 105.1436 | -0.93153 | 106.5792 | -0.86123 | 104.3822 | -0.86123 | 102.2113 | -0.91396 | 104.938 | -0.84365 | 1396.6981 | -0.93153 | 1386.2838 | -0.93153 | 1388.2346 | -0.96669 | 1386.7431 | -0.96669 | 1384.7896 | -0.98426 | 1384.7896 | -0.98426 | 1374.2498
| -0.79092 | 105.1642 | -0.89638 | 106.6406 | -0.82608 | 105.6981 | -0.82608 | 102.4602 | -0.87881 | 104.2998 | -0.8085 | 1396.8123 | -0.89638 | 1388.6932 | -0.89638 | 1388.2346 | -0.93153 | 1386.7431 | -0.93153 | 1384.6745 | -0.94911 | 1383.7537 | -0.94911 | 1374.0162
| -0.75577 | 105.1436 | -0.86123 | 106.8656 | -0.79092 | 106.2106 | -0.79092 | 102.6468 | -0.84365 | 104.1762 | -0.77335 | 1396.9265 | -0.86123 | 1388.2346 | -0.86123 | 1386.5135 | -0.89638 | 1386.7431 | -0.89638 | 1384.6745 | -0.91396 | 1382.2554 | -0.91396 | 1373.8993
| -0.72062 | 105.1231 | -0.82608 | 106.9678 | -0.75577 | 106.8451 | -0.75577 | 102.916 | -0.8085 | 103.5777 | -0.7382 | 1399.0957 | -0.82608 | 1388.2346 | -0.82608 | 1386.3987 | -0.86123 | 1385.9392 | -0.86123 | 1386.6283 | -0.87881 | 1383.8689 | -0.87881 | 1374.0162
| -0.68547 | 105.2258 | -0.79092 | 107.3763 | -0.72062 | 106.7428 | -0.72062 | 103.4124 | -0.77335 | 103.5984 | -0.70304 | 1400.0091 | -0.79092 | 1391.5558 | -0.79092 | 1387.5465 | -0.82608 | 1385.8243 | -0.82608 | 1387.4318 | -0.84365 | 1385.4795 | -0.84365 | 1379.1342
| -0.65032 | 105.3696 | -0.75577 | 107.5804 | -0.68547 | 106.7633 | -0.68547 | 103.4124 | -0.7382 | 104.5264 | -0.66789 | 1401.0367 | -0.75577 | 1392.4707 | -0.75577 | 1388.0053 | -0.79092 | 1385.9392 | -0.79092 | 1389.4953 | -0.8085 | 1385.5944 | -0.8085 | 1379.3658
| -0.61516 | 105.4107 | -0.72062 | 107.7844 | -0.65032 | 106.8451 | -0.65032 | 103.908 | -0.70304 | 104.5676 | -0.63274 | 1401.4935 | -0.72062 | 1392.4707 | -0.72062 | 1388.4639 | -0.75577 | 1386.3987 | -0.75577 | 1390.2969 | -0.77335 | 1385.5944 | -0.77335 | 1379.3658
| -0.58001 | 105.4312 | -0.68547 | 108.192 | -0.61516 | 106.0467 | -0.61516 | 104.8969 | -0.66789 | 104.8969 | -0.59759 | 1402.7498 | -0.68547 | 1393.1567 | -0.68547 | 1390.7548 | -0.72062 | 1386.3987 | -0.72062 | 1390.2969 | -0.7382 | 1387.2023 | -0.7382 | 1379.3658
| -0.54486 | 105.4312 | -0.65032 | 108.2531 | -0.58001 | 105.9852 | -0.58001 | 104.8969 | -0.63274 | 105.082 | -0.56244 | 1402.7498 | -0.65032 | 1393.1567 | -0.65032 | 1391.2125 | -0.68547 | 1387.0875 | -0.68547 | 1391.4414 | -0.70304 | 1387.5465 | -0.70304 | 1377.1618
| -0.50971 | 105.4312 | -0.61516 | 108.2939 | -0.54486 | 106.2311 | -0.54486 | 104.6705 | -0.59759 | 105.5339 | -0.52728 | 1403.7781 | -0.61516 | 1393.1567 | -0.61516 | 1392.0133 | -0.65032 | 1389.8389 | -0.65032 | 1392.242 | -0.66789 | 1389.2662 | -0.66789 | 1376.5805
| -0.47455 | 105.4312 | -0.58001 | 108.4364 | -0.50971 | 105.8827 | -0.50971 | 105.5544 | -0.56244 | 106.2106 | -0.49213 | 1405.7215 | -0.58001 | 1393.1567 | -0.58001 | 1392.1277 | -0.61516 | 1391.0981 | -0.61516 | 1394.528 | -0.63274 | 1389.2662 | -0.63274 | 1377.1618
| -0.4394 | 105.5544 | -0.54486 | 108.4364 | -0.47455 | 105.8622 | -0.47455 | 104.938 | -0.52728 | 106.9269 | -0.45698 | 1406.5222 | -0.54486 | 1391.8989 | -0.54486 | 1393.0424 | -0.58001 | 1391.6702 | -0.58001 | 1394.8708 | -0.59759 | 1390.9837 | -0.59759 | 1379.7132
| -0.40425 | 105.7391 | -0.50971 | 107.866 | -0.4394 | 106.2926 | -0.4394 | 104.6705 | -0.49213 | 106.8656 | -0.42183 | 1406.8655 | -0.50971 | 1393.1567 | -0.50971 | 1393.271 | -0.54486 | 1392.4707 | -0.54486 | 1394.8708 | -0.56244 | 1391.4414 | -0.56244 | 1381.5629
| -0.3691 | 107.5804 | -0.47455 | 107.9271 | -0.40425 | 106.395 | -0.40425 | 104.8969 | -0.45698 | 107.0086 | -0.38667 | 1406.5222 | -0.47455 | 1394.1853 | -0.47455 | 1395.2135 | -0.50971 | 1392.6994 | -0.50971 | 1394.8708 | -0.52728 | 1392.242 | -0.52728 | 1384.6745
| -0.33395 | 107.7436 | -0.4394 | 107.8252 | -0.3691 | 106.354 | -0.3691 | 104.7117 | -0.42183 | 106.8247 | -0.35152 | 1406.5222 | -0.4394 | 1394.985 | -0.4394 | 1396.013 | -0.47455 | 1393.3853 | -0.47455 | 1394.8708 | -0.49213 | 1392.3564 | -0.49213 | 1384.7896
| -0.29879 | 108.4974 | -0.40425 | 107.9271 | -0.33395 | 106.9269 | -0.33395 | 104.3616 | -0.38667 | 106.313 | -0.31637 | 1406.8655 | -0.40425 | 1395.3277 | -0.40425 | 1396.8123 | -0.4394 | 1394.4138 | -0.4394 | 1396.013 | -0.45698 | 1393.7282 | -0.45698 | 1384.7896
| -0.26364 | 108.4974 | -0.3691 | 108.192 | -0.29879 | 106.8656 | -0.29879 | 104.4646 | -0.35152 | 105.9442 | -0.28122 | 1407.5523 | -0.3691 | 1396.6981 | -0.3691 | 1396.9265 | -0.40425 | 1393.3853 | -0.40425 | 1394.6423 | -0.42183 | 1394.4138 | -0.42183 | 1385.0196
| -0.22849 | 107.5192 | -0.33395 | 108.6195 | -0.26364 | 106.8656 | -0.26364 | 104.8969 | -0.31637 | 106.2106 | -0.24607 | 1407.7814 | -0.33395 | 1397.0407 | -0.33395 | 1396.9265 | -0.3691 | 1394.8708 | -0.3691 | 1397.0407 | -0.38667 | 1395.7845 | -0.38667 | 1385.1346
| -0.19334 | 107.3763 | -0.29879 | 108.4974 | -0.22849 | 107.1108 | -0.22849 | 104.8352 | -0.28122 | 106.4359 | -0.21091 | 1408.2395 | -0.29879 | 1395.8987 | -0.29879 | 1396.2413 | -0.33395 | 1395.3277 | -0.33395 | 1397.269 | -0.35152 | 1395.7845 | -0.35152 | 1385.9392
| -0.15818 | 107.5192 | -0.26364 | 108.6602 | -0.19334 | 107.2947 | -0.19334 | 105.2875 | -0.24607 | 106.3335 | -0.17576 | 1409.7294 | -0.26364 | 1395.8987 | -0.26364 | 1393.1567 | -0.29879 | 1394.985 | -0.29879 | 1395.8987 | -0.31637 | 1397.3832 | -0.31637 | 1388.3493
| -0.12303 | 107.3763 | -0.22849 | 108.3142 | -0.15818 | 107.3967 | -0.15818 | 105.4107 | -0.21091 | 106.4564 | -0.14061 | 1410.4177 | -0.22849 | 1397.1548 | -0.22849 | 1392.0133 | -0.26364 | 1394.985 | -0.26364 | 1393.8425 | -0.28122 | 1398.0682 | -0.28122 | 1390.5259
| -0.087881 | 107.7232 | -0.19334 | 108.6195 | -0.12303 | 108.1513 | -0.12303 | 105.4518 | -0.17576 | 106.8656 | -0.10546 | 1410.8768 | -0.19334 | 1397.0407 | -0.19334 | 1392.9281 | -0.22849 | 1396.1272 | -0.22849 | 1394.1853 | -0.24607 | 1396.5839 | -0.24607 | 1388.3493
| -0.052728 | 108.355 | -0.15818 | 109.2294 | -0.087881 | 108.2124 | -0.087881 | 105.657 | -0.14061 | 106.8656 | -0.070304 | 1410.8768 | -0.15818 | 1393.3853 | -0.15818 | 1393.1567 | -0.19334 | 1394.985 | -0.19334 | 1396.013 | -0.21091 | 1394.4138 | -0.21091 | 1388.3493
| -0.017576 | 108.192 | -0.12303 | 109.7165 | -0.052728 | 109.5136 | -0.052728 | 105.8417 | -0.10546 | 107.0086 | -0.035152 | 1410.8768 | -0.12303 | 1394.985 | -0.12303 | 1398.5249 | -0.15818 | 1392.242 | -0.15818 | 1396.013 | -0.17576 | 1394.071 | -0.17576 | 1383.984
| 0.017576 | 108.4771 | -0.087881 | 110.3449 | -0.017576 | 109.9802 | -0.017576 | 106.2516 | -0.070304 | 107.1313 | 0 | 1411.336 | -0.087881 | 1395.6703 | -0.087881 | 1399.4382 | -0.12303 | 1390.9837 | -0.12303 | 1396.013 | -0.14061 | 1392.6994 | -0.14061 | 1383.984
| 0.052728 | 108.9043 | -0.052728 | 110.9116 | 0.017576 | 109.7165 | 0.017576 | 106.9269 | -0.035152 | 107.7028 | 0.035152 | 1410.9916 | -0.052728 | 1396.013 | -0.052728 | 1399.8949 | -0.087881 | 1390.6403 | -0.087881 | 1396.4697 | -0.10546 | 1390.6403 | -0.10546 | 1381.6783
| 0.087881 | 108.5381 | -0.017576 | 111.0329 | 0.052728 | 108.7212 | 0.052728 | 107.6824 | 0 | 107.2947 | 0.070304 | 1410.5324 | -0.017576 | 1396.2413 | -0.017576 | 1399.4382 | -0.052728 | 1390.9837 | -0.052728 | 1396.8123 | -0.070304 | 1390.1824 | -0.070304 | 1381.6783
| 0.12303 | 108.0902 | 0.017576 | 110.9318 | 0.087881 | 108.4771 | 0.087881 | 108.3142 | 0.035152 | 107.3355 | 0.10546 | 1409.7294 | 0.017576 | 1400.2374 | 0.017576 | 1399.8949 | -0.017576 | 1392.9281 | -0.017576 | 1397.269 | -0.035152 | 1390.6403 | -0.035152 | 1387.8906
| 0.15818 | 107.5192 | 0.052728 | 110.9116 | 0.12303 | 108.5788 | 0.12303 | 108.3957 | 0.070304 | 107.7232 | 0.14061 | 1410.4177 | 0.052728 | 1401.4935 | 0.052728 | 1400.6942 | 0.017576 | 1397.1548 | 0.017576 | 1398.2966 | 0 | 1394.985 | 0 | 1388.12
| 0.19334 | 106.9473 | 0.087881 | 111.0329 | 0.15818 | 108.4771 | 0.15818 | 108.3142 | 0.10546 | 108.1105 | 0.17576 | 1410.4177 | 0.087881 | 1403.5495 | 0.087881 | 1400.6942 | 0.052728 | 1399.3241 | 0.052728 | 1398.7532 | 0.035152 | 1396.3555 | 0.035152 | 1388.12
| 0.22849 | 106.7224 | 0.12303 | 111.1137 | 0.19334 | 108.7212 | 0.19334 | 108.355 | 0.14061 | 108.3346 | 0.21091 | 1410.4177 | 0.12303 | 1404.6924 | 0.12303 | 1401.4935 | 0.087881 | 1401.3793 | 0.087881 | 1399.6666 | 0.070304 | 1397.7257 | 0.070304 | 1393.1567
| 0.26364 | 106.5382 | 0.15818 | 111.3966 | 0.22849 | 108.9043 | 0.22849 | 108.3957 | 0.17576 | 109.2497 | 0.24607 | 1412.9451 | 0.15818 | 1406.0646 | 0.15818 | 1402.4071 | 0.12303 | 1401.4935 | 0.12303 | 1401.8361 | 0.10546 | 1399.6666 | 0.10546 | 1393.271
| 0.29879 | 105.9237 | 0.19334 | 111.9215 | 0.26364 | 109.7774 | 0.26364 | 108.2531 | 0.21091 | 109.4121 | 0.28122 | 1412.9451 | 0.19334 | 1406.179 | 0.19334 | 1404.3495 | 0.15818 | 1401.3793 | 0.15818 | 1402.2929 | 0.14061 | 1400.6942 | 0.14061 | 1393.1567
| 0.33395 | 105.616 | 0.22849 | 112.0224 | 0.29879 | 109.4121 | 0.29879 | 108.2124 | 0.24607 | 109.2497 | 0.31637 | 1411.6806 | 0.22849 | 1404.6924 | 0.22849 | 1403.4353 | 0.19334 | 1401.3793 | 0.19334 | 1402.2929 | 0.17576 | 1401.4935 | 0.17576 | 1393.271
| 0.3691 | 105.5544 | 0.26364 | 111.9417 | 0.33395 | 109.4121 | 0.33395 | 108.7619 | 0.28122 | 109.0465 | 0.35152 | 1409.6147 | 0.26364 | 1401.8361 | 0.26364 | 1401.4935 | 0.22849 | 1402.864 | 0.22849 | 1402.1787 | 0.21091 | 1401.6077 | 0.21091 | 1395.6703
| 0.40425 | 105.616 | 0.29879 | 111.5582 | 0.3691 | 109.27 | 0.3691 | 108.7823 | 0.31637 | 109.1888 | 0.38667 | 1408.9269 | 0.29879 | 1396.9265 | 0.29879 | 1396.013 | 0.26364 | 1402.5214 | 0.26364 | 1400.58 | 0.24607 | 1401.4935 | 0.24607 | 1395.6703
| 0.4394 | 105.5749 | 0.33395 | 111.437 | 0.40425 | 109.27 | 0.40425 | 109.0059 | 0.35152 | 109.1481 | 0.42183 | 1407.4379 | 0.33395 | 1394.6423 | 0.33395 | 1395.2135 | 0.29879 | 1399.7808 | 0.29879 | 1398.5249 | 0.28122 | 1400.58 | 0.28122 | 1392.6994
| 0.47455 | 105.616 | 0.3691 | 110.7295 | 0.4394 | 108.9043 | 0.4394 | 108.7823 | 0.38667 | 108.9856 | 0.45698 | 1406.4078 | 0.3691 | 1393.3853 | 0.3691 | 1394.4138 | 0.33395 | 1396.013 | 0.33395 | 1397.0407 | 0.31637 | 1398.8674 | 0.31637 | 1392.1277
| 0.50971 | 106.0262 | 0.40425 | 110.4866 | 0.47455 | 108.3346 | 0.47455 | 108.4974 | 0.42183 | 109.473 | 0.49213 | 1403.4353 | 0.40425 | 1393.271 | 0.40425 | 1393.7282 | 0.3691 | 1393.9567 | 0.3691 | 1395.6703 | 0.35152 | 1396.013 | 0.35152 | 1390.9837
| 0.54486 | 106.2926 | 0.4394 | 110.4866 | 0.50971 | 107.866 | 0.50971 | 107.9679 | 0.45698 | 109.6557 | 0.52728 | 1402.1787 | 0.4394 | 1392.0133 | 0.4394 | 1393.1567 | 0.40425 | 1392.4707 | 0.40425 | 1394.528 | 0.38667 | 1393.1567 | 0.38667 | 1388.5786
| 0.58001 | 106.5996 | 0.47455 | 109.818 | 0.54486 | 107.4376 | 0.54486 | 107.7028 | 0.49213 | 109.676 | 0.56244 | 1401.4935 | 0.47455 | 1391.8989 | 0.47455 | 1392.8138 | 0.4394 | 1390.4114 | 0.4394 | 1393.271 | 0.42183 | 1393.1567 | 0.42183 | 1387.2023
| 0.61516 | 107.4784 | 0.50971 | 109.7571 | 0.58001 | 106.9882 | 0.58001 | 107.5396 | 0.52728 | 109.9802 | 0.59759 | 1401.4935 | 0.50971 | 1391.8989 | 0.50971 | 1389.9534 | 0.47455 | 1390.4114 | 0.47455 | 1391.7846 | 0.45698 | 1392.4707 | 0.45698 | 1384.6745
| 0.65032 | 107.56 | 0.54486 | 109.676 | 0.61516 | 106.5792 | 0.61516 | 106.5792 | 0.56244 | 109.7368 | 0.63274 | 1401.4935 | 0.54486 | 1390.9837 | 0.54486 | 1388.5786 | 0.50971 | 1389.8389 | 0.50971 | 1391.6702 | 0.49213 | 1392.1277 | 0.49213 | 1384.0992
| 0.68547 | 108.1105 | 0.58001 | 109.3512 | 0.65032 | 106.5792 | 0.65032 | 105.9852 | 0.59759 | 109.1278 | 0.66789 | 1401.4935 | 0.58001 | 1389.2662 | 0.58001 | 1387.0875 | 0.54486 | 1389.037 | 0.54486 | 1391.6702 | 0.52728 | 1391.7846 | 0.52728 | 1383.8689
| 0.72062 | 107.56 | 0.61516 | 108.6602 | 0.68547 | 106.5792 | 0.68547 | 105.8417 | 0.63274 | 108.6602 | 0.70304 | 1399.3241 | 0.61516 | 1389.2662 | 0.61516 | 1389.1516 | 0.58001 | 1388.3493 | 0.58001 | 1391.6702 | 0.56244 | 1389.037 | 0.56244 | 1383.8689
| 0.75577 | 107.213 | 0.65032 | 108.5178 | 0.72062 | 106.5792 | 0.72062 | 105.5749 | 0.66789 | 108.3753 | 0.7382 | 1400.8083 | 0.65032 | 1391.327 | 0.65032 | 1389.9534 | 0.61516 | 1387.5465 | 0.61516 | 1391.5558 | 0.59759 | 1388.0053 | 0.59759 | 1384.0992
| 0.79092 | 106.661 | 0.68547 | 108.4771 | 0.75577 | 107.3967 | 0.75577 | 105.0408 | 0.70304 | 108.3346 | 0.77335 | 1400.8083 | 0.68547 | 1390.5259 | 0.68547 | 1390.0679 | 0.65032 | 1388.3493 | 0.65032 | 1391.8989 | 0.63274 | 1385.4795 | 0.63274 | 1382.4861
| 0.82608 | 105.5955 | 0.72062 | 108.5381 | 0.79092 | 107.3967 | 0.79092 | 105.5749 | 0.7382 | 107.7232 | 0.8085 | 1400.2374 | 0.72062 | 1391.5558 | 0.72062 | 1390.0679 | 0.68547 | 1388.5786 | 0.68547 | 1389.1516 | 0.66789 | 1385.4795 | 0.66789 | 1382.4861
| 0.86123 | 105.5955 | 0.75577 | 108.6602 | 0.82608 | 107.3967 | 0.82608 | 105.3285 | 0.77335 | 107.4376 | 0.84365 | 1400.0091 | 0.75577 | 1392.5851 | 0.75577 | 1390.0679 | 0.72062 | 1388.5786 | 0.72062 | 1388.8078 | 0.70304 | 1386.5135 | 0.70304 | 1382.4861
| 0.89638 | 105.5955 | 0.79092 | 108.7619 | 0.86123 | 107.8456 | 0.86123 | 105.2875 | 0.8085 | 107.0086 | 0.87881 | 1401.9503 | 0.79092 | 1391.5558 | 0.79092 | 1390.0679 | 0.75577 | 1391.4414 | 0.75577 | 1388.3493 | 0.7382 | 1387.5465 | 0.7382 | 1383.4082
| 0.93153 | 106.661 | 0.82608 | 109.3309 | 0.89638 | 108.5585 | 0.89638 | 105.7801 | 0.84365 | 106.5996 | 0.91396 | 1400.8083 | 0.82608 | 1391.327 | 0.82608 | 1386.8579 | 0.79092 | 1392.1277 | 0.79092 | 1388.0053 | 0.77335 | 1388.4639 | 0.77335 | 1384.6745
| 0.96669 | 107.0904 | 0.86123 | 109.9193 | 0.93153 | 108.029 | 0.93153 | 106.0467 | 0.87881 | 106.5996 | 0.94911 | 1400.2374 | 0.86123 | 1391.2125 | 0.86123 | 1387.7759 | 0.82608 | 1390.8692 | 0.82608 | 1387.5465 | 0.8085 | 1389.8389 | 0.8085 | 1384.6745
| 1.0018 | 107.2334 | 0.89638 | 110.7902 | 0.96669 | 108.5788 | 0.96669 | 105.9647 | 0.91396 | 106.9269 | 0.98426 | 1397.7257 | 0.89638 | 1387.317 | 0.89638 | 1386.6283 | 0.86123 | 1390.4114 | 0.86123 | 1385.9392 | 0.84365 | 1389.8389 | 0.84365 | 1383.4082
| 1.037 | 107.3763 | 0.93153 | 110.7902 | 1.0018 | 108.029 | 1.0018 | 106.4973 | 0.94911 | 106.5178 | 1.0194 | 1396.9265 | 0.93153 | 1385.5944 | 0.93153 | 1383.8689 | 0.89638 | 1388.2346 | 0.89638 | 1385.9392 | 0.87881 | 1388.12 | 0.87881 | 1383.0625
| 1.0721 | 107.3151 | 0.96669 | 110.2639 | 1.037 | 107.5396 | 1.037 | 107.2538 | 0.98426 | 107.0086 | 1.0546 | 1395.0992 | 0.96669 | 1385.9392 | 0.96669 | 1383.8689 | 0.93153 | 1384.0992 | 0.93153 | 1384.4444 | 0.91396 | 1386.9727 | 0.91396 | 1380.2917
| 1.1073 | 107.3151 | 1.0018 | 109.676 | 1.0721 | 107.0904 | 1.0721 | 107.6008 | 1.0194 | 107.2742 | 1.0897 | 1396.9265 | 1.0018 | 1385.3645 | 1.0018 | 1384.3294 | 0.96669 | 1383.8689 | 0.96669 | 1383.8689 | 0.94911 | 1385.2496 | 0.94911 | 1375.7658
| 1.1424 | 107.3967 | 1.037 | 109.27 | 1.1073 | 107.0495 | 1.1073 | 107.0086 | 1.0546 | 107.9271 | 1.1249 | 1398.8674 | 1.037 | 1386.5135 | 1.037 | 1384.3294 | 1.0018 | 1382.2554 | 1.0018 | 1383.4082 | 0.98426 | 1381.2164 | 0.98426 | 1375.5328
| 1.1776 | 107.3151 | 1.0721 | 108.9652 | 1.1424 | 107.0904 | 1.1424 | 106.8451 | 1.0897 | 109.1278 | 1.16 | 1399.2099 | 1.0721 | 1388.0053 | 1.0721 | 1385.2496 | 1.037 | 1382.2554 | 1.037 | 1382.7167 | 1.0194 | 1381.2164 | 1.0194 | 1372.3782
| 1.2128 | 107.1517 | 1.1073 | 108.6399 | 1.1776 | 107.5396 | 1.1776 | 106.4564 | 1.1249 | 109.1278 | 1.1952 | 1399.5524 | 1.1073 | 1388.2346 | 1.1073 | 1385.7094 | 1.0721 | 1383.8689 | 1.0721 | 1382.6014 | 1.0546 | 1382.4861 | 1.0546 | 1373.7825
| 1.2479 | 107.56 | 1.1424 | 108.6399 | 1.2128 | 108.0698 | 1.2128 | 105.1231 | 1.16 | 109.0669 | 1.2303 | 1399.5524 | 1.1424 | 1389.037 | 1.1424 | 1386.3987 | 1.1073 | 1384.9046 | 1.1073 | 1383.6386 | 1.0897 | 1383.6386 | 1.0897 | 1375.7658
| 1.2831 | 107.9271 | 1.1776 | 109.0872 | 1.2479 | 108.3346 | 1.2479 | 105.0408 | 1.1952 | 108.4974 | 1.2655 | 1399.7808 | 1.1776 | 1389.2662 | 1.1776 | 1386.0541 | 1.1424 | 1387.2023 | 1.1424 | 1384.9046 | 1.1249 | 1384.5595 | 1.1249 | 1378.0907
| 1.3182 | 109.818 | 1.2128 | 109.6963 | 1.2831 | 108.7212 | 1.2831 | 105.0408 | 1.2303 | 108.2124 | 1.3006 | 1400.9225 | 1.2128 | 1389.1516 | 1.2128 | 1385.3645 | 1.1776 | 1388.2346 | 1.1776 | 1386.3987 | 1.16 | 1384.6745 | 1.16 | 1380.4074
| 1.3534 | 110.4056 | 1.2479 | 110.0815 | 1.3182 | 108.7212 | 1.3182 | 105.7596 | 1.2655 | 107.6008 | 1.3358 | 1403.4353 | 1.2479 | 1388.0053 | 1.2479 | 1388.2346 | 1.2128 | 1388.8078 | 1.2128 | 1388.5786 | 1.1952 | 1385.2496 | 1.1952 | 1380.7542
| 1.3885 | 110.7093 | 1.2831 | 110.1828 | 1.3534 | 108.7212 | 1.3534 | 106.1492 | 1.3006 | 106.9882 | 1.3709 | 1403.6638 | 1.2831 | 1390.4114 | 1.2831 | 1390.0679 | 1.2479 | 1389.9534 | 1.2479 | 1387.4318 | 1.2303 | 1385.2496 | 1.2303 | 1380.6386
| 1.4237 | 111.0733 | 1.3182 | 110.0815 | 1.3885 | 108.7212 | 1.3885 | 107.2947 | 1.3358 | 106.7633 | 1.4061 | 1404.3495 | 1.3182 | 1391.4414 | 1.3182 | 1390.5259 | 1.2831 | 1390.4114 | 1.2831 | 1389.3807 | 1.2655 | 1385.5944 | 1.2655 | 1380.0604
| 1.4588 | 110.9722 | 1.3534 | 110.2841 | 1.4237 | 108.7212 | 1.4237 | 107.3559 | 1.3709 | 106.9882 | 1.4412 | 1401.2651 | 1.3534 | 1391.4414 | 1.3534 | 1391.4414 | 1.3182 | 1390.7548 | 1.3182 | 1390.7548 | 1.3006 | 1386.6283 | 1.3006 | 1379.8289
| 1.494 | 111.0733 | 1.3885 | 110.2841 | 1.4588 | 108.9652 | 1.4588 | 107.3559 | 1.4061 | 107.1721 | 1.4764 | 1401.2651 | 1.3885 | 1393.3853 | 1.3885 | 1391.5558 | 1.3534 | 1390.9837 | 1.3534 | 1393.8425 | 1.3358 | 1388.3493 | 1.3358 | 1379.8289
| 1.5291 | 110.9318 | 1.4237 | 110.9116 | 1.494 | 109.7165 | 1.494 | 107.3559 | 1.4412 | 108.4771 | 1.5115 | 1404.3495 | 1.4237 | 1394.7565 | 1.4237 | 1392.1277 | 1.3885 | 1392.1277 | 1.3885 | 1394.4138 | 1.3709 | 1390.2969 | 1.3709 | 1379.8289
| 1.5643 | 110.9318 | 1.4588 | 111.538 | 1.5291 | 110.77 | 1.5291 | 107.458 | 1.4764 | 108.6602 | 1.5467 | 1405.6071 | 1.4588 | 1394.7565 | 1.4588 | 1393.6139 | 1.4237 | 1391.5558 | 1.4237 | 1394.528 | 1.4061 | 1389.4953 | 1.4061 | 1382.7167
| 1.5994 | 110.7902 | 1.494 | 111.7399 | 1.5643 | 110.77 | 1.5643 | 107.213 | 1.5115 | 108.7619 | 1.5818 | 1406.0646 | 1.494 | 1394.7565 | 1.494 | 1393.9567 | 1.4588 | 1392.1277 | 1.4588 | 1392.8138 | 1.4412 | 1388.9224 | 1.4412 | 1385.2496
| 1.6346 | 110.7902 | 1.5291 | 112.1837 | 1.5994 | 111.639 | 1.5994 | 107.458 | 1.5467 | 108.7619 | 1.617 | 1406.0646 | 1.5291 | 1396.3555 | 1.5291 | 1394.7565 | 1.494 | 1392.0133 | 1.494 | 1394.4138 | 1.4764 | 1390.4114 | 1.4764 | 1385.2496
| 1.6697 | 110.4866 | 1.5643 | 113.0297 | 1.6346 | 112.224 | 1.6346 | 107.9068 | 1.5818 | 108.8839 | 1.6522 | 1406.2934 | 1.5643 | 1396.3555 | 1.5643 | 1398.2966 | 1.5291 | 1392.0133 | 1.5291 | 1396.5839 | 1.5115 | 1389.8389 | 1.5115 | 1385.1346
| 1.7049 | 110.9116 | 1.5994 | 113.5928 | 1.6697 | 112.2643 | 1.6697 | 108.3346 | 1.617 | 108.8026 | 1.6873 | 1405.264 | 1.5994 | 1396.3555 | 1.5994 | 1401.1509 | 1.5643 | 1392.4707 | 1.5643 | 1397.4974 | 1.5467 | 1388.9224 | 1.5467 | 1385.0196
| 1.74 | 111.2754 | 1.6346 | 114.7971 | 1.7049 | 111.9618 | 1.7049 | 109.6963 | 1.6522 | 108.8026 | 1.7225 | 1402.6356 | 1.6346 | 1396.2413 | 1.6346 | 1400.8083 | 1.5994 | 1394.6423 | 1.5994 | 1397.9541 | 1.5818 | 1391.5558 | 1.5818 | 1385.0196
| 1.7752 | 111.6995 | 1.6697 | 115.0176 | 1.74 | 111.9618 | 1.74 | 110.3044 | 1.6873 | 108.9856 | 1.7576 | 1402.9783 | 1.6697 | 1395.3277 | 1.6697 | 1398.7532 | 1.6346 | 1396.2413 | 1.6346 | 1397.9541 | 1.617 | 1393.0424 | 1.617 | 1385.2496
| 1.8103 | 112.224 | 1.7049 | 115.0176 | 1.7752 | 111.8004 | 1.7752 | 111.3562 | 1.7225 | 109.6354 | 1.7928 | 1406.2934 | 1.7049 | 1391.8989 | 1.7049 | 1397.6116 | 1.6697 | 1398.7532 | 1.6697 | 1397.9541 | 1.6522 | 1393.1567 | 1.6522 | 1386.8579
| 1.8455 | 112.224 | 1.74 | 114.6568 | 1.8103 | 111.7197 | 1.8103 | 111.5178 | 1.7576 | 109.9599 | 1.8279 | 1407.8959 | 1.74 | 1389.7244 | 1.74 | 1394.6423 | 1.7049 | 1398.2966 | 1.7049 | 1394.071 | 1.6873 | 1394.4138 | 1.6873 | 1387.317
| 1.8806 | 112.2442 | 1.7752 | 114.2154 | 1.8455 | 111.7197 | 1.8455 | 111.3562 | 1.7928 | 111.0127 | 1.8631 | 1410.1882 | 1.7752 | 1389.1516 | 1.7752 | 1391.7846 | 1.74 | 1398.2966 | 1.74 | 1393.6139 | 1.7225 | 1396.5839 | 1.7225 | 1387.317
| 1.9158 | 112.5465 | 1.8103 | 114.115 | 1.8806 | 111.7197 | 1.8806 | 110.9924 | 1.8279 | 111.6188 | 1.8982 | 1410.1882 | 1.8103 | 1387.2023 | 1.8103 | 1390.7548 | 1.7752 | 1398.7532 | 1.7752 | 1394.985 | 1.7576 | 1394.4138 | 1.7576 | 1388.12
| 1.9509 | 112.6472 | 1.8455 | 113.633 | 1.9158 | 111.7601 | 1.9158 | 110.3651 | 1.8631 | 112.0224 | 1.9334 | 1410.1882 | 1.8455 | 1387.6612 | 1.8455 | 1391.4414 | 1.8103 | 1400.3516 | 1.8103 | 1398.4107 | 1.7928 | 1394.4138 | 1.7928 | 1388.12
| 1.9861 | 113.1303 | 1.8806 | 113.3716 | 1.9509 | 112.0627 | 1.9509 | 110.0815 | 1.8982 | 112.9089 | 1.9685 | 1410.1882 | 1.8806 | 1393.271 | 1.8806 | 1394.985 | 1.8455 | 1401.3793 | 1.8455 | 1398.9816 | 1.8279 | 1395.7845 | 1.8279 | 1388.8078
| 2.0213 | 113.2309 | 1.9158 | 113.1303 | 1.9861 | 112.1837 | 1.9861 | 110.0004 | 1.9334 | 112.9089 | 2.0037 | 1410.1882 | 1.9158 | 1402.864 | 1.9158 | 1397.3832 | 1.8806 | 1402.0645 | 1.8806 | 1399.8949 | 1.8631 | 1395.8987 | 1.8631 | 1389.3807
| 2.0564 | 113.2107 | 1.9509 | 113.4923 | 2.0213 | 112.4659 | 2.0213 | 109.8788 | 1.9685 | 112.9089 | 2.0388 | 1409.9588 | 1.9509 | 1407.2089 | 1.9509 | 1397.3832 | 1.9158 | 1402.0645 | 1.9158 | 1399.8949 | 1.8982 | 1395.8987 | 1.8982 | 1389.7244
| 2.0916 | 113.2309 | 1.9861 | 114.0347 | 2.0564 | 113.2912 | 2.0564 | 110.3854 | 2.0037 | 112.2442 | 2.074 | 1410.1882 | 1.9861 | 1407.2089 | 1.9861 | 1403.7781 | 1.9509 | 1399.8949 | 1.9509 | 1400.6942 | 1.9334 | 1396.013 | 1.9334 | 1391.327
| 2.1267 | 113.2107 | 2.0213 | 114.2556 | 2.0916 | 113.5526 | 2.0916 | 110.5069 | 2.0388 | 112.0627 | 2.1091 | 1411.4509 | 2.0213 | 1407.2089 | 2.0213 | 1404.2352 | 1.9861 | 1399.0957 | 1.9861 | 1400.6942 | 1.9685 | 1396.9265 | 1.9685 | 1392.1277
| 2.1619 | 113.0699 | 2.0564 | 115.3982 | 2.1267 | 113.6933 | 2.1267 | 110.5069 | 2.074 | 111.5178 | 2.1443 | 1412.6 | 2.0564 | 1404.4638 | 2.0564 | 1401.7219 | 2.0213 | 1399.0957 | 2.0213 | 1399.4382 | 2.0037 | 1397.1548 | 2.0037 | 1392.1277
| 2.197 | 113.1303 | 2.0916 | 115.7585 | 2.1619 | 114.1753 | 2.1619 | 110.7093 | 2.1091 | 111.1137 | 2.1794 | 1414.2112 | 2.0916 | 1400.9225 | 2.0916 | 1395.2135 | 2.0564 | 1397.7257 | 2.0564 | 1399.0957 | 2.0388 | 1397.1548 | 2.0388 | 1392.242
| 2.2322 | 113.1303 | 2.1267 | 115.7585 | 2.197 | 114.3358 | 2.197 | 110.8711 | 2.1443 | 111.1137 | 2.2146 | 1414.4416 | 2.1267 | 1400.58 | 2.1267 | 1399.4382 | 2.0916 | 1396.1272 | 2.0916 | 1399.2099 | 2.074 | 1397.1548 | 2.074 | 1392.242
| 2.2673 | 112.5666 | 2.1619 | 115.7585 | 2.2322 | 114.3559 | 2.2322 | 110.8711 | 2.1794 | 111.982 | 2.2497 | 1415.2485 | 2.1619 | 1400.9225 | 2.1619 | 1401.7219 | 2.1267 | 1401.7219 | 2.1267 | 1399.7808 | 2.1091 | 1397.7257 | 2.1091 | 1392.242
| 2.3025 | 112.4256 | 2.197 | 116.5584 | 2.2673 | 114.5164 | 2.2673 | 110.9924 | 2.2146 | 113.0095 | 2.2849 | 1418.1371 | 2.197 | 1403.2068 | 2.197 | 1401.7219 | 2.1619 | 1404.1209 | 2.1619 | 1400.2374 | 2.1443 | 1397.7257 | 2.1443 | 1393.9567
| 2.3376 | 112.5062 | 2.2322 | 116.2986 | 2.3025 | 114.376 | 2.3025 | 111.1744 | 2.2497 | 113.1102 | 2.32 | 1418.1371 | 2.2322 | 1404.6924 | 2.2322 | 1403.4353 | 2.197 | 1401.2651 | 2.197 | 1401.8361 | 2.1794 | 1398.0682 | 2.1794 | 1394.4138
| 2.3728 | 112.2038 | 2.2673 | 116.3585 | 2.3376 | 114.376 | 2.3376 | 111.5582 | 2.2849 | 113.1303 | 2.3552 | 1417.6741 | 2.2673 | 1408.5832 | 2.2673 | 1404.8067 | 2.2322 | 1404.1209 | 2.2322 | 1402.864 | 2.2146 | 1399.5524 | 2.2146 | 1395.5561
| 2.4079 | 111.8811 | 2.3025 | 116.3585 | 2.3728 | 114.2556 | 2.3728 | 112.3853 | 2.32 | 113.1504 | 2.3903 | 1418.2529 | 2.3025 | 1411.336 | 2.3025 | 1405.264 | 2.2673 | 1405.0354 | 2.2673 | 1405.0354 | 2.2497 | 1400.58 | 2.2497 | 1396.8123
| 2.4431 | 112.6673 | 2.3376 | 116.5584 | 2.4079 | 114.1351 | 2.4079 | 112.6069 | 2.3552 | 113.5124 | 2.4255 | 1418.3687 | 2.3376 | 1412.1402 | 2.3376 | 1407.2089 | 2.3025 | 1405.0354 | 2.3025 | 1407.0944 | 2.2849 | 1401.4935 | 2.2849 | 1397.6116
| 2.4782 | 112.6673 | 2.3728 | 116.2986 | 2.4431 | 113.6129 | 2.4431 | 112.6069 | 2.3903 | 113.5124 | 2.4607 | 1418.3687 | 2.3728 | 1414.9026 | 2.3728 | 1406.7511 | 2.3376 | 1406.0646 | 2.3376 | 1408.354 | 2.32 | 1404.1209 | 2.32 | 1400.9225
| 2.5134 | 112.6673 | 2.4079 | 116.2386 | 2.4782 | 113.8741 | 2.4782 | 112.8083 | 2.4255 | 113.8339 | 2.4958 | 1416.6336 | 2.4079 | 1414.9026 | 2.4079 | 1406.7511 | 2.3728 | 1410.762 | 2.3728 | 1408.1249 | 2.3552 | 1405.0354 | 2.3552 | 1401.1509
| 2.5485 | 112.6673 | 2.4431 | 116.1386 | 2.5134 | 113.4923 | 2.5134 | 112.8083 | 2.4607 | 114.1351 | 2.531 | 1418.832 | 2.4431 | 1414.9026 | 2.4431 | 1406.5222 | 2.4079 | 1409.7294 | 2.4079 | 1410.0735 | 2.3903 | 1405.3784 | 2.3903 | 1402.6356
| 2.5837 | 112.6673 | 2.4782 | 115.7185 | 2.5485 | 112.7277 | 2.5485 | 112.6472 | 2.4958 | 114.1351 | 2.5661 | 1418.3687 | 2.4782 | 1412.1402 | 2.4782 | 1405.4927 | 2.4431 | 1410.0735 | 2.4431 | 1408.1249 | 2.4255 | 1409.8441 | 2.4255 | 1403.0925
| 2.6188 | 112.224 | 2.5134 | 115.258 | 2.5837 | 112.929 | 2.5837 | 112.1837 | 2.531 | 114.0147 | 2.6013 | 1420.9213 | 2.5134 | 1408.5832 | 2.5134 | 1404.5781 | 2.4782 | 1410.0735 | 2.4782 | 1406.179 | 2.4607 | 1409.8441 | 2.4607 | 1402.9783
| 2.654 | 112.103 | 2.5485 | 115.0377 | 2.6188 | 112.8888 | 2.6188 | 112.0627 | 2.5661 | 114.3358 | 2.6364 | 1423.3677 | 2.5485 | 1409.8441 | 2.5485 | 1406.4078 | 2.5134 | 1408.4686 | 2.5134 | 1407.4379 | 2.4958 | 1407.4379 | 2.4958 | 1403.6638
| 2.6891 | 111.5582 | 2.5837 | 115.1579 | 2.654 | 112.5666 | 2.654 | 111.7197 | 2.6013 | 114.3358 | 2.6716 | 1423.7181 | 2.5837 | 1412.0253 | 2.5837 | 1407.5523 | 2.5485 | 1406.8655 | 2.5485 | 1407.4379 | 2.531 | 1407.5523 | 2.531 | 1403.6638
| 2.7243 | 111.4774 | 2.6188 | 115.258 | 2.6891 | 112.7881 | 2.6891 | 111.7197 | 2.6364 | 114.0347 | 2.7067 | 1423.7181 | 2.6188 | 1412.0253 | 2.6188 | 1408.354 | 2.5837 | 1408.5832 | 2.5837 | 1408.9269 | 2.5661 | 1407.7814 | 2.5661 | 1403.0925
| 2.7594 | 111.3966 | 2.654 | 115.4383 | 2.7243 | 112.9089 | 2.7243 | 111.7197 | 2.6716 | 114.0347 | 2.7419 | 1424.0686 | 2.654 | 1411.7955 | 2.654 | 1408.354 | 2.6188 | 1409.8441 | 2.6188 | 1410.762 | 2.6013 | 1407.5523 | 2.6013 | 1401.6077
| 2.7946 | 111.336 | 2.6891 | 115.4583 | 2.7594 | 112.9089 | 2.7594 | 112.1434 | 2.7067 | 114.2355 | 2.777 | 1423.1343 | 2.6891 | 1412.0253 | 2.6891 | 1408.2395 | 2.654 | 1407.7814 | 2.654 | 1408.9269 | 2.6364 | 1409.2708 | 2.6364 | 1401.6077
| 2.8298 | 111.3966 | 2.7243 | 115.3382 | 2.7946 | 112.9089 | 2.7946 | 112.2442 | 2.7419 | 113.7334 | 2.8122 | 1423.0176 | 2.7243 | 1412.0253 | 2.7243 | 1407.3234 | 2.6891 | 1408.5832 | 2.6891 | 1411.7955 | 2.6716 | 1409.1562 | 2.6716 | 1399.4382
| 2.8649 | 111.5784 | 2.7594 | 115.258 | 2.8298 | 112.9089 | 2.8298 | 112.3651 | 2.777 | 113.5928 | 2.8473 | 1423.0176 | 2.7594 | 1412.6 | 2.7594 | 1407.3234 | 2.7243 | 1412.0253 | 2.7243 | 1413.0601 | 2.7067 | 1407.7814 | 2.7067 | 1399.0957
| 2.9001 | 111.134 | 2.7946 | 115.0978 | 2.8649 | 112.5062 | 2.8649 | 112.3046 | 2.8122 | 113.5928 | 2.8825 | 1423.4845 | 2.7946 | 1413.0601 | 2.7946 | 1407.6668 | 2.7594 | 1410.3029 | 2.7594 | 1410.762 | 2.7419 | 1411.2212 | 2.7419 | 1399.4382
| 2.9352 | 111.134 | 2.8298 | 114.8773 | 2.9001 | 112.3651 | 2.9001 | 112.3046 | 2.8473 | 114.0347 | 2.9176 | 1424.3025 | 2.8298 | 1415.1332 | 2.8298 | 1408.2395 | 2.7946 | 1410.3029 | 2.7946 | 1410.762 | 2.777 | 1410.6472 | 2.777 | 1400.2374
| 2.9704 | 111.982 | 2.8649 | 114.737 | 2.9352 | 112.1837 | 2.9352 | 112.2442 | 2.8825 | 113.9946 | 2.9528 | 1428.1759 | 2.8649 | 1415.1332 | 2.8649 | 1413.5204 | 2.8298 | 1412.0253 | 2.8298 | 1410.8768 | 2.8122 | 1404.3495 | 2.8122 | 1400.8083
| 3.0055 | 112.1434 | 2.9001 | 113.9544 | 2.9704 | 112.1837 | 2.9704 | 112.1232 | 2.9176 | 114.0749 | 2.9879 | 1428.1759 | 2.9001 | 1416.2871 | 2.9001 | 1414.9026 | 2.8649 | 1413.1751 | 2.8649 | 1410.762 | 2.8473 | 1410.6472 | 2.8473 | 1401.2651
| 3.0407 | 112.2442 | 2.9352 | 113.9544 | 3.0055 | 112.345 | 3.0055 | 111.4572 | 2.9528 | 114.0749 | 3.0231 | 1428.1759 | 2.9352 | 1416.7492 | 2.9352 | 1417.5585 | 2.9001 | 1411.7955 | 2.9001 | 1410.8768 | 2.8825 | 1416.5181 | 2.8825 | 1401.7219
| 3.0758 | 112.2845 | 2.9704 | 113.5928 | 3.0407 | 112.345 | 3.0407 | 110.9722 | 2.9879 | 113.854 | 3.0582 | 1426.764 | 2.9704 | 1416.8647 | 2.9704 | 1416.0562 | 2.9352 | 1411.7955 | 2.9352 | 1419.8758 | 2.9176 | 1416.5181 | 2.9176 | 1401.9503
| 3.111 | 112.8485 | 3.0055 | 113.9544 | 3.0758 | 112.4659 | 3.0758 | 110.9722 | 3.0231 | 113.6531 | 3.0934 | 1425.7076 | 3.0055 | 1417.0959 | 3.0055 | 1415.0179 | 2.9704 | 1412.6 | 2.9704 | 1425.7076 | 2.9528 | 1418.4845 | 2.9528 | 1401.9503
| 3.1461 | 112.1434 | 3.0407 | 114.396 | 3.111 | 112.5465 | 3.111 | 110.8913 | 3.0582 | 112.8284 | 3.1285 | 1423.7181 | 3.0407 | 1417.0959 | 3.0407 | 1414.7873 | 3.0055 | 1412.0253 | 3.0055 | 1418.9479 | 2.9879 | 1422.6677 | 2.9879 | 1403.7781
| 3.1813 | 111.8408 | 3.0758 | 114.4362 | 3.1461 | 113.1504 | 3.1461 | 110.8913 | 3.0934 | 112.4256 | 3.1637 | 1422.4346 | 3.0758 | 1417.0959 | 3.0758 | 1413.2902 | 3.0407 | 1412.0253 | 3.0407 | 1412.1402 | 3.0231 | 1419.5277 | 3.0231 | 1404.921
| 3.2164 | 111.336 | 3.111 | 114.4362 | 3.1813 | 113.0297 | 3.1813 | 110.952 | 3.1285 | 112.2643 | 3.1989 | 1417.9056 | 3.111 | 1416.4026 | 3.111 | 1412.83 | 3.0758 | 1411.7955 | 3.0758 | 1410.6472 | 3.0582 | 1417.9056 | 3.0582 | 1405.7215
| 3.2516 | 111.4168 | 3.1461 | 115.0577 | 3.2164 | 113.3716 | 3.2164 | 111.4572 | 3.1637 | 112.0627 | 3.234 | 1419.5277 | 3.1461 | 1414.9026 | 3.1461 | 1411.5657 | 3.111 | 1412.6 | 3.111 | 1409.2708 | 3.0934 | 1417.9056 | 3.0934 | 1405.1497
| 3.2867 | 111.336 | 3.1813 | 115.2981 | 3.2516 | 114.0749 | 3.2516 | 111.4572 | 3.1989 | 112.224 | 3.2692 | 1417.3272 | 3.1813 | 1411.7955 | 3.1813 | 1409.5001 | 3.1461 | 1411.5657 | 3.1461 | 1407.7814 | 3.1285 | 1421.7357 | 3.1285 | 1405.7215
| 3.3219 | 111.8408 | 3.2164 | 115.4383 | 3.2867 | 113.7736 | 3.2867 | 111.4572 | 3.234 | 112.2643 | 3.3043 | 1415.5946 | 3.2164 | 1411.336 | 3.2164 | 1409.2708 | 3.1813 | 1411.5657 | 3.1813 | 1409.2708 | 3.1637 | 1419.5277 | 3.1637 | 1404.5781
| 3.357 | 112.5263 | 3.2516 | 115.5984 | 3.3219 | 113.633 | 3.3219 | 111.5986 | 3.2692 | 112.5666 | 3.3395 | 1413.5204 | 3.2516 | 1409.0415 | 3.2516 | 1408.1249 | 3.2164 | 1410.9916 | 3.2164 | 1409.2708 | 3.1989 | 1414.9026 | 3.1989 | 1402.4071
| 3.3922 | 112.5263 | 3.2867 | 115.5984 | 3.357 | 113.5325 | 3.357 | 111.6995 | 3.3043 | 112.627 | 3.3746 | 1409.7294 | 3.2867 | 1408.2395 | 3.2867 | 1407.4379 | 3.2516 | 1408.5832 | 3.2516 | 1405.7215 | 3.234 | 1414.5568 | 3.234 | 1400.2374
| 3.4273 | 111.6591 | 3.3219 | 115.5184 | 3.3922 | 112.9693 | 3.3922 | 112.3853 | 3.3395 | 112.7076 | 3.4098 | 1407.3234 | 3.3219 | 1404.4638 | 3.3219 | 1405.4927 | 3.2867 | 1405.9502 | 3.2867 | 1404.4638 | 3.2692 | 1408.2395 | 3.2692 | 1398.9816
| 3.4625 | 112.2643 | 3.357 | 115.1979 | 3.4273 | 112.486 | 3.4273 | 112.4659 | 3.3746 | 113.0095 | 3.4449 | 1407.2089 | 3.357 | 1400.58 | 3.357 | 1396.6981 | 3.3219 | 1402.1787 | 3.3219 | 1403.2068 | 3.3043 | 1402.7498 | 3.3043 | 1396.8123
| 3.4976 | 111.8004 | 3.3922 | 114.6167 | 3.4625 | 112.4659 | 3.4625 | 112.3853 | 3.4098 | 112.627 | 3.4801 | 1407.2089 | 3.3922 | 1397.4974 | 3.3922 | 1396.3555 | 3.357 | 1399.7808 | 3.357 | 1399.0957 | 3.3395 | 1402.5214 | 3.3395 | 1392.9281
| 3.5328 | 111.6591 | 3.4273 | 114.4763 | 3.4976 | 112.4054 | 3.4976 | 112.4659 | 3.4449 | 113.633 | 3.5152 | 1406.7511 | 3.4273 | 1396.6981 | 3.4273 | 1395.6703 | 3.3922 | 1397.6116 | 3.3922 | 1397.0407 | 3.3746 | 1398.0682 | 3.3746 | 1391.6702
| 3.5679 | 111.8811 | 3.4625 | 114.6167 | 3.5328 | 112.2845 | 3.5328 | 112.345 | 3.4801 | 113.6933 | 3.5504 | 1406.8655 | 3.4625 | 1396.6981 | 3.4625 | 1395.6703 | 3.4273 | 1396.1272 | 3.4273 | 1394.7565 | 3.4098 | 1397.0407 | 3.4098 | 1391.6702
| 3.6031 | 111.8811 | 3.4976 | 114.6167 | 3.5679 | 112.2643 | 3.5679 | 111.3764 | 3.5152 | 113.6129 | 3.5855 | 1406.7511 | 3.4976 | 1395.8987 | 3.4976 | 1395.8987 | 3.4625 | 1395.2135 | 3.4625 | 1394.2995 | 3.4449 | 1395.8987 | 3.4449 | 1390.4114
| 3.6383 | 111.5784 | 3.5328 | 114.4963 | 3.6031 | 112.2845 | 3.6031 | 110.8104 | 3.5504 | 113.251 | 3.6207 | 1404.8067 | 3.5328 | 1396.9265 | 3.5328 | 1396.2413 | 3.4976 | 1392.6994 | 3.4976 | 1395.0992 | 3.4801 | 1391.4414 | 3.4801 | 1389.1516
| 3.6734 | 111.2956 | 3.5679 | 114.6367 | 3.6383 | 112.5867 | 3.6383 | 110.8104 | 3.5855 | 113.3314 | 3.6558 | 1404.1209 | 3.5679 | 1396.9265 | 3.5679 | 1395.4419 | 3.5328 | 1392.0133 | 3.5328 | 1396.3555 | 3.5152 | 1394.7565 | 3.5152 | 1388.4639
| 3.7086 | 111.2956 | 3.6031 | 114.6768 | 3.6734 | 112.6874 | 3.6734 | 110.8104 | 3.6207 | 112.4256 | 3.691 | 1404.1209 | 3.6031 | 1394.7565 | 3.6031 | 1395.2135 | 3.5679 | 1392.0133 | 3.5679 | 1396.3555 | 3.5504 | 1395.8987 | 3.5504 | 1385.9392
| 3.7437 | 110.9924 | 3.6383 | 114.6568 | 3.7086 | 112.7479 | 3.7086 | 110.8711 | 3.6558 | 112.345 | 3.7261 | 1403.321 | 3.6383 | 1394.1853 | 3.6383 | 1395.0992 | 3.6031 | 1389.3807 | 3.6031 | 1396.3555 | 3.5855 | 1390.9837 | 3.5855 | 1385.3645
| 3.7789 | 110.9924 | 3.6734 | 114.6768 | 3.7437 | 112.5867 | 3.7437 | 111.1137 | 3.691 | 112.345 | 3.7613 | 1403.0925 | 3.6734 | 1395.7845 | 3.6734 | 1394.6423 | 3.6383 | 1389.3807 | 3.6383 | 1395.7845 | 3.6207 | 1390.9837 | 3.6207 | 1384.3294
| 3.814 | 111.134 | 3.7086 | 114.6768 | 3.7789 | 112.4256 | 3.7789 | 111.4976 | 3.7261 | 112.345 | 3.7964 | 1403.321 | 3.7086 | 1394.7565 | 3.7086 | 1395.4419 | 3.6734 | 1389.3807 | 3.6734 | 1395.0992 | 3.6558 | 1389.4953 | 3.6558 | 1383.293
| 3.8492 | 111.538 | 3.7437 | 114.0749 | 3.814 | 112.3248 | 3.814 | 111.5582 | 3.7613 | 112.7881 | 3.8316 | 1403.321 | 3.7437 | 1395.5561 | 3.7437 | 1395.7845 | 3.7086 | 1391.327 | 3.7086 | 1393.3853 | 3.691 | 1389.4953 | 3.691 | 1384.0992
| 3.8843 | 111.6793 | 3.7789 | 113.4722 | 3.8492 | 111.9417 | 3.8492 | 111.4976 | 3.7964 | 112.8888 | 3.8667 | 1403.0925 | 3.7789 | 1396.2413 | 3.7789 | 1398.0682 | 3.7437 | 1394.4138 | 3.7437 | 1392.9281 | 3.7261 | 1390.7548 | 3.7261 | 1384.0992
| 3.9195 | 111.6793 | 3.814 | 113.5727 | 3.8843 | 111.8408 | 3.8843 | 111.4976 | 3.8316 | 113.1705 | 3.9019 | 1404.4638 | 3.814 | 1396.2413 | 3.814 | 1398.6391 | 3.7789 | 1394.985 | 3.7789 | 1392.6994 | 3.7613 | 1392.0133 | 3.7613 | 1384.9046
| 3.9546 | 111.336 | 3.8492 | 113.6129 | 3.9195 | 111.7803 | 3.9195 | 111.3158 | 3.8667 | 113.2309 | 3.937 | 1406.6367 | 3.8492 | 1397.269 | 3.8492 | 1398.6391 | 3.814 | 1394.985 | 3.814 | 1394.7565 | 3.7964 | 1392.242 | 3.7964 | 1385.9392
| 3.9898 | 111.3562 | 3.8843 | 113.7937 | 3.9546 | 111.6188 | 3.9546 | 111.3158 | 3.9019 | 113.2309 | 3.9722 | 1407.8959 | 3.8843 | 1396.8123 | 3.8843 | 1399.3241 | 3.8492 | 1394.985 | 3.8492 | 1395.8987 | 3.8316 | 1393.8425 | 3.8316 | 1386.7431
| 4.0249 | 111.4976 | 3.9195 | 114.0147 | 3.9898 | 111.2754 | 3.9898 | 110.7295 | 3.937 | 112.8284 | 4.0074 | 1409.9588 | 3.9195 | 1397.6116 | 3.9195 | 1399.7808 | 3.8843 | 1396.013 | 3.8843 | 1398.2966 | 3.8667 | 1395.3277 | 3.8667 | 1387.8906
| 4.0601 | 110.9722 | 3.9546 | 113.9142 | 4.0249 | 110.6081 | 4.0249 | 110.7295 | 3.9722 | 112.8284 | 4.0425 | 1411.5657 | 3.9546 | 1401.4935 | 3.9546 | 1400.9225 | 3.9195 | 1396.013 | 3.9195 | 1400.0091 | 3.9019 | 1395.8987 | 3.9019 | 1389.3807
| 4.0952 | 110.77 | 3.9898 | 113.7736 | 4.0601 | 111.1946 | 4.0601 | 110.3651 | 4.0074 | 112.8083 | 4.0777 | 1412.3701 | 3.9898 | 1402.1787 | 3.9898 | 1401.6077 | 3.9546 | 1398.1824 | 3.9546 | 1400.6942 | 3.937 | 1396.2413 | 3.937 | 1390.5259
| 4.1304 | 110.8509 | 4.0249 | 113.5325 | 4.0952 | 111.5986 | 4.0952 | 110.3651 | 4.0425 | 112.8284 | 4.1128 | 1413.9809 | 4.0249 | 1404.921 | 4.0249 | 1404.0066 | 3.9898 | 1399.6666 | 3.9898 | 1401.8361 | 3.9722 | 1396.2413 | 3.9722 | 1391.2125
| 4.1655 | 110.4259 | 4.0601 | 113.5928 | 4.1304 | 112.2038 | 4.1304 | 109.8788 | 4.0777 | 112.8083 | 4.148 | 1414.9026 | 4.0601 | 1406.2934 | 4.0601 | 1404.1209 | 4.0249 | 1400.3516 | 4.0249 | 1402.7498 | 4.0074 | 1396.3555 | 4.0074 | 1390.9837
| 4.2007 | 110.1828 | 4.0952 | 113.633 | 4.1655 | 112.768 | 4.1655 | 109.8788 | 4.1128 | 112.486 | 4.1831 | 1414.096 | 4.0952 | 1405.9502 | 4.0952 | 1401.6077 | 4.0601 | 1401.8361 | 4.0601 | 1402.6356 | 4.0425 | 1397.4974 | 4.0425 | 1392.0133
| 4.2358 | 110.2031 | 4.1304 | 114.737 | 4.2007 | 112.7076 | 4.2007 | 110.1423 | 4.148 | 111.639 | 4.2183 | 1414.096 | 4.1304 | 1407.4379 | 4.1304 | 1404.3495 | 4.0952 | 1401.6077 | 4.0952 | 1401.4935 | 4.0777 | 1399.5524 | 4.0777 | 1393.271
| 4.271 | 110.5271 | 4.1655 | 115.258 | 4.2358 | 112.2038 | 4.2358 | 110.6283 | 4.1831 | 111.0531 | 4.2534 | 1413.4053 | 4.1655 | 1409.1562 | 4.1655 | 1404.6924 | 4.1304 | 1400.6942 | 4.1304 | 1401.4935 | 4.1128 | 1397.8399 | 4.1128 | 1393.4996
| 4.3061 | 110.6081 | 4.2007 | 115.258 | 4.271 | 112.0425 | 4.271 | 111.2956 | 4.2183 | 110.8913 | 4.2886 | 1414.9026 | 4.2007 | 1406.5222 | 4.2007 | 1403.321 | 4.1655 | 1400.6942 | 4.1655 | 1401.8361 | 4.148 | 1399.5524 | 4.148 | 1393.7282
| 4.3413 | 111.0127 | 4.2358 | 114.6768 | 4.3061 | 111.982 | 4.3061 | 112.0829 | 4.2534 | 110.8306 | 4.3237 | 1415.8254 | 4.2358 | 1408.1249 | 4.2358 | 1404.6924 | 4.2007 | 1400.6942 | 4.2007 | 1402.6356 | 4.1831 | 1398.9816 | 4.1831 | 1394.528
| 4.3764 | 111.4572 | 4.271 | 114.5966 | 4.3413 | 111.6793 | 4.3413 | 112.1837 | 4.2886 | 111.1744 | 4.3589 | 1417.6741 | 4.271 | 1409.0415 | 4.271 | 1405.1497 | 4.2358 | 1401.6077 | 4.2358 | 1403.8923 | 4.2183 | 1400.4658 | 4.2183 | 1396.1272
| 4.4116 | 111.6793 | 4.3061 | 114.1953 | 4.3764 | 111.4168 | 4.3764 | 112.2643 | 4.3237 | 111.6995 | 4.394 | 1420.108 | 4.3061 | 1408.6977 | 4.3061 | 1404.6924 | 4.271 | 1404.0066 | 4.271 | 1407.0944 | 4.2534 | 1400.4658 | 4.2534 | 1398.6391
| 4.4468 | 112.103 | 4.3413 | 113.5727 | 4.4116 | 112.0425 | 4.4116 | 111.9013 | 4.3589 | 112.2845 | 4.4292 | 1421.5029 | 4.3413 | 1408.8123 | 4.3413 | 1405.6071 | 4.3061 | 1405.264 | 4.3061 | 1407.2089 | 4.2886 | 1402.9783 | 4.2886 | 1400.9225
| 4.4819 | 112.103 | 4.3764 | 113.251 | 4.4468 | 112.0425 | 4.4468 | 111.134 | 4.394 | 113.1303 | 4.4643 | 1422.2015 | 4.3764 | 1410.0735 | 4.3764 | 1405.8358 | 4.3413 | 1405.264 | 4.3413 | 1408.5832 | 4.3237 | 1402.6356 | 4.3237 | 1401.9503
| 4.5171 | 111.4976 | 4.4116 | 113.1705 | 4.4819 | 111.3764 | 4.4819 | 110.3854 | 4.4292 | 113.2309 | 4.4995 | 1422.2015 | 4.4116 | 1411.1064 | 4.4116 | 1406.0646 | 4.3764 | 1406.4078 | 4.3764 | 1409.5001 | 4.3589 | 1404.3495 | 4.3589 | 1402.9783
| 4.5522 | 111.4168 | 4.4468 | 113.0297 | 4.5171 | 111.3158 | 4.5171 | 110.2233 | 4.4643 | 113.4521 | 4.5346 | 1422.2015 | 4.4468 | 1410.9916 | 4.4468 | 1408.1249 | 4.4116 | 1407.3234 | 4.4116 | 1410.3029 | 4.394 | 1404.0066 | 4.394 | 1403.4353
| 4.5874 | 111.3764 | 4.4819 | 113.0297 | 4.5522 | 111.6591 | 4.5522 | 110.041 | 4.4995 | 113.1303 | 4.5698 | 1420.108 | 4.4819 | 1410.9916 | 4.4819 | 1408.1249 | 4.4468 | 1406.4078 | 4.4468 | 1410.8768 | 4.4292 | 1405.6071 | 4.4292 | 1404.8067
| 4.6225 | 111.4976 | 4.5171 | 113.1303 | 4.5874 | 111.8408 | 4.5874 | 109.9802 | 4.5346 | 112.929 | 4.6049 | 1419.8758 | 4.5171 | 1411.4509 | 4.5171 | 1408.5832 | 4.4819 | 1407.4379 | 4.4819 | 1411.2212 | 4.4643 | 1406.8655 | 4.4643 | 1403.4353
| 4.6577 | 111.639 | 4.5522 | 113.2711 | 4.6225 | 112.0627 | 4.6225 | 110.1828 | 4.5698 | 112.3248 | 4.6401 | 1419.5277 | 4.5522 | 1411.4509 | 4.5522 | 1410.3029 | 4.5171 | 1408.1249 | 4.5171 | 1411.4509 | 4.4995 | 1408.354 | 4.4995 | 1403.4353
| 4.6928 | 111.9013 | 4.5874 | 113.7535 | 4.6577 | 112.4457 | 4.6577 | 110.3854 | 4.6049 | 111.9215 | 4.6752 | 1419.5277 | 4.5874 | 1410.9916 | 4.5874 | 1410.3029 | 4.5522 | 1408.4686 | 4.5522 | 1411.9104 | 4.5346 | 1408.8123 | 4.5346 | 1403.321
| 4.728 | 112.7076 | 4.6225 | 114.2957 | 4.6928 | 113.2107 | 4.6928 | 110.6081 | 4.6401 | 111.639 | 4.7104 | 1419.6437 | 4.6225 | 1412.4851 | 4.6225 | 1408.5832 | 4.5874 | 1409.0415 | 4.5874 | 1411.9104 | 4.5698 | 1410.0735 | 4.5698 | 1403.321
| 4.7631 | 112.7881 | 4.6577 | 114.9174 | 4.728 | 113.2107 | 4.728 | 111.0935 | 4.6752 | 111.5784 | 4.7455 | 1421.1539 | 4.6577 | 1412.4851 | 4.6577 | 1408.0104 | 4.6225 | 1408.4686 | 4.6225 | 1411.2212 | 4.6049 | 1409.0415 | 4.6049 | 1403.5495
| 4.7983 | 112.7881 | 4.6928 | 115.3582 | 4.7631 | 113.1102 | 4.7631 | 111.538 | 4.7104 | 111.5986 | 4.7807 | 1421.3865 | 4.6928 | 1412.4851 | 4.6928 | 1407.4379 | 4.6577 | 1408.0104 | 4.6577 | 1411.5657 | 4.6401 | 1408.354 | 4.6401 | 1403.5495
| 4.8334 | 111.861 | 4.728 | 115.4383 | 4.7983 | 112.9693 | 4.7983 | 111.9013 | 4.7455 | 111.9215 | 4.8159 | 1420.805 | 4.728 | 1412.6 | 4.728 | 1407.2089 | 4.6928 | 1406.5222 | 4.6928 | 1411.5657 | 4.6752 | 1408.2395 | 4.6752 | 1403.6638
| 4.8686 | 111.4774 | 4.7631 | 115.1579 | 4.8334 | 112.9089 | 4.8334 | 112.3046 | 4.7807 | 112.6069 | 4.851 | 1422.6677 | 4.7631 | 1412.83 | 4.7631 | 1406.2934 | 4.728 | 1406.5222 | 4.728 | 1410.762 | 4.7104 | 1409.7294 | 4.7104 | 1403.8923
| 4.9037 | 111.4774 | 4.7983 | 114.5766 | 4.8686 | 112.7076 | 4.8686 | 112.3046 | 4.8159 | 112.929 | 4.8862 | 1424.0686 | 4.7983 | 1412.83 | 4.7983 | 1407.5523 | 4.7631 | 1406.6367 | 4.7631 | 1410.762 | 4.7455 | 1408.8123 | 4.7455 | 1404.1209
| 4.9389 | 111.7399 | 4.8334 | 114.5766 | 4.9037 | 112.5263 | 4.9037 | 112.2442 | 4.851 | 113.2107 | 4.9213 | 1424.6534 | 4.8334 | 1417.3272 | 4.8334 | 1409.1562 | 4.7983 | 1408.9269 | 4.7983 | 1412.4851 | 4.7807 | 1409.7294 | 4.7807 | 1404.1209
| 4.974 | 112.8687 | 4.8686 | 114.5766 | 4.9389 | 112.5263 | 4.9389 | 111.6793 | 4.8862 | 113.6531 | 4.9565 | 1424.1856 | 4.8686 | 1417.3272 | 4.8686 | 1412.0253 | 4.8334 | 1408.9269 | 4.8334 | 1413.5204 | 4.8159 | 1409.9588 | 4.8159 | 1403.8923
| 5.0092 | 113.2309 | 4.9037 | 114.9375 | 4.974 | 112.5465 | 4.974 | 111.336 | 4.9213 | 113.8741 | 4.9916 | 1425.2388 | 4.9037 | 1417.7899 | 4.9037 | 1414.4416 | 4.8686 | 1409.2708 | 4.8686 | 1415.0179 | 4.851 | 1410.6472 | 4.851 | 1404.8067
| 5.0443 | 113.2711 | 4.9389 | 115.0577 | 5.0092 | 111.9417 | 5.0092 | 110.4259 | 4.9565 | 113.7736 | 5.0268 | 1424.7704 | 4.9389 | 1417.7899 | 4.9389 | 1415.71 | 4.9037 | 1409.7294 | 4.9037 | 1418.0213 | 4.8862 | 1410.9916 | 4.8862 | 1405.1497
| 5.0795 | 113.2711 | 4.974 | 115.0777 | 5.0443 | 111.7197 | 5.0443 | 111.336 | 4.9916 | 113.2711 | 5.0619 | 1424.5364 | 4.974 | 1418.3687 | 4.974 | 1418.1371 | 4.9389 | 1410.3029 | 4.9389 | 1419.5277 | 4.9213 | 1411.4509 | 4.9213 | 1405.6071
| 5.1146 | 113.3515 | 5.0092 | 114.7771 | 5.0795 | 111.1137 | 5.0795 | 111.336 | 5.0268 | 112.8284 | 5.0971 | 1425.9422 | 5.0092 | 1421.7357 | 5.0092 | 1416.4026 | 4.974 | 1412.0253 | 4.974 | 1419.7598 | 4.9565 | 1412.0253 | 4.9565 | 1406.179
| 5.1498 | 114.0548 | 5.0443 | 114.5966 | 5.1146 | 111.0329 | 5.1146 | 111.3764 | 5.0619 | 112.8284 | 5.1322 | 1430.5384 | 5.0443 | 1423.9518 | 5.0443 | 1419.4117 | 5.0092 | 1414.2112 | 5.0092 | 1421.6193 | 4.9916 | 1415.4792 | 4.9916 | 1409.5001
| 5.1849 | 115.4783 | 5.0795 | 114.0548 | 5.1498 | 110.6081 | 5.1498 | 111.2552 | 5.0971 | 112.6874 | 5.1674 | 1432.6754 | 5.0795 | 1425.2388 | 5.0795 | 1419.4117 | 5.0443 | 1417.0959 | 5.0443 | 1421.3865 | 5.0268 | 1417.4428 | 5.0268 | 1410.1882
| 5.2201 | 115.5984 | 5.1146 | 113.4119 | 5.1849 | 110.5271 | 5.1849 | 111.0329 | 5.1322 | 113.09 | 5.2025 | 1436.5013 | 5.1146 | 1426.2942 | 5.1146 | 1419.8758 | 5.0795 | 1418.2529 | 5.0795 | 1421.3865 | 5.0619 | 1418.832 | 5.0619 | 1413.6355
| 5.2553 | 115.5984 | 5.1498 | 112.9089 | 5.2201 | 111.0329 | 5.2201 | 110.5271 | 5.1674 | 113.09 | 5.2377 | 1438.4282 | 5.1498 | 1427.9403 | 5.1498 | 1422.4346 | 5.1146 | 1420.6888 | 5.1146 | 1422.318 | 5.0971 | 1419.0638 | 5.0971 | 1415.3638
| 5.2904 | 115.7185 | 5.1849 | 112.9089 | 5.2553 | 112.1837 | 5.2553 | 110.4056 | 5.2025 | 113.2309 | 5.2728 | 1439.5164 | 5.1849 | 1429.4738 | 5.1849 | 1426.1768 | 5.1498 | 1422.901 | 5.1498 | 1425.5904 | 5.1322 | 1420.9213 | 5.1322 | 1418.7162
| 5.3256 | 115.3582 | 5.2201 | 112.9089 | 5.2904 | 112.5666 | 5.2904 | 110.0004 | 5.2377 | 113.0095 | 5.308 | 1439.5164 | 5.2201 | 1429.4738 | 5.2201 | 1427.8225 | 5.1849 | 1423.8349 | 5.1849 | 1427.2341 | 5.1674 | 1423.3677 | 5.1674 | 1419.2957
| 5.3607 | 115.218 | 5.2553 | 113.432 | 5.3256 | 112.7479 | 5.3256 | 110.0004 | 5.2728 | 112.8888 | 5.3431 | 1439.0324 | 5.2553 | 1429.4738 | 5.2553 | 1430.42 | 5.2201 | 1426.4116 | 5.2201 | 1429.4738 | 5.2025 | 1426.4116 | 5.2025 | 1418.1371
| 5.3959 | 115.3181 | 5.2904 | 113.854 | 5.3607 | 113.0498 | 5.3607 | 110.3246 | 5.308 | 112.2038 | 5.3783 | 1439.0324 | 5.2904 | 1431.2496 | 5.2904 | 1431.131 | 5.2553 | 1426.764 | 5.2553 | 1429.8284 | 5.2377 | 1427.7048 | 5.2377 | 1419.2957
| 5.431 | 115.8386 | 5.3256 | 113.854 | 5.3959 | 113.2309 | 5.3959 | 110.8913 | 5.3431 | 112.2038 | 5.4134 | 1440.1224 | 5.3256 | 1430.3016 | 5.3256 | 1430.1833 | 5.2904 | 1427.4694 | 5.2904 | 1429.9467 | 5.2728 | 1428.2937 | 5.2728 | 1420.5726
| 5.4662 | 116.4784 | 5.3607 | 114.1351 | 5.431 | 114.115 | 5.431 | 111.134 | 5.3783 | 111.5784 | 5.4486 | 1441.7025 | 5.3607 | 1430.8939 | 5.3607 | 1431.3682 | 5.3256 | 1427.9403 | 5.3256 | 1429.9467 | 5.308 | 1427.8225 | 5.308 | 1421.0376
| 5.5013 | 117.0377 | 5.3959 | 114.8573 | 5.4662 | 114.717 | 5.4662 | 111.6591 | 5.4134 | 111.5582 | 5.4837 | 1444.8843 | 5.3959 | 1433.1517 | 5.3959 | 1431.7243 | 5.3607 | 1428.7654 | 5.3607 | 1429.9467 | 5.3431 | 1428.7654 | 5.3431 | 1422.901
| 5.5365 | 117.0377 | 5.431 | 116.2586 | 5.5013 | 115.4583 | 5.5013 | 112.2845 | 5.4486 | 111.5582 | 5.5189 | 1445.6228 | 5.431 | 1435.4216 | 5.431 | 1431.7243 | 5.3959 | 1430.1833 | 5.3959 | 1429.9467 | 5.3783 | 1429.2376 | 5.3783 | 1424.0686
| 5.5716 | 116.4185 | 5.4662 | 116.5384 | 5.5365 | 115.6985 | 5.5365 | 112.2845 | 5.4837 | 112.103 | 5.554 | 1449.7142 | 5.4662 | 1438.3075 | 5.4662 | 1435.1821 | 5.431 | 1430.3016 | 5.431 | 1429.9467 | 5.4134 | 1430.3016 | 5.4134 | 1424.7704
| 5.6068 | 115.9786 | 5.5013 | 117.0577 | 5.5716 | 115.7385 | 5.5716 | 113.1705 | 5.5189 | 112.3046 | 5.5892 | 1452.0939 | 5.5013 | 1440.4864 | 5.5013 | 1437.2228 | 5.4662 | 1430.42 | 5.4662 | 1432.3185 | 5.4486 | 1431.2496 | 5.4486 | 1424.7704
| 5.6419 | 115.4183 | 5.5365 | 117.0976 | 5.6068 | 116.2386 | 5.6068 | 113.7937 | 5.554 | 112.9894 | 5.6244 | 1453.7327 | 5.5365 | 1442.6785 | 5.5365 | 1440.365 | 5.5013 | 1434.9427 | 5.5013 | 1436.7417 | 5.4837 | 1432.3185 | 5.4837 | 1426.4116
| 5.6771 | 115.4183 | 5.5716 | 116.838 | 5.6419 | 116.6383 | 5.6419 | 113.854 | 5.5892 | 113.3917 | 5.6595 | 1454.9994 | 5.5716 | 1443.7796 | 5.5716 | 1441.5807 | 5.5365 | 1437.3432 | 5.5365 | 1440.0011 | 5.5189 | 1434.7035 | 5.5189 | 1429.0014
| 5.7122 | 115.4583 | 5.6068 | 117.2772 | 5.6771 | 116.6183 | 5.6771 | 113.8942 | 5.6244 | 113.3917 | 5.6947 | 1456.2714 | 5.6068 | 1444.0248 | 5.6068 | 1443.4122 | 5.5716 | 1440.365 | 5.5716 | 1442.0682 | 5.554 | 1435.7812 | 5.554 | 1432.4374
| 5.7474 | 115.8786 | 5.6419 | 117.7759 | 5.7122 | 116.8979 | 5.7122 | 114.0347 | 5.6595 | 114.3559 | 5.7298 | 1456.5265 | 5.6419 | 1445.3764 | 5.6419 | 1446.61 | 5.6068 | 1442.4342 | 5.6068 | 1442.8006 | 5.5892 | 1436.5013 | 5.5892 | 1433.8672
| 5.7825 | 115.8786 | 5.6771 | 117.6762 | 5.7474 | 116.4984 | 5.7474 | 113.854 | 5.6947 | 115.0577 | 5.765 | 1457.1652 | 5.6771 | 1445.3764 | 5.6771 | 1447.6001 | 5.6419 | 1444.7614 | 5.6419 | 1443.7796 | 5.6244 | 1437.9456 | 5.6244 | 1435.1821
| 5.8177 | 116.2786 | 5.7122 | 118.1547 | 5.7825 | 116.9179 | 5.7825 | 114.0548 | 5.7298 | 115.0777 | 5.8001 | 1460.8973 | 5.7122 | 1449.4648 | 5.7122 | 1447.6001 | 5.6771 | 1443.6571 | 5.6771 | 1446.8572 | 5.6595 | 1439.8799 | 5.6595 | 1435.1821
| 5.8528 | 116.2586 | 5.7474 | 118.3738 | 5.8177 | 116.6183 | 5.8177 | 114.1552 | 5.765 | 115.9986 | 5.8353 | 1462.847 | 5.7474 | 1449.4648 | 5.7474 | 1448.2204 | 5.7122 | 1443.2898 | 5.7122 | 1446.8572 | 5.6947 | 1443.2898 | 5.6947 | 1435.4216
| 5.888 | 116.5584 | 5.7825 | 119.2696 | 5.8528 | 116.4585 | 5.8528 | 115.4383 | 5.8001 | 116.2986 | 5.8704 | 1462.847 | 5.7825 | 1449.5895 | 5.7825 | 1448.9665 | 5.7474 | 1445.9926 | 5.7474 | 1447.9722 | 5.7298 | 1443.6571 | 5.7298 | 1436.5013
| 5.9231 | 116.7981 | 5.8177 | 119.2696 | 5.888 | 116.9778 | 5.888 | 115.9186 | 5.8353 | 116.2986 | 5.9056 | 1462.847 | 5.8177 | 1451.5914 | 5.8177 | 1452.8492 | 5.7825 | 1445.0073 | 5.7825 | 1448.3447 | 5.765 | 1445.6228 | 5.765 | 1436.9822
| 5.9583 | 117.6164 | 5.8528 | 119.389 | 5.9231 | 116.9778 | 5.9231 | 116.1786 | 5.8704 | 116.3785 | 5.9407 | 1461.8055 | 5.8528 | 1453.2275 | 5.8528 | 1454.1121 | 5.8177 | 1448.842 | 5.8177 | 1449.2155 | 5.8001 | 1446.61 | 5.8001 | 1436.9822
| 5.9935 | 117.6164 | 5.888 | 119.2099 | 5.9583 | 116.4585 | 5.9583 | 116.2186 | 5.9056 | 117.1774 | 5.9759 | 1461.5457 | 5.888 | 1453.48 | 5.888 | 1452.8492 | 5.8528 | 1454.2387 | 5.8528 | 1453.2275 | 5.8353 | 1447.2284 | 5.8353 | 1437.7045
| 6.0286 | 117.8158 | 5.9231 | 118.7522 | 5.9935 | 116.5384 | 5.9935 | 116.2186 | 5.9407 | 117.5565 | 6.011 | 1456.9095 | 5.9231 | 1454.6188 | 5.9231 | 1453.48 | 5.888 | 1454.2387 | 5.888 | 1449.5895 | 5.8704 | 1448.3447 | 5.8704 | 1439.7587
| 6.0638 | 117.8158 | 5.9583 | 118.7124 | 6.0286 | 116.7981 | 6.0286 | 115.9386 | 5.9759 | 117.5565 | 6.0462 | 1455.8893 | 5.9583 | 1454.2387 | 5.9583 | 1452.3454 | 5.9231 | 1454.2387 | 5.9231 | 1451.4659 | 5.9056 | 1451.4659 | 5.9056 | 1441.8244
| 6.0989 | 116.3585 | 5.9935 | 118.792 | 6.0638 | 116.6583 | 6.0638 | 115.7986 | 6.011 | 118.314 | 6.0813 | 1454.8725 | 5.9935 | 1453.48 | 5.9935 | 1449.2155 | 5.9583 | 1453.8591 | 5.9583 | 1451.4659 | 5.9407 | 1453.3537 | 5.9407 | 1446.2394
| 6.1341 | 116.3585 | 6.0286 | 118.8318 | 6.0989 | 116.9379 | 6.0989 | 115.7986 | 6.0462 | 117.5565 | 6.1165 | 1454.1121 | 6.0286 | 1452.9752 | 6.0286 | 1449.7142 | 5.9935 | 1453.6063 | 5.9935 | 1450.2137 | 5.9759 | 1451.8425 | 5.9759 | 1446.2394
| 6.1692 | 116.3585 | 6.0638 | 118.8517 | 6.1341 | 117.2573 | 6.1341 | 115.9986 | 6.0813 | 117.2772 | 6.1516 | 1456.399 | 6.0638 | 1451.0897 | 6.0638 | 1448.9665 | 6.0286 | 1450.2137 | 6.0286 | 1448.7176 | 6.011 | 1451.4659 | 6.011 | 1446.61
| 6.2044 | 116.5983 | 6.0989 | 118.8517 | 6.1692 | 117.5765 | 6.1692 | 116.0386 | 6.1165 | 116.5384 | 6.1868 | 1457.6771 | 6.0989 | 1451.0897 | 6.0989 | 1449.2155 | 6.0638 | 1450.2137 | 6.0638 | 1448.7176 | 6.0462 | 1449.9639 | 6.0462 | 1447.6001
| 6.2395 | 117.6762 | 6.1341 | 118.8517 | 6.2044 | 118.3539 | 6.2044 | 116.0986 | 6.1516 | 116.0386 | 6.2219 | 1461.9355 | 6.1341 | 1452.8492 | 6.1341 | 1449.839 | 6.0989 | 1447.4762 | 6.0989 | 1448.0963 | 6.0813 | 1444.7614 | 6.0813 | 1445.8693
| 6.2747 | 118.573 | 6.1692 | 119.0707 | 6.2395 | 118.4933 | 6.2395 | 116.0986 | 6.1868 | 116.0386 | 6.2571 | 1463.8924 | 6.1692 | 1453.7327 | 6.1692 | 1451.9682 | 6.1341 | 1447.3523 | 6.1341 | 1448.7176 | 6.1165 | 1444.7614 | 6.1165 | 1444.7614
| 6.3098 | 120.045 | 6.2044 | 119.6674 | 6.2747 | 118.573 | 6.2747 | 115.8186 | 6.2219 | 116.1986 | 6.2922 | 1465.0733 | 6.2044 | 1454.4921 | 6.2044 | 1453.6063 | 6.1692 | 1447.3523 | 6.1692 | 1451.0897 | 6.1516 | 1444.8843 | 6.1516 | 1445.0073
| 6.345 | 120.8197 | 6.2395 | 119.9457 | 6.3098 | 118.7721 | 6.3098 | 115.8186 | 6.2571 | 116.4185 | 6.3274 | 1464.6791 | 6.2395 | 1456.7818 | 6.2395 | 1457.0373 | 6.2044 | 1450.4637 | 6.2044 | 1453.2275 | 6.1868 | 1446.4864 | 6.1868 | 1444.1475
| 6.3801 | 121.0182 | 6.2747 | 120.1444 | 6.345 | 118.7721 | 6.345 | 116.4185 | 6.2922 | 117.2174 | 6.3625 | 1464.2855 | 6.2747 | 1457.9334 | 6.2747 | 1457.8052 | 6.2395 | 1451.3404 | 6.2395 | 1454.8725 | 6.2219 | 1448.2204 | 6.2219 | 1444.1475
| 6.4153 | 121.2761 | 6.3098 | 120.0848 | 6.3801 | 118.0151 | 6.3801 | 116.9778 | 6.3274 | 117.4568 | 6.3977 | 1464.0234 | 6.3098 | 1458.0616 | 6.3098 | 1458.1899 | 6.2747 | 1455.8893 | 6.2747 | 1455.6347 | 6.2571 | 1452.0939 | 6.2571 | 1444.6385
| 6.4504 | 121.772 | 6.345 | 119.7071 | 6.4153 | 117.9553 | 6.4153 | 117.1375 | 6.3625 | 117.397 | 6.4329 | 1464.9418 | 6.345 | 1459.3471 | 6.345 | 1457.421 | 6.3098 | 1457.1652 | 6.3098 | 1456.0166 | 6.2922 | 1454.7456 | 6.2922 | 1445.9926
| 6.4856 | 121.9703 | 6.3801 | 119.6873 | 6.4504 | 117.9553 | 6.4504 | 117.1375 | 6.3977 | 117.6363 | 6.468 | 1464.6791 | 6.3801 | 1460.3796 | 6.3801 | 1457.9334 | 6.345 | 1458.5751 | 6.345 | 1456.0166 | 6.3274 | 1455.1263 | 6.3274 | 1447.7241
| 6.5207 | 121.8315 | 6.4153 | 119.7071 | 6.4856 | 118.4535 | 6.4856 | 117.2174 | 6.4329 | 117.8357 | 6.5032 | 1464.9418 | 6.4153 | 1459.3471 | 6.4153 | 1459.0895 | 6.3801 | 1458.5751 | 6.3801 | 1456.144 | 6.3625 | 1456.0166 | 6.3625 | 1450.4637
| 6.5559 | 121.772 | 6.4504 | 120.0053 | 6.5207 | 119.1701 | 6.5207 | 116.5783 | 6.468 | 118.3937 | 6.5383 | 1466.7879 | 6.4504 | 1458.7036 | 6.4504 | 1459.3471 | 6.4153 | 1460.1211 | 6.4153 | 1455.6347 | 6.3977 | 1456.0166 | 6.3977 | 1451.3404
| 6.591 | 121.7125 | 6.4856 | 120.7403 | 6.5559 | 119.8463 | 6.5559 | 116.5983 | 6.5032 | 118.3937 | 6.5735 | 1467.8485 | 6.4856 | 1461.1565 | 6.4856 | 1460.2503 | 6.4504 | 1460.7678 | 6.4504 | 1456.144 | 6.4329 | 1455.2534 | 6.4329 | 1452.4713
| 6.6262 | 120.9586 | 6.5207 | 122.0496 | 6.591 | 121.5737 | 6.591 | 116.5783 | 6.5383 | 118.4933 | 6.6086 | 1467.8485 | 6.5207 | 1461.5457 | 6.5207 | 1460.2503 | 6.4856 | 1458.3182 | 6.4856 | 1458.4467 | 6.468 | 1456.7818 | 6.468 | 1454.3654
| 6.6613 | 120.8197 | 6.5559 | 122.4064 | 6.6262 | 122.1884 | 6.6262 | 117.0577 | 6.5735 | 118.4933 | 6.6438 | 1468.6466 | 6.5559 | 1461.5457 | 6.5559 | 1463.2386 | 6.5207 | 1458.3182 | 6.5207 | 1458.7036 | 6.5032 | 1457.6771 | 6.5032 | 1452.9752
| 6.6965 | 120.9785 | 6.591 | 123.852 | 6.6613 | 123.0007 | 6.6613 | 117.7161 | 6.6086 | 118.4734 | 6.6789 | 1469.18 | 6.591 | 1464.5478 | 6.591 | 1463.2386 | 6.5559 | 1458.9608 | 6.5559 | 1459.3471 | 6.5383 | 1459.2183 | 6.5383 | 1452.9752
| 6.7316 | 121.3952 | 6.6262 | 124.2082 | 6.6965 | 122.763 | 6.6965 | 118.7124 | 6.6438 | 118.2144 | 6.7141 | 1470.3841 | 6.6262 | 1466.2592 | 6.6262 | 1463.2386 | 6.591 | 1458.3182 | 6.591 | 1458.9608 | 6.5735 | 1459.4759 | 6.5735 | 1454.4921
| 6.7668 | 121.8117 | 6.6613 | 124.4851 | 6.7316 | 123.0403 | 6.7316 | 121.5142 | 6.6789 | 118.4734 | 6.7492 | 1474.1653 | 6.6613 | 1466.7879 | 6.6613 | 1463.8924 | 6.6262 | 1457.421 | 6.6262 | 1460.5089 | 6.6086 | 1461.6755 | 6.6086 | 1455.2534
| 6.802 | 122.1686 | 6.6965 | 124.6434 | 6.7668 | 122.8819 | 6.7668 | 121.8712 | 6.7141 | 118.7721 | 6.7844 | 1479.1084 | 6.6965 | 1467.0527 | 6.6965 | 1466.7879 | 6.6613 | 1460.6383 | 6.6613 | 1461.2861 | 6.6438 | 1461.6755 | 6.6438 | 1459.3471
| 6.8371 | 123.0601 | 6.7316 | 124.4851 | 6.802 | 122.763 | 6.802 | 122.1289 | 6.7492 | 119.9854 | 6.8195 | 1482.7383 | 6.7316 | 1471.8631 | 6.7316 | 1467.4503 | 6.6965 | 1465.7315 | 6.6965 | 1464.1544 | 6.6789 | 1461.6755 | 6.6789 | 1460.3796
| 6.8723 | 124.2675 | 6.7668 | 124.3071 | 6.8371 | 122.1884 | 6.8371 | 121.8712 | 6.7844 | 121.891 | 6.8547 | 1483.0196 | 6.7668 | 1473.3505 | 6.7668 | 1469.8483 | 6.7316 | 1470.1161 | 6.7316 | 1467.1851 | 6.7141 | 1462.0655 | 6.7141 | 1461.8055
| 6.9074 | 125.9679 | 6.802 | 124.2082 | 6.8723 | 122.0496 | 6.8723 | 122.0893 | 6.8195 | 122.6045 | 6.8898 | 1484.2896 | 6.802 | 1472.944 | 6.802 | 1470.25 | 6.7668 | 1471.5936 | 6.7668 | 1470.3841 | 6.7492 | 1464.9418 | 6.7492 | 1461.8055
| 6.9426 | 125.9679 | 6.8371 | 123.9114 | 6.9074 | 121.5539 | 6.9074 | 121.7919 | 6.8547 | 123.0007 | 6.925 | 1484.7144 | 6.8371 | 1474.3014 | 6.8371 | 1471.4589 | 6.802 | 1471.8631 | 6.802 | 1470.1161 | 6.7844 | 1466.3913 | 6.7844 | 1461.8055
| 6.9777 | 125.9679 | 6.8723 | 123.5551 | 6.9426 | 121.5539 | 6.9426 | 121.6729 | 6.8898 | 122.763 | 6.9601 | 1488.7135 | 6.8723 | 1475.9395 | 6.8723 | 1476.2135 | 6.8371 | 1472.403 | 6.8371 | 1470.3841 | 6.8195 | 1467.7157 | 6.8195 | 1463.108
| 7.0129 | 124.8016 | 6.9074 | 123.3373 | 6.9777 | 122.0496 | 6.9777 | 121.1372 | 6.925 | 123.1987 | 6.9953 | 1490.1572 | 6.9074 | 1475.9395 | 6.9074 | 1479.1084 | 6.8723 | 1471.7283 | 6.8723 | 1472.6734 | 6.8547 | 1468.7799 | 6.8547 | 1466.9203
| 7.048 | 122.1289 | 6.9426 | 123.4165 | 7.0129 | 122.0694 | 7.0129 | 120.7403 | 6.9601 | 123.2977 | 7.0304 | 1491.0275 | 6.9426 | 1480.3591 | 6.9426 | 1482.7383 | 6.9074 | 1476.3507 | 6.9074 | 1472.944 | 6.8898 | 1469.5808 | 6.8898 | 1467.9813
| 7.0832 | 120.8594 | 6.9777 | 123.5155 | 7.048 | 122.8423 | 7.048 | 120.7403 | 6.9953 | 123.2977 | 7.0656 | 1491.0275 | 6.9777 | 1480.9169 | 6.9777 | 1483.7244 | 6.9426 | 1476.6251 | 6.9426 | 1474.9827 | 6.925 | 1471.0554 | 6.925 | 1469.18
| 7.1183 | 120.2636 | 7.0129 | 124.0499 | 7.0832 | 123.3571 | 7.0832 | 120.7403 | 7.0304 | 123.0997 | 7.1007 | 1490.592 | 7.0129 | 1480.7774 | 7.0129 | 1483.5833 | 6.9777 | 1477.8638 | 6.9777 | 1478.8313 | 6.9601 | 1471.7283 | 6.9601 | 1471.0554
| 7.1535 | 120.3629 | 7.048 | 124.6236 | 7.1183 | 124.4851 | 7.1183 | 121.4547 | 7.0656 | 123.1987 | 7.1359 | 1490.0125 | 7.048 | 1481.1963 | 7.048 | 1483.0196 | 7.0129 | 1478.0018 | 7.0129 | 1480.7774 | 6.9953 | 1472.1329 | 6.9953 | 1473.2149
| 7.1886 | 120.9189 | 7.0832 | 125.9284 | 7.1535 | 125.6517 | 7.1535 | 121.9703 | 7.1007 | 122.228 | 7.171 | 1489.4343 | 7.0832 | 1482.7383 | 7.0832 | 1483.0196 | 7.048 | 1478.4162 | 7.048 | 1480.7774 | 7.0304 | 1474.0293 | 7.0304 | 1473.6218
| 7.2238 | 121.4547 | 7.1183 | 127.173 | 7.1886 | 127.1532 | 7.1886 | 122.5847 | 7.1359 | 122.2082 | 7.2062 | 1487.9947 | 7.1183 | 1482.7383 | 7.1183 | 1482.5977 | 7.0832 | 1479.1084 | 7.0832 | 1480.7774 | 7.0656 | 1475.1192 | 7.0656 | 1474.8463
| 7.2589 | 121.4547 | 7.1535 | 127.9629 | 7.2238 | 127.1532 | 7.2238 | 123.4759 | 7.171 | 122.3469 | 7.2414 | 1487.9947 | 7.1535 | 1483.3013 | 7.1535 | 1480.9169 | 7.1183 | 1480.0807 | 7.1183 | 1479.6635 | 7.1007 | 1476.6251 | 7.1007 | 1475.8026
| 7.2941 | 121.6927 | 7.1886 | 128.8316 | 7.2589 | 127.1532 | 7.2589 | 124.4258 | 7.2062 | 122.5055 | 7.2765 | 1487.8512 | 7.1886 | 1485.8505 | 7.1886 | 1479.9415 | 7.1535 | 1478.9698 | 7.1535 | 1479.5247 | 7.1359 | 1475.529 | 7.1359 | 1474.8463
| 7.3292 | 120.9983 | 7.2238 | 128.8316 | 7.2941 | 126.2445 | 7.2941 | 125.8889 | 7.2414 | 123.0007 | 7.3117 | 1490.447 | 7.2238 | 1484.2896 | 7.2238 | 1478.6929 | 7.1886 | 1479.6635 | 7.1886 | 1478.278 | 7.171 | 1475.529 | 7.171 | 1475.3923
| 7.3644 | 120.502 | 7.2589 | 128.654 | 7.3292 | 123.852 | 7.3292 | 126.4224 | 7.2765 | 123.7135 | 7.3468 | 1491.1728 | 7.2589 | 1483.3013 | 7.2589 | 1478.5545 | 7.2238 | 1480.0807 | 7.2238 | 1478.1399 | 7.2062 | 1475.9395 | 7.2062 | 1475.529
| 7.3995 | 120.2835 | 7.2941 | 125.8691 | 7.3644 | 123.0205 | 7.3644 | 126.4224 | 7.3117 | 124.3071 | 7.382 | 1495.8686 | 7.2941 | 1486.1354 | 7.2941 | 1478.6929 | 7.2589 | 1478.8313 | 7.2589 | 1478.278 | 7.2414 | 1477.3125 | 7.2414 | 1475.2557
| 7.4347 | 120.3033 | 7.3292 | 124.4258 | 7.3995 | 121.6134 | 7.3995 | 126.2248 | 7.3468 | 126.2248 | 7.4171 | 1498.1006 | 7.3292 | 1488.4257 | 7.3292 | 1481.1963 | 7.2941 | 1478.8313 | 7.2941 | 1481.476 | 7.2765 | 1477.3125 | 7.2765 | 1475.2557
| 7.4698 | 120.1444 | 7.3644 | 124.2082 | 7.4347 | 121.6134 | 7.4347 | 123.7926 | 7.382 | 126.9162 | 7.4523 | 1500.5033 | 7.3644 | 1492.4846 | 7.3644 | 1488.8575 | 7.3292 | 1482.5977 | 7.3292 | 1485.2818 | 7.3117 | 1478.0018 | 7.3117 | 1476.6251
| 7.505 | 120.2835 | 7.3995 | 124.1884 | 7.4698 | 121.8315 | 7.4698 | 121.6531 | 7.4171 | 127.8247 | 7.4874 | 1499.5997 | 7.3995 | 1496.9077 | 7.3995 | 1489.4343 | 7.3644 | 1487.2779 | 7.3644 | 1491.1728 | 7.3468 | 1480.9169 | 7.3468 | 1478.4162
| 7.5401 | 120.3828 | 7.4347 | 124.4456 | 7.505 | 121.8315 | 7.505 | 119.8264 | 7.4523 | 127.8247 | 7.5226 | 1499.2992 | 7.4347 | 1498.2501 | 7.4347 | 1490.3021 | 7.3995 | 1491.4637 | 7.3995 | 1494.8339 | 7.382 | 1483.7244 | 7.382 | 1479.2471
| 7.5753 | 120.3033 | 7.4698 | 124.94 | 7.5401 | 122.01 | 7.5401 | 119.8264 | 7.4874 | 127.6075 | 7.5577 | 1497.0565 | 7.4698 | 1498.5494 | 7.4698 | 1489.29 | 7.4347 | 1493.5097 | 7.4347 | 1496.3134 | 7.4171 | 1486.8489 | 7.4171 | 1481.3361
| 7.6105 | 120.4424 | 7.505 | 125.0389 | 7.5753 | 122.6045 | 7.5753 | 119.9059 | 7.5226 | 124.3862 | 7.5929 | 1499.2992 | 7.505 | 1498.5494 | 7.505 | 1488.8575 | 7.4698 | 1494.539 | 7.4698 | 1496.1651 | 7.4523 | 1490.7371 | 7.4523 | 1482.5977
| 7.6456 | 121.3158 | 7.5401 | 124.7225 | 7.6105 | 122.3469 | 7.6105 | 121.2761 | 7.5577 | 123.2779 | 7.628 | 1496.9077 | 7.5401 | 1496.4619 | 7.5401 | 1486.4205 | 7.505 | 1494.3917 | 7.505 | 1494.8339 | 7.4874 | 1494.3917 | 7.4874 | 1482.8789
| 7.6808 | 121.3555 | 7.5753 | 124.94 | 7.6456 | 121.8712 | 7.6456 | 121.2761 | 7.5929 | 121.8514 | 7.6632 | 1496.6104 | 7.5753 | 1492.3385 | 7.5753 | 1486.2779 | 7.5401 | 1494.3917 | 7.5401 | 1493.5097 | 7.5226 | 1494.0974 | 7.5226 | 1482.5977
| 7.7159 | 121.3555 | 7.6105 | 124.3269 | 7.6808 | 121.9108 | 7.6808 | 122.3073 | 7.628 | 121.8514 | 7.6983 | 1494.2445 | 7.6105 | 1485.566 | 7.6105 | 1486.1354 | 7.5753 | 1491.1728 | 7.5753 | 1492.6308 | 7.5577 | 1492.9234 | 7.5577 | 1482.5977
| 7.7511 | 121.9108 | 7.6456 | 124.1686 | 7.7159 | 121.8315 | 7.7159 | 122.1091 | 7.6632 | 121.9307 | 7.7335 | 1494.2445 | 7.6456 | 1490.592 | 7.6456 | 1483.8656 | 7.6105 | 1489.29 | 7.6105 | 1488.282 | 7.5929 | 1490.8822 | 7.5929 | 1479.5247
| 7.7862 | 121.9108 | 7.6808 | 123.8916 | 7.7511 | 121.4547 | 7.7511 | 121.2761 | 7.6983 | 122.9413 | 7.7686 | 1491.9008 | 7.6808 | 1489.5787 | 7.6808 | 1482.5977 | 7.6456 | 1486.2779 | 7.6456 | 1486.1354 | 7.628 | 1489.29 | 7.628 | 1476.4879
| 7.8214 | 121.3158 | 7.7159 | 123.5353 | 7.7862 | 121.6531 | 7.7862 | 121.6927 | 7.7335 | 122.9413 | 7.8038 | 1493.9503 | 7.7159 | 1487.1348 | 7.7159 | 1482.5977 | 7.6808 | 1486.2779 | 7.6808 | 1483.7244 | 7.6632 | 1487.8512 | 7.6632 | 1476.0765
| 7.8565 | 121.0579 | 7.7511 | 123.6343 | 7.8214 | 122.1487 | 7.8214 | 121.772 | 7.7686 | 123.0007 | 7.8389 | 1494.0974 | 7.7511 | 1484.0069 | 7.7511 | 1480.2198 | 7.7159 | 1485.7082 | 7.7159 | 1483.0196 | 7.6983 | 1485.2818 | 7.6983 | 1476.0765
| 7.8917 | 121.0579 | 7.7862 | 124.1686 | 7.8565 | 122.5847 | 7.8565 | 120.9785 | 7.8038 | 123.0799 | 7.8741 | 1495.4246 | 7.7862 | 1485.1399 | 7.7862 | 1480.2198 | 7.7511 | 1483.3013 | 7.7511 | 1481.1963 | 7.7335 | 1483.8656 | 7.7335 | 1478.8313
| 7.9268 | 121.3555 | 7.8214 | 124.2675 | 7.8917 | 123.0997 | 7.8917 | 121.415 | 7.8389 | 122.7828 | 7.9092 | 1496.9077 | 7.8214 | 1485.1399 | 7.8214 | 1484.4311 | 7.7862 | 1482.0362 | 7.7862 | 1480.0807 | 7.7686 | 1483.4423 | 7.7686 | 1479.2471
| 7.962 | 121.6134 | 7.8565 | 124.2675 | 7.9268 | 123.3373 | 7.9268 | 121.6531 | 7.8741 | 122.5451 | 7.9444 | 1496.9077 | 7.8565 | 1485.4239 | 7.8565 | 1490.8822 | 7.8214 | 1482.8789 | 7.8214 | 1480.3591 | 7.8038 | 1481.896 | 7.8038 | 1479.2471
| 7.9971 | 122.8819 | 7.8917 | 124.7225 | 7.962 | 123.753 | 7.962 | 121.5142 | 7.9092 | 122.8423 | 7.9795 | 1498.999 | 7.8917 | 1489.1457 | 7.8917 | 1491.3182 | 7.8565 | 1482.1765 | 7.8565 | 1481.7559 | 7.8389 | 1482.1765 | 7.8389 | 1481.7559
| 8.0323 | 122.8819 | 7.9268 | 125.0784 | 7.9971 | 123.7926 | 7.9971 | 121.8117 | 7.9444 | 122.8423 | 8.0147 | 1498.999 | 7.9268 | 1490.8822 | 7.9268 | 1492.3385 | 7.8917 | 1484.7144 | 7.8917 | 1481.7559 | 7.8741 | 1483.8656 | 7.8741 | 1482.0362
| 8.0674 | 122.8819 | 7.962 | 125.0784 | 8.0323 | 123.7926 | 8.0323 | 122.7036 | 7.9795 | 122.6045 | 8.0499 | 1495.5726 | 7.962 | 1489.4343 | 7.962 | 1490.8822 | 7.9268 | 1484.8561 | 7.9268 | 1482.5977 | 7.9092 | 1483.8656 | 7.9092 | 1481.7559
| 8.1026 | 123.0799 | 7.9971 | 125.3947 | 8.0674 | 123.4957 | 8.0674 | 123.0601 | 8.0147 | 122.8423 | 8.085 | 1495.5726 | 7.9971 | 1488.8575 | 7.9971 | 1484.7144 | 7.962 | 1489.29 | 7.962 | 1482.5977 | 7.9444 | 1484.0069 | 7.9444 | 1484.4311
| 8.1377 | 123.3967 | 8.0323 | 125.6122 | 8.1026 | 123.0799 | 8.1026 | 123.1393 | 8.0499 | 122.6838 | 8.1202 | 1498.3997 | 8.0323 | 1489.4343 | 8.0323 | 1484.5727 | 7.9971 | 1489.29 | 7.9971 | 1482.3168 | 7.9795 | 1484.7144 | 7.9795 | 1485.4239
| 8.1729 | 123.7926 | 8.0674 | 125.3947 | 8.1377 | 122.664 | 8.1377 | 123.1393 | 8.085 | 122.6045 | 8.1553 | 1502.3165 | 8.0674 | 1492.9234 | 8.0674 | 1485.7082 | 8.0323 | 1488.7135 | 8.0323 | 1482.1765 | 8.0147 | 1485.1399 | 8.0147 | 1483.1604
| 8.208 | 123.852 | 8.1026 | 124.8411 | 8.1729 | 122.664 | 8.1729 | 123.0007 | 8.1202 | 122.9809 | 8.1905 | 1504.5935 | 8.1026 | 1493.6565 | 8.1026 | 1491.755 | 8.0674 | 1489.29 | 8.0674 | 1483.3013 | 8.0499 | 1483.7244 | 8.0499 | 1482.8789
| 8.2432 | 124.2082 | 8.1377 | 124.7225 | 8.208 | 123.1789 | 8.208 | 122.1686 | 8.1553 | 123.3571 | 8.2256 | 1509.3368 | 8.1377 | 1495.4246 | 8.1377 | 1495.2768 | 8.1026 | 1490.1572 | 8.1026 | 1488.8575 | 8.085 | 1483.4423 | 8.085 | 1484.8561
| 8.2783 | 124.228 | 8.1729 | 125.0191 | 8.2432 | 123.951 | 8.2432 | 121.3952 | 8.1905 | 123.6145 | 8.2608 | 1511.1864 | 8.1729 | 1498.6991 | 8.1729 | 1501.409 | 8.1377 | 1493.8034 | 8.1377 | 1491.4637 | 8.1202 | 1483.4423 | 8.1202 | 1484.998
| 8.3135 | 124.3269 | 8.208 | 125.2761 | 8.2783 | 124.129 | 8.2783 | 121.038 | 8.2256 | 123.8916 | 8.2959 | 1515.5314 | 8.208 | 1502.6194 | 8.208 | 1502.6194 | 8.1729 | 1495.8686 | 8.1729 | 1499.2992 | 8.1553 | 1485.7082 | 8.1553 | 1485.8505
| 8.3486 | 124.3862 | 8.2432 | 126.126 | 8.3135 | 124.129 | 8.3135 | 121.1968 | 8.2608 | 124.0103 | 8.3311 | 1516.6241 | 8.2432 | 1506.7294 | 8.2432 | 1504.2892 | 8.208 | 1501.8625 | 8.208 | 1501.2579 | 8.1905 | 1489.5787 | 8.1905 | 1489.8678
| 8.3838 | 124.4456 | 8.2783 | 126.126 | 8.3486 | 123.951 | 8.3486 | 121.9307 | 8.2959 | 123.7333 | 8.3662 | 1517.4061 | 8.2783 | 1508.4148 | 8.2783 | 1508.1079 | 8.2432 | 1502.3165 | 8.2432 | 1502.4679 | 8.2256 | 1497.6526 | 8.2256 | 1494.2445
| 8.419 | 124.2082 | 8.3135 | 126.2445 | 8.3838 | 123.951 | 8.3838 | 122.9809 | 8.3311 | 122.8819 | 8.4014 | 1519.2883 | 8.3135 | 1509.7985 | 8.3135 | 1508.1079 | 8.2783 | 1504.5935 | 8.2783 | 1505.5077 | 8.2608 | 1499.7501 | 8.2608 | 1495.7206
| 8.4541 | 123.6541 | 8.3486 | 126.2643 | 8.419 | 123.951 | 8.419 | 123.0205 | 8.3662 | 122.8026 | 8.4365 | 1523.2336 | 8.3486 | 1510.2607 | 8.3486 | 1508.8756 | 8.3135 | 1506.2709 | 8.3135 | 1505.8128 | 8.2959 | 1501.5601 | 8.2959 | 1497.8018
| 8.4893 | 123.8916 | 8.3838 | 126.2841 | 8.4541 | 125.197 | 8.4541 | 124.228 | 8.4014 | 123.0997 | 8.4717 | 1521.0202 | 8.3838 | 1510.4148 | 8.3838 | 1508.8756 | 8.3486 | 1508.4148 | 8.3486 | 1508.5684 | 8.3311 | 1503.6813 | 8.3311 | 1499.2992
| 8.5244 | 124.1884 | 8.419 | 127.4495 | 8.4893 | 126.4619 | 8.4893 | 124.8411 | 8.4365 | 123.3769 | 8.5068 | 1519.2883 | 8.419 | 1511.65 | 8.419 | 1508.7219 | 8.3838 | 1510.4148 | 8.3838 | 1509.3368 | 8.3662 | 1503.9852 | 8.3662 | 1500.0512
| 8.5596 | 123.3373 | 8.4541 | 127.9629 | 8.5244 | 126.6594 | 8.5244 | 124.8411 | 8.4717 | 124.0697 | 8.542 | 1518.817 | 8.4541 | 1511.8046 | 8.4541 | 1506.1181 | 8.419 | 1513.0435 | 8.419 | 1509.3368 | 8.4014 | 1503.9852 | 8.4014 | 1499.2992
| 8.5947 | 123.6541 | 8.4893 | 128.1999 | 8.5596 | 127.252 | 8.5596 | 124.8609 | 8.5068 | 124.1488 | 8.5771 | 1515.9994 | 8.4893 | 1510.2607 | 8.4893 | 1503.8332 | 8.4541 | 1515.8433 | 8.4541 | 1511.8046 | 8.4365 | 1508.8756 | 8.4365 | 1505.0504
| 8.6299 | 123.5947 | 8.5244 | 128.6342 | 8.5947 | 127.252 | 8.5947 | 125.9877 | 8.542 | 124.2477 | 8.6123 | 1511.65 | 8.5244 | 1508.2614 | 8.5244 | 1500.6541 | 8.4893 | 1513.8196 | 8.4893 | 1512.8885 | 8.4717 | 1509.7985 | 8.4717 | 1507.648
| 8.665 | 123.5947 | 8.5596 | 129.187 | 8.6299 | 127.2915 | 8.6299 | 126.8767 | 8.5771 | 124.5445 | 8.6474 | 1510.4148 | 8.5596 | 1508.2614 | 8.5596 | 1496.1651 | 8.5244 | 1506.5765 | 8.5244 | 1507.1885 | 8.5068 | 1509.7985 | 8.5068 | 1507.0354
| 8.7002 | 123.5947 | 8.5947 | 129.3449 | 8.665 | 127.3705 | 8.665 | 127.173 | 8.6123 | 124.6236 | 8.6826 | 1511.65 | 8.5947 | 1502.4679 | 8.5947 | 1495.8686 | 8.5596 | 1506.5765 | 8.5596 | 1505.9654 | 8.542 | 1504.1372 | 8.542 | 1503.8332
| 8.7353 | 124.8214 | 8.6299 | 129.187 | 8.7002 | 127.3902 | 8.7002 | 127.7062 | 8.6474 | 124.6236 | 8.7177 | 1510.8776 | 8.6299 | 1501.8625 | 8.6299 | 1496.1651 | 8.5947 | 1506.1181 | 8.5947 | 1498.6991 | 8.5771 | 1504.1372 | 8.5771 | 1495.7206
| 8.7705 | 125.3157 | 8.665 | 127.7457 | 8.7353 | 127.1137 | 8.7353 | 127.6075 | 8.6826 | 125.2959 | 8.7529 | 1511.9593 | 8.665 | 1502.4679 | 8.665 | 1497.6526 | 8.6299 | 1505.3552 | 8.6299 | 1498.2501 | 8.6123 | 1501.8625 | 8.6123 | 1495.4246
| 8.8056 | 126.0272 | 8.7002 | 127.5482 | 8.7705 | 127.2125 | 8.7705 | 127.4692 | 8.7177 | 126.4619 | 8.7881 | 1514.2858 | 8.7002 | 1506.1181 | 8.7002 | 1502.9225 | 8.665 | 1505.0504 | 8.665 | 1498.6991 | 8.6474 | 1500.805 | 8.6474 | 1494.0974
| 8.8408 | 126.7977 | 8.7353 | 127.5482 | 8.8056 | 127.1137 | 8.8056 | 126.6594 | 8.7529 | 126.6792 | 8.8232 | 1515.3755 | 8.7353 | 1508.2614 | 8.7353 | 1506.1181 | 8.7002 | 1506.7294 | 8.7002 | 1500.0512 | 8.6826 | 1500.805 | 8.6826 | 1494.0974
| 8.8759 | 126.8965 | 8.7705 | 127.5482 | 8.8408 | 126.857 | 8.8408 | 126.2841 | 8.7881 | 127.3902 | 8.8584 | 1516.7804 | 8.7705 | 1509.7985 | 8.7705 | 1508.4148 | 8.7353 | 1507.3416 | 8.7353 | 1504.898 | 8.7177 | 1502.0138 | 8.7177 | 1495.8686
| 8.9111 | 127.015 | 8.8056 | 127.7852 | 8.8759 | 126.1458 | 8.8759 | 125.9877 | 8.8232 | 127.568 | 8.8935 | 1518.1894 | 8.8056 | 1509.6446 | 8.8056 | 1508.4148 | 8.7705 | 1507.8012 | 8.7705 | 1505.3552 | 8.7529 | 1503.3776 | 8.7529 | 1498.5494
| 8.9462 | 127.2125 | 8.8408 | 127.7852 | 8.9111 | 125.5924 | 8.9111 | 126.126 | 8.8584 | 127.568 | 8.9287 | 1518.1894 | 8.8408 | 1509.6446 | 8.8408 | 1508.4148 | 8.8056 | 1507.8012 | 8.8056 | 1507.0354 | 8.7881 | 1504.5935 | 8.7881 | 1504.7458
| 8.9814 | 127.173 | 8.8759 | 127.7062 | 8.9462 | 125.5924 | 8.9462 | 126.126 | 8.8935 | 126.6792 | 8.9638 | 1520.8625 | 8.8759 | 1509.6446 | 8.8759 | 1508.8756 | 8.8408 | 1508.2614 | 8.8408 | 1508.4148 | 8.8232 | 1506.2709 | 8.8232 | 1505.6602
| 9.0165 | 127.252 | 8.9111 | 127.2323 | 8.9814 | 126.0074 | 8.9814 | 126.0469 | 8.9287 | 126.5014 | 8.999 | 1521.6516 | 8.9111 | 1512.2688 | 8.9111 | 1509.6446 | 8.8759 | 1508.2614 | 8.8759 | 1509.6446 | 8.8584 | 1506.7294 | 8.8584 | 1507.1885
| 9.0517 | 128.0222 | 8.9462 | 127.4297 | 9.0165 | 126.8372 | 9.0165 | 126.0469 | 8.9638 | 126.5409 | 9.0341 | 1526.892 | 8.9462 | 1512.8885 | 8.9462 | 1514.2858 | 8.9111 | 1508.8756 | 8.9111 | 1509.6446 | 8.8935 | 1506.8824 | 8.8935 | 1506.8824
| 9.0868 | 129.029 | 8.9814 | 128.1011 | 9.0517 | 127.4297 | 9.0517 | 126.0469 | 8.999 | 126.9162 | 9.0693 | 1529.935 | 8.9814 | 1517.2496 | 8.9814 | 1516.4678 | 8.9462 | 1510.4148 | 8.9462 | 1511.9593 | 8.9287 | 1505.3552 | 8.9287 | 1504.7458
| 9.122 | 129.7595 | 9.0165 | 129.6608 | 9.0868 | 128.2986 | 9.0868 | 126.0469 | 9.0341 | 126.9162 | 9.1044 | 1532.835 | 9.0165 | 1519.4455 | 9.0165 | 1520.3897 | 8.9814 | 1514.5969 | 8.9814 | 1515.3755 | 8.9638 | 1507.9545 | 8.9638 | 1503.3776
| 9.1571 | 130.5491 | 9.0517 | 130.1345 | 9.122 | 128.3183 | 9.122 | 126.3236 | 9.0693 | 126.5409 | 9.1396 | 1532.835 | 9.0517 | 1520.7048 | 9.0517 | 1521.3358 | 9.0165 | 1516.6241 | 9.0165 | 1518.3462 | 8.999 | 1510.4148 | 8.999 | 1503.3776
| 9.1923 | 130.2332 | 9.0868 | 130.5886 | 9.1571 | 128.496 | 9.1571 | 128.0814 | 9.1044 | 125.8691 | 9.1747 | 1533.32 | 9.0868 | 1520.7048 | 9.0868 | 1522.4419 | 9.0517 | 1521.3358 | 9.0517 | 1520.7048 | 9.0341 | 1515.5314 | 9.0341 | 1504.898
| 9.2275 | 130.5491 | 9.122 | 130.5886 | 9.1923 | 128.5947 | 9.1923 | 128.5947 | 9.1396 | 125.5529 | 9.2099 | 1535.2646 | 9.122 | 1520.7048 | 9.122 | 1523.0752 | 9.0868 | 1522.2838 | 9.0868 | 1520.7048 | 9.0693 | 1518.5031 | 9.0693 | 1505.3552
| 9.2626 | 130.7662 | 9.1571 | 130.8649 | 9.2275 | 128.9501 | 9.2275 | 128.7724 | 9.1747 | 125.5529 | 9.245 | 1536.4024 | 9.1571 | 1522.9168 | 9.1571 | 1524.8208 | 9.122 | 1525.4572 | 9.122 | 1521.8095 | 9.1044 | 1518.66 | 9.1044 | 1506.1181
| 9.2978 | 130.7662 | 9.1923 | 130.9242 | 9.2626 | 129.1672 | 9.2626 | 128.7724 | 9.2099 | 125.6715 | 9.2802 | 1543.7774 | 9.1923 | 1523.0752 | 9.1923 | 1524.8208 | 9.1571 | 1524.9799 | 9.1571 | 1523.5506 | 9.1396 | 1519.2883 | 9.1396 | 1509.3368
| 9.3329 | 130.9834 | 9.2275 | 130.5491 | 9.2978 | 129.4633 | 9.2978 | 128.7724 | 9.245 | 126.6989 | 9.3153 | 1546.259 | 9.2275 | 1529.7743 | 9.2275 | 1525.1389 | 9.1923 | 1525.4572 | 9.1923 | 1523.8679 | 9.1747 | 1519.4455 | 9.1747 | 1511.8046
| 9.3681 | 130.9834 | 9.2626 | 131.1216 | 9.3329 | 129.6015 | 9.3329 | 128.7527 | 9.2802 | 128.417 | 9.3505 | 1548.2527 | 9.2626 | 1530.7389 | 9.2626 | 1527.8509 | 9.2275 | 1528.8117 | 9.2275 | 1524.1853 | 9.2099 | 1522.2838 | 9.2099 | 1513.3538
| 9.4032 | 130.6873 | 9.2978 | 131.4177 | 9.3681 | 131.398 | 9.3681 | 128.6342 | 9.3153 | 129.0685 | 9.3856 | 1551.592 | 9.2978 | 1534.94 | 9.2978 | 1532.9967 | 9.2626 | 1532.1891 | 9.2626 | 1528.8117 | 9.245 | 1519.4455 | 9.245 | 1517.4061
| 9.4384 | 129.7595 | 9.3329 | 131.9113 | 9.4032 | 132.2273 | 9.4032 | 128.7527 | 9.3505 | 129.029 | 9.4208 | 1554.2783 | 9.3329 | 1537.869 | 9.3329 | 1539.9949 | 9.2978 | 1535.7519 | 9.2978 | 1531.8665 | 9.2802 | 1525.7757 | 9.2802 | 1521.178
| 9.4735 | 126.6792 | 9.3681 | 132.8988 | 9.4384 | 132.4643 | 9.4384 | 129.3252 | 9.3856 | 129.0685 | 9.4559 | 1554.7835 | 9.3681 | 1541.965 | 9.3681 | 1540.8148 | 9.3329 | 1536.728 | 9.3329 | 1535.7519 | 9.3153 | 1526.7324 | 9.3153 | 1525.4572
| 9.5087 | 126.4224 | 9.4032 | 133.2346 | 9.4735 | 132.5828 | 9.4735 | 130.1543 | 9.4208 | 129.029 | 9.4911 | 1555.4577 | 9.4032 | 1544.1076 | 9.4032 | 1542.9528 | 9.3681 | 1538.1955 | 9.3681 | 1540.1588 | 9.3505 | 1531.8665 | 9.3505 | 1529.6138
|  |  | 9.4384 | 133.61 | 9.5087 | 132.9581 | 9.5087 | 130.4701 | 9.4559 | 128.5355 | 9.5262 | 1556.1328 | 9.4384 | 1549.2523 | 9.4384 | 1545.0995 | 9.4032 | 1543.2825 | 9.4032 | 1540.8148 | 9.3856 | 1535.7519 | 9.3856 | 1530.578
|  |  | 9.4735 | 133.9657 | 9.5438 | 132.9581 | 9.5438 | 131.5757 | 9.4911 | 128.5355 |  |  | 9.4735 | 1551.592 | 9.4735 | 1548.9189 | 9.4384 | 1546.4249 | 9.4384 | 1543.2825 | 9.4208 | 1536.8908 | 9.4208 | 1532.3505
|  |  | 9.5087 | 133.8669 |  |  |  |  | 9.5262 | 129.2462 |  |  | 9.5087 | 1556.3017 | 9.5087 | 1553.1014 | 9.4735 | 1549.0856 | 9.4735 | 1544.7687 | 9.4559 | 1542.9528 | 9.4559 | 1534.4535
|  |  | 9.5438 | 133.9657 |  |  |  |  | 9.5614 | 130.1345 |  |  | 9.5438 | 1556.6396 | 9.5438 | 1554.2783 | 9.5087 | 1551.2572 | 9.5087 | 1546.4249 | 9.4911 | 1545.9275 | 9.4911 | 1537.0537
|  |  | 9.579 | 133.9855 |  |  |  |  | 9.5966 | 130.8255 |  |  | 9.579 | 1557.3162 | 9.579 | 1553.2694 | 9.5438 | 1551.592 | 9.5438 | 1547.2549 | 9.5262 | 1549.9198 | 9.5262 | 1538.1955
|  |  | 9.6141 | 133.4322 | [8 empty] | 9.6141 | 1557.3162 | 9.6141 | 1553.1014 | 9.579 | 1554.9519 | 9.579 | 1550.2538 | 9.5614 | 1550.0868 | 9.5614 | 1542.4587
[17 empty] | 9.6141 | 1555.2891 | 9.6141 | 1551.4246 | 9.5966 | 1552.0947 | 9.5966 | 1543.9425
[17 empty] | 9.6493 | 1555.2891 | 9.6493 | 1552.5978 | 9.6317 | 1552.4301 | 9.6317 | 1545.9275
[21 empty] | 9.6669 | 1552.0947 | 9.6669 | 1548.7523

## repeatability
The step wedge was scanned 8 times consecutively and an ROI with the coordinates x1=8.49, x2=10.9, y1=6.42, y2=6.98 was applied to each image in order to obtain the mean dose value in this ROI.
cGy
| 14:50 |  | 15:20 |  | 15:50 |  | 15th august testing immediately after turn on | x 8.78 12.47 y 6.52 6.93 | 11th sept | 25 minutes later after being on | 8.62 11.24 6.38 6.88
| 559.658 |  | 557.8652 |  | 557.4287 |  | 561.6745 |  | 562.1696 | 562.2925
| 556.3621 |  | 556.4022 |  | 556.6499 |  | 564.5898 |  | 559.8571 | 559.4542
| 555.7725 |  | 555.7559 |  | 554.7845 |  | 564.0859 |  | 559.0304 | 559.455
| 554.8055 |  | 555.4498 |  | 553.9613 |  | 563.6658 |  | 558.5612 | 558.3285
| 553.8083 |  | 554.7375 |  | 554.4346 |  | 562.8902 |  | 558.6137 | 559.7588
| 554.165 |  | 553.7738 |  | 553.6778 |  | 562.577 |  | 560.6678 | 558.2825
| 553.7615 |  | 554.2863 |  | 553.2975 |  | 562.5483 |  |  | 558.9726
| 554.7724 |  |  |  | 553.3293 |  | 566.3022 |  |  | 558.8134
mean | 555.3882 | mean | 555.4672 | mean | 554.6955 |  | 563.5417 |  | 559.8166 | 559.4197
SD | 1.9525 | SD | 1.3846 | SD | 1.5474 |  | 1.4569 |  | 1.4074 | 1.2774
coefficient of variation | 0.35155 |  | 0.24927 |  | 0.27897 |  | 0.25853 |  | 0.2514 | 0.22834
number of measurements | 8 |  | 7 |  | 8
| Time taken | Number of measurements | mean (cGy) | Standard deviation | Coefficient of variation %
1 | 14:50 | 8 | 555.3882 | 1.9525 | 0.35155
2 | 15:20 | 7 | 555.4672 | 1.3846 | 0.24927
3 | 15:50 | 8 | 554.6955 | 1.5474 | 0.27897
4 | 11th Aug | 8 | 563.5417 | 1.4569 | 0.25853
| 11th sept | 6 | 559.8166 | 1.4074 | 0.2514 | these two sets were done so that we could measure if the value was the same after a certain length of time and not just the repeatability range.
| 30 minutes later | 8 | 559.4197 | 1.2774 | 0.22834

## positional shifts
measurement location | dose | profile | measurement |  | measurement location | dose | profile | measurement
peak at ~7.5cm | 1Gy | 11.5cm up | 7.3 | 0.4 | peak at ~4cm | 14Gy | 10cm up
peak at ~-5cm | 1Gy | 11.5cm up | -4.4 |  | peak at ~-5cm | 14Gy | 10cm up | -4.6
peak at ~7.5cm | 1Gy | 4.5cm up | 7.2 | 0.3 | peak at ~4cm | 14Gy | 4.5cm up
peak at ~-5cm | 1Gy | 4.5cm up | -4.5 |  | peak at ~-5cm | 14Gy | 4.5cm up
peak at ~7.5cm | 1Gy | middle | 6.9 | 0 | peak at ~4cm | 14Gy | middle
peak at ~-5cm | 1Gy | middle | -4.7 |  | peak at ~-5cm | 14Gy | middle
peak at ~7.5cm | 1Gy | 4.5cm down | 7.2 | 0.3 | peak at ~4cm | 14Gy | 4.5cm down | 3.5
peak at ~-5cm | 1Gy | 4.5cm down | -4.5 |  | peak at ~-5cm | 14Gy | 4.5cm down | -4.7
peak at ~7.5cm | 1Gy | 11.5cm down | 7.5 | 0.6 | peak at ~4cm | 14Gy | 10cm down | 3.6
peak at ~-5cm | 1Gy | 11.5cm down | -4.2 |  | peak at ~-5cm | 14Gy | 10cm down | -4.6
0.21679

## variation in dose profiles
middle 1Gy | 4.5cm up 1Gy | 4.5cm down 1Gy | 11.5cm up 1Gy | 11.5cm down 1Gy | middle 14Gy | 4.5cm up 14Gy | 10cm up 14Gy | 4.5cm down 14Gy | 10cm down 14Gy
108.365 | 110.861 | 109.1035 | 106.779 | 107.366 | 1410.9151 | 1398.8675 | 1400.0091 | 1397.8779 | 1386.774
1Gy at centre of profile | 14Gy variation at centre of profile
108.4949 | 1398.8887 | mean
1.5972 | 8.5685 | std
1.4721 | 0.61252 | CoV

## measuring linearity of uniformy
roi | 0.02 | 19.07
| 8.44 | 17.61
cm | cGy |  | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy | cm | cGy
-9.5087 | 251.7004 |  | -9.5087 | 435.319 | -9.5087 | 621.2492 | -9.5087 | 966.2418 | -9.5087 | 1250.6402 | -9.5087 | 1542.9528 | -9.4911 | 1935.7645
-9.4735 | 250.6074 |  | -9.4735 | 434.0799 | -9.4735 | 621.7699 | -9.4735 | 965.3583 | -9.4735 | 1250.2652 | -9.4735 | 1544.2728 | -9.4559 | 1935.7645
-9.4384 | 250.6074 |  | -9.4384 | 432.7196 | -9.4384 | 621.7699 | -9.4384 | 963.9468 | -9.4384 | 1243.3846 | -9.4384 | 1543.7774 | -9.4208 | 1936.277
-9.4032 | 250.8435 |  | -9.4032 | 431.6909 | -9.4032 | 622.4069 | -9.4032 | 963.4181 | -9.4032 | 1243.3846 | -9.4032 | 1542.1295 | -9.3856 | 1936.277
-9.3681 | 250.6074 |  | -9.3681 | 431.0744 | -9.3681 | 623.1026 | -9.3681 | 963.1538 | -9.3681 | 1242.7585 | -9.3681 | 1542.1295 | -9.3505 | 1935.5082
-9.3329 | 251.0206 |  | -9.3329 | 431.0334 | -9.3329 | 623.683 | -9.3329 | 963.33 | -9.3329 | 1242.7585 | -9.3329 | 1539.831 | -9.3153 | 1936.277
-9.2978 | 251.2274 |  | -9.2978 | 430.5817 | -9.2978 | 624.264 | -9.2978 | 963.7705 | -9.2978 | 1242.0071 | -9.2978 | 1534.94 | -9.2802 | 1933.4602
-9.2626 | 250.9616 |  | -9.2626 | 430.6228 | -9.2626 | 624.0897 | -9.2626 | 962.2736 | -9.2626 | 1242.508 | -9.2626 | 1531.3829 | -9.245 | 1930.3937
-9.2275 | 250.4599 |  | -9.2275 | 430.5407 | -9.2275 | 623.1026 | -9.2275 | 961.0428 | -9.2275 | 1237.8719 | -9.2275 | 1528.4912 | -9.2099 | 1927.0793
-9.1923 | 249.694 |  | -9.1923 | 429.0647 | -9.1923 | 623.683 | -9.1923 | 960.1646 | -9.1923 | 1237.1196 | -9.1923 | 1523.0752 | -9.1747 | 1924.7894
-9.1571 | 248.3133 |  | -9.1571 | 428.2053 | -9.1571 | 623.2186 | -9.1571 | 958.2354 | -9.1571 | 1236.2418 | -9.1571 | 1520.8625 | -9.1396 | 1914.4053
-9.122 | 247.9322 |  | -9.122 | 427.3878 | -9.122 | 622.9286 | -9.122 | 955.5229 | -9.122 | 1232.4772 | -9.122 | 1518.1894 | -9.1044 | 1908.3634
-9.0868 | 247.3761 |  | -9.0868 | 426.0818 | -9.0868 | 622.9286 | -9.0868 | 953.3398 | -9.0868 | 1231.0959 | -9.0868 | 1517.7193 | -9.0693 | 1903.5994
-9.0517 | 247.2006 |  | -9.0517 | 425.5927 | -9.0517 | 622.9286 | -9.0517 | 953.6888 | -9.0517 | 1229.3371 | -9.0517 | 1515.2197 | -9.0341 | 1897.3568
-9.0165 | 246.4995 |  | -9.0165 | 425.4705 | -9.0165 | 621.5384 | -9.0165 | 953.3398 | -9.0165 | 1226.9488 | -9.0165 | 1512.8885 | -8.999 | 1891.1436
-8.9814 | 245.5959 |  | -8.9814 | 425.5112 | -8.9814 | 619.5749 | -8.9814 | 953.3398 | -8.9814 | 1224.0556 | -8.9814 | 1512.4237 | -8.9638 | 1888.6666
-8.9462 | 244.1721 |  | -8.9462 | 425.2261 | -8.9462 | 618.2506 | -8.9462 | 953.1653 | -8.9462 | 1219.6482 | -8.9462 | 1511.9593 | -8.9287 | 1883.2337
-8.9111 | 243.5057 |  | -8.9111 | 423.5992 | -8.9111 | 617.5034 | -8.9111 | 949.2446 | -8.9111 | 1215.7396 | -8.9111 | 1506.4237 | -8.8935 | 1871.948
-8.8759 | 242.956 |  | -8.8759 | 423.3555 | -8.8759 | 614.2392 | -8.8759 | 946.0273 | -8.8759 | 1213.0891 | -8.8759 | 1503.9852 | -8.8584 | 1866.8292
-8.8408 | 242.956 |  | -8.8408 | 422.0572 | -8.8408 | 610.3132 | -8.8408 | 944.898 | -8.8408 | 1209.1724 | -8.8408 | 1498.2501 | -8.8232 | 1859.7946
-8.8056 | 242.5515 |  | -8.8056 | 421.3684 | -8.8056 | 607.824 | -8.8056 | 943.8559 | -8.8056 | 1208.4137 | -8.8056 | 1491.9008 | -8.7881 | 1855.2074
-8.7705 | 242.667 |  | -8.7705 | 421.1255 | -8.7705 | 607.3163 | -8.7705 | 943.5954 | -8.7705 | 1205.8835 | -8.7705 | 1491.3182 | -8.7529 | 1852.8
-8.7353 | 243.3031 |  | -8.7353 | 421.0445 | -8.7353 | 607.4291 | -8.7353 | 942.1198 | -8.7353 | 1205.8835 | -8.7353 | 1488.1383 | -8.7177 | 1849.6775
-8.7002 | 243.4188 |  | -8.7002 | 421.1255 | -8.7002 | 606.8653 | -8.7002 | 940.6446 | -8.7002 | 1203.9845 | -8.7002 | 1486.4205 | -8.6826 | 1848.4786
-8.665 | 242.4938 |  | -8.665 | 421.0445 | -8.665 | 606.2459 | -8.665 | 938.6493 | -8.665 | 1203.6045 | -8.665 | 1481.6159 | -8.6474 | 1844.4115
-8.6299 | 241.7727 |  | -8.6299 | 420.3972 | -8.6299 | 605.7958 | -8.6299 | 935.7866 | -8.6299 | 1195.5926 | -8.6299 | 1479.6635 | -8.6123 | 1837.0305
-8.5947 | 240.4496 |  | -8.5947 | 419.0236 | -8.5947 | 604.9529 | -8.5947 | 935.0058 | -8.5947 | 1191.0819 | -8.5947 | 1476.7625 | -8.5771 | 1834.8962
-8.5596 | 239.6753 |  | -8.5596 | 416.2039 | -8.5596 | 604.0552 | -8.5596 | 933.7043 | -8.5596 | 1189.9148 | -8.5596 | 1476.2135 | -8.542 | 1834.1856 | difference between edge of profile and centre
-8.5244 | 238.7025 |  | -8.5244 | 415.4405 | -8.5244 | 602.0407 | -8.5244 | 930.9267 | -8.5244 | 1188.8751 | -8.5244 | 1473.6218 | -8.5068 | 1832.7657 | centre dose value | edge dose value | absolute difference | percentage diff
-8.4893 | 238.5025 |  | -8.4893 | 414.718 | -8.4893 | 600.479 | -8.4893 | 930.9267 | -8.4893 | 1188.8751 | -8.4893 | 1473.0794 | -8.4717 | 1828.5165 | 220 | 252 | 32 | 14.5455
-8.4541 | 238.5882 |  | -8.4541 | 414.1164 | -8.4541 | 598.533 | -8.4541 | 930.9267 | -8.4541 | 1189.5252 | -8.4541 | 1471.0554 | -8.4365 | 1829.4594 | 390 | 435 | 45 | 11.5385
-8.419 | 238.103 |  | -8.419 | 412.8347 | -8.419 | 596.9259 | -8.419 | 929.8845 | -8.419 | 1188.2242 | -8.419 | 1470.6524 | -8.4014 | 1828.5165 | 557 | 621 | 64 | 11.4901
-8.3838 | 238.103 |  | -8.3838 | 411.5952 | -8.3838 | 593.8358 | -8.3838 | 929.7108 | -8.3838 | 1184.9575 | -8.3838 | 1469.3135 | -8.3662 | 1824.048 | 849 | 966 | 117 | 13.7809
-8.3486 | 238.2741 |  | -8.3486 | 411.5552 | -8.3486 | 593.6157 | -8.3486 | 928.755 | -8.3486 | 1182.5935 | -8.3486 | 1469.3135 | -8.3311 | 1829.2236 | 1083 | 1251 | 168 | 15.5125
-8.3135 | 238.3598 |  | -8.3135 | 411.5952 | -8.3135 | 592.4618 | -8.3135 | 927.3642 | -8.3135 | 1181.9351 | -8.3135 | 1469.3135 | -8.2959 | 1829.6953 | 1400 | 1543 | 143 | 10.2143
-8.2783 | 238.1885 |  | -8.2783 | 411.5952 | -8.2783 | 592.1875 | -8.2783 | 924.4055 | -8.2783 | 1181.9351 | -8.2783 | 1468.6466 | -8.2608 | 1829.2236 | 1745 | 1936 | 191 | 10.9456
-8.2432 | 238.1315 |  | -8.2432 | 411.1958 | -8.2432 | 592.2972 | -8.2432 | 923.7957 | -8.2432 | 1181.9351 | -8.2432 | 1468.6466 | -8.2256 | 1831.5838
-8.208 | 238.103 |  | -8.208 | 411.1559 | -8.208 | 591.7488 | -8.208 | 924.057 | -8.208 | 1181.2761 | -8.208 | 1470.1161 | -8.1905 | 1833.4754
-8.1729 | 237.8748 |  | -8.1729 | 411.5952 | -8.1729 | 592.2423 | -8.1729 | 923.7957 | -8.1729 | 1180.3522 | -8.1729 | 1470.25 | -8.1553 | 1833.4754
-8.1377 | 238.103 |  | -8.1377 | 411.6351 | -8.1377 | 592.2972 | -8.1377 | 922.0521 | -8.1377 | 1180.7483 | -8.1377 | 1470.25 | -8.1202 | 1835.1331
-8.1026 | 238.874 |  | -8.1026 | 412.2747 | -8.1026 | 594.0559 | -8.1026 | 922.0521 | -8.1026 | 1181.0123 | -8.1026 | 1470.25 | -8.085 | 1833.4754
-8.0674 | 240.2774 |  | -8.0674 | 413.1149 | -8.0674 | 595.0477 | -8.0674 | 922.0521 | -8.0674 | 1181.2761 | -8.0674 | 1467.4503 | -8.0499 | 1830.8753
-8.0323 | 240.4783 |  | -8.0323 | 414.1966 | -8.0323 | 595.2131 | -8.0323 | 922.1393 | -8.0323 | 1181.0123 | -8.0323 | 1465.468 | -8.0147 | 1833.7121
-7.9971 | 240.2774 |  | -7.9971 | 414.7581 | -7.9971 | 595.5442 | -7.9971 | 922.0521 | -7.9971 | 1181.5398 | -7.9971 | 1463.3692 | -7.9795 | 1837.2678
-7.962 | 240.2774 |  | -7.962 | 415.2397 | -7.962 | 597.0919 | -7.962 | 921.8776 | -7.962 | 1182.1985 | -7.962 | 1460.7678 | -7.9444 | 1839.1686
-7.9268 | 240.1913 |  | -7.9268 | 415.9226 | -7.9268 | 596.7599 | -7.9268 | 923.6214 | -7.9268 | 1182.5935 | -7.9268 | 1460.7678 | -7.9092 | 1841.3106
-7.8917 | 239.618 |  | -7.8917 | 416.8475 | -7.8917 | 596.8706 | -7.8917 | 923.2728 | -7.8917 | 1184.6953 | -7.8917 | 1460.7678 | -7.8741 | 1847.281
-7.8565 | 239.7899 |  | -7.8565 | 417.9347 | -7.8565 | 597.8674 | -7.8565 | 924.5796 | -7.8565 | 1189.0052 | -7.8565 | 1461.9355 | -7.8389 | 1850.1573
-7.8214 | 240.2487 |  | -7.8214 | 419.1043 | -7.8214 | 598.2001 | -7.8214 | 924.9279 | -7.8214 | 1190.3042 | -7.8214 | 1463.8924 | -7.8038 | 1851.3578
-7.7862 | 239.8185 |  | -7.7862 | 419.2658 | -7.7862 | 599.4218 | -7.7862 | 925.015 | -7.7862 | 1189.9148 | -7.7862 | 1464.6791 | -7.7686 | 1851.3578
-7.7511 | 240.2487 |  | -7.7511 | 420.2355 | -7.7511 | 599.7554 | -7.7511 | 925.7114 | -7.7511 | 1190.5636 | -7.7511 | 1462.3258 | -7.7335 | 1854.4847
-7.7159 | 240.3922 |  | -7.7159 | 420.3163 | -7.7159 | 599.4218 | -7.7159 | 927.3642 | -7.7159 | 1188.745 | -7.7159 | 1464.6791 | -7.6983 | 1856.4129
-7.6808 | 240.9093 |  | -7.6808 | 420.3568 | -7.6808 | 598.4221 | -7.6808 | 925.015 | -7.6808 | 1188.4847 | -7.6808 | 1466.6556 | -7.6632 | 1855.6895
-7.6456 | 240.9093 |  | -7.6456 | 420.5185 | -7.6456 | 597.4796 | -7.6456 | 925.7114 | -7.6456 | 1188.3545 | -7.6456 | 1469.3135 | -7.628 | 1855.4484
-7.6105 | 240.9093 |  | -7.6105 | 420.9636 | -7.6105 | 597.3688 | -7.6105 | 925.3632 | -7.6105 | 1188.745 | -7.6105 | 1471.3244 | -7.5929 | 1855.4484
-7.5753 | 240.5932 |  | -7.5753 | 422.1382 | -7.5753 | 596.4281 | -7.5753 | 925.015 | -7.5753 | 1187.9636 | -7.5753 | 1476.2135 | -7.5577 | 1855.2074
-7.5401 | 240.1053 |  | -7.5401 | 422.3815 | -7.5401 | 596.4281 | -7.5401 | 925.015 | -7.5401 | 1187.9636 | -7.5401 | 1478.1399 | -7.5226 | 1855.2074
-7.505 | 239.6466 |  | -7.505 | 423.2743 | -7.505 | 597.4242 | -7.505 | 925.015 | -7.505 | 1187.8333 | -7.505 | 1479.3858 | -7.4874 | 1855.2074
-7.4698 | 239.6466 |  | -7.4698 | 423.6804 | -7.4698 | 598.0338 | -7.4698 | 924.057 | -7.4698 | 1188.3545 | -7.4698 | 1480.4985 | -7.4523 | 1858.5858
-7.4347 | 239.5893 |  | -7.4347 | 423.9243 | -7.4347 | 598.3111 | -7.4347 | 924.4055 | -7.4347 | 1188.6149 | -7.4347 | 1482.0362 | -7.4171 | 1858.5858
-7.3995 | 238.3883 |  | -7.3995 | 424.1275 | -7.3995 | 598.4221 | -7.3995 | 924.9279 | -7.3995 | 1192.3753 | -7.3995 | 1483.7244 | -7.382 | 1858.3442
-7.3644 | 238.3883 |  | -7.3644 | 424.2902 | -7.3644 | 599.9779 | -7.3644 | 924.057 | -7.3644 | 1193.9225 | -7.3644 | 1485.1399 | -7.3468 | 1857.6196
-7.3292 | 239.1599 |  | -7.3292 | 424.2088 | -7.3292 | 601.6499 | -7.3292 | 924.1441 | -7.3292 | 1195.7208 | -7.3292 | 1486.4205 | -7.3117 | 1857.6196
-7.2941 | 239.1599 |  | -7.2941 | 424.1275 | -7.2941 | 602.2642 | -7.2941 | 924.9279 | -7.2941 | 1197.5115 | -7.2941 | 1483.8656 | -7.2765 | 1854.9665
-7.2589 | 240.3922 |  | -7.2589 | 424.6562 | -7.2589 | 603.3269 | -7.2589 | 924.9279 | -7.2589 | 1198.7857 | -7.2589 | 1483.0196 | -7.2414 | 1854.7256
-7.2238 | 240.9093 |  | -7.2238 | 424.0869 | -7.2238 | 603.2709 | -7.2238 | 924.057 | -7.2238 | 1199.8021 | -7.2238 | 1480.7774 | -7.2062 | 1851.1176
-7.1886 | 240.9093 |  | -7.1886 | 424.6562 | -7.1886 | 602.9351 | -7.1886 | 921.7031 | -7.1886 | 1200.1826 | -7.1886 | 1481.0566 | -7.171 | 1849.9174
-7.1535 | 241.5135 |  | -7.1535 | 425.8372 | -7.1535 | 602.0407 | -7.1535 | 921.7031 | -7.1535 | 1200.1826 | -7.1535 | 1482.7383 | -7.1359 | 1847.5204
-7.1183 | 241.9456 |  | -7.1183 | 427.3061 | -7.1183 | 601.092 | -7.1183 | 921.8776 | -7.1183 | 1199.2942 | -7.1183 | 1482.8789 | -7.1007 | 1846.563
-7.0832 | 242.061 |  | -7.0832 | 427.061 | -7.0832 | 600.2562 | -7.0832 | 921.6159 | -7.0832 | 1199.04 | -7.0832 | 1485.1399 | -7.0656 | 1843.6952
-7.048 | 241.7151 |  | -7.048 | 427.3878 | -7.048 | 599.533 | -7.048 | 921.8776 | -7.048 | 1198.5312 | -7.048 | 1487.2779 | -7.0304 | 1846.563
-7.0129 | 241.7151 |  | -7.0129 | 427.3061 | -7.0129 | 599.811 | -7.0129 | 922.9242 | -7.0129 | 1197.1284 | -7.0129 | 1487.7077 | -6.9953 | 1846.563
-6.9777 | 240.4783 |  | -6.9777 | 426.7345 | -6.9777 | 599.533 | -6.9777 | 923.7957 | -6.9777 | 1193.1495 | -6.9777 | 1487.5644 | -6.9601 | 1841.5489
-6.9426 | 239.9045 |  | -6.9426 | 425.5927 | -6.9426 | 598.4775 | -6.9426 | 921.0922 | -6.9426 | 1192.2461 | -6.9426 | 1487.5644 | -6.925 | 1839.4064
-6.9074 | 239.3889 |  | -6.9074 | 424.9411 | -6.9074 | 597.0919 | -6.9074 | 920.3064 | -6.9074 | 1191.5997 | -6.9074 | 1486.1354 | -6.8898 | 1838.4555
-6.8723 | 239.8759 |  | -6.8723 | 424.9818 | -6.8723 | 595.6546 | -6.8723 | 921.0922 | -6.8723 | 1190.1744 | -6.8723 | 1486.1354 | -6.8547 | 1835.3701
-6.8371 | 239.9045 |  | -6.8371 | 424.9411 | -6.8371 | 595.6546 | -6.8371 | 921.7904 | -6.8371 | 1186.9199 | -6.8371 | 1485.8505 | -6.8195 | 1838.4555
-6.802 | 239.8472 |  | -6.802 | 424.9818 | -6.802 | 595.8755 | -6.802 | 921.7904 | -6.802 | 1184.3018 | -6.802 | 1485.566 | -6.7844 | 1837.5053
-6.7668 | 239.8472 |  | -6.7668 | 424.6969 | -6.7668 | 596.9259 | -6.7668 | 923.0985 | -6.7668 | 1182.9881 | -6.7668 | 1482.4572 | -6.7492 | 1838.4555
-6.7316 | 239.5893 |  | -6.7316 | 424.0869 | -6.7316 | 596.3175 | -6.7316 | 923.0985 | -6.7316 | 1181.1442 | -6.7316 | 1479.8025 | -6.7141 | 1838.4555
-6.6965 | 239.8472 |  | -6.6965 | 422.1382 | -6.6965 | 596.4834 | -6.6965 | 924.6667 | -6.6965 | 1180.7483 | -6.6965 | 1477.8638 | -6.6789 | 1838.6931
-6.6613 | 239.2744 |  | -6.6613 | 419.9525 | -6.6613 | 596.1517 | -6.6613 | 924.6667 | -6.6613 | 1180.7483 | -6.6613 | 1475.2557 | -6.6438 | 1837.7427
-6.6262 | 238.8454 |  | -6.6262 | 417.7735 | -6.6262 | 596.1517 | -6.6262 | 924.2313 | -6.6262 | 1181.5398 | -6.6262 | 1474.7099 | -6.6086 | 1844.8893
-6.591 | 238.8168 |  | -6.591 | 416.9682 | -6.591 | 595.5994 | -6.591 | 923.8828 | -6.591 | 1182.0668 | -6.591 | 1472.403 | -6.5735 | 1847.281
-6.5559 | 238.5882 |  | -6.5559 | 416.4854 | -6.5559 | 595.4338 | -6.5559 | 923.4471 | -6.5559 | 1182.0668 | -6.5559 | 1471.8631 | -6.5383 | 1849.9174
-6.5207 | 238.16 |  | -6.5207 | 415.3602 | -6.5207 | 595.7098 | -6.5207 | 919.6948 | -6.5207 | 1182.0668 | -6.5207 | 1469.7145 | -6.5032 | 1851.3578
-6.4856 | 237.3621 |  | -6.4856 | 414.0763 | -6.4856 | 595.2683 | -6.4856 | 918.7331 | -6.4856 | 1180.2201 | -6.4856 | 1469.5808 | -6.468 | 1852.0787
-6.4504 | 237.2767 |  | -6.4504 | 413.3551 | -6.4504 | 594.8271 | -6.4504 | 917.8582 | -6.4504 | 1179.5593 | -6.4504 | 1469.5808 | -6.4329 | 1846.563
-6.4153 | 236.9922 |  | -6.4153 | 413.235 | -6.4153 | 594.6067 | -6.4153 | 915.6678 | -6.4153 | 1179.8237 | -6.4153 | 1469.5808 | -6.3977 | 1845.3672
-6.3801 | 237.3621 |  | -6.3801 | 412.8747 | -6.3801 | 594.4965 | -6.3801 | 914.3513 | -6.3801 | 1180.8803 | -6.3801 | 1469.9821 | -6.3625 | 1844.8893
-6.345 | 237.2198 |  | -6.345 | 412.2747 | -6.345 | 593.8908 | -6.345 | 913.7363 | -6.345 | 1182.0668 | -6.345 | 1469.9821 | -6.3274 | 1844.1727
-6.3098 | 237.4759 |  | -6.3098 | 412.4346 | -6.3098 | 594.4965 | -6.3098 | 913.9999 | -6.3098 | 1183.3825 | -6.3098 | 1472.8087 | -6.2922 | 1845.3672
-6.2747 | 237.4759 |  | -6.2747 | 412.4346 | -6.2747 | 594.2211 | -6.2747 | 913.7363 | -6.2747 | 1183.1196 | -6.2747 | 1469.5808 | -6.2571 | 1847.0416
-6.2395 | 235.8282 |  | -6.2395 | 412.5146 | -6.2395 | 593.8358 | -6.2395 | 913.9999 | -6.2395 | 1183.1196 | -6.2395 | 1466.2592 | -6.2219 | 1842.9793
-6.2044 | 234.1316 |  | -6.2044 | 414.0763 | -6.2044 | 593.8358 | -6.2044 | 913.9999 | -6.2044 | 1179.5593 | -6.2044 | 1466.2592 | -6.1868 | 1841.0725
-6.1692 | 233.5397 |  | -6.1692 | 415.1193 | -6.1692 | 593.3957 | -6.1692 | 913.9999 | -6.1692 | 1179.5593 | -6.1692 | 1466.1272 | -6.1516 | 1839.8822
-6.1341 | 233.4552 |  | -6.1341 | 414.8785 | -6.1341 | 592.1326 | -6.1341 | 913.9999 | -6.1341 | 1178.633 | -6.1341 | 1463.8924 | -6.1165 | 1839.4064
-6.0989 | 233.2019 |  | -6.0989 | 414.9186 | -6.0989 | 593.3957 | -6.0989 | 913.5605 | -6.0989 | 1177.4402 | -6.0989 | 1463.8924 | -6.0813 | 1838.6931
-6.0638 | 233.4834 |  | -6.0638 | 414.9587 | -6.0638 | 592.1875 | -6.0638 | 910.6549 | -6.0638 | 1176.9095 | -6.0638 | 1463.6307 | -6.0462 | 1835.3701
-6.0286 | 234.9789 |  | -6.0286 | 413.9962 | -6.0286 | 592.0778 | -6.0286 | 909.3309 | -6.0286 | 1174.5162 | -6.0286 | 1462.1956 | -6.011 | 1838.2178
-5.9935 | 234.4138 |  | -5.9935 | 412.3946 | -5.9935 | 589.1239 | -5.9935 | 908.8007 | -5.9935 | 1171.4474 | -5.9935 | 1460.7678 | -5.9759 | 1838.2178
-5.9583 | 234.8375 |  | -5.9583 | 412.1147 | -5.9583 | 588.8512 | -5.9583 | 908.0931 | -5.9583 | 1165.0097 | -5.9583 | 1459.0895 | -5.9407 | 1835.1331
-5.9231 | 234.3573 |  | -5.9231 | 410.7966 | -5.9231 | 587.7618 | -5.9231 | 906.7647 | -5.9231 | 1161.7766 | -5.9231 | 1455.5076 | -5.9056 | 1828.0453
-5.888 | 233.8778 |  | -5.888 | 410.7567 | -5.888 | 587.2179 | -5.888 | 906.7647 | -5.888 | 1158.8061 | -5.888 | 1454.1121 | -5.8704 | 1822.875
-5.8528 | 233.4834 |  | -5.8528 | 409.9988 | -5.8528 | 587.0005 | -5.8528 | 904.9897 | -5.8528 | 1156.9129 | -5.8528 | 1453.7327 | -5.8353 | 1822.4061
-5.8177 | 232.4712 |  | -5.8177 | 409.0426 | -5.8177 | 585.7522 | -5.8177 | 904.9897 | -5.8177 | 1151.2231 | -5.8177 | 1452.8492 | -5.8001 | 1817.7278
-5.7825 | 231.378 |  | -5.7825 | 407.7298 | -5.7825 | 585.5353 | -5.7825 | 904.9897 | -5.7825 | 1146.4738 | -5.7825 | 1451.8425 | -5.765 | 1807.5031
-5.7474 | 230.8187 |  | -5.7474 | 407.4913 | -5.7474 | 585.3186 | -5.7474 | 902.7645 | -5.7474 | 1143.3512 | -5.7474 | 1451.8425 | -5.7298 | 1804.7307
-5.7122 | 230.344 |  | -5.7122 | 407.1338 | -5.7122 | 584.8313 | -5.7122 | 901.0683 | -5.7122 | 1140.9074 | -5.7122 | 1450.714 | -5.6947 | 1803.1167
-5.6771 | 229.8421 |  | -5.6771 | 406.7367 | -5.6771 | 584.1822 | -5.6771 | 898.8294 | -5.6771 | 1138.6 | -5.6771 | 1449.091 | -5.6595 | 1801.045
-5.6419 | 229.8421 |  | -5.6419 | 405.3485 | -5.6419 | 583.3719 | -5.6419 | 897.1228 | -5.6419 | 1137.9216 | -5.6419 | 1448.5932 | -5.6244 | 1800.5851
-5.6068 | 229.3966 |  | -5.6068 | 404.0023 | -5.6068 | 581.3784 | -5.6068 | 895.592 | -5.6068 | 1137.6502 | -5.6068 | 1445.0073 | -5.5892 | 1798.5182
-5.5716 | 229.1184 |  | -5.5716 | 402.1846 | -5.5716 | 579.8747 | -5.5716 | 893.7869 | -5.5716 | 1136.7006 | -5.5716 | 1441.3373 | -5.554 | 1798.5182
-5.5365 | 229.1741 |  | -5.5365 | 401.3167 | -5.5365 | 577.4667 | -5.5365 | 892.7922 | -5.5365 | 1136.1581 | -5.5365 | 1440.2437 | -5.5189 | 1795.082
-5.5013 | 229.2575 |  | -5.5013 | 400.5679 | -5.5013 | 576.4532 | -5.5013 | 890.1633 | -5.5013 | 1135.2089 | -5.5013 | 1438.549 | -5.4837 | 1794.8533
-5.4662 | 229.3688 |  | -5.4662 | 400.2528 | -5.4662 | 576.3466 | -5.4662 | 885.0625 | -5.4662 | 1131.1447 | -5.4662 | 1437.825 | -5.4486 | 1794.1675
-5.431 | 229.2019 |  | -5.431 | 399.6623 | -5.431 | 575.3352 | -5.431 | 884.5142 | -5.431 | 1130.0622 | -5.431 | 1437.2228 | -5.4134 | 1791.8846
-5.3959 | 229.5914 |  | -5.3959 | 399.6623 | -5.3959 | 574.8568 | -5.3959 | 884.3313 | -5.3959 | 1129.386 | -5.3959 | 1437.2228 | -5.3783 | 1791.2007
-5.3607 | 229.0906 |  | -5.3607 | 399.0331 | -5.3607 | 574.0073 | -5.3607 | 884.7884 | -5.3607 | 1130.0622 | -5.3607 | 1436.5013 | -5.3431 | 1788.2421
-5.3256 | 228.7294 |  | -5.3256 | 398.9152 | -5.3256 | 573.1062 | -5.3256 | 884.7884 | -5.3256 | 1130.0622 | -5.3256 | 1434.4643 | -5.308 | 1782.1233
-5.2904 | 228.7016 |  | -5.2904 | 398.9938 | -5.2904 | 573.1062 | -5.2904 | 885.4279 | -5.2904 | 1130.1975 | -5.2904 | 1431.4869 | -5.2728 | 1781.4456
-5.2553 | 228.7294 |  | -5.2553 | 399.3083 | -5.2553 | 573.1062 | -5.2553 | 887.799 | -5.2553 | 1129.386 | -5.2553 | 1430.7754 | -5.2377 | 1780.3171
-5.2201 | 228.8127 |  | -5.2201 | 399.151 | -5.2201 | 574.0073 | -5.2201 | 887.9811 | -5.2201 | 1126.9538 | -5.2201 | 1430.5384 | -5.2025 | 1780.0916
-5.1849 | 228.7016 |  | -5.1849 | 400.3315 | -5.1849 | 574.6974 | -5.1849 | 887.799 | -5.1849 | 1124.5256 | -5.1849 | 1428.2937 | -5.1674 | 1781.2198
-5.1498 | 228.0081 |  | -5.1498 | 400.686 | -5.1498 | 574.6974 | -5.1498 | 886.3406 | -5.1498 | 1122.7747 | -5.1498 | 1427.9403 | -5.1322 | 1783.0276
-5.1146 | 228.0081 |  | -5.1146 | 401.0013 | -5.1146 | 574.2726 | -5.1146 | 882.4093 | -5.1146 | 1121.1607 | -5.1146 | 1425.7076 | -5.0971 | 1782.5753
-5.0795 | 228.0635 |  | -5.0795 | 401.159 | -5.0795 | 573.3711 | -5.0795 | 881.4926 | -5.0795 | 1120.4889 | -5.0795 | 1421.5029 | -5.0619 | 1783.2538
-5.0443 | 227.925 |  | -5.0443 | 401.9083 | -5.0443 | 573.2651 | -5.0443 | 880.9422 | -5.0443 | 1119.1467 | -5.0443 | 1421.6193 | -5.0268 | 1784.8384
-5.0092 | 227.565 |  | -5.0092 | 402.4214 | -5.0092 | 573.2651 | -5.0092 | 878.5537 | -5.0092 | 1116.3339 | -5.0092 | 1421.5029 | -4.9916 | 1784.8384
-4.974 | 227.3989 |  | -4.974 | 403.33 | -4.974 | 573.4241 | -4.974 | 880.9422 | -4.974 | 1115.3982 | -4.974 | 1421.6193 | -4.9565 | 1781.8974
-4.9389 | 226.7633 |  | -4.9389 | 403.6463 | -4.9389 | 572.9473 | -4.9389 | 882.1344 | -4.9389 | 1115.6654 | -4.9389 | 1422.318 | -4.9213 | 1776.4893
-4.9037 | 226.3493 |  | -4.9037 | 403.6859 | -4.9037 | 573.4241 | -4.9037 | 882.6841 | -4.9037 | 1117.4044 | -4.9037 | 1422.5511 | -4.8862 | 1772.2277
-4.8686 | 226.57 |  | -4.8686 | 403.2905 | -4.8686 | 574.6974 | -4.8686 | 883.2335 | -4.8686 | 1119.6834 | -4.8686 | 1422.5511 | -4.851 | 1761.318
-4.8334 | 226.8185 |  | -4.8334 | 403.0929 | -4.8334 | 574.6974 | -4.8334 | 884.1485 | -4.8334 | 1121.0263 | -4.8334 | 1423.6013 | -4.8159 | 1752.4988
-4.7983 | 226.8185 |  | -4.7983 | 400.7254 | -4.7983 | 574.5381 | -4.7983 | 883.2335 | -4.7983 | 1121.2951 | -4.7983 | 1424.1856 | -4.7807 | 1748.9929
-4.7631 | 226.6252 |  | -4.7631 | 399.1903 | -4.7631 | 574.4319 | -4.7631 | 884.4228 | -4.7631 | 1120.892 | -4.7631 | 1424.1856 | -4.7455 | 1751.4019
-4.728 | 227.1777 |  | -4.728 | 397.9726 | -4.728 | 574.485 | -4.728 | 885.0625 | -4.728 | 1120.892 | -4.728 | 1425.7076 | -4.7104 | 1752.2793
-4.6928 | 227.0395 |  | -4.6928 | 397.8941 | -4.6928 | 574.5381 | -4.6928 | 886.4318 | -4.6928 | 1122.5055 | -4.6928 | 1428.1759 | -4.6752 | 1758.2224
-4.6577 | 227.0395 |  | -4.6577 | 398.2866 | -4.6577 | 574.3257 | -4.6577 | 886.4318 | -4.6577 | 1123.9866 | -4.6577 | 1429.4738 | -4.6401 | 1767.315
-4.6225 | 226.9566 |  | -4.6225 | 398.5615 | -4.6225 | 574.3788 | -4.6225 | 887.89 | -4.6225 | 1124.7952 | -4.6225 | 1429.9467 | -4.6049 | 1769.0987
-4.5874 | 227.0395 |  | -4.5874 | 399.4656 | -4.5874 | 575.6012 | -4.5874 | 887.89 | -4.5874 | 1127.2239 | -4.5874 | 1429.0014 | -4.5698 | 1788.0148
-4.5522 | 226.9566 |  | -4.5522 | 399.7804 | -4.5522 | 577.3065 | -4.5522 | 886.4318 | -4.5522 | 1128.9804 | -4.5522 | 1427.9403 | -4.5346 | 1770.662
-4.5171 | 226.9566 |  | -4.5171 | 399.623 | -4.5171 | 579.1244 | -4.5171 | 885.1539 | -4.5171 | 1127.2239 | -4.5171 | 1426.0595 | -4.4995 | 1769.0987
-4.4819 | 228.0081 |  | -4.4819 | 399.7804 | -4.4819 | 579.3922 | -4.4819 | 882.9588 | -4.4819 | 1127.0888 | -4.4819 | 1425.2388 | -4.4643 | 1768.2065
-4.4468 | 228.2576 |  | -4.4468 | 400.6466 | -4.4468 | 579.3922 | -4.4468 | 881.1257 | -4.4468 | 1127.2239 | -4.4468 | 1422.318 | -4.4292 | 1768.2065
-4.4116 | 227.5926 |  | -4.4116 | 400.6466 | -4.4116 | 577.3065 | -4.4116 | 881.1257 | -4.4116 | 1130.0622 | -4.4116 | 1420.5726 | -4.394 | 1770.2151
-4.3764 | 226.929 |  | -4.3764 | 401.2772 | -4.3764 | 577.1464 | -4.3764 | 881.2175 | -4.3764 | 1129.386 | -4.3764 | 1420.4564 | -4.3589 | 1768.2065
-4.3413 | 226.929 |  | -4.3413 | 401.2772 | -4.3413 | 575.9738 | -4.3413 | 881.7677 | -4.3413 | 1130.0622 | -4.3413 | 1418.832 | -4.3237 | 1768.2065
-4.3061 | 226.9842 |  | -4.3061 | 400.5679 | -4.3061 | 575.9738 | -4.3061 | 883.5081 | -4.3061 | 1129.927 | -4.3061 | 1418.832 | -4.2886 | 1767.9836
-4.271 | 227.1224 |  | -4.271 | 398.7973 | -4.271 | 575.9738 | -4.271 | 884.3313 | -4.271 | 1126.8188 | -4.271 | 1417.9056 | -4.2534 | 1764.6453
-4.2358 | 227.3436 |  | -4.2358 | 397.7371 | -4.2358 | 576.3999 | -4.2358 | 882.4093 | -4.2358 | 1126.8188 | -4.2358 | 1415.3638 | -4.2183 | 1764.6453
-4.2007 | 227.4266 |  | -4.2007 | 397.0313 | -4.2007 | 575.9738 | -4.2007 | 882.3177 | -4.2007 | 1126.4138 | -4.2007 | 1413.8657 | -4.1831 | 1763.979
-4.1655 | 227.233 |  | -4.1655 | 396.9529 | -4.1655 | 575.8141 | -4.1655 | 882.4093 | -4.1655 | 1123.0439 | -4.1655 | 1412.2552 | -4.148 | 1758.885
-4.1304 | 227.2053 |  | -4.1304 | 396.4828 | -4.1304 | 575.8141 | -4.1304 | 882.5925 | -4.1304 | 1122.1019 | -4.1304 | 1412.0253 | -4.1128 | 1759.548
-4.0952 | 226.6804 |  | -4.0952 | 396.0521 | -4.0952 | 575.8673 | -4.0952 | 882.5925 | -4.0952 | 1121.6984 | -4.0952 | 1412.1402 | -4.0777 | 1763.535
-4.0601 | 226.5148 |  | -4.0601 | 396.4045 | -4.0601 | 573.0003 | -4.0601 | 882.6841 | -4.0601 | 1116.7352 | -4.0601 | 1412.6 | -4.0425 | 1763.0912
-4.0249 | 226.6528 |  | -4.0249 | 396.4045 | -4.0249 | 568.2032 | -4.0249 | 883.9655 | -4.0249 | 1114.4636 | -4.0249 | 1412.1402 | -4.0074 | 1764.8675
-3.9898 | 226.7909 |  | -3.9898 | 396.5611 | -3.9898 | 567.1022 | -3.9898 | 883.9655 | -3.9898 | 1111.0015 | -3.9898 | 1412.4851 | -3.9722 | 1765.0898
-3.9546 | 226.7909 |  | -3.9546 | 396.5611 | -3.9546 | 567.1022 | -3.9546 | 883.6911 | -3.9546 | 1110.0721 | -3.9546 | 1413.2902 | -3.937 | 1766.6469
-3.9195 | 226.9842 |  | -3.9195 | 396.5611 | -3.9195 | 567.9408 | -3.9195 | 883.4166 | -3.9195 | 1110.0721 | -3.9195 | 1413.9809 | -3.9019 | 1763.0912
-3.8843 | 227.0119 |  | -3.8843 | 397.6586 | -3.8843 | 569.3064 | -3.8843 | 884.3313 | -3.8843 | 1110.0721 | -3.8843 | 1413.9809 | -3.8667 | 1763.0912
-3.8492 | 226.7633 |  | -3.8492 | 398.758 | -3.8492 | 569.9378 | -3.8492 | 884.3313 | -3.8492 | 1106.2351 | -3.8492 | 1413.5204 | -3.8316 | 1762.8694
-3.814 | 226.4872 |  | -3.814 | 398.5222 | -3.814 | 570.5173 | -3.814 | 884.3313 | -3.814 | 1106.8951 | -3.814 | 1413.9809 | -3.7964 | 1761.0966
-3.7789 | 226.3493 |  | -3.7789 | 398.5222 | -3.7789 | 569.9378 | -3.7789 | 883.325 | -3.7789 | 1108.879 | -3.7789 | 1413.5204 | -3.7613 | 1765.7569
-3.7437 | 226.0186 |  | -3.7437 | 397.9726 | -3.7437 | 567.6262 | -3.7437 | 883.325 | -3.7437 | 1108.879 | -3.7437 | 1412.1402 | -3.7261 | 1761.0966
-3.7086 | 225.578 |  | -3.7086 | 397.2273 | -3.7086 | 566.788 | -3.7086 | 882.8673 | -3.7086 | 1110.2048 | -3.7086 | 1410.8768 | -3.691 | 1761.0966
-3.6734 | 225.4405 |  | -3.6734 | 396.6786 | -3.6734 | 566.788 | -3.6734 | 878.0937 | -3.6734 | 1112.997 | -3.6734 | 1410.8768 | -3.6558 | 1760.6539
-3.6383 | 225.4405 |  | -3.6383 | 396.6786 | -3.6383 | 566.7357 | -3.6383 | 875.6989 | -3.6383 | 1115.6654 | -3.6383 | 1410.5324 | -3.6207 | 1763.3131
-3.6031 | 225.1656 |  | -3.6031 | 395.9347 | -3.6031 | 566.7357 | -3.6031 | 877.6336 | -3.6031 | 1119.415 | -3.6031 | 1409.7294 | -3.5855 | 1760.6539
-3.5679 | 225.1656 |  | -3.5679 | 394.4885 | -3.5679 | 566.5264 | -3.5679 | 875.2377 | -3.5679 | 1121.6984 | -3.5679 | 1407.5523 | -3.5504 | 1761.9826
-3.5328 | 225.3305 |  | -3.5328 | 394.0982 | -3.5328 | 566.6833 | -3.5328 | 873.5761 | -3.5328 | 1122.6401 | -3.5328 | 1407.0944 | -3.5152 | 1761.5395
-3.4976 | 225.3305 |  | -3.4976 | 393.0844 | -3.4976 | 566.9974 | -3.4976 | 873.3914 | -3.4976 | 1122.6401 | -3.4976 | 1406.7511 | -3.4801 | 1761.5395
-3.4625 | 225.2205 |  | -3.4625 | 392.9675 | -3.4625 | 565.9512 | -3.4625 | 871.2647 | -3.4625 | 1122.6401 | -3.4625 | 1404.4638 | -3.4449 | 1758.6641
-3.4273 | 225.0007 |  | -3.4273 | 392.5002 | -3.4273 | 565.2722 | -3.4273 | 870.339 | -3.4273 | 1120.0861 | -3.4273 | 1403.0925 | -3.4098 | 1757.1192
-3.3922 | 224.2051 |  | -3.3922 | 392.2278 | -3.3922 | 565.3766 | -3.3922 | 870.0612 | -3.3922 | 1117.2705 | -3.3922 | 1403.0925 | -3.3746 | 1754.2564
-3.357 | 224.0407 |  | -3.357 | 391.2945 | -3.357 | 565.2722 | -3.357 | 869.8759 | -3.357 | 1113.1302 | -3.357 | 1401.3793 | -3.3395 | 1751.8405
-3.3219 | 223.6575 |  | -3.3219 | 390.285 | -3.3219 | 564.8548 | -3.3219 | 869.6907 | -3.3219 | 1112.0651 | -3.3219 | 1401.6077 | -3.3043 | 1748.7742
-3.2867 | 223.2474 |  | -3.2867 | 389.5482 | -3.2867 | 564.6984 | -3.2867 | 869.6907 | -3.2867 | 1108.217 | -3.2867 | 1402.9783 | -3.2692 | 1750.9634
-3.2516 | 222.1563 |  | -3.2516 | 389.277 | -3.2516 | 564.6984 | -3.2516 | 869.1347 | -3.2516 | 1104.9173 | -3.2516 | 1402.9783 | -3.234 | 1750.9634
-3.2164 | 221.6393 |  | -3.2164 | 389.9747 | -3.2164 | 564.6984 | -3.2164 | 868.3931 | -3.2164 | 1103.471 | -3.2164 | 1403.0925 | -3.1989 | 1750.5252
-3.1813 | 220.7704 |  | -3.1813 | 390.0135 | -3.1813 | 561.7865 | -3.1813 | 865.2371 | -3.1813 | 1102.1593 | -3.1813 | 1403.0925 | -3.1637 | 1750.9634
-3.1461 | 220.5535 |  | -3.1461 | 390.0135 | -3.1461 | 561.2163 | -3.1461 | 864.9584 | -3.1461 | 1101.6355 | -3.1461 | 1402.0645 | -3.1285 | 1749.8682
-3.111 | 220.7704 |  | -3.111 | 390.2462 | -3.111 | 561.2163 | -3.111 | 864.9584 | -3.111 | 1102.0283 | -3.111 | 1402.0645 | -3.0934 | 1748.1184
-3.0758 | 220.4451 |  | -3.0758 | 390.2462 | -3.0758 | 560.7503 | -3.0758 | 864.7725 | -3.0758 | 1102.0283 | -3.0758 | 1402.7498 | -3.0582 | 1745.9355
-3.0407 | 220.2826 |  | -3.0407 | 389.1608 | -3.0407 | 559.8195 | -3.0407 | 860.6787 | -3.0407 | 1101.6355 | -3.0407 | 1402.7498 | -3.0231 | 1739.8495
-3.0055 | 220.0661 |  | -3.0055 | 388.5414 | -3.0055 | 559.2515 | -3.0055 | 862.4475 | -3.0055 | 1101.3738 | -3.0055 | 1402.9783 | -2.9879 | 1739.1997
-2.9704 | 219.0665 |  | -2.9704 | 387.8839 | -2.9704 | 559.3031 | -2.9704 | 861.796 | -2.9704 | 1101.6355 | -2.9704 | 1403.4353 | -2.9528 | 1736.8208
-2.9352 | 218.2851 |  | -2.9352 | 387.1884 | -2.9352 | 559.3547 | -2.9352 | 861.796 | -2.9352 | 1100.1976 | -2.9352 | 1403.4353 | -2.9176 | 1735.5258
-2.9001 | 216.9693 |  | -2.9001 | 386.5323 | -2.9001 | 559.3547 | -2.9001 | 861.796 | -2.9001 | 1100.1976 | -2.9001 | 1403.6638 | -2.8825 | 1736.1731
-2.8649 | 216.6211 |  | -2.8649 | 386.1852 | -2.8649 | 559.6645 | -2.8649 | 861.796 | -2.8649 | 1099.6756 | -2.8649 | 1403.7781 | -2.8473 | 1738.1177
-2.8298 | 216.4337 |  | -2.8298 | 386.1852 | -2.8298 | 559.6645 | -2.8298 | 861.796 | -2.8298 | 1097.4634 | -2.8298 | 1403.8923 | -2.8122 | 1739.8495
-2.7946 | 216.6211 |  | -2.7946 | 386.5323 | -2.7946 | 558.9935 | -2.7946 | 860.5855 | -2.7946 | 1097.4634 | -2.7946 | 1404.8067 | -2.777 | 1743.9751
-2.7594 | 216.9693 |  | -2.7594 | 386.1852 | -2.7594 | 555.5995 | -2.7594 | 860.5855 | -2.7594 | 1098.113 | -2.7594 | 1403.8923 | -2.7419 | 1747.463
-2.7243 | 217.7743 |  | -2.7243 | 385.5301 | -2.7243 | 555.0359 | -2.7243 | 860.0266 | -2.7243 | 1096.5553 | -2.7243 | 1403.2068 | -2.7067 | 1747.463
-2.6891 | 218.1506 |  | -2.6891 | 384.6448 | -2.6891 | 554.7287 | -2.6891 | 860.6787 | -2.6891 | 1096.5553 | -2.6891 | 1403.2068 | -2.6716 | 1743.1051
-2.654 | 218.1506 |  | -2.654 | 384.5678 | -2.654 | 553.9105 | -2.654 | 860.0266 | -2.654 | 1094.0993 | -2.654 | 1402.9783 | -2.6364 | 1743.1051
-2.6188 | 218.1506 |  | -2.6188 | 384.2987 | -2.6188 | 554.3706 | -2.6188 | 858.3489 | -2.6188 | 1091.5284 | -2.6188 | 1402.5214 | -2.6013 | 1741.3674
-2.5837 | 218.0968 |  | -2.5837 | 384.9526 | -2.5837 | 554.4729 | -2.5837 | 857.2298 | -2.5837 | 1092.0414 | -2.5837 | 1402.9783 | -2.5661 | 1741.1504
-2.5485 | 217.9086 |  | -2.5485 | 385.5301 | -2.5485 | 554.8311 | -2.5485 | 853.2159 | -2.5485 | 1091.2722 | -2.5485 | 1402.5214 | -2.531 | 1741.1504
-2.5134 | 217.2106 |  | -2.5134 | 386.031 | -2.5134 | 555.8046 | -2.5134 | 852.562 | -2.5134 | 1089.6104 | -2.5134 | 1401.8361 | -2.4958 | 1741.1504
-2.4782 | 218.2044 |  | -2.4782 | 386.1081 | -2.4782 | 556.8825 | -2.4782 | 852.9357 | -2.4782 | 1089.6104 | -2.4782 | 1401.8361 | -2.4607 | 1741.3674
-2.4431 | 217.9624 |  | -2.4431 | 386.4166 | -2.4431 | 556.8825 | -2.4431 | 852.0014 | -2.4431 | 1089.4829 | -2.4431 | 1397.8399 | -2.4255 | 1743.3225
-2.4079 | 218.043 |  | -2.4079 | 386.1081 | -2.4079 | 556.8825 | -2.4079 | 851.0669 | -2.4079 | 1089.4829 | -2.4079 | 1397.269 | -2.3903 | 1745.9355
-2.3728 | 217.7474 |  | -2.3728 | 386.0695 | -2.3728 | 555.7533 | -2.3728 | 850.3192 | -2.3728 | 1089.6104 | -2.3728 | 1397.269 | -2.3552 | 1745.9355
-2.3376 | 217.0497 |  | -2.3376 | 385.453 | -2.3376 | 554.0127 | -2.3376 | 849.7582 | -2.3376 | 1090.2488 | -2.3376 | 1397.3832 | -2.32 | 1747.2447
-2.3025 | 216.8621 |  | -2.3025 | 386.0695 | -2.3025 | 553.2466 | -2.3025 | 848.9168 | -2.3025 | 1092.9405 | -2.3025 | 1397.8399 | -2.2849 | 1749.4305
-2.2673 | 216.7282 |  | -2.2673 | 385.9539 | -2.2673 | 552.7365 | -2.2673 | 848.5427 | -2.2673 | 1094.4862 | -2.2673 | 1398.1824 | -2.2497 | 1751.1826
-2.2322 | 216.7282 |  | -2.2322 | 385.9539 | -2.2322 | 552.6346 | -2.2322 | 850.0387 | -2.2322 | 1093.4552 | -2.2322 | 1397.3832 | -2.2146 | 1751.1826
-2.197 | 217.0229 |  | -2.197 | 386.5323 | -2.197 | 552.5836 | -2.197 | 849.3843 | -2.197 | 1092.9405 | -2.197 | 1396.2413 | -2.1794 | 1751.4019
-2.1619 | 217.828 |  | -2.1619 | 385.9539 | -2.1619 | 552.5836 | -2.1619 | 848.9168 | -2.1619 | 1091.6566 | -2.1619 | 1394.528 | -2.1443 | 1751.1826
-2.1267 | 217.8817 |  | -2.1267 | 385.299 | -2.1267 | 552.5836 | -2.1267 | 847.514 | -2.1267 | 1091.6566 | -2.1267 | 1392.4707 | -2.1091 | 1745.2816
-2.0916 | 218.043 |  | -2.0916 | 385.299 | -2.0916 | 552.6346 | -2.0916 | 849.3843 | -2.0916 | 1091.1441 | -2.0916 | 1390.8692 | -2.074 | 1745.0638
-2.0564 | 218.1237 |  | -2.0564 | 385.376 | -2.0564 | 551.8196 | -2.0564 | 848.6362 | -2.0564 | 1091.7848 | -2.0564 | 1390.8692 | -2.0388 | 1744.1928
-2.0213 | 218.2582 |  | -2.0213 | 386.3394 | -2.0213 | 550.9041 | -2.0213 | 847.3269 | -2.0213 | 1092.6834 | -2.0213 | 1391.2125 | -2.0037 | 1744.8459
-1.9861 | 218.5274 |  | -1.9861 | 387.5747 | -1.9861 | 550.9041 | -1.9861 | 848.3557 | -1.9861 | 1092.6834 | -1.9861 | 1388.3493 | -1.9685 | 1744.8459
-1.9509 | 219.2554 |  | -1.9509 | 388.4253 | -1.9509 | 550.9041 | -1.9509 | 846.298 | -1.9509 | 1091.2722 | -1.9509 | 1387.4318 | -1.9334 | 1744.6282
-1.9158 | 220.2014 |  | -1.9158 | 391.3722 | -1.9158 | 550.5486 | -1.9158 | 843.5848 | -1.9158 | 1089.3554 | -1.9158 | 1388.3493 | -1.8982 | 1743.7576
-1.8806 | 220.9331 |  | -1.8806 | 391.4111 | -1.8806 | 550.9041 | -1.8806 | 842.1813 | -1.8806 | 1083.7867 | -1.8806 | 1387.4318 | -1.8631 | 1741.3674
-1.8455 | 221.7209 |  | -1.8455 | 391.4111 | -1.8455 | 550.955 | -1.8455 | 842.5556 | -1.8455 | 1078.5523 | -1.8455 | 1386.8579 | -1.8279 | 1739.6328
-1.8103 | 221.5306 |  | -1.8103 | 391.1391 | -1.8103 | 551.2599 | -1.8103 | 842.8363 | -1.8103 | 1076.5806 | -1.8103 | 1386.7431 | -1.7928 | 1739.1997
-1.7752 | 221.5306 |  | -1.7752 | 391.1391 | -1.7752 | 552.6855 | -1.7752 | 843.4913 | -1.7752 | 1076.0897 | -1.7752 | 1386.5135 | -1.7576 | 1735.3101
-1.74 | 221.5849 |  | -1.74 | 391.1391 | -1.74 | 552.176 | -1.74 | 843.4913 | -1.74 | 1074.1344 | -1.74 | 1386.6283 | -1.7225 | 1731.6512
-1.7049 | 222.3197 |  | -1.7049 | 391.1391 | -1.7049 | 552.6855 | -1.7049 | 843.3977 | -1.7049 | 1071.9507 | -1.7049 | 1386.6283 | -1.6873 | 1729.9343
-1.6697 | 222.3197 |  | -1.6697 | 391.1003 | -1.6697 | 552.176 | -1.6697 | 840.6841 | -1.6697 | 1069.5452 | -1.6697 | 1387.4318 | -1.6522 | 1724.3759
-1.6346 | 222.6196 |  | -1.6346 | 388.9672 | -1.6346 | 551.6669 | -1.6346 | 837.7833 | -1.6346 | 1068.3509 | -1.6346 | 1387.317 | -1.617 | 1719.2745
-1.5994 | 222.6196 |  | -1.5994 | 388.9672 | -1.5994 | 549.8381 | -1.5994 | 837.1284 | -1.5994 | 1066.5702 | -1.5994 | 1386.3987 | -1.5818 | 1718.8507
-1.5643 | 221.3403 |  | -1.5643 | 387.1498 | -1.5643 | 549.8381 | -1.5643 | 835.7251 | -1.5643 | 1065.6259 | -1.5643 | 1384.4444 | -1.5467 | 1716.5231
-1.5291 | 220.5806 |  | -1.5291 | 386.1852 | -1.5291 | 550.0918 | -1.5291 | 836.4735 | -1.5291 | 1064.3338 | -1.5291 | 1378.7865 | -1.5115 | 1716.5231
-1.494 | 219.9849 |  | -1.494 | 386.5323 | -1.494 | 549.8381 | -1.494 | 836.0057 | -1.494 | 1062.9326 | -1.494 | 1378.6706 | -1.4764 | 1716.5231
-1.4588 | 219.4984 |  | -1.4588 | 386.1852 | -1.4588 | 550.0918 | -1.4588 | 836.4735 | -1.4588 | 1062.4675 | -1.4588 | 1379.25 | -1.4412 | 1719.0626
-1.4237 | 219.4714 |  | -1.4237 | 385.9539 | -1.4237 | 550.0918 | -1.4237 | 836.4735 | -1.4237 | 1062.6999 | -1.4237 | 1380.176 | -1.4061 | 1719.6986
-1.3885 | 219.1204 |  | -1.3885 | 384.8756 | -1.3885 | 550.3963 | -1.3885 | 836.1928 | -1.3885 | 1062.9326 | -1.3885 | 1380.176 | -1.3709 | 1722.4596
-1.3534 | 218.9047 |  | -1.3534 | 384.8756 | -1.3534 | 547.8135 | -1.3534 | 836.4735 | -1.3534 | 1064.3338 | -1.3534 | 1380.8698 | -1.3358 | 1722.6723
-1.3182 | 218.9586 |  | -1.3182 | 383.6455 | -1.3182 | 548.9261 | -1.3182 | 838.6254 | -1.3182 | 1064.3338 | -1.3182 | 1381.5629 | -1.3006 | 1725.6556
-1.2831 | 218.8777 |  | -1.2831 | 382.9929 | -1.2831 | 549.5846 | -1.2831 | 840.4033 | -1.2831 | 1067.1623 | -1.2831 | 1380.8698 | -1.2655 | 1728.8628
-1.2479 | 218.9586 |  | -1.2479 | 382.9929 | -1.2479 | 551.6669 | -1.2479 | 840.3098 | -1.2479 | 1067.5183 | -1.2479 | 1380.8698 | -1.2303 | 1729.9343
-1.2128 | 218.9586 |  | -1.2128 | 382.226 | -1.2128 | 551.2599 | -1.2128 | 840.3098 | -1.2128 | 1067.5183 | -1.2128 | 1383.984 | -1.1952 | 1730.3632
-1.1776 | 219.3093 |  | -1.1776 | 382.4943 | -1.1776 | 551.2599 | -1.1776 | 839.4676 | -1.1776 | 1070.5048 | -1.1776 | 1382.14 | -1.16 | 1731.6512
-1.1424 | 219.6334 |  | -1.1424 | 382.4943 | -1.1424 | 551.6669 | -1.1424 | 839.374 | -1.1424 | 1075.2326 | -1.1424 | 1383.7537 | -1.1249 | 1732.5109
-1.1073 | 219.5254 |  | -1.1073 | 383.2999 | -1.1073 | 551.2599 | -1.1073 | 838.9997 | -1.1073 | 1074.378 | -1.1073 | 1382.7167 | -1.0897 | 1737.6852
-1.0721 | 218.6352 |  | -1.0721 | 385.1835 | -1.0721 | 551.2599 | -1.0721 | 838.719 | -1.0721 | 1071.9507 | -1.0721 | 1387.7759 | -1.0546 | 1737.9014
-1.037 | 218.6082 |  | -1.037 | 386.648 | -1.037 | 551.6669 | -1.037 | 840.6841 | -1.037 | 1077.4417 | -1.037 | 1383.7537 | -1.0194 | 1738.5503
-1.0018 | 218.3928 |  | -1.0018 | 388.6962 | -1.0018 | 551.4125 | -1.0018 | 842.3684 | -1.0018 | 1071.9507 | -1.0018 | 1385.7094 | -0.98426 | 1740.4997
-0.96669 | 218.3928 |  | -0.96669 | 386.3394 | -0.96669 | 551.2091 | -0.96669 | 843.5848 | -0.96669 | 1067.9937 | -0.96669 | 1385.9392 | -0.94911 | 1741.8015
-0.93153 | 218.6352 |  | -0.93153 | 386.3009 | -0.93153 | 552.5836 | -0.93153 | 845.9238 | -0.93153 | 1071.5885 | -0.93153 | 1387.7759 | -0.91396 | 1742.4531
-0.89638 | 218.9856 |  | -0.89638 | 388.8123 | -0.89638 | 555.1895 | -0.89638 | 846.1109 | -0.89638 | 1080.9099 | -0.89638 | 1389.1516 | -0.87881 | 1742.2359
-0.86123 | 218.9856 |  | -0.86123 | 388.8123 | -0.86123 | 554.9847 | -0.86123 | 846.2045 | -0.86123 | 1068.7086 | -0.86123 | 1389.8389 | -0.84365 | 1741.5845
-0.82608 | 218.9856 |  | -0.82608 | 390.5178 | -0.82608 | 554.9847 | -0.82608 | 846.2045 | -0.82608 | 1073.0404 | -0.82608 | 1391.7846 | -0.8085 | 1740.2829
-0.79092 | 218.2851 |  | -0.79092 | 390.5178 | -0.79092 | 556.8312 | -0.79092 | 845.3625 | -0.79092 | 1083.9123 | -0.79092 | 1389.4953 | -0.77335 | 1740.2829
-0.75577 | 218.0699 |  | -0.75577 | 390.6343 | -0.75577 | 554.9847 | -0.75577 | 844.2398 | -0.75577 | 1080.9099 | -0.75577 | 1390.1824 | -0.7382 | 1741.8015
-0.72062 | 218.339 |  | -0.72062 | 390.5178 | -0.72062 | 554.9847 | -0.72062 | 843.772 | -0.72062 | 1078.1816 | -0.72062 | 1390.7548 | -0.70304 | 1742.2359
-0.68547 | 218.6082 |  | -0.68547 | 390.5178 | -0.68547 | 556.9339 | -0.68547 | 843.772 | -0.68547 | 1080.9099 | -0.68547 | 1389.2662 | -0.66789 | 1741.8015
-0.65032 | 219.1204 |  | -0.65032 | 391.1391 | -0.65032 | 557.551 | -0.65032 | 847.0463 | -0.65032 | 1082.4078 | -0.65032 | 1389.4953 | -0.63274 | 1741.1504
-0.61516 | 219.7145 |  | -0.61516 | 391.1391 | -0.61516 | 558.7356 | -0.61516 | 850.7865 | -0.61516 | 1082.4078 | -0.61516 | 1391.0981 | -0.59759 | 1739.8495
-0.58001 | 220.391 |  | -0.58001 | 394.4104 | -0.58001 | 559.5096 | -0.58001 | 855.5503 | -0.58001 | 1082.6581 | -0.58001 | 1391.5558 | -0.56244 | 1738.334
-0.54486 | 221.2046 |  | -0.54486 | 394.801 | -0.54486 | 559.6129 | -0.54486 | 856.2969 | -0.54486 | 1087.9555 | -0.54486 | 1387.7759 | -0.52728 | 1735.0945
-0.50971 | 221.2589 |  | -0.50971 | 394.4104 | -0.50971 | 559.5096 | -0.50971 | 856.6701 | -0.50971 | 1089.2279 | -0.50971 | 1389.8389 | -0.49213 | 1735.0945
-0.47455 | 221.6665 |  | -0.47455 | 395.8174 | -0.47455 | 559.5096 | -0.47455 | 858.2557 | -0.47455 | 1085.045 | -0.47455 | 1392.0133 | -0.45698 | 1734.0171
-0.4394 | 221.6665 |  | -0.4394 | 394.7619 | -0.4394 | 558.5809 | -0.4394 | 860.4924 | -0.4394 | 1082.0327 | -0.4394 | 1393.9567 | -0.42183 | 1733.5865
-0.40425 | 221.2046 |  | -0.40425 | 392.9675 | -0.40425 | 558.5809 | -0.40425 | 860.8649 | -0.40425 | 1075.9671 | -0.40425 | 1394.2995 | -0.38667 | 1733.5865
-0.3691 | 220.6348 |  | -0.3691 | 392.9675 | -0.3691 | 559.3547 | -0.3691 | 859.0947 | -0.3691 | 1075.9671 | -0.3691 | 1394.528 | -0.35152 | 1735.3101
-0.33395 | 220.7975 |  | -0.33395 | 390.4402 | -0.33395 | 559.9228 | -0.33395 | 857.6962 | -0.33395 | 1077.1954 | -0.33395 | 1399.2099 | -0.31637 | 1742.2359
-0.29879 | 220.0932 |  | -0.29879 | 389.587 | -0.29879 | 561.9421 | -0.29879 | 856.2969 | -0.29879 | 1081.7829 | -0.29879 | 1398.2966 | -0.28122 | 1744.6282
-0.26364 | 219.5794 |  | -0.26364 | 391.3722 | -0.26364 | 561.9421 | -0.26364 | 850.3192 | -0.26364 | 1082.1577 | -0.26364 | 1399.5524 | -0.24607 | 1744.8459
-0.22849 | 219.9579 |  | -0.22849 | 391.9554 | -0.22849 | 560.5951 | -0.22849 | 850.3192 | -0.22849 | 1083.6611 | -0.22849 | 1399.5524 | -0.21091 | 1744.8459
-0.19334 | 220.1743 |  | -0.19334 | 388.7736 | -0.19334 | 558.2202 | -0.19334 | 849.1973 | -0.19334 | 1083.4101 | -0.19334 | 1398.6391 | -0.17576 | 1742.4531
-0.15818 | 219.6875 |  | -0.15818 | 389.0833 | -0.15818 | 556.5743 | -0.15818 | 851.4407 | -0.15818 | 1080.2878 | -0.15818 | 1397.269 | -0.14061 | 1735.0945
-0.12303 | 219.6605 |  | -0.12303 | 387.4202 | -0.12303 | 556.3689 | -0.12303 | 851.4407 | -0.12303 | 1079.791 | -0.12303 | 1395.4419 | -0.10546 | 1733.5865
-0.087881 | 219.2823 |  | -0.087881 | 387.4202 | -0.087881 | 553.553 | -0.087881 | 848.8233 | -0.087881 | 1078.1816 | -0.087881 | 1386.1689 | -0.070304 | 1733.5865
-0.052728 | 218.8238 |  | -0.052728 | 388.5801 | -0.052728 | 554.5752 | -0.052728 | 846.9528 | -0.052728 | 1074.9882 | -0.052728 | 1386.0541 | -0.035152 | 1733.8018
-0.017576 | 218.3928 |  | -0.017576 | 392.2667 | -0.017576 | 556.9853 | -0.017576 | 846.1109 | -0.017576 | 1073.8909 | -0.017576 | 1384.9046 | 0 | 1733.8018
0.017576 | 218.6082 |  | 0.017576 | 393.2792 | 0.017576 | 557.6539 | 0.017576 | 845.1754 | 0.017576 | 1071.9507 | 0.017576 | 1385.3645 | 0.035152 | 1733.8018
0.052728 | 218.6352 |  | 0.052728 | 393.2792 | 0.052728 | 557.6539 | 0.052728 | 845.1754 | 0.052728 | 1069.4255 | 0.052728 | 1384.0992 | 0.070304 | 1731.007
0.087881 | 218.6082 |  | 0.087881 | 394.3324 | 0.087881 | 560.2847 | 0.087881 | 846.1109 | 0.087881 | 1067.5183 | 0.087881 | 1383.984 | 0.10546 | 1727.5786
0.12303 | 219.2823 |  | 0.12303 | 391.7221 | 0.12303 | 560.4916 | 0.12303 | 845.1754 | 0.12303 | 1067.1623 | 0.12303 | 1382.9473 | 0.14061 | 1727.7925
0.15818 | 219.7145 |  | 0.15818 | 388.4253 | 0.15818 | 560.7503 | 0.15818 | 841.4327 | 0.15818 | 1067.1623 | 0.15818 | 1382.9473 | 0.17576 | 1728.4345
0.19334 | 219.9579 |  | 0.19334 | 386.6095 | 0.19334 | 560.2847 | 0.19334 | 840.7776 | 0.19334 | 1066.5702 | 0.19334 | 1381.7938 | 0.21091 | 1729.077
0.22849 | 220.906 |  | 0.22849 | 385.8768 | 0.22849 | 556.4203 | 0.22849 | 838.3447 | 0.22849 | 1065.3904 | 0.22849 | 1381.5629 | 0.24607 | 1728.0065
0.26364 | 220.906 |  | 0.26364 | 385.7612 | 0.26364 | 556.4203 | 0.26364 | 836.3799 | 0.26364 | 1064.8028 | 0.26364 | 1381.7938 | 0.28122 | 1729.7199
0.29879 | 220.906 |  | 0.29879 | 385.3375 | 0.29879 | 560.4399 | 0.29879 | 836.3799 | 0.29879 | 1063.6321 | 0.29879 | 1383.0625 | 0.31637 | 1727.5786
0.33395 | 221.0145 |  | 0.33395 | 385.3375 | 0.33395 | 560.4399 | 0.33395 | 838.9997 | 0.33395 | 1064.8028 | 0.33395 | 1383.0625 | 0.35152 | 1726.9372
0.3691 | 221.449 |  | 0.3691 | 385.3375 | 0.3691 | 555.6507 | 0.3691 | 837.0348 | 0.3691 | 1067.9937 | 0.3691 | 1383.0625 | 0.38667 | 1727.1509
0.40425 | 221.5577 |  | 0.40425 | 391.4888 | 0.40425 | 555.6507 | 0.40425 | 840.3098 | 0.40425 | 1069.9046 | 0.40425 | 1383.8689 | 0.42183 | 1728.2205
0.4394 | 221.5577 |  | 0.4394 | 391.4888 | 0.4394 | 560.4399 | 0.4394 | 845.0819 | 0.4394 | 1069.9046 | 0.4394 | 1380.8698 | 0.45698 | 1731.4364
0.47455 | 221.5034 |  | 0.47455 | 391.4888 | 0.47455 | 555.1895 | 0.47455 | 841.6198 | 0.47455 | 1070.1445 | 0.47455 | 1379.8289 | 0.49213 | 1735.3101
0.50971 | 220.1743 |  | 0.50971 | 390.9061 | 0.50971 | 548.2179 | 0.50971 | 838.6254 | 0.50971 | 1071.5885 | 0.50971 | 1380.8698 | 0.52728 | 1735.3101
0.54486 | 219.6875 |  | 0.54486 | 390.6343 | 0.54486 | 557.2938 | 0.54486 | 845.7367 | 0.54486 | 1069.9046 | 0.54486 | 1380.8698 | 0.56244 | 1731.2217
0.58001 | 219.6875 |  | 0.58001 | 388.851 | 0.58001 | 559.1482 | 0.58001 | 843.117 | 0.58001 | 1069.5452 | 0.58001 | 1380.8698 | 0.59759 | 1731.6512
0.61516 | 220.8789 |  | 0.61516 | 388.2706 | 0.61516 | 550.5486 | 0.61516 | 838.1576 | 0.61516 | 1071.7092 | 0.61516 | 1381.2164 | 0.63274 | 1731.2217
0.65032 | 221.9657 |  | 0.65032 | 389.5482 | 0.65032 | 550.5486 | 0.65032 | 841.6198 | 0.65032 | 1070.3846 | 0.65032 | 1382.9473 | 0.66789 | 1731.6512
0.68547 | 223.2474 |  | 0.68547 | 390.5955 | 0.68547 | 557.3452 | 0.68547 | 849.0103 | 0.68547 | 1067.637 | 0.68547 | 1383.5234 | 0.70304 | 1735.0945
0.72062 | 223.8491 |  | 0.72062 | 389.1608 | 0.72062 | 555.4457 | 0.72062 | 842.1813 | 0.72062 | 1071.7092 | 0.72062 | 1385.2496 | 0.7382 | 1736.1731
0.75577 | 223.9312 |  | 0.75577 | 389.1608 | 0.75577 | 552.4817 | 0.75577 | 842.1813 | 0.75577 | 1073.2831 | 0.75577 | 1385.7094 | 0.77335 | 1736.389
0.79092 | 224.397 |  | 0.79092 | 389.4707 | 0.79092 | 555.4457 | 0.79092 | 853.4028 | 0.79092 | 1072.7979 | 0.79092 | 1387.8906 | 0.8085 | 1736.6049
0.82608 | 225.468 |  | 0.82608 | 389.1608 | 0.82608 | 555.5482 | 0.82608 | 863.2849 | 0.82608 | 1074.5 | 0.82608 | 1385.7094 | 0.84365 | 1736.8208
0.86123 | 226.4045 |  | 0.86123 | 389.4707 | 0.86123 | 553.8083 | 0.86123 | 851.4407 | 0.86123 | 1074.622 | 0.86123 | 1385.7094 | 0.87881 | 1737.469
0.89638 | 226.5148 |  | 0.89638 | 390.1686 | 0.89638 | 554.4218 | 0.89638 | 851.2538 | 0.89638 | 1075.7221 | 0.89638 | 1384.6745 | 0.91396 | 1737.469
0.93153 | 226.8461 |  | 0.93153 | 390.4014 | 0.93153 | 558.1172 | 0.93153 | 850.2257 | 0.93153 | 1076.5806 | 0.93153 | 1387.7759 | 0.94911 | 1736.389
0.96669 | 227.15 |  | 0.96669 | 391.2945 | 0.96669 | 554.4218 | 0.96669 | 850.2257 | 0.96669 | 1077.0723 | 0.96669 | 1387.7759 | 0.98426 | 1736.389
1.0018 | 227.4266 |  | 1.0018 | 391.9943 | 1.0018 | 557.3452 | 1.0018 | 850.2257 | 1.0018 | 1077.4417 | 1.0018 | 1390.4114 | 1.0194 | 1736.389
1.037 | 227.4543 |  | 1.037 | 392.9286 | 1.037 | 563.4485 | 1.037 | 863.3779 | 1.037 | 1077.5649 | 1.037 | 1389.1516 | 1.0546 | 1737.0369
1.0721 | 227.7311 |  | 1.0721 | 393.7471 | 1.0721 | 558.1172 | 1.0721 | 875.6066 | 1.0721 | 1081.0345 | 1.0721 | 1390.4114 | 1.0897 | 1739.6328
1.1073 | 228.1744 |  | 1.1073 | 395.7782 | 1.1073 | 560.9574 | 1.1073 | 875.6066 | 1.1073 | 1081.0345 | 1.1073 | 1393.0424 | 1.1249 | 1739.8495
1.1424 | 228.424 |  | 1.1424 | 396.8353 | 1.1424 | 564.5941 | 1.1424 | 871.2647 | 1.1424 | 1081.0345 | 1.1424 | 1393.9567 | 1.16 | 1739.6328
1.1776 | 228.7016 |  | 1.1776 | 396.5219 | 1.1776 | 561.0609 | 1.1776 | 862.7267 | 1.1776 | 1080.5365 | 1.1776 | 1396.013 | 1.1952 | 1739.6328
1.2128 | 229.3409 |  | 1.2128 | 397.541 | 1.2128 | 557.9112 | 1.2128 | 851.1604 | 1.2128 | 1080.5365 | 1.2128 | 1401.8361 | 1.2303 | 1739.4162
1.2479 | 229.5914 |  | 1.2479 | 398.2081 | 1.2479 | 565.6899 | 1.2479 | 842.6491 | 1.2479 | 1080.5365 | 1.2479 | 1401.8361 | 1.2655 | 1739.6328
1.2831 | 229.7585 |  | 1.2831 | 397.7763 | 1.2831 | 566.631 | 1.2831 | 839.4676 | 1.2831 | 1084.919 | 1.2831 | 1401.3793 | 1.3006 | 1740.9335
1.3182 | 230.5394 |  | 1.3182 | 398.2081 | 1.3182 | 565.5333 | 1.3182 | 839.8419 | 1.3182 | 1084.919 | 1.3182 | 1400.6942 | 1.3358 | 1739.8495
1.3534 | 230.9864 |  | 1.3534 | 399.0331 | 1.3534 | 565.5333 | 1.3534 | 844.1462 | 1.3534 | 1084.919 | 1.3534 | 1398.2966 | 1.3709 | 1739.6328
1.3885 | 231.2661 |  | 1.3885 | 399.0724 | 1.3885 | 565.5333 | 1.3885 | 855.2703 | 1.3885 | 1085.8021 | 1.3885 | 1393.7282 | 1.4061 | 1735.0945
1.4237 | 232.0223 |  | 1.4237 | 399.1117 | 1.4237 | 569.622 | 1.4237 | 862.2614 | 1.4237 | 1080.5365 | 1.4237 | 1390.4114 | 1.4412 | 1733.8018
1.4588 | 232.1625 |  | 1.4588 | 400.8831 | 1.4588 | 570.2538 | 1.4588 | 877.6336 | 1.4588 | 1077.8114 | 1.4588 | 1386.0541 | 1.4764 | 1736.1731
1.494 | 232.6678 |  | 1.494 | 401.3561 | 1.494 | 570.2538 | 1.494 | 878.0017 | 1.494 | 1085.8021 | 1.494 | 1385.0196 | 1.5115 | 1736.1731
1.5291 | 232.5554 |  | 1.5291 | 402.5794 | 1.5291 | 570.2538 | 1.5291 | 878.0017 | 1.5291 | 1088.2097 | 1.5291 | 1385.1346 | 1.5467 | 1740.9335
1.5643 | 232.724 |  | 1.5643 | 404.5959 | 1.5643 | 569.5168 | 1.5643 | 873.5761 | 1.5643 | 1089.738 | 1.5643 | 1385.5944 | 1.5818 | 1741.5845
1.5994 | 232.6959 |  | 1.5994 | 402.9349 | 1.5994 | 569.5168 | 1.5994 | 865.4229 | 1.5994 | 1097.3336 | 1.5994 | 1393.4996 | 1.617 | 1743.3225
1.6346 | 232.2747 |  | 1.6346 | 402.1451 | 1.6346 | 569.5168 | 1.6346 | 865.2371 | 1.6346 | 1093.3264 | 1.6346 | 1393.4996 | 1.6522 | 1743.54
1.6697 | 232.0223 |  | 1.6697 | 401.9083 | 1.6697 | 574.1665 | 1.6697 | 865.4229 | 1.6697 | 1093.3264 | 1.6697 | 1393.8425 | 1.6873 | 1742.6704
1.7049 | 230.5673 |  | 1.7049 | 402.3425 | 1.7049 | 574.1665 | 1.7049 | 865.4229 | 1.7049 | 1093.5839 | 1.7049 | 1395.4419 | 1.7225 | 1740.4997
1.74 | 230.3719 |  | 1.74 | 402.9744 | 1.74 | 576.1335 | 1.74 | 873.1142 | 1.74 | 1097.3336 | 1.74 | 1395.4419 | 1.7576 | 1740.2829
1.7752 | 230.2882 |  | 1.7752 | 403.2905 | 1.7752 | 574.1665 | 1.7752 | 874.4072 | 1.7752 | 1097.983 | 1.7752 | 1393.6139 | 1.7928 | 1740.2829
1.8103 | 230.3719 |  | 1.8103 | 403.2905 | 1.8103 | 576.1335 | 1.8103 | 872.3747 | 1.8103 | 1098.5032 | 1.8103 | 1394.071 | 1.8279 | 1740.2829
1.8455 | 232.0503 |  | 1.8455 | 403.251 | 1.8455 | 572.5239 | 1.8455 | 869.5054 | 1.8455 | 1099.1542 | 1.8455 | 1404.921 | 1.8631 | 1744.1928
1.8806 | 233.005 |  | 1.8806 | 402.8163 | 1.8806 | 572.9473 | 1.8806 | 873.6685 | 1.8806 | 1099.6756 | 1.8806 | 1401.9503 | 1.8982 | 1745.7175
1.9158 | 233.005 |  | 1.9158 | 401.5533 | 1.9158 | 572.9473 | 1.9158 | 872.3747 | 1.9158 | 1100.4587 | 1.9158 | 1398.0682 | 1.9334 | 1745.9355
1.9509 | 232.4712 |  | 1.9509 | 401.711 | 1.9509 | 574.2196 | 1.9509 | 872.0048 | 1.9509 | 1100.4587 | 1.9509 | 1403.6638 | 1.9685 | 1748.337
1.9861 | 232.0503 |  | 1.9861 | 402.1056 | 1.9861 | 574.5381 | 1.9861 | 872.8369 | 1.9861 | 1101.8973 | 1.9861 | 1404.5781 | 2.0037 | 1748.9929
2.0213 | 231.7981 |  | 2.0213 | 403.4091 | 2.0213 | 574.9099 | 2.0213 | 873.0218 | 2.0213 | 1103.471 | 2.0213 | 1402.2929 | 2.0388 | 1750.5252
2.0564 | 231.658 |  | 2.0564 | 404.6751 | 2.0564 | 575.0694 | 2.0564 | 872.0048 | 2.0564 | 1104.9173 | 2.0564 | 1404.4638 | 2.074 | 1750.9634
2.0916 | 231.658 |  | 2.0916 | 406.8161 | 2.0916 | 575.7076 | 2.0916 | 872.1898 | 2.0916 | 1105.9713 | 2.0916 | 1406.8655 | 2.1091 | 1752.4988
2.1267 | 232.3589 |  | 2.1267 | 408.3262 | 2.1267 | 575.3352 | 2.1267 | 870.802 | 2.1267 | 1110.4703 | 2.1267 | 1402.2929 | 2.1443 | 1755.3565
2.1619 | 233.2301 |  | 2.1619 | 409.5604 | 2.1619 | 576.4532 | 2.1619 | 872.9294 | 2.1619 | 1113.2634 | 2.1619 | 1406.8655 | 2.1794 | 1758.2224
2.197 | 233.8496 |  | 2.197 | 410.9163 | 2.197 | 576.773 | 2.197 | 876.8048 | 2.197 | 1115.3982 | 2.197 | 1411.2212 | 2.2146 | 1757.3397
2.2322 | 234.2727 |  | 2.2322 | 410.0387 | 2.2322 | 577.5735 | 2.2322 | 878.6457 | 2.2322 | 1115.6654 | 2.2322 | 1406.0646 | 2.2497 | 1755.3565
2.2673 | 234.5267 |  | 2.2673 | 408.6445 | 2.2673 | 577.7337 | 2.2673 | 878.6457 | 2.2673 | 1115.6654 | 2.2673 | 1402.864 | 2.2849 | 1754.9163
2.3025 | 234.7245 |  | 2.3025 | 409.5604 | 2.3025 | 579.0173 | 2.3025 | 880.9422 | 2.3025 | 1110.2048 | 2.3025 | 1411.2212 | 2.32 | 1751.6212
2.3376 | 234.8092 |  | 2.3376 | 409.6799 | 2.3376 | 580.4649 | 2.3376 | 881.3092 | 2.3376 | 1105.7076 | 2.3376 | 1411.1064 | 2.3552 | 1749.6493
2.3728 | 235.6298 |  | 2.3728 | 408.3262 | 2.3728 | 581.6474 | 2.3728 | 882.9588 | 2.3728 | 1103.9965 | 2.3728 | 1411.1064 | 2.3903 | 1751.6212
2.4079 | 235.0071 |  | 2.4079 | 407.8093 | 2.4079 | 583.8039 | 2.4079 | 883.7826 | 2.4079 | 1103.9965 | 2.4079 | 1412.6 | 2.4255 | 1753.8167
2.4431 | 235.2051 |  | 2.4431 | 408.1274 | 2.4431 | 584.1822 | 2.4431 | 886.3406 | 2.4431 | 1105.7076 | 2.4431 | 1411.1064 | 2.4607 | 1757.3397
2.4782 | 235.1203 |  | 2.4782 | 407.4913 | 2.4782 | 584.2903 | 2.4782 | 884.6056 | 2.4782 | 1107.5557 | 2.4782 | 1401.8361 | 2.4958 | 1759.769
2.5134 | 234.781 |  | 2.5134 | 406.8161 | 2.5134 | 584.2903 | 2.5134 | 887.9811 | 2.5134 | 1108.6141 | 2.5134 | 1395.2135 | 2.531 | 1766.2018
2.5485 | 234.5267 |  | 2.5485 | 407.2529 | 2.5485 | 584.2903 | 2.5485 | 891.1615 | 2.5485 | 1111.9321 | 2.5485 | 1393.8425 | 2.5661 | 1766.2018
2.5837 | 234.8092 |  | 2.5837 | 407.69 | 2.5837 | 584.3444 | 2.5837 | 893.9676 | 2.5837 | 1113.1302 | 2.5837 | 1392.8138 | 2.6013 | 1765.9793
2.6188 | 234.8092 |  | 2.6188 | 408.6445 | 2.6188 | 584.9937 | 2.6188 | 898.8294 | 2.6188 | 1113.1302 | 2.6188 | 1392.9281 | 2.6364 | 1763.979
2.654 | 234.2727 |  | 2.654 | 410.3976 | 2.654 | 585.427 | 2.654 | 900.8 | 2.654 | 1112.8638 | 2.654 | 1393.8425 | 2.6716 | 1763.0912
2.6891 | 234.2727 |  | 2.6891 | 411.9149 | 2.6891 | 586.6203 | 2.6891 | 903.3883 | 2.6891 | 1112.8638 | 2.6891 | 1397.3832 | 2.7067 | 1763.0912
2.7243 | 235.0071 |  | 2.7243 | 411.9548 | 2.7243 | 589.2331 | 2.7243 | 903.3883 | 2.7243 | 1112.0651 | 2.7243 | 1399.2099 | 2.7419 | 1766.6469
2.7594 | 234.9506 |  | 2.7594 | 411.9149 | 2.7594 | 589.9975 | 2.7594 | 903.6554 | 2.7594 | 1111.9321 | 2.7594 | 1404.8067 | 2.777 | 1768.4295
2.7946 | 234.8375 |  | 2.7946 | 411.1958 | 2.7946 | 590.3254 | 2.7946 | 903.6554 | 2.7946 | 1112.0651 | 2.7946 | 1407.5523 | 2.8122 | 1769.3219
2.8298 | 234.9223 |  | 2.8298 | 410.637 | 2.8298 | 590.1068 | 2.8298 | 903.6554 | 2.8298 | 1112.1981 | 2.8298 | 1404.8067 | 2.8473 | 1767.9836
2.8649 | 234.8375 |  | 2.8649 | 410.5971 | 2.8649 | 586.5117 | 2.8649 | 899.9051 | 2.8649 | 1111.5331 | 2.8649 | 1400.0091 | 2.8825 | 1765.5344
2.9001 | 234.2727 |  | 2.9001 | 410.4774 | 2.9001 | 584.02 | 2.9001 | 893.6965 | 2.9001 | 1110.3375 | 2.9001 | 1394.8708 | 2.9176 | 1760.2114
2.9352 | 232.9769 |  | 2.9352 | 410.3976 | 2.9352 | 574.2726 | 2.9352 | 893.6965 | 2.9352 | 1108.217 | 2.9352 | 1392.9281 | 2.9528 | 1757.1192
2.9704 | 232.6678 |  | 2.9704 | 408.3262 | 2.9704 | 570.5173 | 2.9704 | 888.2542 | 2.9704 | 1106.631 | 2.9704 | 1392.242 | 2.9879 | 1755.5767
3.0055 | 232.724 |  | 3.0055 | 407.9683 | 3.0055 | 570.3065 | 3.0055 | 888.2542 | 3.0055 | 1103.0772 | 3.0055 | 1391.5558 | 3.0231 | 1752.9379
3.0407 | 232.6678 |  | 3.0407 | 407.9683 | 3.0407 | 570.5173 | 3.0407 | 890.9801 | 3.0407 | 1101.7664 | 3.0407 | 1389.8389 | 3.0582 | 1751.6212
3.0758 | 232.8083 |  | 3.0758 | 407.9683 | 3.0758 | 574.5912 | 3.0758 | 889.0729 | 3.0758 | 1100.5893 | 3.0758 | 1391.5558 | 3.0934 | 1748.7742
3.111 | 232.6678 |  | 3.111 | 408.0876 | 3.111 | 578.5891 | 3.111 | 885.6105 | 3.111 | 1098.2431 | 3.111 | 1393.8425 | 3.1285 | 1748.7742
3.1461 | 231.8541 |  | 3.1461 | 408.8037 | 3.1461 | 581.1633 | 3.1461 | 884.6056 | 3.1461 | 1096.8146 | 3.1461 | 1390.7548 | 3.1637 | 1748.7742
3.1813 | 231.434 |  | 3.1813 | 408.1274 | 3.1813 | 582.509 | 3.1813 | 882.7757 | 3.1813 | 1096.8146 | 3.1813 | 1386.6283 | 3.1989 | 1748.7742
3.2164 | 231.378 |  | 3.2164 | 408.1274 | 3.2164 | 582.7785 | 3.2164 | 882.6841 | 3.2164 | 1097.074 | 3.2164 | 1389.7244 | 3.234 | 1744.8459
3.2516 | 231.322 |  | 3.2516 | 408.3262 | 3.2516 | 582.4012 | 3.2516 | 884.057 | 3.2516 | 1097.983 | 3.2516 | 1390.0679 | 3.2692 | 1742.0187
3.2867 | 231.686 |  | 3.2867 | 407.9683 | 3.2867 | 582.4012 | 3.2867 | 883.8741 | 3.2867 | 1098.2431 | 3.2867 | 1392.242 | 3.3043 | 1740.7166
3.3219 | 231.686 |  | 3.3219 | 407.5311 | 3.3219 | 582.7785 | 3.3219 | 880.6669 | 3.3219 | 1100.5893 | 3.3219 | 1390.0679 | 3.3395 | 1740.7166
3.357 | 232.4151 |  | 3.357 | 407.0146 | 3.357 | 582.4012 | 3.357 | 881.676 | 3.357 | 1104.6541 | 3.357 | 1386.3987 | 3.3746 | 1740.4997
3.3922 | 232.4151 |  | 3.3922 | 405.4674 | 3.3922 | 581.3784 | 3.3922 | 878.3697 | 3.3922 | 1106.8951 | 3.3922 | 1399.2099 | 3.4098 | 1736.6049
3.4273 | 232.387 |  | 3.4273 | 405.4674 | 3.4273 | 582.5629 | 3.4273 | 878.0017 | 3.4273 | 1111.0015 | 3.4273 | 1403.8923 | 3.4449 | 1740.4997
3.4625 | 231.8541 |  | 3.4625 | 405.6656 | 3.4625 | 582.1857 | 3.4625 | 877.6336 | 3.4625 | 1111.0015 | 3.4625 | 1393.1567 | 3.4801 | 1740.9335
3.4976 | 231.574 |  | 3.4976 | 405.4674 | 3.4976 | 582.0242 | 3.4976 | 880.5751 | 3.4976 | 1107.6879 | 3.4976 | 1393.1567 | 3.5152 | 1740.7166
3.5328 | 230.9584 |  | 3.5328 | 405.6656 | 3.5328 | 582.3473 | 3.5328 | 882.5925 | 3.5328 | 1107.6879 | 3.5328 | 1399.2099 | 3.5504 | 1738.334
3.5679 | 230.8746 |  | 3.5679 | 406.1018 | 3.5679 | 582.3473 | 3.5679 | 884.2399 | 3.5679 | 1107.9524 | 3.5679 | 1402.0645 | 3.5855 | 1738.334
3.6031 | 230.3998 |  | 3.6031 | 406.3398 | 3.6031 | 582.3473 | 3.6031 | 887.9811 | 3.6031 | 1106.631 | 3.6031 | 1399.7808 | 3.6207 | 1738.334
3.6383 | 230.1487 |  | 3.6383 | 406.697 | 3.6383 | 582.4012 | 3.6383 | 886.8878 | 3.6383 | 1106.763 | 3.6383 | 1400.6942 | 3.6558 | 1737.0369
3.6734 | 230.3998 |  | 3.6734 | 407.2529 | 3.6734 | 582.4012 | 3.6734 | 886.8878 | 3.6734 | 1107.9524 | 3.6734 | 1403.0925 | 3.691 | 1737.0369
3.7086 | 230.6511 |  | 3.7086 | 407.8093 | 3.7086 | 582.4012 | 3.7086 | 887.4346 | 3.7086 | 1107.9524 | 3.7086 | 1403.0925 | 3.7261 | 1739.8495
3.7437 | 230.4556 |  | 3.7437 | 408.4456 | 3.7437 | 582.4012 | 3.7437 | 887.9811 | 3.7437 | 1106.1032 | 3.7437 | 1400.6942 | 3.7613 | 1743.9751
3.7789 | 230.1209 |  | 3.7789 | 408.6843 | 3.7789 | 582.4012 | 3.7789 | 887.4346 | 3.7789 | 1108.879 | 3.7789 | 1399.4382 | 3.7964 | 1745.4996
3.814 | 228.7294 |  | 3.814 | 407.7695 | 3.814 | 582.4551 | 3.814 | 887.799 | 3.814 | 1109.8068 | 3.814 | 1399.4382 | 3.8316 | 1745.4996
3.8492 | 228.2576 |  | 3.8492 | 407.2529 | 3.8492 | 582.4012 | 3.8492 | 888.3452 | 3.8492 | 1110.3375 | 3.8492 | 1403.2068 | 3.8667 | 1747.6814
3.8843 | 227.925 |  | 3.8843 | 407.1735 | 3.8843 | 581.8088 | 3.8843 | 887.799 | 3.8843 | 1110.3375 | 3.8843 | 1403.2068 | 3.9019 | 1746.5899
3.9195 | 227.8972 |  | 3.9195 | 407.1735 | 3.9195 | 581.755 | 3.9195 | 886.6142 | 3.9195 | 1111.799 | 3.9195 | 1405.6071 | 3.937 | 1745.4996
3.9546 | 228.2576 |  | 3.9546 | 407.1735 | 3.9546 | 580.5186 | 3.9546 | 886.8878 | 3.9546 | 1112.7306 | 3.9546 | 1405.8358 | 3.9722 | 1747.6814
3.9898 | 228.7572 |  | 3.9898 | 407.1735 | 3.9898 | 579.3922 | 3.9898 | 888.982 | 3.9898 | 1115.1311 | 3.9898 | 1407.7814 | 4.0074 | 1748.337
4.0249 | 228.5906 |  | 4.0249 | 406.6573 | 4.0249 | 578.3216 | 4.0249 | 892.249 | 4.0249 | 1113.53 | 4.0249 | 1407.0944 | 4.0425 | 1747.8999
4.0601 | 228.8683 |  | 4.0601 | 406.7367 | 4.0601 | 579.0173 | 4.0601 | 892.7922 | 4.0601 | 1115.5318 | 4.0601 | 1409.1562 | 4.0777 | 1748.337
4.0952 | 228.8683 |  | 4.0952 | 405.5863 | 4.0952 | 576.9864 | 4.0952 | 893.4253 | 4.0952 | 1115.5318 | 4.0952 | 1405.264 | 4.1128 | 1748.9929
4.1304 | 229.1462 |  | 4.1304 | 405.5863 | 4.1304 | 578.4286 | 4.1304 | 892.7922 | 4.1304 | 1117.4044 | 4.1304 | 1404.1209 | 4.148 | 1751.6212
4.1655 | 229.5636 |  | 4.1655 | 405.7449 | 4.1655 | 579.1779 | 4.1655 | 889.2548 | 4.1655 | 1116.3339 | 4.1655 | 1403.321 | 4.1831 | 1753.3772
4.2007 | 230.7628 |  | 4.2007 | 406.8558 | 4.2007 | 579.7674 | 4.2007 | 886.7966 | 4.2007 | 1116.869 | 4.2007 | 1405.1497 | 4.2183 | 1754.6963
4.2358 | 231.518 |  | 4.2358 | 407.849 | 4.2358 | 581.8626 | 4.2358 | 886.7966 | 4.2358 | 1116.3339 | 4.2358 | 1405.6071 | 4.2534 | 1759.548
4.271 | 231.574 |  | 4.271 | 408.1671 | 4.271 | 582.1857 | 4.271 | 886.8878 | 4.271 | 1114.0633 | 4.271 | 1405.9502 | 4.2886 | 1760.8752
4.3061 | 230.8466 |  | 4.3061 | 408.3262 | 4.3061 | 582.1857 | 4.3061 | 887.89 | 4.3061 | 1113.9299 | 4.3061 | 1406.98 | 4.3237 | 1761.761
4.3413 | 230.8466 |  | 4.3413 | 408.2069 | 4.3413 | 581.3246 | 4.3413 | 888.1632 | 4.3413 | 1113.7966 | 4.3413 | 1406.7511 | 4.3589 | 1761.0966
4.3764 | 230.9025 |  | 4.3764 | 408.2069 | 4.3764 | 578.7496 | 4.3764 | 888.2542 | 4.3764 | 1117.5383 | 4.3764 | 1406.98 | 4.394 | 1761.5395
4.4116 | 230.9584 |  | 4.4116 | 409.4409 | 4.4116 | 578.7496 | 4.4116 | 888.1632 | 4.4116 | 1120.0861 | 4.4116 | 1408.354 | 4.4292 | 1762.4259
4.4468 | 232.0223 |  | 4.4468 | 409.6003 | 4.4468 | 579.3922 | 4.4468 | 887.9811 | 4.4468 | 1120.7576 | 4.4468 | 1411.2212 | 4.4643 | 1763.979
4.4819 | 232.2467 |  | 4.4819 | 410.2779 | 4.4819 | 579.7674 | 4.4819 | 888.4362 | 4.4819 | 1121.2951 | 4.4819 | 1412.0253 | 4.4995 | 1767.0923
4.5171 | 232.2467 |  | 4.5171 | 410.2779 | 4.5171 | 581.1096 | 4.5171 | 890.8894 | 4.5171 | 1120.7576 | 4.5171 | 1412.9451 | 4.5346 | 1769.9917
4.5522 | 232.0784 |  | 4.5522 | 410.2779 | 4.5522 | 581.8088 | 4.5522 | 891.5242 | 4.5522 | 1122.2364 | 4.5522 | 1414.2112 | 4.5698 | 1769.9917
4.5874 | 232.0784 |  | 4.5874 | 410.1982 | 4.5874 | 582.7246 | 4.5874 | 893.3349 | 4.5874 | 1122.6401 | 4.5874 | 1414.9026 | 4.6049 | 1769.9917
4.6225 | 232.8083 |  | 4.6225 | 410.2779 | 4.6225 | 583.5339 | 4.6225 | 894.3289 | 4.6225 | 1124.93 | 4.6225 | 1416.7492 | 4.6401 | 1775.3661
4.6577 | 232.5274 |  | 4.6577 | 410.3976 | 4.6577 | 583.1561 | 4.6577 | 897.0328 | 4.6577 | 1126.6838 | 4.6577 | 1419.6437 | 4.6752 | 1774.4684
4.6928 | 232.3028 |  | 4.6928 | 410.5173 | 4.6928 | 585.102 | 4.6928 | 898.65 | 4.6928 | 1128.5749 | 4.6928 | 1420.108 | 4.7104 | 1777.8387
4.728 | 232.6116 |  | 4.728 | 410.8365 | 4.728 | 585.427 | 4.728 | 901.8723 | 4.728 | 1133.1759 | 4.728 | 1422.085 | 4.7455 | 1778.9645
4.7631 | 231.686 |  | 4.7631 | 411.076 | 4.7631 | 585.427 | 4.7631 | 903.0319 | 4.7631 | 1133.7179 | 4.7631 | 1424.5364 | 4.7807 | 1781.6715
4.7983 | 231.0982 |  | 4.7983 | 411.2756 | 4.7983 | 585.2103 | 4.7983 | 903.9225 | 4.7983 | 1133.9889 | 4.7983 | 1425.7076 | 4.8159 | 1781.6715
4.8334 | 230.9864 |  | 4.8334 | 411.2756 | 4.8334 | 586.1318 | 4.8334 | 903.9225 | 4.8334 | 1133.9889 | 4.8334 | 1425.9422 | 4.851 | 1784.8384
4.8686 | 231.0982 |  | 4.8686 | 412.5146 | 4.8686 | 587.7618 | 4.8686 | 902.8536 | 4.8686 | 1135.3445 | 4.8686 | 1426.999 | 4.8862 | 1788.6967
4.9037 | 231.406 |  | 4.9037 | 412.7947 | 4.9037 | 588.4152 | 4.9037 | 903.3883 | 4.9037 | 1140.6359 | 4.9037 | 1431.4869 | 4.9213 | 1793.939
4.9389 | 231.9382 |  | 4.9389 | 412.7947 | 4.9389 | 588.1973 | 4.9389 | 903.2101 | 4.9389 | 1141.9935 | 4.9389 | 1431.6056 | 4.9565 | 1803.5776
4.974 | 232.2186 |  | 4.974 | 413.0749 | 4.974 | 588.9058 | 4.974 | 902.9428 | 4.974 | 1141.4505 | 4.974 | 1432.4374 | 4.9916 | 1807.5031
5.0092 | 232.3028 |  | 5.0092 | 412.9548 | 5.0092 | 590.3254 | 5.0092 | 905.0785 | 5.0092 | 1141.4505 | 5.0092 | 1433.7478 | 5.0268 | 1807.5031
5.0443 | 232.4151 |  | 5.0443 | 412.1147 | 5.0443 | 590.5988 | 5.0443 | 908.0931 | 5.0443 | 1140.9074 | 5.0443 | 1438.7906 | 5.0619 | 1807.9659
5.0795 | 232.6397 |  | 5.0795 | 411.6351 | 5.0795 | 591.1461 | 5.0795 | 908.5354 | 5.0795 | 1141.0432 | 5.0795 | 1441.2156 | 5.0971 | 1808.4288
5.1146 | 232.8926 |  | 5.1146 | 411.4354 | 5.1146 | 593.3407 | 5.1146 | 910.8313 | 5.1146 | 1140.093 | 5.1146 | 1442.4342 | 5.1322 | 1810.5144
5.1498 | 233.3989 |  | 5.1498 | 411.1559 | 5.1498 | 594.2211 | 5.1498 | 913.912 | 5.1498 | 1141.4505 | 5.1498 | 1444.2702 | 5.1674 | 1810.9784
5.1849 | 234.1034 |  | 5.1849 | 411.2357 | 5.1849 | 595.2131 | 5.1849 | 916.895 | 5.1849 | 1141.8578 | 5.1849 | 1446.61 | 5.2025 | 1814.6972
5.2201 | 234.3291 |  | 5.2201 | 411.5552 | 5.2201 | 595.7098 | 5.2201 | 918.4707 | 5.2201 | 1144.1658 | 5.2201 | 1447.2284 | 5.2377 | 1815.6288
5.2553 | 235.29 |  | 5.2553 | 415.6815 | 5.2553 | 596.207 | 5.2553 | 920.3064 | 5.2553 | 1146.4738 | 5.2553 | 1447.2284 | 5.2728 | 1818.8956
5.2904 | 236.85 |  | 5.2904 | 418.2572 | 5.2904 | 597.3688 | 5.2904 | 922.3138 | 5.2904 | 1149.4597 | 5.2904 | 1448.7176 | 5.308 | 1821.9374
5.3256 | 237.1913 |  | 5.3256 | 419.7909 | 5.3256 | 597.3688 | 5.3256 | 923.7957 | 5.3256 | 1155.9656 | 5.3256 | 1448.842 | 5.3431 | 1825.6923
5.3607 | 237.7323 |  | 5.3607 | 420.2355 | 5.3607 | 597.3688 | 5.3607 | 923.7957 | 5.3607 | 1158.5358 | 5.3607 | 1449.4648 | 5.3783 | 1831.1114
5.3959 | 238.0174 |  | 5.3959 | 420.9231 | 5.3959 | 598.3111 | 5.3959 | 923.0113 | 5.3959 | 1159.2115 | 5.3959 | 1451.0897 | 5.4134 | 1832.2928
5.431 | 237.9033 |  | 5.431 | 421.004 | 5.431 | 600.479 | 5.431 | 923.0113 | 5.431 | 1161.102 | 5.431 | 1453.1013 | 5.4486 | 1833.2388
5.4662 | 237.7893 |  | 5.4662 | 421.3279 | 5.4662 | 601.5941 | 5.4662 | 922.401 | 5.4662 | 1162.8552 | 5.4662 | 1453.1013 | 5.4837 | 1832.2928
5.5013 | 237.5899 |  | 5.5013 | 420.8826 | 5.5013 | 602.3759 | 5.5013 | 921.4414 | 5.5013 | 1165.8166 | 5.5013 | 1451.0897 | 5.5189 | 1830.8753
5.5365 | 237.5044 |  | 5.5365 | 420.8826 | 5.5365 | 604.0552 | 5.5365 | 921.0049 | 5.5365 | 1169.0382 | 5.5365 | 1450.3387 | 5.554 | 1832.5293
5.5716 | 237.5044 |  | 5.5716 | 419.4273 | 5.5716 | 605.5709 | 5.5716 | 921.0922 | 5.5716 | 1172.3826 | 5.5716 | 1446.7336 | 5.5892 | 1833.7121
5.6068 | 237.2198 |  | 5.6068 | 419.4273 | 5.6068 | 605.9083 | 5.6068 | 921.0922 | 5.6068 | 1174.2498 | 5.6068 | 1444.8843 | 5.6244 | 1832.7657
5.6419 | 236.7648 |  | 5.6419 | 418.7411 | 5.6419 | 606.077 | 5.6419 | 921.4414 | 5.6419 | 1174.2498 | 5.6419 | 1444.8843 | 5.6595 | 1835.6072
5.6771 | 237.2767 |  | 5.6771 | 419.5485 | 5.6771 | 605.9083 | 5.6771 | 924.2313 | 5.6771 | 1171.5811 | 5.6771 | 1446.2394 | 5.6947 | 1837.0305
5.7122 | 238.3598 |  | 5.7122 | 419.5485 | 5.7122 | 605.9083 | 5.7122 | 926.5814 | 5.7122 | 1170.1098 | 5.7122 | 1446.2394 | 5.7298 | 1842.5023
5.7474 | 238.0459 |  | 5.7474 | 419.8313 | 5.7474 | 605.9645 | 5.7474 | 927.1033 | 5.7474 | 1167.8313 | 5.7474 | 1447.3523 | 5.765 | 1837.2678
5.7825 | 238.9311 |  | 5.7825 | 420.0333 | 5.7825 | 607.3163 | 5.7825 | 928.9288 | 5.7825 | 1167.8313 | 5.7825 | 1449.5895 | 5.8001 | 1843.4565
5.8177 | 240.0766 |  | 5.8177 | 420.195 | 5.8177 | 607.9934 | 5.8177 | 928.9288 | 5.8177 | 1168.6361 | 5.8177 | 1454.9994 | 5.8353 | 1846.0845
5.8528 | 240.1053 |  | 5.8528 | 420.559 | 5.8528 | 608.4453 | 5.8528 | 928.755 | 5.8528 | 1171.7147 | 5.8528 | 1456.9095 | 5.8704 | 1840.5962
5.888 | 239.9619 |  | 5.888 | 420.5185 | 5.888 | 609.4633 | 5.888 | 929.1895 | 5.888 | 1173.85 | 5.888 | 1459.992 | 5.9056 | 1840.1202
5.9231 | 239.7899 |  | 5.9231 | 420.559 | 5.9231 | 610.0864 | 5.9231 | 929.9714 | 5.9231 | 1175.0487 | 5.9231 | 1459.992 | 5.9407 | 1841.7872
5.9583 | 239.5893 |  | 5.9583 | 420.7612 | 5.9583 | 609.0672 | 5.9583 | 930.753 | 5.9583 | 1175.9796 | 5.9583 | 1456.9095 | 5.9759 | 1840.3582
5.9935 | 239.532 |  | 5.9935 | 421.004 | 5.9935 | 609.0672 | 5.9935 | 931.4476 | 5.9935 | 1175.9796 | 5.9935 | 1459.4759 | 6.011 | 1840.3582
6.0286 | 239.4461 |  | 6.0286 | 421.3279 | 6.0286 | 610.2565 | 6.0286 | 931.1872 | 6.0286 | 1175.9796 | 6.0286 | 1456.0166 | 6.0462 | 1842.9793
6.0638 | 239.5607 |  | 6.0638 | 422.2193 | 6.0638 | 610.4833 | 6.0638 | 930.8398 | 6.0638 | 1178.5006 | 6.0638 | 1455.6347 | 6.0813 | 1844.1727
6.0989 | 240.1339 |  | 6.0989 | 422.6249 | 6.0989 | 610.7669 | 6.0989 | 931.4476 | 6.0989 | 1178.633 | 6.0989 | 1457.549 | 6.1165 | 1846.8023
6.1341 | 239.5607 |  | 6.1341 | 422.9089 | 6.1341 | 611.7324 | 6.1341 | 931.7081 | 6.1341 | 1180.088 | 6.1341 | 1459.7338 | 6.1516 | 1847.9994
6.1692 | 239.1027 |  | 6.1692 | 422.9089 | 6.1692 | 611.9598 | 6.1692 | 933.1837 | 6.1692 | 1182.4619 | 6.1692 | 1462.456 | 6.1868 | 1850.6374
6.2044 | 239.4461 |  | 6.2044 | 423.9243 | 6.2044 | 612.2443 | 6.2044 | 934.7455 | 6.2044 | 1184.3018 | 6.2044 | 1464.8104 | 6.2219 | 1850.3973
6.2395 | 239.0455 |  | 6.2395 | 422.9089 | 6.2395 | 611.6756 | 6.2395 | 935.2661 | 6.2395 | 1186.7893 | 6.2395 | 1467.3177 | 6.2571 | 1850.6374
6.2747 | 239.0455 |  | 6.2747 | 422.2599 | 6.2747 | 608.6148 | 6.2747 | 934.7455 | 6.2747 | 1187.7029 | 6.2747 | 1467.3177 | 6.2922 | 1845.8454
6.3098 | 239.7039 |  | 6.3098 | 421.4494 | 6.3098 | 607.9934 | 6.3098 | 933.8779 | 6.3098 | 1190.4339 | 6.3098 | 1470.1161 | 6.3274 | 1845.8454
6.345 | 240.4783 |  | 6.345 | 422.9089 | 6.345 | 606.8653 | 6.345 | 935.0058 | 6.345 | 1190.9524 | 6.345 | 1470.3841 | 6.3625 | 1850.3973
6.3801 | 241.053 |  | 6.3801 | 424.1275 | 6.3801 | 607.2035 | 6.3801 | 935.0058 | 6.3801 | 1190.9524 | 6.3801 | 1470.3841 | 6.3977 | 1854.0031
6.4153 | 242.8982 |  | 6.4153 | 426.5304 | 6.4153 | 607.9934 | 6.4153 | 943.1614 | 6.4153 | 1190.0447 | 6.4153 | 1470.25 | 6.4329 | 1854.9665
6.4504 | 243.1873 |  | 6.4504 | 427.592 | 6.4504 | 608.9541 | 6.4504 | 944.3769 | 6.4504 | 1191.2114 | 6.4504 | 1470.921 | 6.468 | 1856.6542
6.4856 | 243.3609 |  | 6.4856 | 427.592 | 6.4856 | 610.0297 | 6.4856 | 945.7667 | 6.4856 | 1191.9876 | 6.4856 | 1471.8631 | 6.5032 | 1862.2158
6.5207 | 243.8532 |  | 6.5207 | 426.8569 | 6.5207 | 611.3915 | 6.5207 | 946.3749 | 6.5207 | 1195.5926 | 6.5207 | 1476.2135 | 6.5383 | 1868.2896
6.5559 | 244.2301 |  | 6.5559 | 426.2041 | 6.5559 | 613.1553 | 6.5559 | 950.2021 | 6.5559 | 1198.0216 | 6.5559 | 1477.588 | 6.5735 | 1872.9254
6.591 | 244.6655 |  | 6.591 | 426.2041 | 6.591 | 615.7258 | 6.591 | 949.9409 | 6.591 | 1201.8307 | 6.591 | 1480.0807 | 6.6086 | 1869.0204
6.6262 | 244.9269 |  | 6.6262 | 427.3061 | 6.6262 | 615.783 | 6.6262 | 950.2021 | 6.6262 | 1204.6176 | 6.6262 | 1480.0807 | 6.6438 | 1871.948
6.6613 | 245.1304 |  | 6.6613 | 428.4916 | 6.6613 | 615.783 | 6.6613 | 950.8988 | 6.6613 | 1205.3772 | 6.6613 | 1480.0807 | 6.6789 | 1876.8426
6.6965 | 245.5086 |  | 6.6965 | 429.4744 | 6.6965 | 617.1015 | 6.6965 | 950.2021 | 6.6965 | 1206.0101 | 6.6965 | 1479.1084 | 6.7141 | 1877.333
6.7316 | 245.5086 |  | 6.7316 | 429.4744 | 6.7316 | 617.1015 | 6.7316 | 949.6798 | 6.7316 | 1205.8835 | 6.7316 | 1478.4162 | 6.7492 | 1877.333
6.7668 | 245.5086 |  | 6.7668 | 428.9828 | 6.7668 | 619.2291 | 6.7668 | 946.1142 | 6.7668 | 1207.7814 | 6.7668 | 1478.4162 | 6.7844 | 1876.8426
6.802 | 245.5086 |  | 6.802 | 428.0008 | 6.802 | 620.3249 | 6.802 | 945.5929 | 6.802 | 1208.7931 | 6.802 | 1479.3858 | 6.8195 | 1876.8426
6.8371 | 245.9455 |  | 6.8371 | 427.7964 | 6.8371 | 621.1336 | 6.8371 | 945.2454 | 6.8371 | 1206.2632 | 6.8371 | 1479.3858 | 6.8547 | 1876.5974
6.8723 | 246.4995 |  | 6.8723 | 428.328 | 6.8723 | 621.5384 | 6.8723 | 943.7691 | 6.8723 | 1206.2632 | 6.8723 | 1481.1963 | 6.8898 | 1876.5974
6.9074 | 247.698 |  | 6.9074 | 429.1466 | 6.9074 | 621.4806 | 6.9074 | 943.4218 | 6.9074 | 1208.2873 | 6.9074 | 1483.8656 | 6.925 | 1877.0878
6.9426 | 247.8737 |  | 6.9426 | 432.2256 | 6.9426 | 620.6136 | 6.9426 | 940.6446 | 6.9426 | 1208.2873 | 6.9426 | 1485.1399 | 6.9601 | 1877.5784
6.9777 | 247.9029 |  | 6.9777 | 432.8431 | 6.9777 | 621.5384 | 6.9777 | 940.6446 | 6.9777 | 1208.2873 | 6.9777 | 1488.282 | 6.9953 | 1877.8237
7.0129 | 247.698 |  | 7.0129 | 432.8431 | 7.0129 | 621.3649 | 7.0129 | 940.6446 | 7.0129 | 1210.31 | 7.0129 | 1491.9008 | 7.0304 | 1878.0691
7.048 | 247.2006 |  | 7.048 | 432.2256 | 7.048 | 620.4403 | 7.048 | 940.8181 | 7.048 | 1209.9308 | 7.048 | 1491.9008 | 7.0656 | 1873.4144
7.0832 | 245.8581 |  | 7.0832 | 427.592 | 7.0832 | 619.8056 | 7.0832 | 940.2976 | 7.0832 | 1209.046 | 7.0832 | 1490.592 | 7.1007 | 1873.4144
7.1183 | 244.8688 |  | 7.1183 | 426.9385 | 7.1183 | 618.7684 | 7.1183 | 938.5625 | 7.1183 | 1206.6428 | 7.1183 | 1489.0016 | 7.1359 | 1873.1699
7.1535 | 244.8688 |  | 7.1535 | 424.1275 | 7.1535 | 618.0206 | 7.1535 | 937.6083 | 7.1535 | 1204.8708 | 7.1535 | 1486.9918 | 7.171 | 1877.333
7.1886 | 245.7124 |  | 7.1886 | 424.0462 | 7.1886 | 617.446 | 7.1886 | 935.6131 | 7.1886 | 1203.4778 | 7.1886 | 1484.4311 | 7.2062 | 1873.1699
7.2238 | 246.8791 |  | 7.2238 | 425.3076 | 7.2238 | 617.5609 | 7.2238 | 933.9647 | 7.2238 | 1202.9711 | 7.2238 | 1485.1399 | 7.2414 | 1878.3146
7.2589 | 247.3176 |  | 7.2589 | 426.5304 | 7.2589 | 617.3886 | 7.2589 | 934.1382 | 7.2589 | 1205.2506 | 7.2589 | 1483.5833 | 7.2765 | 1878.8057
7.2941 | 249.4292 |  | 7.2941 | 427.6738 | 7.2941 | 617.6758 | 7.2941 | 935.4396 | 7.2941 | 1207.7814 | 7.2941 | 1483.3013 | 7.3117 | 1870.7274
7.3292 | 249.7823 |  | 7.3292 | 428.6144 | 7.3292 | 616.5279 | 7.3292 | 936.3071 | 7.3292 | 1208.9195 | 7.3292 | 1484.1482 | 7.3468 | 1870.4834
7.3644 | 249.0177 |  | 7.3644 | 427.8781 | 7.3644 | 615.1535 | 7.3644 | 936.0469 | 7.3644 | 1208.9195 | 7.3644 | 1484.1482 | 7.382 | 1870.7274
7.3995 | 247.1714 |  | 7.3995 | 426.5304 | 7.3995 | 612.2443 | 7.3995 | 934.4853 | 7.3995 | 1210.0572 | 7.3995 | 1483.3013 | 7.4171 | 1871.7038
7.4347 | 245.334 |  | 7.4347 | 426.9385 | 7.4347 | 611.6756 | 7.4347 | 933.2704 | 7.4347 | 1206.5163 | 7.4347 | 1483.5833 | 7.4523 | 1874.6377
7.4698 | 244.2591 |  | 7.4698 | 425.5927 | 7.4698 | 611.4483 | 7.4698 | 933.0101 | 7.4698 | 1202.4643 | 7.4698 | 1484.1482 | 7.4874 | 1877.333
7.505 | 243.0716 |  | 7.505 | 425.552 | 7.505 | 611.5619 | 7.505 | 932.6629 | 7.505 | 1202.4643 | 7.505 | 1485.1399 | 7.5226 | 1877.8237
7.5401 | 242.436 |  | 7.5401 | 425.104 | 7.5401 | 612.1874 | 7.5401 | 933.0969 | 7.5401 | 1202.9711 | 7.5401 | 1485.1399 | 7.5577 | 1877.8237
7.5753 | 242.1186 |  | 7.5753 | 424.6969 | 7.5753 | 614.4105 | 7.5753 | 935.1793 | 7.5753 | 1203.0978 | 7.5753 | 1485.1399 | 7.5929 | 1878.8057
7.6105 | 242.436 |  | 7.6105 | 424.5342 | 7.6105 | 615.2679 | 7.6105 | 938.5625 | 7.6105 | 1204.8708 | 7.6105 | 1487.9947 | 7.628 | 1878.8057
7.6456 | 243.1584 |  | 7.6456 | 425.104 | 7.6456 | 614.1821 | 7.6456 | 940.2976 | 7.6456 | 1206.7693 | 7.6456 | 1487.5644 | 7.6632 | 1879.0513
7.6808 | 243.8242 |  | 7.6808 | 425.3483 | 7.6808 | 614.1821 | 7.6808 | 943.335 | 7.6808 | 1221.6637 | 7.6808 | 1488.282 | 7.6983 | 1886.1943
7.7159 | 244.4042 |  | 7.7159 | 426.5304 | 7.7159 | 615.0963 | 7.7159 | 945.6798 | 7.7159 | 1222.7969 | 7.7159 | 1491.3182 | 7.7335 | 1890.1522
7.7511 | 244.956 |  | 7.7511 | 427.7146 | 7.7511 | 615.0963 | 7.7511 | 948.3746 | 7.7511 | 1225.188 | 7.7511 | 1494.539 | 7.7686 | 1895.1166
7.7862 | 246.0329 |  | 7.7862 | 428.6963 | 7.7862 | 616.5852 | 7.7862 | 949.3316 | 7.7862 | 1225.3138 | 7.7862 | 1500.5033 | 7.8038 | 1899.8503
7.8214 | 247.0252 |  | 7.8214 | 429.0238 | 7.8214 | 616.9294 | 7.8214 | 952.7292 | 7.8214 | 1225.5654 | 7.8214 | 1502.6194 | 7.8389 | 1911.8845
7.8565 | 247.9908 |  | 7.8565 | 429.5154 | 7.8565 | 617.5034 | 7.8565 | 954.2126 | 7.8565 | 1227.2003 | 7.8565 | 1503.3776 | 7.8741 | 1911.8845
7.8917 | 249.5763 |  | 7.8917 | 429.7613 | 7.8917 | 618.4807 | 7.8917 | 954.9112 | 7.8917 | 1228.0803 | 7.8917 | 1503.9852 | 7.9092 | 1910.8776
7.9268 | 249.7823 |  | 7.9268 | 429.5154 | 7.9268 | 618.5382 | 7.9268 | 956.5721 | 7.9268 | 1229.4627 | 7.9268 | 1505.5077 | 7.9444 | 1911.6327
7.962 | 249.8411 |  | 7.962 | 429.0238 | 7.962 | 618.5382 | 7.962 | 957.3597 | 7.962 | 1230.4678 | 7.962 | 1507.648 | 7.9795 | 1913.3964
7.9971 | 249.6057 |  | 7.9971 | 428.7781 | 7.9971 | 620.3249 | 7.9971 | 957.6223 | 7.9971 | 1229.8397 | 7.9971 | 1507.4947 | 8.0147 | 1913.3964
8.0323 | 249.194 |  | 8.0323 | 428.5326 | 8.0323 | 621.2492 | 8.0323 | 957.3597 | 8.0323 | 1226.5716 | 8.0323 | 1507.3416 | 8.0499 | 1921.9957
8.0674 | 246.996 |  | 8.0674 | 427.6329 | 8.0674 | 621.4806 | 8.0674 | 954.4745 | 8.0674 | 1222.7969 | 8.0674 | 1504.5935 | 8.085 | 1924.7894
8.1026 | 246.996 |  | 8.1026 | 427.7555 | 8.1026 | 621.8856 | 8.1026 | 950.8988 | 8.1026 | 1212.3314 | 8.1026 | 1503.3776 | 8.1202 | 1926.5701
8.1377 | 246.996 |  | 8.1377 | 427.1835 | 8.1377 | 622.8126 | 8.1377 | 948.2876 | 8.1377 | 1212.3314 | 8.1377 | 1499.2992 | 8.1553 | 1921.9957
8.1729 | 247.8151 |  | 8.1729 | 427.3469 | 8.1729 | 622.1172 | 8.1729 | 946.2011 | 8.1729 | 1212.3314 | 8.1729 | 1497.2054 | 8.1905 | 1919.2076
8.208 | 249.9295 |  | 8.208 | 428.7372 | 8.208 | 622.2331 | 8.208 | 944.4637 | 8.208 | 1227.5775 | 8.208 | 1495.1291 | 8.2256 | 1919.2076
8.2432 | 250.4599 |  | 8.2432 | 431.732 | 8.2432 | 621.712 | 8.2432 | 944.0296 | 8.2432 | 1231.9749 | 8.2432 | 1497.2054 | 8.2608 | 1928.863
8.2783 | 249.9 |  | 8.2783 | 434.4514 | 8.2783 | 621.0758 | 8.2783 | 944.0296 | 8.2783 | 1234.6109 | 8.2783 | 1496.3134 | 8.2959 | 1928.863
8.3135 | 249.7823 |  | 8.3135 | 436.0222 | 8.3135 | 619.9786 | 8.3135 | 944.0296 | 8.3135 | 1240.128 | 8.3135 | 1502.6194 | 8.3311 | 1929.118
8.3486 | 249.9295 |  | 8.3486 | 436.1878 | 8.3486 | 619.1139 | 8.3486 | 947.5919 | 8.3486 | 1241.005 | 8.3486 | 1504.4414 | 8.3662 | 1932.1817
8.3838 | 249.8411 |  | 8.3838 | 435.8981 | 8.3838 | 618.9412 | 8.3838 | 953.1653 | 8.3838 | 1237.245 | 8.3838 | 1506.7294 | 8.4014 | 1930.9044
8.419 | 250.3125 |  | 8.419 | 434.9883 | 8.419 | 618.9412 | 8.419 | 953.427 | 8.419 | 1235.2382 | 8.419 | 1503.0742 | 8.4365 | 1929.118
8.4541 | 250.7844 |  | 8.4541 | 434.8231 | 8.4541 | 619.7479 | 8.4541 | 953.5143 | 8.4541 | 1233.6069 | 8.4541 | 1503.0742 | 8.4717 | 1932.1817
8.4893 | 254.0446 |  | 8.4893 | 434.9883 | 8.4893 | 619.8056 | 8.4893 | 952.3804 | 8.4893 | 1228.7087 | 8.4893 | 1505.3552 | 8.5068 | 1932.1817
8.5244 | 254.4913 |  | 8.5244 | 435.5258 | 8.5244 | 622.0593 | 8.5244 | 951.0731 | 8.5244 | 1224.6848 | 8.5244 | 1506.7294 | 8.542 | 1932.1817
8.5596 | 254.6702 |  | 8.5596 | 435.6498 | 8.5596 | 623.9734 | 8.5596 | 950.8988 | 8.5596 | 1223.8039 | 8.5596 | 1507.8012 | 8.5771 | 1936.277
8.5947 | 255.1176 |  | 8.5947 | 436.6848 | 8.5947 | 623.9734 | 8.5947 | 951.1602 | 8.5947 | 1223.5522 | 8.5947 | 1511.1864 | 8.6123 | 1939.8702
8.6299 | 254.6702 |  | 8.6299 | 436.6848 | 8.6299 | 623.9734 | 8.6299 | 952.3804 | 8.6299 | 1223.5522 | 8.6299 | 1512.114 | 8.6474 | 1943.4726
8.665 | 254.4913 |  | 8.665 | 435.8981 | 8.665 | 624.5548 | 8.665 | 954.8239 | 8.665 | 1223.5522 | 8.665 | 1512.8885 | 8.6826 | 1946.8258
8.7002 | 254.6403 |  | 8.7002 | 436.4363 | 8.7002 | 625.0785 | 8.7002 | 954.9112 | 8.7002 | 1223.8039 | 8.7002 | 1515.3755 | 8.7177 | 1950
8.7353 | 254.6702 |  | 8.7353 | 435.8981 | 8.7353 | 625.6609 | 8.7353 | 957.885 | 8.7353 | 1224.5589 | 8.7353 | 1517.0931 | 8.7529 | 1950
8.7705 | 254.7596 |  | 8.7705 | 436.5191 | 8.7705 | 626.3607 | 8.7705 | 959.2873 | 8.7705 | 1226.9488 | 8.7705 | 1519.76 | 8.7881 | 1950
8.8056 | 255.1176 |  | 8.8056 | 437.3066 | 8.8056 | 628.6995 | 8.8056 | 959.7258 | 8.8056 | 1229.714 | 8.8056 | 1521.4936 | 8.8232 | 1950
8.8408 | 254.9385 |  | 8.8408 | 438.0535 | 8.8408 | 629.4617 | 8.8408 | 960.6914 | 8.8408 | 1230.5934 | 8.8408 | 1524.0266 | 8.8584 | 1950
8.8759 | 254.6702 |  | 8.8759 | 439.4251 | 8.8759 | 630.1075 | 8.8759 | 965.0935 | 8.8759 | 1232.3516 | 8.8759 | 1525.935 | 8.8935 | 1950
8.9111 | 254.5509 |  | 8.9111 | 439.7164 | 8.9111 | 630.1075 | 8.9111 | 966.0651 | 8.9111 | 1232.6027 | 8.9111 | 1526.5728 | 8.9287 | 1950
8.9462 | 254.5211 |  | 8.9462 | 440.8414 | 8.9462 | 631.46 | 8.9462 | 965.9767 | 8.9462 | 1235.2382 | 8.9462 | 1526.7324 | 8.9638 | 1950
8.9814 | 254.5509 |  | 8.9814 | 442.2191 | 8.9814 | 631.9902 | 8.9814 | 965.9767 | 8.9814 | 1237.4958 | 8.9814 | 1529.6138 | 8.999 | 1950
9.0165 | 255.4759 |  | 9.0165 | 443.3485 | 9.0165 | 632.5798 | 9.0165 | 966.0651 | 9.0165 | 1240.6292 | 9.0165 | 1534.1294 | 9.0341 | 1950
9.0517 | 255.984 |  | 9.0517 | 443.3066 | 9.0517 | 632.4618 | 9.0517 | 967.0378 | 9.0517 | 1242.6333 | 9.0517 | 1536.4024 | 9.0693 | 1950
9.0868 | 256.2832 |  | 9.0868 | 444.2284 | 9.0868 | 632.9339 | 9.0868 | 968.5437 | 9.0868 | 1243.5098 | 9.0868 | 1541.3074 | 9.1044 | 1950
9.122 | 257.1522 |  | 9.122 | 445.7816 | 9.122 | 632.6388 | 9.122 | 969.1646 | 9.122 | 1244.1357 | 9.122 | 1542.4587 | 9.1396 | 1950
9.1571 | 257.2122 |  | 9.1571 | 447.2124 | 9.1571 | 633.2882 | 9.1571 | 971.0308 | 9.1571 | 1249.3901 | 9.1571 | 1547.0888 | 9.1747 | 1950
9.1923 | 257.1822 |  | 9.1923 | 447.0439 | 9.1923 | 635.7151 | 9.1923 | 971.0308 | 9.1923 | 1250.3902 | 9.1923 | 1547.2549 | 9.2099 | 1950
9.2275 | 257.2122 |  | 9.2275 | 447.2124 | 9.2275 | 636.368 | 9.2275 | 970.9418 | 9.2275 | 1252.6396 | 9.2275 | 1547.7536 | 9.245 | 1950
9.2626 | 257.4523 |  | 9.2626 | 448.2244 | 9.2626 | 638.9871 | 9.2626 | 971.2088 | 9.2626 | 1257.385 | 9.2626 | 1550.4209 | 9.2802 | 1950
9.2978 | 257.4523 |  | 9.2978 | 448.2244 | 9.2978 | 638.9871 | 9.2978 | 972.3669 | 9.2978 | 1258.2587 | 9.2978 | 1546.259 | 9.3153 | 1950
9.3329 | 257.7827 |  | 9.3329 | 448.5198 | 9.3329 | 637.1406 | 9.3329 | 971.2088 | 9.3329 | 1257.385 | 9.3329 | 1544.9341 | 9.3505 | 1950
9.3681 | 257.7827 |  | 9.3681 | 448.5198 | 9.3681 | 635.7151 | 9.3681 | 969.7861 | 9.3681 | 1251.3901 | 9.3681 | 1540.1588 | 9.3856 | 1950
9.4032 | 257.0922 |  | 9.4032 | 446.8755 | 9.4032 | 629.6964 | 9.4032 | 964.6523 | 9.4032 | 1245.5125 | 9.4032 | 1527.8509 | 9.4208 | 1950
9.4384 | 253.0637 |  | 9.4384 | 446.5387 | 9.4384 | 621.3071 | 9.4384 | 946.9833 | 9.4384 | 1216.4964 | 9.4384 | 1492.4846 | 9.4559 | 1874.3929
9.4735 | 251.9076 |  | 9.4735 | 441.2585 | 9.4735 | 615.7258 | 9.4735 | 938.1288 | 9.4735 | 1184.9575 | 9.4735 | 1466.9203 | 9.4911 | 1794.6246
9.5087 | 248.2839 |  | 9.5087 | 435.9395 | 9.5087 | 590.5988 | 9.5087 | 830.6758 | 9.5087 | 936.9143 | 9.5087 | 986.9928 | 9.5262 | 1791.8846
9.5438 | 248.0494 |  | 9.5438 | 435.9395 | 9.5438 | 588.1973 | 9.5438 | 810.7581 | 9.5438 | 926.4075 | 9.5438 | 980.7159